## ScopeCosts
|  |  | Total Cost = | 208630.755
PARTS SUMMARY
Lasers
Component | Supplier/Vendor | Price |  | Quotes | Detailed description |  |  | Date Requested | Date Recieved | Notes
Table | Newport, Stanford | 14163.8 |  |  | Optics table and setup
Lasers | Coherent, MPB | 82796.985 |  |  | 1W+ lasers for STORM (488, 560, 647, 750) and 405 for activation
Camera | Hamamatsu | 19475 |  |  | CMOS camera and calibration
Microscope Body | Nikon | 20737.2 |  |  | Microscope body, Objectives
Microscope Stage | Ludl | 30852.2 |  |  | Microscope stage, XY and Piezo Z
Laser Optics | Thorlabs, Semrock, Chroma, Edmund | 10534.5 |  |  | Optics to combine beams
AOTF | Gooch and Housego | 5911.46 |  |  | AOTF to adjust laser power
Focus Lock | Thorlabs | 7063.08 |  |  | Focus lock to keep system in registration
Timing Control | National Instruments | 3106.2 |  |  | National Instruments card and components to synchronize AOTF, camera, and shutters
Computer | Dell | 5105.35 |  |  | Dell Precision Tower (5810), with SSD, RAM and CPU for high speed imaging
Misc Parts | Thorlabs, Amazon | 64.94 |  |  | Other parts (power meters, tools, safety, cables, etc).
Fluidics System | Hamilton, Gilson, Bioptechs | 8128.77 |  |  | Peristaltic pump, valve controllers, tubing, connectors and cables for automated fluid handling
Enclosure | McMaster | 691.27 |  |  | Laser-proof and light proof enclosure to protect system in operation and protect users.
ITEMIZED PARTS
Supplier/Vendor* | Catalog#/Supplier item* | Qty.* | Unit of Measure* | Unit Price* | Item Description* | Total Cost | Quote Requested | Date Requested | Date Recieved | Notes / URL
Computer | 5105.35
Dell | T5810 | 1 | each | 3532.88 | see configuration in quote | 3532.88 | 2017-06-05
Amazon |  | 1 | each | 1572.47 | SAMSUNG 850 EVO 2.5" 4 TB SATA III 3-D Vertical Internal Solid State Drive (SSD) MZ-75E4T0B/AM | 1572.47 | 2017-06-26 |  |  | https://www.amazon.com/gp/product/B01G844OOO/ref=oh_aui_detailpage_o01_s00?ie=UTF8&psc=1
Table | 14163.8 |  |  | (10% discount)
Newport | RS4000-36-12 | 1 | each | 5031 | RS4000 Series Optical Table, 3 ft x 6 ft x 12 in | 5031 | 2017-04-14
Newport | S-2000A-423.5 | 1 | set of 4 | 3340.8 | Set of four S-2000 23.5 in tall standard isolators (legs) | 3340.8 | 2017-04-14
Newport | ERS2010-323.5 | 3 | each | 550.8 | Earthquake Restraints, 23.5 in. tall, mounting hardware to be ordered seperately | 1652.4 | 2017-04-14
Newport | ERS2010-3-KT | 1 | each | 174.6 | Mounting Kit for Earthquake restraints | 174.6 | 2017-04-14
Newport | Shipping | 1 | each | 515 | Shipping costs | 515 | 2017-04-14
Peninsula Crane and Rigging | 408 294 4841 | 1 | each | 2290 | Riggers to setup table - Joe Bauer and Richard White | 2290
Peninsula Crane and Rigging | 408 294 4841 | 1 | each | 1160 | Riggers to install seismic bracing | 1160
Microscope Stage | 30852.2 | - | - | (5% discount) | -
Ludl Elect. Prod. | 96A602 | 1 | each | 11519.7 | Ludl Piezo Z-axis insert, 350um for MAC6000 (use 96S107 Stage) | 11519.7 | 2017-04-25
Ludl Elect. Prod. | 73006063 | 1 | each | 1347.1 | Ludl Piezo Z-axis driver, MAC6000 | 1347.1 | 2017-04-25
Ludl Elect. Prod. | 96S107-N3-LE2 | 1 | each | 12903.85 | Stage, Flat Top pz Nikon Ti2-U St Motors/Inverted/Lin-Enc 50n | 12903.85 | 2017-04-25
Ludl Elect. Prod. | 996052 | 1 | each | 4341.5 | Controller, MAC6000 XY Stage Stepper, Includes Joystick | 4341.5 | 2017-04-25
Ludl Elect. Prod. | 99A074 | 1 | each | 740.05 | Adjustable insert for slide/petri dish – inverted stage | 740.05 | 2017-04-25 |  |  | http://ludl.com/stage-accessories/inventory/--13/
Lasers | 82796.985
Coherent | 1284370 | 1 | each | 4784.6 | OBIS 405nm LX 50mW LASER SYSTEM | 4784.6
Coherent | 1220124 | 1 | each | 8267.37 | OBIS 488nm LX 150 mW Laser System | 8267.37 | 2017-06-13
MPB | 2RU-VFL-P-2000-560-B1R | 1 | each | 22325 | 2000mW at 560nm Laser | 22325 | 2017-04-03
MPB | 2RU-VFL-P-2000-647-B1R | 1 | each | 24652.5 | 2000mW at 647nm Laser | 24652.5 | 2017-04-03
MPB | 2RU-VFL-P-500-750-B1R | 1 | each | 22400 | 500mW at 750 nm Laser | 22400
eMachineShop | MPB_mount.ems | 1 | set of 3 | 261.79 | Custom machined plate for MPB laser mount | 261.79 | 2017-05-26
eMachineShop | Coheren_mount.ems | 1 | set of 1 | 105.725 | Custom macined plate for OBIS laser mount | 105.725 | 2017-05-26
Camera | 19475 |  |  | (5% discount)
Hamamatsu via Nikon | C11440-22CUPLUS-KIT | 1 | each | 17575 | ORCA-Flash4.0 V2+ CMOS Camera Kit | 17575 | 2017-04-03
Hamamatsu via Nikon | CAMRA-1001-000-KIT | 1 | each | 1900 | Camera Link Interface Kit for C11440-22CU | 1900 | 2017-04-03
Laser Optics | 10534.5
Newport Optics | GBS-AR14 | 0 | each | 6210 | Refractive Beam Shaper, 430-700nm | 0
AVR (Semrock) | Di03-R405-t3-25D | 1 | each | 665 | 405 nm laser BrightLine® single-edge super-resolution laser dichroic beamsplitter, 25mm round, 3mm thick | 665 |  |  |  | https://www.semrock.com/FilterDetails.aspx?id=Di03-R405-t3-25x36 | combine beams | Get 25mm round 3mm
AVR (Semrock) | Di03-R488-t3-25D | 0 | each | 665 | 488 nm laser BrightLine® single-edge super-resolution laser dichroic beamsplitter, 25mm round, 3mm thick | 0 | 2017-06-16 |  |  | https://www.semrock.com/FilterDetails.aspx?id=Di03-R488-t3-25x36 | combine beams | Get 25mm round 3mm
AVR (Semrock) | Di03-R561-t3-25D | 1 | each | 665 | 561 nm laser BrightLine® single-edge super-resolution laser dichroic beamsplitter, 25mm round, 3mm thick | 517.75 | 2017-04-14 |  |  | https://www.semrock.com/FilterDetails.aspx?id=Di03-R561-t3-25x36 | combine beams | Get 25mm round 3mm
AVR (Semrock) | Di03-R635-t3-25D | 1 | each | 665 | 635 nm laser BrightLine® single-edge super-resolution laser dichroic beamsplitter, 25mm round, 3mm thick | 517.75 | 2017-04-14 |  |  | https://www.semrock.com/FilterDetails.aspx?id=Di03-R635-t3-25x36 | combine beams | Get 25mm round 3mm
Thorlabs | LA1433-A-ML | 3 | each | 42.3 | Ø1" Mounted N-BK7 Plano-Convex Lenses (AR Coating: 350 - 700 nm) f=150mm | 126.9 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=LA1433-A-ML | for laser zoom system to adjust beam size
Thorlabs | LA1509-A-ML | 3 | each | 43.6 | Ø1" Mounted N-BK7 Plano-Convex Lenses (AR Coating: 350 - 700 nm) f=100mm | 130.8 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=LA1509-A-ML | for laser zoom system to adjust beam size
Thorlabs | LA1433-B-ML | 1 | each | 42.3 | Ø1" Mounted N-BK7 Plano-Convex Lenses (AR Coating: 650 - 1050 nm) f =150mm | 42.3 |  |  |  |  | for laser zoom system to adjust beam size
Thorlabs | LA1509-B-ML | 1 | each | 43.6 | Ø1" Mounted N-BK7 Plano-Convex Lenses (AR Coating: 650 - 1050 nm) f = 100mm | 43.6 |  |  |  |  | for laser zoom system to adjust beam size
Thorlabs | LD2297-A | 3 | each | 41.21 | N-SF11 Bi-Concave Lens, Ø25.4 mm, f = -25.0 mm, ARC: 350-700 nm | 123.63 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=LD2297-A | for laser zoom system to adjust beam size
Thorlabs | LD2297-B | 1 | each | 41.21 | N-SF11 Bi-Concave Lens, Ø25.4 mm, f = -25.0 mm, ARC:650 - 1050 nm | 41.21 |  |  |  |  | for laser zoom system to adjust beam size
Thorlabs | WPH05M-488 | 1 | each | 418.2 | Ø1/2" Mounted Zero-Order Half-Wave Plate, Ø1" Mount, 488 nm | 418.2 |  |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=WPH05M-488 | To adjust beam polarization
Thorlabs | WPH05M-561 | 1 | each | 418.2 | Ø1/2" Mounted Zero-Order Half-Wave Plate, Ø1" Mount, 561 nm | 418.2 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=WPH05M-561 | To adjust beam polarization
Thorlabs | WPH05M-633 | 1 | each | 418.2 | Ø1/2" Mounted Zero-Order Half-Wave Plate, Ø1" Mount, 633 nm | 418.2 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=WPH05M-633 | To adjust beam polarization
Thorlabs | WPH05M-780 | 1 | each | 418.2 | Ø1/2" Mounted Zero-Order Half-Wave Plate, Ø1" Mount, 780 nm | 418.2 |  |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=WPH05M-780 | To adjust beam polarization
Thorlabs | CRM1 | 3 | each | 78.94 | Cage Rotation Mount for Ø1" Optics, SM1 Threaded, 8-32 Tap | 236.82 | 2017-05-03 |  | backordered | https://www.thorlabs.com/thorproduct.cfm?partnumber=CRM1 | for mounting half-wave plates
Thorlabs | CP02 | 15 | each | 16 | CP02 - SM1-Threaded 30 mm Cage Plate, 0.35" Thick, 2 Retaining Rings, 8-32 Tap | 240 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CP02 | to mount laser zoom
Thorlabs | ER10 | 12 | each | 16.44 | ER10 - Cage Assembly Rod, 10" (254.0 mm) Long, Ø6 mm | 197.28 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=ER10 | for laser zoom
Thorlabs | CF125-P5 | 8 | 5-pack | 41.5 | CF125-P5 - Short Clamping Fork, 1.24" Counterbored Slot, Universal, 5 Pack | 332 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CF125-P5 | for anchoring mounts
Thorlabs | KM100-E02 | 13 | each | 103.5 | KM100-E02 - Kinematic Mirror Mount for Ø1" Optics with Visible Laser Quality Mirror | 1345.5 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=KM100-E02 | mirrors
Thorlabs | KM200-E02 | 2 | each | 198 | KM200-E02 - Kinematic Mirror Mount for Ø2" Optics with Visible Laser Quality Mirror | 396 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=KM200-E02 | mirrors for expanded beam
Thorlabs | PH1.5E | 6 | each | 23 | PH1.5E - Ø1/2" Pedestal Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=1.69" | 138 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=PH1.5E | for mounting laser-zoom rail systems
Thorlabs | TR1.5-P5 | 2 | each | 22.37 | TR1.5-P5 - Ø1/2" Optical Post, SS, 8-32 Setscrew, 1/4"-20 Tap, L = 1.5", 5 Pack | 44.74 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=TR1.5-P5 | for mounting laser-zoom rail systems
Thorlabs | PH082E | 3 | each | 22.5 | PH082E - Ø1/2" Pedestal Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=1.00" | 67.5 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=PH082E | for mounting beam expander
Thorlabs | TR1 | 3 | each | 4.74 | TR1 - Ø1/2" Optical Post, SS, 8-32 Setscrew, 1/4"-20 Tap, L = 1" | 14.22 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=TR1 | for mounting beam expander
Edmund Optics | 49-395 | 1 | each | 145 | 50mm Dia. x 350mm FL, VIS-NIR, Achromatic Lens | 145 | 2017-05-02 |  |  | http://www.edmundoptics.com/optics/optical-lenses/achromatic-lenses/vis-nir-coated-achromatic-lenses/49395/ | for beam expander
Edmund Optics | 49-352 | 2 | each | 94 | 25mm Dia. x 30mm FL, VIS-NIR Coated, Achromatic Lens | 188 | 2017-05-02 |  |  | http://www.edmundoptics.com/optics/optical-lenses/achromatic-lenses/vis-nir-coated-achromatic-lenses/49352/ | for beam expander
Edmund Optics | 49-392 | 1 | each | 145 | 50mm Dia. x 200mm FL, VIS-NIR Coated, Achromatic Lens | 145 | 2017-05-02 |  |  | http://www.edmundoptics.com/optics/optical-lenses/achromatic-lenses/vis-nir-coated-achromatic-lenses/49392/ | for beam expander
Thorlabs | SM1L05-P5 | 1 | 5-pk | 59.8 | SM1 Lens Tube, 0.50" Thread Depth, SM1RR Retaining Ring, 5 Pack | 59.8 | 2017-05-03 |  |  |  | 2 for edmund 25mm achromatic
Thorlabs | ER12 | 8 | each | 16.44 | ER12 - Cage Assembly Rod, 12" (304.8 mm) Long, Ø6 mm | 131.52 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=ER12 | for beam expander
Thorlabs | CDA1 | 4 | each | 30 | CDA1 - Drop-In 30 mm to 60 mm Cage System Adapter, Qty. 1 | 120 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CDA1 | attach 1" optics in beam expander
Thorlabs | ER1-P4 | 2 | 4-pack | 18.77 | ER1-P4 - Cage Assembly Rod, 1" (25.4 mm) Long, Ø6 mm, 4 Pack | 37.54 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=ER1-P4 | attach 1" optics in beam expander and connect periscope
Thorlabs | ER05-P4 | 3 | 4-pack | 18.77 | ER05-P4 - Cage Assembly Rod, 0.5" Long, Ø6 mm, 4 Pack | 56.31 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=ER05-P4 | connect periscope
Thorlabs | ER025 | 8 | each | 5.05 | ER025 - Cage Assembly Rod, 1/4" Long, Ø6 mm | 40.4 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=ER025 | height flexibility for periscope
Thorlabs | CXY1 | 2 | each | 169.05 | CXY1 - 30 mm Cage XY Translator for Ø1" Optics with Quick-Release Mounting Carriage | 338.1 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CXY1 | attach 1" optics in beam expander
Thorlabs | XRN25C | 1 | each | 362.7 | XRN25C - Compact 25 mm Travel Linear Translation Stage, Rear Micrometer, 1/4"-20 Taps | 362.7 | 2017-05-03 |  | backordered | https://www.thorlabs.com/thorproduct.cfm?partnumber=XR25C | For TIRF
Thorlabs | PH3 | 1 | each | 9.17 | PH3 - Ø1/2" Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L = 3" | 9.17 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=PH3 | For mounting TIRF
Thorlabs | TR3 | 1 | each | 5.43 | TR3 - Ø1/2" Optical Post, SS, 8-32 Setscrew, 1/4"-20 Tap, L = 3" | 5.43 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=TR3 | For mounting TIRF
Thorlabs | CXY2 | 2 | each | 185 | CXY2 - 60 mm Cage System Translating Lens Mount for Ø2" Optics | 370 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CXY2 | for beam expander and back lens
Thorlabs | SM2L05 | 2 | each | 24.75 | SM2L05 - SM2 Lens Tube, 0.5" Thread Depth, One Retaining Ring Included | 49.5 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=SM2L05 | for mounting 2" diameter lenses
Thorlabs | LCPMA1 | 2 | each | 35 | LCPMA1 - Snap-On 60 mm Cage Mounting Bracket, #8 (M4) Slot | 70 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=LCPMA1 | for mounting periscope
Thorlabs | ERSCA | 1 | 4-pack | 42.14 | ERSCA-P4 - Rod Adapter for Ø6 mm ER Rods, 4 Pack | 42.14 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=ERSCA-P4 | for connecting periscope parts
Thorlabs | ER2 | 4 | each | 5.94 | ER2 - Cage Assembly Rod, 2" (50.8 mm) Long, Ø6 mm | 23.76 | 2017-05-03 |  |  |  | for periscope
Thorlabs | CV1530 | 1 | each | 179 | CV1530 - Vertical 30 mm Cage Clamp for Ø1.5" Posts | 179 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CV1530 | for holding periscope
Thorlabs | P6 | 1 | each | 50 | P6 - Ø1.5" Mounting Post, 1/4"-20 Taps, L = 6" | 50 | 2017-05-03 |  |  | https://www.thorlabs.com/newgrouppage9.cfm?objectgroup_id=1313&pn=P6#175 | mount periscope
Thorlabs | BBE2-E02 | 2 | each | 185 | BBE2-E02 - 2" Broadband Dielectric Elliptical Mirror, 400 - 750 nm | 370 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=BBE2-E02 | for periscope
Thorlabs | ND20A | 4 | each | 68.34 | ND20A - Reflective Ø25 mm ND Filter, SM1-Threaded Mount, Optical Density: 2.0 | 273.36 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=ND20A | reduce laser power for aligning beam
Thorlabs | RS1.5P | 10 | each | 23.75 | RS1.5P - Ø1" Pedestal Pillar Post, 1/4"-20 Taps, L = 1.5" | 237.5 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=RS1.5P | Mounts for 2" mirrors and for laser plates (to match 2" mounts for 1" mirrors)
Thorlabs | PB4 | 1 | each | 12.4 | PB4 - Studded Pedestal Base Adapter, 1/4"-20 Thread | 12.4 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=PB4 | mount periscope
Thorlabs | PF175 | 1 | each | 17.47 | PF175 - Clamping Fork for Ø1.5" Pedestal Post or Post Pedestal Base Adapter, Universal | 17.47 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=PF175 | mount periscope
Thorlabs | KCB2E | 1 | each | 225 | KCB2E - Right-Angle Kinematic Elliptical Mirror Mount, 60 mm Cage System and SM2 Compatible, 8-32 and 1/4"-20 Mounting Holes | 225 |  |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=KCB2E | for periscope
Thorlabs | LCP01 | 1 | each | 37.1 | 60 mm Cage Plate, SM2 Threaded, 0.5" Thick, 8-32 Tap (Two SM2RR Retaining Rings Included) | 37.1 |  |  |  |  | for beam expansion cage and periscope
Thorlabs | SM1V05 | 1 | each | 30.25 | SM1V05 - Ø1" Adjustable Lens Tube, 0.31" Travel Range | 30.25 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=SM1V05 | for beam expansion after AOTF before beam-shaper
Thorlabs | SM1L10 | 1 | each | 14.25 | SM1L10 - SM1 Lens Tube, 1.00" Thread Depth, One Retaining Ring Included | 14.25 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=SM1L10 | for beam expansion after AOTF before beam-shaper
AOTF | 5911.46
Gooch and Housego | 97-03309-01 | 1 | each | 2450 | AOTF, 488-800 nm | 2450 | 2017-04-14 |  | 2017-04-25
Gooch and Housego | 97-03926-12 | 1 | each | 2875 | DDS driver (8-chn) for AOTF | 2875 | 2017-04-14 |  | 2017-04-25
Gooch and Housego | 38-00028-01 | 1 | each | 35 | 24V Power Supply | 35 | 2017-04-14 |  | 2017-04-25
Gooch and Housego | Software CD | 1 | each | 0 | Software CD | 0 | 2017-04-14 |  | 2017-04-25
Gooch and Housego | 31-81822-02 | 1 | each | 95 | 40 pin MDR cable, 2 meter | 95 | 2017-04-14 |  | 2017-04-25
Gooch and Housego | 86-00136-02 | 1 | each | 195 | temperature sensor | 195 | 2017-04-14 |  | 2017-04-25
Thorlabs | T3283 | 1 | each | 6.8 | BNC Adapter - Straight Adapter (F-F) | 6.8 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=T3283 | for joining SMA cables to make long SMA
Thorlabs | CA2848 | 2 | each | 17.5 | CA2848 - SMA Coaxial Cable, SMA Male to BNC Male, 48" (1219 mm) | 35 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CA2848 | SMA cables to join
Thorlabs | PB4 | 1 | each | 12.4 | PB4 - Studded Pedestal Base Adapter, 1/4"-20 Thread | 12.4 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=PB4 | mount AOTF
Thorlabs | PF175 | 1 | each | 17.47 | PF175 - Clamping Fork for Ø1.5" Pedestal Post or Post Pedestal Base Adapter, Universal | 17.47 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=PF175 | mount AOTF
Thorlabs | TR8 | 1 | each | 8.12 | TR8 - Ø1/2" Optical Post, SS, 8-32 Setscrew, 1/4"-20 Tap, L = 8" | 8.12 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=TR8 | For mounting AOTF
Thorlabs | RA90 | 2 | each | 9.76 | RA90 - Right-Angle Clamp for Ø1/2" Posts, 3/16" Hex | 19.52 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=RA90 | For mounting AOTF
eMachineShop | ATOF_mount.ems | 1 | each | 157.16 | aluminum mounting plate for AOTF, custom machined | 157.16 | 2017-05-26
Amazon |  | 1 | each | 4.99 | AmazonBasics USB 2.0 Cable - A-Male to B-Male - 6 Feet (1.8 Meters) | 4.99 |  |  |  | AmazonBasics USB 2.0 Cable - A-Male to B-Male - 6 Feet (1.8 Meters) | to run AOTF
Focus Lock | 7063.08
Thorlabs | SM1A30 | 1 | each | 60.8 | Nikon Eclipse Ti Microscope Epi-Fluorescence Port Adapter, Internal SM1 Threads, 30 and 60 mm Cage Compatibility | 60.8 | 2017-05-03 |  |  |  | for attaching focus lock system to scope
Thorlabs | LDC210C | 1 | each | 1040 | Benchtop LD Current Controller, ±1 A HV | 1040 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=LDC210C | power supply for IR laser
Thorlabs | LDM9LP | 1 | each | 617 | LD/TEC Mount for Thorlabs Fiber-Pigtailed Laser Diodes | 617 | 2017-05-03 |  |  | https://www.thorlabs.com/newgrouppage9.cfm?objectgroup_id=4839&pn=LDM9LP | Mount and temp control for IR laser
Thorlabs | LP940-SF30 | 1 | each | 690 | 940 nm 30 mW Thorlabs Fiber-Pigtailed Laser Diode | 690 | 2017-05-03 |  | backordered |  | IR laser fiber |  | 980 nm, 15 mW, E Pin Code, SM Fiber-Pigtailed Laser Diode, FC/PC
Thorlabs | NENIR10A | 1 | each | 57 | Ø25 mm NIR Absorptive ND Filter, SM1-Threaded Mount, OD: 1.0 | 57 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=NENIR10A | modulate laser power
Thorlabs | NENIR20A | 1 | each | 57 | Ø25 mm NIR Absorptive ND Filter, SM1-Threaded Mount, OD: 2.0 | 57 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=NENIR20A | modulate laser power
Thorlabs | DCC1545M | 1 | each | 355 | High Resolution USB2.0 CMOS Camera, 1280 x 1024, Monochrome Sensor | 355 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=DCC1545M | IR camera
Thorlabs | PAFA-X-4-B | 1 | each | 515 | Achromatic FiberPort, FC/PC & FC/APC, f = 4.00 mm, 650 - 1050 nm, Ø0.87 mm Waist | 515 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=PAFA-X-4-B | Fiber port
Thorlabs | CP08FP | 1 | each | 28 | FiberPort and LaserPort Adapter for 30 mm Cage System, Enhanced Clamping | 28 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CP08FP | Mount for fiber port
Thorlabs | LA1608-B-ML | 1 | each | 43.6 | Ø1" N-BK7 Plano-Convex Lens, SM1-Threaded Mount, f = 75.0 mm, ARC: 650-1050 nm | 43.6 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=LA1608-B-ML | Focus spots onto camera
Thorlabs | LA1172-B-ML | 1 | each | 41.9 | Ø1" N-BK7 Plano-Convex Lens, SM1-Threaded Mount, f = 400.0 mm, ARC: 650-1050 nm | 41.9 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=LA1172-B-ML | Focus spots onto objective (might try different lenses)
Thorlabs | CXY1 | 1 | each | 169.05 | 30 mm Cage XY Translator for Ø1" Optics with Quick-Release Mounting Carriage | 169.05 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CXY1 | Mount and adjust lens to focus on objective
Thorlabs | CT1 | 1 | each | 418 | 1/2" Translation Stage for Cage System | 418 | 2017-05-03 |  |  | https://www.thorlabs.com/newgrouppage9.cfm?objectgroup_id=1895 | Offset control for final lens focusing onto camera
Thorlabs | CT101 | 1 | each | 43.1 | Optic Mount for use with CT1 in 30 mm Cage System | 43.1 | 2017-05-03 |  |  | https://www.thorlabs.com/newgrouppage9.cfm?objectgroup_id=1895 | Offset control for final lens focusing onto camera
Thorlabs | BB1-E03 | 2 | each | 75.1 | BB1-E03 - Ø1" Broadband Dielectric Mirror, 750-1100 nm | 150.2 | 2017-05-03 |  |  | http://www.thorlabs.com/thorproduct.cfm?partnumber=BB1-E03 | IR mirrors
Thorlabs | KCB1C | 2 | each | 139.94 | Right-Angle Kinematic Mirror Mount with Tapped Cage Rod Holes, 30 mm Cage System and SM1 Compatible, 8-32 and 1/4"-20 Mounting Holes | 279.88 | 2017-05-03 |  |  | https://www.thorlabs.com/newgrouppage9.cfm?objectgroup_id=6813&pn=KCB1C | Mirror mounts
Thorlabs | C6WA | 2 | each | 41.2 | C6W Platform Adapter Kit | 82.4 | 2017-05-03 |  |  |  | for rotating beam splitter
Thorlabs | C6W | 2 | each | 61.2 | 30 mm Cage Cube, Ø6 mm Through Holes | 122.4 | 2017-05-03 |  |  |  | for rotating beam splitter
Thorlabs | B4CRP | 2 | each | 283 | 30 mm Cage Cube Precision Kinematic Rotation Platform | 566 | 2017-05-03 |  |  |  | for rotating beam splitter
Thorlabs | BS017 | 2 | each | 190 | 50:50 Non-Polarizing Beamsplitter Cube, 700-1100 nm, 20 mm | 380 | 2017-05-03 |  |  |  | for rotating beam splitter
Thorlabs | PM3 | 2 | each | 18 | Small Adjustable Clamping Arm, 6-32 Threaded Post | 36 | 2017-05-03 |  |  |  | for rotating beam splitter
Thorlabs | ER2-P4 | 6 | 4-pack | 22.57 | Cage Assembly Rod, 2" (50.8 mm) Long, Ø6 mm, 4 Pack | 135.42 | 2017-05-03 |  |  | http://www.thorlabs.com/thorproduct.cfm?partnumber=ER2-P4 | cage system for IR optics
Thorlabs | ER6-P4 | 1 | 4-pack | 31.42 | ER6-P4 - Cage Assembly Rod, 6" (152.4 mm) Long, Ø6 mm, 4 Pack | 31.42 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=ER6-P4 | cage connector to camera
Thorlabs | ERSCA | 12 | each | 10.8 | ERSCA - Rod Adapter for Ø6 mm ER Rods | 129.6 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=ERSCA
Thorlabs | CP02 | 2 | each | 16 | CP02 - SM1-Threaded 30 mm Cage Plate, 0.35" Thick, 2 Retaining Rings, 8-32 Tap | 32 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CP02
Thorlabs | CM1-BP145B3 | 1 | each | 268 | 30 mm Cage Cube-Mounted Pellicle Beamsplitter, 45:55 (R:T), 1 - 2 µm | 268 | 2017-05-03 |  |  | http://www.thorlabs.com/thorproduct.cfm?partnumber=CM1-BP145B3 | To send light to camera, pellicle to reduce back-scatter from final beam splitter
Thorlabs | NENIR60A | 1 | each | 57 | NENIR60A - Ø25 mm NIR Absorptive ND Filter, SM1-Threaded Mount, OD: 6.0 | 57 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=NENIR60A | to deflect back-scatter away from camera
Thorlabs | SM1L03T | 1 | each | 48 | SM1L03T - 10° Angled Optic Mount for 1" Optics with SM1 Threading | 48 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=SM1L03T | to deflect back-scatter away from camera
Thorlabs | SM1CP2 | 1 | each | 17 | SM1CP2 - Externally SM1-Threaded End Cap | 17 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=SM1CP2 | to deflect back-scatter away from camera
Thorlabs | SM1A9 | 1 | each | 18.75 | SM1A9 - Adapter with External C-Mount Threads and Internal SM1 Threads | 18.75 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=SM1A9 | to mount IR camera
Thorlabs | PH4E | 3 | each | 24.5 | PH4E - Ø1/2" Pedestal Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=4.19" | 73.5 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=PH4E
Thorlabs | TR6 | 3 | each | 6.77 | TR6 - Ø1/2" Optical Post, SS, 8-32 Setscrew, 1/4"-20 Tap, L = 6" | 20.31 | 2017-05-03 |  |  | https://www.thorlabs.com/newgrouppage9.cfm?objectgroup_id=1266&pn=TR6#237
AVR (Semrock) | FF01-935/170-25 | 1 | each | 384.75 | 935/170 nm BrightLine® single-band bandpass filter | 384.75 | 2017-04-14 |  |  | https://www.semrock.com/FilterDetails.aspx?id=FF01-935/170-25 | To prevent room light getting to IR camera
AVR (Semrock) |  | 1 | each | 95 | Custom sizing fee for filter | 95 | 2017-04-14
Other Parts | 64.94
Amazon |  | 2 | each | 19.99 | Sabrent 13 Port High Speed USB 2.0 Hub with Power Adapter And 2 Control Switches (HB-U14P) | 39.98 | 2017-06-20 |  |  | http://www.amazon.com/Sabrent-Adapter-Control-Switches-HB-U14P/dp/B00HL7Z46K/ref=sr_1_15?ie=UTF8&qid=1461345814&sr=8-15&keywords=USB+hub
Amazon |  | 1 | 2-pack | 9.97 | AmazonBasics 6-Outlet Surge Protector Power Strip 2-Pack, 200 Joule - Black | 9.97 |  |  |  | http://www.amazon.com/AmazonBasics-6-Outlet-Surge-Protector-2-Pack/dp/B014EKQ5AA/ref=sr_1_3?ie=UTF8&qid=1461346017&sr=8-3&keywords=power+strip
Amazon |  | 1 | 2-pack | 14.99 | 10-Feet Power Cord Extension - NEMA 5-15R to NEMA 5-15P | 14.99 | 2017-06-20 |  |  | 2 of StarTech.com PAC10110 10-Feet Power Cord Extension - NEMA 5-15R to NEMA 5-15P
Timing Control | 3106.2 | Quote Requested
National Instruments | 781049-01 | 1 | each | 1508.4 | NI PCIe-6353 | 1508.4 | 2017-04-14 |  |  | http://sine.ni.com/nips/cds/view/p/lang/en/nid/207410
National Instruments | 192061-02 | 2 | each | 138.6 | NI Cables - Shielded, for PCIe-6353 | 277.2 | 2017-04-14 |  |  | http://sine.ni.com/nips/cds/view/p/lang/en/nid/207410
National Instruments | 777664-01 | 1 | each | 282.6 | CA-1000 Connector Accessory Enclosure | 282.6 | 2017-04-14 |  |  | http://sine.ni.com/nips/cds/view/p/lang/en/nid/1835
National Instruments | 184721-01 | 1 | each | 63 | Strain Relief Connector Panelette for CA-1000 Style Enclosure | 63 | 2017-04-14 |  |  | http://sine.ni.com/nips/cds/view/p/lang/en/nid/10713
National Instruments | 184737-01 | 4 | each | 49.5 | BNC Connector Panelette, CA-1000 Style Enclosure | 198 | 2017-04-14 |  |  | http://sine.ni.com/nips/cds/view/p/lang/en/nid/10713
National Instruments | 777145-02 | 2 | each | 113.4 | CB-68LPR I/O Connector Block | 226.8 | 2017-04-14 |  |  | http://sine.ni.com/nips/cds/view/p/lang/en/nid/1186
National Instruments | 184483-01 | 3 | each | 12.6 | Blank Connector Panelette for the CA-1000 Style Enclosure | 37.8 | 2017-04-14 |  |  | http://sine.ni.com/nips/cds/view/p/lang/en/nid/10688
National Instruments | 782606-01 | 1 | each | 365.4 | Low-Cost Multifunction DAQ, 16-bit, 50 kS/s, 8 AI, 2 AO, 13 DIO | 365.4 | 2017-04-14 |  |  | http://sine.ni.com/nips/cds/view/p/lang/en/nid/212384 | For focus drive
National Instruments | 782909-01 | 1 | each | 19.8 | High Speed USB Cable - Micro to STD Male (1m) | 19.8 | 2017-04-14 |  |  | http://sine.ni.com/nips/cds/view/p/lang/en/nid/212384 | for focus drive
Thorlabs | CA2672 | 2 | each | 44.2 | CA2672 - SMC Coaxial Cable, SMC Female to BNC Male, 72" (1829 mm) | 88.4 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CA2672 | connect 405 and 488 laser to National instruments panel
Thorlabs | CA3136 | 2 | each | 16 | CA3136 - RG-174 BNC Coaxial Cable, BNC Male to BNC Male, 36" (914 mm) | 32 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=CA3136 | connect IR laser to National Instruments panel
Thorlabs | T3283 | 1 | each | 6.8 | BNC Adapter - Straight Adapter (F-F) | 6.8 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=T3283 | for Peizo-Z output (might not need)
Thorlabs | 2249-C-60 | 1 | each | 18.5 | 2249-C-60 - RG-58 BNC Coaxial Cable, BNC Male to BNC Male, 60" (1524 mm) |  | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=2249-C-60 | for Peizo-Z output
Microscope Body | 20737.2 | Quote Requested
Nikon | MEA53200 | 1 | each | 7038 | Inverted Microscope Eclipse Ti-U Main Body,with: left and rightside ports, intermediate magnification changer 1.5X, coaxial coarse/finefocus w/tension adjustment on coarse and focus stop on coarse, laser safety interlock signal output; Light-distribution: 100 Eyepiece, 100Left, 100 Right (*), and Auxiliary position (*) Note: (*) prism can beexchanged for optional prisms; Eye20/Left80, Eye20/Right80 | 7038 | 2017-04-03
Nikon | MEV51005 | 2 | each | 750.55 | TI-FLC Epi-FL Filter Turret Six position turret, ergonomicallypositioned light shut-off shutter, "Noise Terminator" included for highsignal-to-noise fluorecence images | 1501.1 | 2017-04-03
Nikon | MQD42070 | 1 | each | 532.1 | 0.7x C-Mount Relay Lens | 532.1 | 2017-04-03
Nikon | MED53200 | 1 | each | 938.4 | Ti-BSUK70 70mm Stage Up Kit | 938.4 | 2017-04-03
Nikon | MRD01605 | 1 | each | 5078.75 | CFI Plan Apo Lambda 60x oil 0.13mm, F.O.V. 25mm, DIC | 5078.75 | 2017-04-03
Nikon | MRH00101 | 1 | each | 800 | CFI60 Plan Fluor 10x Objective Lens, N.A. 0.3, W.D. 16mm, F.O.V. 25mm, DIC | 800 | 2017-04-03
Nikon | MEP53315 | 1 | each | 1164.5 | Ti-ND6 Sextuple DIC Nosepiece | 1164.5 | 2017-04-03
Nikon | MEB55800 | 1 | each | 290.7 | Ti-T-B Eyepiece Base Unit (to mount front of stage) | 290.7 | 2017-04-03
Nikon |  | 1 | each | 159.8 | T-DH Dia-Illuminator Base Kit (replaces tower arm, to mount back of stage) | 159.8 | 2017-04-03
Shapeways | none | 1 | each | 17.05 | Custom 3D printing of Nikon Filter Holder 1 -LARGE SIZE (UPDATE PRICE) | 17.05 | 2017-06-20 |  |  |  | to hold IR filter
Chroma | ZT405/488/561/647/752rpc-UF2 | 1 | each | 1100 | Penta-band polychroic mirror 32 x 44 x 2mm | 1100 | 2018-07-03 |  |  | 20% discount!
Chroma | ZET405/488/561/647-656/752 | 1 | each | 680 | Penta-band emission/barrier/notch filter, 25mm diameter, to be mounted in 91032 Nikon laser/TIRF filtercube for Nikon Ti | 680 | 2018-07-03 |  |  | 20% discount!
Chroma | 91054 | 2 | each | 318.4 | Nikon Ti-2 Filter Cube | 636.8 | 2018-07-03 |  |  | 20% discount!
Chroma | 900DCSP | 1 | each | 800 | 900DCSP shortpass dichroic mirror, 32 x 44 x 2 (CHECK THAT WE ACTUALLY SIZED UP HERE) | 800 | 2017-04-14 |  |  | 20% discount | Reflects IR beam for focus lock
Edmund Optics | #64-334 | 1 | each | 215 | 850nm 25mm Diameter, OD 4 Shortpass Filter (UPDATE TO 32 mm Diameter) | 215 | 2017-05-02 |  |  | http://www.edmundoptics.com/optics/optical-filters/shortpass-edge-filters/high-performance-od-4-shortpass-filters/64334/ | to block backscatter IR from hcam
Fluidics System | 8128.77
Hamilton | 4833 -01 | 1 | each | 93 | Power Supply. UNIV24VDC.1/7A | 93 | 2017-04-12 |  | 2017-04-25 |  | Call or email Hamilton and order, didn't go through website
Hamilton | 7750-16 | 1 | 6/pk | 63.9 | 25 gauge, Kel-F Hub NDL, custom length (4 in), point style 3, 6/PK | 63.9 | 2017-04-12 |  |  | https://www.hamiltoncompany.com/products/syringes-and-needles/needles/luer-lock-needles/kel-f-hub-needles/25-gauge-KelF-Hub-NDL-custom-length-0375-to-12-in-point-style-2-3-or-4-6PK
Hamilton | 36782 | 2 | each | 172 | HVXM 8-7 VALVE | 344 | 2017-04-12 |  | 2017-04-25
Hamilton | 77902 | 1 | each | 43 | MVP Output to MVP Input Cable | 43 | 2017-04-12 |  | 2017-04-25 | http://www.hamiltoncompany.com/products/MVP-Output-to-MVP-Input-Cable
Hamilton | 36798 | 1 | each | 950 | MVP RS-232, PC1, w/o Valve | 950 | 2017-04-12 |  | 2017-04-25 | http://www.hamiltoncompany.com/products/MVPRS232PC1WO-VALVE
Gilson | F155006 | 1 | each | 2438 | Peristaltic pump | 2438 | 2017-04-24 |  |  |  | On Gilson quote
Gilson | 361832 | 1 | each | 436 | 361832 361832: 508 INTERFACE MODULE, 110-220 VOLT | 436 | 2017-04-24 |  |  |  | On Gilson quote
Gilson | F117933 | 1 | pk-12 | 41 | PVC Manifold Tubing, 0.38 mm ID (0.015"), orange/green collars | 41 | 2017-04-24 |  |  |  | On Gilson quote
Gilson | 21061820 | 1 | each | 0 | SFW,GILSON SERVER FOR XP AND 7 | 0 | 2017-04-24 |  |  |  | On Gilson quote
Bioptechs | 03060319-2-NH | 1 | each | 2686.86 | FCS2 Chamber Flow Cell | 2686.86 | 2017-04-24 |  |  | http://www.bioptechs.com/Products/Prices/prices.html | Not needed yet (for flow)
Bioptechs | 060319-2-2611 | 1 | each | 970.26 | Zeiss (K), Marzhauser, ASI, Ludl & Prior Adapter | 970.26 | 2017-04-24 |  |  | http://www.bioptechs.com/Products/Prices/prices.html | Not needed yet (for flow)
Bioptechs | 1907 -250 | 0 | set-5 | 37.32 | Blank Gasket cut using die# 443792, 0.25mm Thick – 5/pack – 1907-250 | 0
Bioptechs | 1907 -750 | 0 | set-5 | 37.32 | Blank Gasket cut using die# 443792, 0.75mm Thick – 5/pack – 1907-250 | 0
Bioptechs | 1907 -750 | 0 | set-5 | 37.32 | Blank Gasket cut using die# 527762, 0.75mm Thick – 5/pack – 1907-250 | 0 |  |  |  | http://216.92.84.132/wordpress/product/chamber-accessories/ | Not needed yet (for flow)
Bioptechs | 1907 -750 | 0 | set-5 | 37.32 | Blank Gasket cut using die# 449379-D, 0.75mm Thick – 5/pack – 1907-250 | 0 |  |  |  | http://216.92.84.132/wordpress/product/chamber-accessories/ | Not needed yet (for flow)
Bioptechs | 40-1313 -03192 | 0 | 500pk | 522 | 40mm Coverslips | 0 |  |  |  |  | Not needed yet (for flow)
Bioptechs | Discount | 0 | each | -432.88 | Lab startup discount | 0
Amazon | Pcard | 1 | each | 22.95 | Keyspan by Tripp Lite USA-19HS High-Speed USB Serial Adapter, PC, MAC, supports Cisco Break Sequence | 22.95 |  |  |  | http://www.amazon.com/Tripp-Lite-USA-19HS-High-Speed-supports/dp/B0000VYJRY/ref=sr_1_1?ie=UTF8&qid=1453835083&sr=8-1&keywords=tripp-lite+keyspan+usa-19hs
Amazon | Pcard | 1 | each | 9.99 | Anker Ultra Slim 4-Port USB 3.0 Data Hub | 9.99 |  |  |  | http://www.amazon.com/AmazonBasics-Port-2-5A-power-adapter/dp/B00DQFGH80
Amazon | Pcard | 1 | each | 3.45 | StarTech 6-Feet Straight Through Serial Cable - DB9 F/F (MXT100FF) | 3.45 |  |  |  | http://www.amazon.com/StarTech-6-Feet-Straight-Through-Serial/dp/B00066HP2E
Amazon | Pcard | 3 | each | 5.79 | AmazonBasics USB 3.0 Extension Cable - A-Male to A-Female - 9.8 Feet | 17.37 |  |  |  | http://www.amazon.com/AmazonBasics-USB-3-0-Extension-Cable/dp/B00NH134L6
Amazon | Pcard | 1 | 2-pack | 8.99 | CableCreation [2-PACK] 3 Feet 18 AWG Universal Power Cord for NEMA 5-15P to IEC320C13 Cable, 0.915M / Black | 8.99 |  |  |  | http://www.amazon.com/SF-Cable-Universal-Power-IEC320/dp/B006W0IHPW?ie=UTF8&psc=1&redirect=true&ref_=ox_sc_act_title_1&smid=A33MASNQGU8ZOV
IDEX/Upchurch | P -675 | 0 | each | 9.23 | 1/4-28 Female to Male Luer Assy | 0 |  |  |  | https://www.idex-hs.com/checkout/cart/configure/id/46408/
IDEX | XP -202 | 0 | each | 2.41 | Flangeless Fitting Delrin®, Red 1/4-28 Flat-Bottom, for1/16" OD | 0 |  |  |  | https://www.idex-hs.com/flangeless-fitting-delrinr-red-1-4-24-flat-bottom-for1-16-od.html
McMaster | 5583k52 | 0 | each | 212 | Abrasion-Resistant White ETFE Tubing, .020" ID, 1/16" OD, 50ft | 0
Amazon | pcard | 0 | assorted | 42.44 | Assorted tubing from Amazon | 0
eBay | CNC3040T | 1 | each | 445.01 | 3 AXIS USB Router Engraver Engraving Drilling Milling Machine, 300 x 400 mm (maybe not right size) | 445.01 |  |  |  | https://www.ebay.com/itm/CNC3040T-3-AXIS-USB-Router-Engraver-Engraving-Drilling-Milling-Machine-300x400/163143304813?epid=5029963288&hash=item25fc19466d:g:OMQAAOSwojdbNZ7H:sc:FedExHomeDelivery!94305!US!-1
Laser Safety | 0
Microscope Enclosure | 691.27 | [8 empty] | http://www.mcmaster.com/#t-slotted-framing/=13k1al5
Late Additions
Hamilton | 7750-16 | 1 | 6/pk | 71 | 25 gauge, Kel-F Hub NDL, custom length (5 in), point style 3, 6/PK | 71 | 2017-07-14 |  |  | https://www.hamiltoncompany.com/products/syringes-and-needles/needles/luer-lock-needles/kel-f-hub-needles/25-gauge-KelF-Hub-NDL-custom-length-0375-to-12-in-point-style-2-3-or-4-6PK
Thorlabs | LB1 | 2 | each | 49.5 | Beam Block, 400 - 700 nm, 10 W Max Avg. Power, CW Only, Includes TR3 Post | 90 | 2017-07-18
Thorlabs | ER2-P4 | 2 | 4-pack | 23.18 | Cage Assembly Rod, 2" (50.8 mm) Long, Ø6 mm, 4 Pack | 23.18 | 2017-07-18
Thorlabs | SM1CP2 | 3 | each | 17.4 | SM1CP2 - Externally SM1-Threaded End Cap | 52.2 | 2017-07-18
Thorlabs | ERSCA-P4 | 1 | 4-pack | 42.14 | ERSCA - Rod Adapter for Ø6 mm ER Rods | 42.14 | 2017-07-18
Thorlabs | XE25L09 | 4 | each | 15 | 9" Long Construction Rail | 60 | 2017-07-18
Thorlabs | IDA20-P5 | 1 | 5-pack | 213 | IDA20-P5 - Post-Mountable Standard Iris, 20.0 mm Max Aperture, 8-32 Threaded Stud, Pack of 5 | 213 | ?
Thorlabs | SM1D12 | 1 | each | 57 | SM1D12 - SM1 Lever-Actuated Iris Diaphragm (Ø0.8 - Ø12 mm) | 57 | ?
Thorlabs | SM2D25D | 1 | each | 82.75 | SM2D25D - Ring-Actuated SM2 Iris Diaphragm | 82.75 | ?
Thorlabs | RS1.5P | 6 | each | 23.75 | RS1.5P - Ø1" Pedestal Pillar Post, 1/4"-20 Taps, L = 1.5" | 142.5 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=RS1.5P | Mounts for 2" mirrors and for laser plates (to match 2" mounts for 1" mirrors)
Thorlabs | BE1R | 6 | each | 16.4 | Ø1.25" Magnetic Studded Pedestal Base Adapter, 1/4"-20 Thread | 98.4 | 2017-05-03 |  |  | https://www.thorlabs.com/thorproduct.cfm?partnumber=RS1.5P | Mounts for 2" mirrors and for laser plates (to match 2" mounts for 1" mirrors)