Braille Sorter Installation and Software Instruction Manual

Ramesh Utharala - Merten Grou p, EM B L - August 2020

# Installation Guide

**NOTE: You DO NOT need to have full-LabVIEW version installed to run the executable, however if LabVIEW is not installed you will need to install LabVIEW Run-Time Engine 2013.  An .exe is included in the Braille Sorter.zip to install LabVIEW Run-Time Engine 2013.

## To Download Files

- Download Braille Sorter.zip and extract to the desired location.

## To Launch the Braille Sorter Executable

- In Braille Sorter directory navigate to the Executables folder.

- Double-click Braille Sorter.exe application file to launch the program.

- Once opened the program will begin running immediately.  To correctly configure the software please refer to Running the Braille SorterExecutable.

## To Install LabVIEW Run-Time Engine 2013

- Download the run-time engine from National instruments website (https://www.ni.com/en-ca/support/downloads/software-products/download.labview-runtime.html#359538).

- Double click LVRTE2013std_downloader executable.  The LabVIEW Runtime Engine 2013 Downloader appears and begins downloading.

- When the pop-up dialog box appears click ok and you will be prompted to unzip the files to your desired location.  Use Browse to choose the desired location and leave the When done unzipping open: .|setup.exe box checked.

- Click okay to the dialog box confirming files have been successfully unzipped and if necessary allow the National Instruments Installer to make changes to your computer. This will launch the NI LabVIEW Run-Time Engine 2013 Installer.

- Follow the instructions to install LabVIEW Run-Time Engine 2013.

## To Connect Braille Unit with computer

USB-port of the Braille controller (see file on wiring diagram for details) should be connected to USB-port of the computer. Then all the Braille pins should go down.

## Running the Braille Sorter Executable

Tab 1: Manual Sorting Configuration

1. First open the Braille Sorter.exe. Before pressing the LabVIEW run button, it is important to give desired input in ‘Waste cluster’ and ‘Sorting mode’. That means user should know which pins are used for waste purpose and based on that select one of the avilable 16 pins.

2. Sorting mode is to define the type of experiment the user is going to operate. In this drop-down menu, user has preference to choose cell, embryo, or color sorting options.

3. After pressing LabVIEW run button, select the right ‘Com Port’ and press ‘CONNECT BD’ then BD status should be on (LED lights up).

4. Once BD is connected, priming should be performed to check whether all the liquids are flowing into the chip or not. This can be performed by clicking on green LED buttons in the ‘sample cluster’ that are labelled from 1 to 16. Each number corresponds to the 4-pins in the Braille display (i.e. each column of Braille display)

5. Once Data is selected, click on ‘ADD data’ button such that the newly added data can be seen in ‘Data Indicator’ on the right as a form of matrix.

6.  ‘SAVE Data’ button is used to save all the information in ‘Data Indicator’ and the user has to give file name.

7. Previously saved file can be uploaded into the ‘Data Indicator’ using ‘LOAD Data’ button. Once clicked, the user has to browse and select a file from the computer.

8. Moreover, if there is a mistake in the configuration, it can be deleted by pressing ‘DELETE Data’. This option will only be useful to delete the latest configuration that was uploaded into the ‘Data Indicator’.

9. The next button ‘RESET BD’ is used to reset braille display such that all the pins go down.

10. Finally, ‘Primining cluster’ is used to prime the valves prior to start of the experiment to remove the air from the chip and to makesure chip is filled with buffer or some sort of medium.

Tab 2: Fluorescence Droplet Sorting

1. Input parameters such as acquisition frame rate, gain, exposure time are adjusted for camera by looking at the ‘Live Video’ until the cell is detected.

2. Region of Interest (ROI) in the form of a rectangle should be created on Live Data so that the analysis is be performed only inside the ROI

3. ‘Live Streaming’ is a plot with number of data points on X-axis and green-pixel intensity values on Y-axis. The dotted yellow line is used to set the threshold for detecting cells. That means all the data above threshold is considered for analysis.

4. ‘Peak Width’ is another paramater to make a sorting decision, it is basically defined as the required number of data points above the threshold. For example, one can chose 2 or more data points above the threshold to switch the valves and sort a particular cell.

5. ‘Number of data points’ is used to resets the Live Streaming once it is reaches the input number.

6. ‘FREEZE’ is used to stop the data update at the current time point. Pressed again to collect the data again.

7. ‘RESET’ to clear the current plot.

8. ‘Frame rate’ shows the current acquisition frame rate. It is better to have higher frame rate so that cell can be detected.

Tab 3: Cell sorting

1.  Input parameters such as brightness, gain, autoexposure, shutter modes and values are adjusted for the camera by looking at the ‘Image Out’ until the cell is detected. User should make a clear difference between positive and negative embryo.

2. Region of Interest (ROI) in the form a rectangle should be created on Live Data so that the analysis can be performed only inside the ROI

3. ‘Live Streaming’ is a plot with number of data points on X-axis and green-pixel intensity values on Y-axis. The dotted yellow line is used to set the threshold for detecting embryos. That means all the data above threshold is considered for analysis.

4. ‘Peak Width’ is another paramater to make a sorting decision, it is basically defined as the number of data points above the threshold. In this case if we have 2 or more data above the threshold and once embryo passed the ROI then activating valves for specific amount of time to sort that particular embryo.

5. ‘Number of data points’ is used to resets the Live Streaming once it is reaches the input number.

6. ‘FREEZE’ is used to stop the data update at the current time point. Pressed again to collect the data again.

7. ‘RESET’ to clear the current plot.

8. In addition, a special feature is added to stop the flow inside for a short amount of time (1 to 2 sec) to analyse embryos in more detail. In order to achieve this, change boolean to False state then select the stopping configuration by clicking on all the valves that are involved in the experiment. ‘Stopping time’ is to define how long both the valves are closed before going back to default waste configuration.

9. ‘Actual width of embryo’ is the calculated as the  width of embryo

10. ‘Frame rate’ shows the current acquisition frame rate.

Tab 4. multiway Sorting

1. Input parameters such as brightness, gain, shutter values are adjusted for camera by looking at the ‘Live Video’ until the droplet can be seen nicely.

2. Input RGB values for each color should be determined and given as input here. Sorting decision is based on these RGB ranges.

3. ‘Time (s)’ to define how long the valves should be opened. ‘Decision’ is the output value that shows ‘color value’ of plug, which fall under the particular color category.

4. ‘Frame rate’ shows the current acquisition frame rate.

Sorting modes

- The fluorescene sorting section, which can beused to sort cells, embryos, and fluorescent droplets.

2. Press ‘sort on/off’ so that LED lights up to activate the sorting mode. In addition, press Start-FL-based SORT to start the actual sorting operation.

3. Other input parameters ‘Sample Time (ms)’ is the valve opening time once the right object is detected.

‘Sorting way’ is to select whether the sorting is the 2-way, 3- way,… or the 8-way. ‘Delay time (ms) is necessary sometimes to adjust the sorting opearion. Usually delay times from 1 ms to 25 ms are used. This is the waiting time after detecting the object and before switching the valves.

4. The output is shown once a ‘Cell Detected’ LED glows and the ‘Number of sorts’ shows the number of objects that are sorted at that momemt.

5. ‘Stop FL-based Sort’ is used to immediately stop the current sorting procedure. Then press STOP-BD to completely turn off Braille display.

6. ‘Color Sorting’, which mainly used to sort color droplets.

7. Press ‘Start Color SORT’ to activate the color sorting.

8. Once the droplet is detected, ‘Droplet Detected’ LED turn on for few ms.

9. ‘Stop Color Sort’ is used to immediately stop the current sorting procedure.Then press STOP-BD to completely turn off Braille display.