## Fluence Calculations
Modified from Bolton LP Spreadsheet May 2016 by Natalie Hull | [7 empty] | Photorepair lamp 222nm (modified by BPF for Summer 2019, OSU CEGE) | [7 empty] | Determination of the "Petri Factor"
INSTRUCTIONS AND NOTES: All inputs are shaded yellow
Make sure radiometer calibration data is up to date on Weighting tab | [18 empty] | x | -10 | -9.75 | -9.5 | -9.25 | -9 | -8.75 | -8.5 | -8.25 | -8 | -7.75 | -7.5 | -7.25 | -7 | -6.75 | -6.5 | -6.25 | -6 | -5.75 | -5.5 | -5.25 | -5 | -4.75 | -4.5 | -4.25 | -4 | -3.75 | -3.5 | -3.25 | -3 | -2.75 | -2.5 | -2.25 | -2 | -1.75 | -1.5 | -1.25 | -1 | -0.75 | -0.5 | -0.25 | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.25 | 1.5 | 1.75 | 2 | 2.25 | 2.5 | 2.75 | 3 | 3.25 | 3.5 | 3.75 | 4 | 4.25 | 4.5 | 4.75 | 5 | 5.25 | 5.5 | 5.75 | 6 | 6.25 | 6.5 | 6.75 | 7 | 7.25 | 7.5 | 7.75 | 8 | 8.25 | 8.5 | 8.75 | 9 | 9.25 | 9.5 | 9.75 | 10 |  |  |  | x | y | Ratio | x | y | Ratio
Enter lamp spectra from Spectra tab in Weighting tab. Normalize radiometer factor to peak output wavelength, based on Spectra tab information | [10 empty] | Steps: | [7 empty] | y
| Or normalize to 254 for germicidal comparisons (useful for MP UV). | [9 empty] | 1. Draw a 0.5 cm x 0.5 cm grid and place the center of the grid at the center of the collimated beam | [7 empty] | -10 | [82 empty] | 0 | 0 | 0 | -10 | #DIV/0! | -10 | 0 | #DIV/0!
Set up a 'quasi' collimated beam apparatus. A 'collimating tube' may be used, but it is preferable to use circular 'masks' to define the beam. Make sure that | [10 empty] | and at the height of the sample surface/ radiometer detector plane. | [7 empty] | -9.75 | [82 empty] | 0 | 0 | 0 | -9.75 | #DIV/0! | -9.75 | 0 | #DIV/0!
safety measures are taken to protect workers from exposure to the UV from the lamp. EYE PROTECTION IS AN ABSOLUTE REQUIREMENT. | [10 empty] | 2. Measure the UV irradiance (with a radiometer) every 0.5 cm in the x and y directions and input readings into columns S and X. | [7 empty] | -9.5 | [82 empty] | 0 | 0 | 0 | -9.5 | #DIV/0! | -9.5 | 0 | #DIV/0!
Place the detector head of the UV radiometer on a horizontal surface, containing a 0.5 cm x 0.5 cm grid, such that the 'calibration plane' (see the Calibration | [10 empty] | 3. Take readings out to 0.5 cm beyond the radius of the sample dish (the distance required will be shaded in grey once you input the diameter) | [7 empty] | -9.25 | [82 empty] | 0 | 0 | 0 | -9.25 | #DIV/0! | -9.25 | 0 | #DIV/0!
Sheet provided by the manufacturer of the Radiometer) is at the level of where the top of the solution will be during exposures to the UV. | [10 empty] | 4. The Petri factor is calculated in Cell DL90. | [7 empty] | -9 | [82 empty] | 0 | 0 | 0 | -9 | #DIV/0! | -9 | 0 | #DIV/0!
Zero the radiometer with the cap on, and measure the output at the center of the dish, ensuring to use the radiometer factor for the lamp peak wavelength | [18 empty] | -8.75 | [82 empty] | 0 | 0 | 0 | -8.75 | #DIV/0! | -8.75 | 0 | #DIV/0!
Determine the Petri Factor (See column Q) | [18 empty] | -8.5 | [82 empty] | 0 | 0 | 0 | -8.5 | #DIV/0! | -8.5 | 0 | #DIV/0!
Make sure that the spectrophotometer is balanced with distilled water in the same cuvette. Then measure the absorption coefficient (1 cm absorbance) | [10 empty] | UV Distribution across the Petri Dish for a low pressure UV Lamp | [7 empty] | -8.25 | [82 empty] | 0 | 0 | 0 | -8.25 | #DIV/0! | -8.25 | 0 | #DIV/0!
from 200-400nm for the sample to be irradiated and paste into Weighting tab. | [18 empty] | -8 | [82 empty] | 0 | 0 | 0 | -8 | #DIV/0! | -8 | 0 | #DIV/0!
Insert the solution volume added to the Petri dish | [12 empty] | Meter |  |  |  | Meter |  | -7.75 | [82 empty] | 0 | 0 | 0 | -7.75 | #DIV/0! | -7.75 | 0 | #DIV/0!
Insert the stirrer volume | [10 empty] | x | y | Reading | Ratio | x | y | Reading | Ratio | -7.5 | [82 empty] | 0 | 0 | 0 | -7.5 | #DIV/0! | -7.5 | 0 | #DIV/0!
Insert the distance from the center of the UV lamp to the surface of the water in the Petri Dish | [18 empty] | -7.25 | [82 empty] | 0 | 0 | 0 | -7.25 | #DIV/0! | -7.25 | 0 | #DIV/0!
Insert the desired Fluences (UV Doses) | [10 empty] | 0 | -10 |  | #DIV/0! | -10 | 0 |  | #DIV/0! | -7 | [82 empty] | 0 | 0 | 0 | -7 | #DIV/0! | -7 | 0 | #DIV/0!
Choose a spectral weighting basis (ie germicidal, DNA absorbance, action spectrum, or normalized to 254 for MP) and apply the formula in the weighting tab | [10 empty] | 0 | -9.5 |  | #DIV/0! | -9.5 | 0 |  | #DIV/0! | -6.75 | [82 empty] | 0 | 0 | 0 | -6.75 | #DIV/0! | -6.75 | 0 | #DIV/0!
Remove the radiometer detector and place a Petri Dish (or other container), containing the cell suspension, on a stirring motor placed so that the top of | [10 empty] | 0 | -9 |  | #DIV/0! | -9 | 0 |  | #DIV/0! | -6.5 | [82 empty] | 0 | 0 | 0 | -6.5 | #DIV/0! | -6.5 | 0 | #DIV/0!
the solution is at the same level as that of the calibration plane of the detector head. Add a very small stir bar and make sure that the stirring rate is such that | [10 empty] | 0 | -8.5 |  | #DIV/0! | -8.5 | 0 |  | #DIV/0! | -6.25 | [82 empty] | 0 | 0 | 0 | -6.25 | #DIV/0! | -6.25 | 0 | #DIV/0!
there is no vortex. | [10 empty] | 0 | -8 |  | #DIV/0! | -8 | 0 |  | #DIV/0! | -6 | [82 empty] | 0 | 0 | 0 | -6 | #DIV/0! | -6 | 0 | #DIV/0!
Expose samples in the UV beam for the times calculated | [10 empty] | 0 | -7.5 |  | #DIV/0! | -7.5 | 0 |  | #DIV/0! | -5.75 | [82 empty] | 0 | 0 | 0 | -5.75 | #DIV/0! | -5.75 | 0 | #DIV/0!
For a control, or zero dose determination, a sample solution is added to the Petri dish, set into the collimated beam apparatus and left stirring for a | [10 empty] | 0 | -7 |  | #DIV/0! | -7 | 0 |  | #DIV/0! | -5.5 | [82 empty] | 0 | 0 | 0 | -5.5 | #DIV/0! | -5.5 | 0 | #DIV/0!
time equal to the average of the exposure times used. The resulting counts are used as the 0 dose concentration | [10 empty] | 0 | -6.5 |  | #DIV/0! | -6.5 | 0 |  | #DIV/0! | -5.25 | [82 empty] | 0 | 0 | 0 | -5.25 | #DIV/0! | -5.25 | 0 | #DIV/0!
Note: the exposure times should be at least 30 s. If they are calculated to be shorter, arrange the irradiation platform further away from the UV lamp so | [10 empty] | 0 | -6 |  | #DIV/0! | -6 | 0 |  | #DIV/0! | -5 | [82 empty] | 0 | 0 | 0 | -5 | #DIV/0! | -5 | 0 | #DIV/0!
that the irradiance will be smaller. If this is done, the Petri factor must be re-measured | [10 empty] | 0 | -5.5 |  | #DIV/0! | -5.5 | 0 |  | #DIV/0! | -4.75 | [82 empty] | 0 | 0 | 0 | -4.75 | #DIV/0! | -4.75 | 0 | #DIV/0!
[11 empty] | 0 | -5 |  | #DIV/0! | -5 | 0 |  | #DIV/0! | -4.5 | [82 empty] | 0 | 0 | 0 | -4.5 | #DIV/0! | -4.5 | 0 | #DIV/0!
|  |  |  | λ peak = | 436 | from Spectra tab |  |  |  |  | 0 | -4.5 |  | #DIV/0! | -4.5 | 0 |  | #DIV/0! | -4.25 | [82 empty] | 0 | 0 | 0 | -4.25 | #DIV/0! | -4.25 | 0 | #DIV/0!
|  |  |  | λ avg = | 447.7086 | from Spectra tab |  |  |  |  | 0 | -4 |  | #DIV/0! | -4 | 0 |  | #DIV/0! | -4 | [82 empty] | 0 | 0 | 0 | -4 | #DIV/0! | -4 | 0 | #DIV/0!
[11 empty] | 0 | -3.5 |  | #DIV/0! | -3.5 | 0 |  | #DIV/0! | -3.75 | [82 empty] | 0 | 0 | 0 | -3.75 | #DIV/0! | -3.75 | 0 | #DIV/0!
|  |  |  | Divergence Factor = | 0.96098 | =d/(d+l). Target >0.9 |  |  |  |  | 0 | -3 |  | #DIV/0! | -3 | 0 |  | #DIV/0! | -3.5 | [82 empty] | 0 | 0 | 0 | -3.5 | #DIV/0! | -3.5 | 0 | #DIV/0!
|  |  |  | Reflection Factor at λpeak = | 0.97855 | from Reflection Factor tab |  |  |  |  | 0 | -2.5 |  | #DIV/0! | -2.5 | 0 |  | #DIV/0! | -3.25 | [82 empty] | 0 | 0 | 0 | -3.25 | #DIV/0! | -3.25 | 0 | #DIV/0!
|  |  |  | Petri Factor = | 1 | calculated in cell DL90. Target >0.9 |  |  |  |  | 0 | -2 |  | #DIV/0! | -2 | 0 |  | #DIV/0! | -3 | [82 empty] | 0 | 0 | 0 | -3 | #DIV/0! | -3 | 0 | #DIV/0!
|  |  |  | Spectral Water Factor = | 0.37395 | = Incident / Average Intensity. Target >0.9 |  |  |  |  | 0 | -1.5 |  | #DIV/0! | -1.5 | 0 |  | #DIV/0! | -2.75 | [38 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [39 empty] | #DIV/0! | 0 | 0 | -2.75 | #DIV/0! | -2.75 | 0 | #DIV/0!
[11 empty] | 0 | -1 |  | #DIV/0! | -1 | 0 |  | #DIV/0! | -2.5 | [35 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [36 empty] | #DIV/0! | 0 | 0 | -2.5 | #DIV/0! | -2.5 | 0 | #DIV/0!
|  |  |  | solution volume added to Petri dish = | 10 | mL |  |  |  |  | 0 | -0.5 |  | #DIV/0! | -0.5 | 0 |  | #DIV/0! | -2.25 | [34 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [35 empty] | #DIV/0! | 0 | 0 | -2.25 | #DIV/0! | -2.25 | 0 | #DIV/0!
|  |  |  | stirrer volume = | 0 | mL |  |  |  |  | 0 | 0 |  | #DIV/0! | 0 | 0 |  | #DIV/0! | -2 | [33 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [34 empty] | #DIV/0! | 0 | 0 | -2 | #DIV/0! | -2 | 0 | #DIV/0!
|  |  |  | Petri dish internal diameter = | 5.6 | cm |  |  |  |  | 0 | 0.5 |  | #DIV/0! | 0.5 | 0 |  | #DIV/0! | -1.75 | [32 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [33 empty] | #DIV/0! | 0 | 0 | -1.75 | #DIV/0! | -1.75 | 0 | #DIV/0!
|  |  |  | total volume in Petri dish = | 10 | mL |  |  |  |  | 0 | 1 |  | #DIV/0! | 1 | 0 |  | #DIV/0! | -1.5 | [31 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [32 empty] | #DIV/0! | 0 | 0 | -1.5 | #DIV/0! | -1.5 | 0 | #DIV/0!
|  |  |  | water path length = | 0.40601 | cm |  |  |  |  | 0 | 1.5 |  | #DIV/0! | 1.5 | 0 |  | #DIV/0! | -1.25 | [30 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [31 empty] | #DIV/0! | 0 | 0 | -1.25 | #DIV/0! | -1.25 | 0 | #DIV/0!
|  |  |  | distance from UV lamp to top of water surface = | 10 | cm |  |  |  |  | 0 | 2 |  | #DIV/0! | 2 | 0 |  | #DIV/0! | -1 | [30 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [31 empty] | #DIV/0! | 0 | 0 | -1 | #DIV/0! | -1 | 0 | #DIV/0!
|  |  |  | Absorbance scan = | Import Data! | use Absorbance tab |  |  |  |  | 0 | 2.5 |  | #DIV/0! | 2.5 | 0 |  | #DIV/0! | -0.75 | [30 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [31 empty] | #DIV/0! | 0 | 0 | -0.75 | #DIV/0! | -0.75 | 0 | #DIV/0!
|  |  |  | absorption coefficient at 365 nm = | 0.10982 | cm-1 |  |  |  |  | 0 | 3 |  | #DIV/0! | 3 | 0 |  | #DIV/0! | -0.5 | [29 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [30 empty] | #DIV/0! | 0 | 0 | -0.5 | #DIV/0! | -0.5 | 0 | #DIV/0!
|  |  |  | UVT % at 365 = | 77.6569 | make sure this isn't <90% for disinfection studies |  |  |  |  | 0 | 3.5 |  | #DIV/0! | 3.5 | 0 |  | #DIV/0! | -0.25 | [29 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [30 empty] | #DIV/0! | 0 | 0 | -0.25 | #DIV/0! | -0.25 | 0 | #DIV/0!
[11 empty] | 0 | 4 |  | #DIV/0! | 4 | 0 |  | #DIV/0! | 0 | [29 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [30 empty] | #DIV/0! | 0 | 0 | 0 | #DIV/0! | 0 | 0 | #DIV/0!
|  |  |  | Radiometer Factor at λpeak = | 4.9698e-06 | [5 empty] | 0 | 4.5 |  | #DIV/0! | 4.5 | 0 |  | #DIV/0! | 0.25 | [29 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [30 empty] | #DIV/0! | 0 | 0 | 0.25 | #DIV/0! | 0.25 | 0 | #DIV/0!
|  |  |  | Radiometer Reading Center of Dish = | 0.14319 | mW/cm2 |  |  | 127.4253 | µW/cm2 | 0 | 5 |  | #DIV/0! | 5 | 0 |  | #DIV/0! | 0.5 | [29 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [30 empty] | #DIV/0! | 0 | 0 | 0.5 | #DIV/0! | 0.5 | 0 | #DIV/0!
|  |  |  | Incident Irradiance= | 0.42241 | mW/cm2 |  |  |  |  | 0 | 5.5 |  | #DIV/0! | 5.5 | 0 |  | #DIV/0! | 0.75 | [30 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [31 empty] | #DIV/0! | 0 | 0 | 0.75 | #DIV/0! | 0.75 | 0 | #DIV/0!
|  |  |  | Average Irradiance= | 0.15796 | mW/cm2 |  |  |  |  | 0 | 6 |  | #DIV/0! | 6 | 0 |  | #DIV/0! | 1 | [30 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [31 empty] | #DIV/0! | 0 | 0 | 1 | #DIV/0! | 1 | 0 | #DIV/0!
[11 empty] | 0 | 6.5 |  | #DIV/0! | 6.5 | 0 |  | #DIV/0! | 1.25 | [30 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [31 empty] | #DIV/0! | 0 | 0 | 1.25 | #DIV/0! | 1.25 | 0 | #DIV/0!
|  |  |  | Weighting Factor = | Photolyase Absorbance | choose and apply formula on Weighting tab, record weighting type here |  |  |  |  | 0 | 7 |  | #DIV/0! | 7 | 0 |  | #DIV/0! | 1.5 | [31 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [32 empty] | #DIV/0! | 0 | 0 | 1.5 | #DIV/0! | 1.5 | 0 | #DIV/0!
|  |  |  | Weighted Average Irradiance= | 0.03253 | mW/cm2 |  |  |  |  | 0 | 7.5 |  | #DIV/0! | 7.5 | 0 |  | #DIV/0! | 1.75 | [32 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [33 empty] | #DIV/0! | 0 | 0 | 1.75 | #DIV/0! | 1.75 | 0 | #DIV/0!
[11 empty] | 0 | 8 |  | #DIV/0! | 8 | 0 |  | #DIV/0! | 2 | [33 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [34 empty] | #DIV/0! | 0 | 0 | 2 | #DIV/0! | 2 | 0 | #DIV/0!
|  | Seconds for a Fluence (UV Dose) of | 1 | mJ/cm2 = | 30.7411 | change formula to use weighted average irradiance, if applicable |  |  |  |  | 0 | 8.5 |  | #DIV/0! | 8.5 | 0 |  | #DIV/0! | 2.25 | [34 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [35 empty] | #DIV/0! | 0 | 0 | 2.25 | #DIV/0! | 2.25 | 0 | #DIV/0!
[11 empty] | 0 | 9 |  | #DIV/0! | 9 | 0 |  | #DIV/0! | 2.5 | [35 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [36 empty] | #DIV/0! | 0 | 0 | 2.5 | #DIV/0! | 2.5 | 0 | #DIV/0!
[9 empty] | min | seconds | 0 | 9.5 |  | #DIV/0! | 9.5 | 0 |  | #DIV/0! | 2.75 | [38 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [39 empty] | #DIV/0! | 0 | 0 | 2.75 | #DIV/0! | 2.75 | 0 | #DIV/0!
|  |  |  | Time for a Fluence (UV Dose) of | 100 | mJ/cm2 = | 3074.1121 | s = | 51 | 14.1121 | 0 | 10 |  | #DIV/0! | 10 | 0 |  | #DIV/0! | 3 | [82 empty] | 0 | 0 | 0 | 3 | #DIV/0! | 3 | 0 | #DIV/0!
|  |  |  | Time for a Fluence (UV Dose) of | 300 | mJ/cm2 = | 9222.3363 | s = | 153 | 42.3363 | [8 empty] | 3.25 | [82 empty] | 0 | 0 | 0 | 3.25 | #DIV/0! | 3.25 | 0 | #DIV/0!
|  |  |  | Time for a Fluence (UV Dose) of | 600 | mJ/cm2 = | 18444.6727 | s = | 307 | 24.6727 | [8 empty] | 3.5 | [82 empty] | 0 | 0 | 0 | 3.5 | #DIV/0! | 3.5 | 0 | #DIV/0!
|  |  |  | Time for a Fluence (UV Dose) of | 1200 | mJ/cm2 = | 36889.3453 | s = | 614 | 49.3453 | [8 empty] | 3.75 | [82 empty] | 0 | 0 | 0 | 3.75 | #DIV/0! | 3.75 | 0 | #DIV/0!
|  |  |  | Time for a Fluence (UV Dose) of | 2400 | mJ/cm2 = | 73778.6907 | s = | 1229 | 38.6907 | [8 empty] | 4 | [82 empty] | 0 | 0 | 0 | 4 | #DIV/0! | 4 | 0 | #DIV/0!
|  |  |  | Time for a Fluence (UV Dose) of | 5000 | mJ/cm2 = | 153705.6055 | s = | 2561 | 45.6055 | [8 empty] | 4.25 | [82 empty] | 0 | 0 | 0 | 4.25 | #DIV/0! | 4.25 | 0 | #DIV/0!
|  |  |  | Time for a Fluence (UV Dose) of |  | mJ/cm2 = | 0 | s = | 0 | 0 | [8 empty] | 4.5 | [82 empty] | 0 | 0 | 0 | 4.5 | #DIV/0! | 4.5 | 0 | #DIV/0!
[7 empty] | 0 |  | 0 | 0 | [8 empty] | 4.75 | [82 empty] | 0 | 0 | 0 | 4.75 | #DIV/0! | 4.75 | 0 | #DIV/0!
[19 empty] | 5 | [82 empty] | 0 | 0 | 0 | 5 | #DIV/0! | 5 | 0 | #DIV/0!
[19 empty] | 5.25 | [82 empty] | 0 | 0 | 0 | 5.25 | #DIV/0! | 5.25 | 0 | #DIV/0!
[19 empty] | 5.5 | [82 empty] | 0 | 0 | 0 | 5.5 | #DIV/0! | 5.5 | 0 | #DIV/0!
[19 empty] | 5.75 | [82 empty] | 0 | 0 | 0 | 5.75 | #DIV/0! | 5.75 | 0 | #DIV/0!
[19 empty] | 6 | [82 empty] | 0 | 0 | 0 | 6 | #DIV/0! | 6 | 0 | #DIV/0!
[19 empty] | 6.25 | [82 empty] | 0 | 0 | 0 | 6.25 | #DIV/0! | 6.25 | 0 | #DIV/0!
[19 empty] | 6.5 | [82 empty] | 0 | 0 | 0 | 6.5 | #DIV/0! | 6.5 | 0 | #DIV/0!
[19 empty] | 6.75 | [82 empty] | 0 | 0 | 0 | 6.75 | #DIV/0! | 6.75 | 0 | #DIV/0!
[19 empty] | 7 | [82 empty] | 0 | 0 | 0 | 7 | #DIV/0! | 7 | 0 | #DIV/0!
[19 empty] | 7.25 | [82 empty] | 0 | 0 | 0 | 7.25 | #DIV/0! | 7.25 | 0 | #DIV/0!
[19 empty] | 7.5 | [82 empty] | 0 | 0 | 0 | 7.5 | #DIV/0! | 7.5 | 0 | #DIV/0!
[19 empty] | 7.75 | [82 empty] | 0 | 0 | 0 | 7.75 | #DIV/0! | 7.75 | 0 | #DIV/0!
[19 empty] | 8 | [82 empty] | 0 | 0 | 0 | 8 | #DIV/0! | 8 | 0 | #DIV/0!
[19 empty] | 8.25 | [82 empty] | 0 | 0 | 0 | 8.25 | #DIV/0! | 8.25 | 0 | #DIV/0!
[19 empty] | 8.5 | [82 empty] | 0 | 0 | 0 | 8.5 | #DIV/0! | 8.5 | 0 | #DIV/0!
[19 empty] | 8.75 | [82 empty] | 0 | 0 | 0 | 8.75 | #DIV/0! | 8.75 | 0 | #DIV/0!
[19 empty] | 9 | [82 empty] | 0 | 0 | 0 | 9 | #DIV/0! | 9 | 0 | #DIV/0!
[19 empty] | 9.25 | [82 empty] | 0 | 0 | 0 | 9.25 | #DIV/0! | 9.25 | 0 | #DIV/0!
[19 empty] | 9.5 | [82 empty] | 0 | 0 | 0 | 9.5 | #DIV/0! | 9.5 | 0 | #DIV/0!
[19 empty] | 9.75 | [82 empty] | 0 | 0 | 0 | 9.75 | #DIV/0! | 9.75 | 0 | #DIV/0!
[19 empty] | 10 | [82 empty] | 0 | 0 | 0 | 10 | #DIV/0! | 10 | 0 | #DIV/0!
[102 empty] | #DIV/0! | 0
[101 empty] | Petri factor = | #DIV/0!
Note on G38: Jim Bolton: Measure carefully with a plastic mm ruler or better with a pair of calipers.
Note on G40: Jim Bolton: The water path length may be entered directly.

## Absorbance
Paste the absorbance spectra of the microorganism sample below.
UV VIS Wavelength | Absorbance
500 | 0.059853
499 | 0.060144
498 | 0.060364
497 | 0.060621
496 | 0.06082
495 | 0.061042
494 | 0.061305
493 | 0.061562
492 | 0.061767
491 | 0.061972
490 | 0.06214
489 | 0.06241
488 | 0.062589
487 | 0.062885
486 | 0.063119
485 | 0.063332
484 | 0.063579
483 | 0.063809
482 | 0.064094
481 | 0.064315
480 | 0.064547
479 | 0.064852
478 | 0.065074
477 | 0.065347
476 | 0.065586
475 | 0.065874
474 | 0.066132
473 | 0.066395
472 | 0.066687
471 | 0.066955
470 | 0.067181
469 | 0.067443
468 | 0.06776
467 | 0.068022
466 | 0.068321
465 | 0.068561
464 | 0.068877
463 | 0.069167
462 | 0.069395
461 | 0.069682
460 | 0.07004
459 | 0.070321
458 | 0.070561
457 | 0.070886
456 | 0.071182
455 | 0.071493
454 | 0.071744
453 | 0.072119
452 | 0.072458
451 | 0.072726
450 | 0.072975
449 | 0.073311
448 | 0.07364
447 | 0.073921
446 | 0.074286
445 | 0.074543
444 | 0.074925
443 | 0.075215
442 | 0.075587
441 | 0.075871
440 | 0.076182
439 | 0.076525
438 | 0.076871
437 | 0.077187
436 | 0.077549
435 | 0.077891
434 | 0.078229
433 | 0.078573
432 | 0.078977
431 | 0.079352
430 | 0.079663
429 | 0.079985
428 | 0.080396
427 | 0.080748
426 | 0.081147
425 | 0.081515
424 | 0.081917
423 | 0.082278
422 | 0.082681
421 | 0.08302
420 | 0.083397
419 | 0.083772
418 | 0.084164
417 | 0.084562
416 | 0.084934
415 | 0.085378
414 | 0.085677
413 | 0.08615
412 | 0.086527
411 | 0.086948
410 | 0.087359
409 | 0.087762
408 | 0.088228
407 | 0.088556
406 | 0.088985
405 | 0.089445
404 | 0.089852
403 | 0.090274
402 | 0.090714
401 | 0.091145
400 | 0.091572
399 | 0.091996
398 | 0.092411
397 | 0.092894
396 | 0.093358
395 | 0.093759
394 | 0.094268
393 | 0.094768
392 | 0.095238
391 | 0.095691
390 | 0.096133
389 | 0.096662
388 | 0.097198
387 | 0.097638
386 | 0.098199
385 | 0.098622
384 | 0.099182
383 | 0.099713
382 | 0.10025
381 | 0.10073
380 | 0.10121
379 | 0.10175
378 | 0.10235
377 | 0.10289
376 | 0.10341
375 | 0.10391
374 | 0.10449
373 | 0.10503
372 | 0.10564
371 | 0.10622
370 | 0.10671
369 | 0.10734
368 | 0.10799
367 | 0.1086
366 | 0.10913
365 | 0.10982
364 | 0.11045
363 | 0.11106
362 | 0.11163
361 | 0.11226
360 | 0.11291
359 | 0.1136
358 | 0.11417
357 | 0.11489
356 | 0.11556
355 | 0.11617
354 | 0.11683
353 | 0.11751
352 | 0.11825
351 | 0.1189
350 | 0.11961
349 | 0.12046
348 | 0.12184
347 | 0.12191
346 | 0.12321
345 | 0.12359
344 | 0.12448
343 | 0.12538
342 | 0.12577
341 | 0.12698
340 | 0.12752
339 | 0.12846
338 | 0.12947
337 | 0.13005
336 | 0.13101
335 | 0.13179
334 | 0.13267
333 | 0.13361
332 | 0.13448
331 | 0.13548
330 | 0.13616
329 | 0.1372
328 | 0.1381
327 | 0.13913
326 | 0.14023
325 | 0.14098
324 | 0.14219
323 | 0.14309
322 | 0.14422
321 | 0.1452
320 | 0.14638
319 | 0.1475
318 | 0.14851
317 | 0.14977
316 | 0.15099
315 | 0.15217
314 | 0.1534
313 | 0.15468
312 | 0.15596
311 | 0.15734
310 | 0.15888
309 | 0.16026
308 | 0.16181
307 | 0.16346
306 | 0.16504
305 | 0.16691
304 | 0.16882
303 | 0.17076
302 | 0.17276
301 | 0.17505
300 | 0.17729
299 | 0.17961
298 | 0.18216
297 | 0.18479
296 | 0.18747
295 | 0.19019
294 | 0.19297
293 | 0.19573
292 | 0.1983
291 | 0.20092
290 | 0.20334
289 | 0.20582
288 | 0.2083
287 | 0.21081
286 | 0.21325
285 | 0.21543
284 | 0.21765
283 | 0.21973
282 | 0.22163
281 | 0.22355
280 | 0.22529
279 | 0.22712
278 | 0.22865
277 | 0.23018
276 | 0.23166
275 | 0.23299
274 | 0.23422
273 | 0.23535
272 | 0.23631
271 | 0.23731
270 | 0.23818
269 | 0.23894
268 | 0.23951
267 | 0.24007
266 | 0.24053
265 | 0.24097
264 | 0.24116
263 | 0.24133
262 | 0.24138
261 | 0.24142
260 | 0.24125
259 | 0.24104
258 | 0.24084
257 | 0.24052
256 | 0.24032
255 | 0.24009
254 | 0.23997
253 | 0.2397
252 | 0.23951
251 | 0.23949
250 | 0.23953
249 | 0.23977
248 | 0.24018
247 | 0.24077
246 | 0.24159
245 | 0.24271
244 | 0.24422
243 | 0.24599
242 | 0.24809
241 | 0.25064
240 | 0.25351
239 | 0.25695
238 | 0.26066
237 | 0.26465
236 | 0.26905
235 | 0.27368
234 | 0.2785
233 | 0.28338
232 | 0.2884
231 | 0.29322
230 | 0.29806
229 | 0.30281
228 | 0.30735
227 | 0.31201
226 | 0.31652
225 | 0.32121
224 | 0.32567
223 | 0.33036
222 | 0.33496
221 | 0.3395
220 | 0.34412
219 | 0.3488
218 | 0.35365
217 | 0.35878
216 | 0.36439
215 | 0.37099
214 | 0.3786
213 | 0.38794
212 | 0.40027
211 | 0.41669
210 | 0.43939
209 | 0.47133
208 | 0.51697
207 | 0.58286
206 | 0.67836
205 | 0.81652
204 | 1.0143
203 | 1.2903
202 | 1.6678
201 | 2.1831
200 | 2.8706

## Reflection Factor
Recommended set of data from Jim Bolton, personal communication, June 2016
Wavelength
| range | [11 empty] | Recommended absorption coefficients
| 170-197 nm | a1 = | 7.2264e-05 | b1 = | -0.0399 | c1 = | 7.1409 | d1 = | -412.8021 | [5 empty] | 1/10
| 198-206 nm |  |  | b2 = | 0.0071158 | c2 = | -2.9461 | d2 = | 302.1054 |  |  |  | wavelength |  | absorption
| 207-320 nm | a3 = | -2.921e-07 | b3 = | 0.00028708 | c3 = | -0.099558 | d3 = | 8.0526 |  |  |  | / nm | a / cm-1 | length / cm
| 320-400 nm | a4 = | 9.325e-07 | b4 = | -0.00088013 | c4 = | 0.26227 | d4 = | -28.3699 |  |  |  | 172 | 575.44 | 0.0017378
[5 empty] | 0.00063175 |  |  | fraction | 1/10 |  |  |  | 184.9 | 1.8017 | 0.55502
see hidden columns for calcuations
| wavelength | [6 empty] | absorbed in | absorption | Refractive | Reflection | Reflection | 253.7 | 0.00031814 | 3143.2527
| / nm | a / cm-1 | log(a) | log(a) fit | deviation | devation2 | UVT | 10 cm | length / cm | Index | coeficient | Factor
| 170 | 1174.9 | 3.07 | 3.0739 | -0.0039311 | 1.5454e-05 | 0% | 1 | 0.00085114
| 170.5 | 977.24 | 2.99 | 2.9933 | -0.0032729 | 1.0712e-05 | 0% | 1 | 0.0010233
| 171 | 831.76 | 2.92 | 2.9111 | 0.0088503 | 7.8328e-05 | 0% | 1 | 0.0012023 |  |  |  | For absorption coefficients from 170 to 190 nm:
| 171.5 | 691.83 | 2.84 | 2.8276 | 0.012392 | 0.00015357 | 0% | 1 | 0.0014454 |  |  |  | J. L. Weeks, G. M. A. C. Meaburn and S. Gordon, Absorption Coefficients of Liquid Water
| 172 | 575.44 | 2.76 | 2.7427 | 0.017294 | 0.00029907 | 0% | 1 | 0.0017378 |  |  |  | and aqueous solutions in the far ultraviolet, Radiation Research, 19(3), 559-567 (1963)
| 172.5 | 467.74 | 2.67 | 2.6565 | 0.013505 | 0.00018238 | 0% | 1 | 0.0021379
| 173 | 380.19 | 2.58 | 2.569 | 0.010959 | 0.00012011 | 0% | 1 | 0.0026303 |  |  |  | For absorption coefficients from 196 to 320 nm:
| 173.5 | 309.03 | 2.49 | 2.4804 | 0.0096152 | 9.2452e-05 | 0% | 1 | 0.0032359 |  |  |  | T. I. Quickenden and J. A. Irvin, "The ultraviolet absorption spectrum
| 174 | 251.19 | 2.4 | 2.3906 | 0.009416 | 8.8661e-05 | 6.457e-250% | 1 | 0.0039811 |  |  |  | of liquid water," J. Chem Phys., 72, 4416--4428, (1980).
| 174.5 | 204.17 | 2.31 | 2.2997 | 0.010294 | 0.00010596 | 6.761e-203% | 1 | 0.0048979
| 175 | 165.96 | 2.22 | 2.2078 | 0.012228 | 0.00014953 | 1.096e-164% | 1 | 0.0060255 |  |  |  | For absorption coefficients from 320 to 400 nm:
| 175.5 | 134.9 | 2.13 | 2.1149 | 0.015141 | 0.00022924 | 1.259e-133% | 1 | 0.0074129 |  |  |  | H. Buiteveld, J. M. H. Hakvoort and M. Donze, The optical properties of pure water,
| 176 | 107.15 | 2.03 | 2.021 | 0.0089513 | 8.0126e-05 | 7.079e-106% | 1 | 0.0093327 |  |  |  | in SPIE Proceedings on Ocean Optics XII, edited by J. S. Jaffe, 2258, 174--183, (1994).
| 176.5 | 87.1 | 1.94 | 1.9263 | 0.01368 | 0.00018714 | 7.943e-86% | 1 | 0.011481
| 177 | 67.61 | 1.83 | 1.8308 | -0.00080658 | 6.5056e-07 | 2.455e-66% | 1 | 0.014791
| 177.5 | 54.95 | 1.74 | 1.7345 | 0.0054348 | 2.9537e-05 | 1.122e-53% | 1 | 0.018198
| 178 | 44.67 | 1.65 | 1.6375 | 0.012477 | 0.00015569 | 2.138e-43% | 1 | 0.022386
| 178.5 | 34.67 | 1.54 | 1.5399 | 6.5173e-05 | 4.2475e-09 | 2.138e-33% | 1 | 0.028843
| 179 | 28.84 | 1.46 | 1.4416 | 0.018358 | 0.00033701 | 1.445e-27% | 1 | 0.034674
| 179.5 | 22.387 | 1.35 | 1.3428 | 0.0071567 | 5.1218e-05 | 4.102e-21% | 1 | 0.044669 |  |  |  | For refractive indices from 182 to 400 nm:
| 180 | 17.783 | 1.25 | 1.2435 | 0.006457 | 4.1693e-05 | 1.648e-16% | 1 | 0.056233 |  |  |  | M. Daimon and A. Masumura, "Measurement of the refractive index of distilled water from
| 180.5 | 14.125 | 1.15 | 1.1438 | 0.0061704 | 3.8074e-05 | 7.499e-13% | 1 | 0.070796 |  |  |  | the near-infrared region to the ultraviolet region", Applied Optics, 46(18), 3811-3820 (2007)
| 181 | 11.382 | 1.0562 | 1.0437 | 0.012515 | 0.00015662 | 4.15e-10% | 1 | 0.087858
| 181.5 | 8.913 | 0.95002 | 0.94326 | 0.0067651 | 4.5766e-05 | 1.222e-07% | 1 | 0.1122
| 182 | 6.761 | 0.83001 | 0.84254 | -0.012527 | 0.00015693 | 1.734e-05% | 1 | 0.14791 | 1.4677 | 0.035922 | 0.96408
| 182.5 | 5.37 | 0.72997 | 0.7416 | -0.011621 | 0.00013504 | 0.0004266% | 1 | 0.18622 | 1.4659 | 0.035699 | 0.9643
| 183 | 4.074 | 0.61002 | 0.64048 | -0.030464 | 0.00092803 | 0.008433% | 1 | 0.24546 | 1.4642 | 0.035482 | 0.96452
| 183.5 | 3.311 | 0.51996 | 0.53926 | -0.019301 | 0.00037254 | 0.04887% | 1 | 0.30202 | 1.4625 | 0.035271 | 0.96473
| 184 | 2.692 | 0.43008 | 0.43798 | -0.007902 | 6.2442e-05 | 0.2032% | 1 | 0.37147 | 1.4608 | 0.035067 | 0.96493
| 184.5 | 2.138 | 0.33001 | 0.33669 | -0.0066808 | 4.4633e-05 | 0.7278% | 1 | 0.46773 | 1.4592 | 0.034869 | 0.96513
| 184.9 | 1.8017 | 0.25569 | 0.25569 | 0 | 3.0815e-33 | 1.579% | 1 | 0.55502 | 1.458 | 0.034714 | 0.96529
| 185 | 1.66 | 0.22011 | 0.23545 | -0.015341 | 0.00023534 | 2.188% | 1 | 0.60241 | 1.4576 | 0.034676 | 0.96532
| 185.5 | 1.259 | 0.10003 | 0.13431 | -0.034287 | 0.0011756 | 5.508% | 1 | 0.79428 | 1.4561 | 0.034488 | 0.96551
| 186 | 1.023 | 0.0098756 | 0.033334 | -0.023458 | 0.00055029 | 9.484% | 1 | 0.97752 | 1.4546 | 0.034306 | 0.96569
| 186.5 | 0.813 | -0.089909 | -0.067433 | -0.022476 | 0.00050518 | 15.38% | 1 | 1.23 | 1.4532 | 0.034128 | 0.96587
| 187 | 0.661 | -0.1798 | -0.16793 | -0.011864 | 0.00014076 | 21.83% | 1 | 1.5129 | 1.4518 | 0.033954 | 0.96605
| 187.5 | 0.525 | -0.27984 | -0.26812 | -0.011725 | 0.00013748 | 29.85% | 0.99999 | 1.9048 | 1.4504 | 0.033786 | 0.96621
| 188 | 0.417 | -0.37986 | -0.36792 | -0.011941 | 0.0001426 | 38.28% | 0.99993 | 2.3981 | 1.4491 | 0.033621 | 0.96638
| 188.5 | 0.331 | -0.48017 | -0.4673 | -0.012871 | 0.00016566 | 46.67% | 0.99951 | 3.0211 | 1.4477 | 0.03346 | 0.96654
| 189 | 0.282 | -0.54975 | -0.5662 | 0.016446 | 0.00027047 | 52.24% | 0.99849 | 3.5461 | 1.4465 | 0.033303 | 0.9667
| 189.5 | 0.224 | -0.64975 | -0.66456 | 0.014804 | 0.00021916 | 59.7% | 0.99425 | 4.4643 | 1.4452 | 0.03315 | 0.96685
| 190 | 0.2 | -0.69897 | -0.76232 | 0.063354 | 0.0040137 | 63.1% | 0.99 | 5 | 1.444 | 0.033 | 0.967
| 196 | 0.0126 | -1.8996 | -1.8697 | -0.029919 | 0.00089514 | 97.14% | 0.25183 | 79.3651 | 1.431 | 0.031435 | 0.96857
| 197 | 0.00852 | -2.0696 | -2.0387 | -0.030895 | 0.00095449 | 98.06% | 0.17814 | 117.3709 | 1.4291 | 0.031208 | 0.96879
| 198 | 0.00567 | -2.2464 | -2.2554 | 0.0089439 | 7.9994e-05 | 98.7% | 0.12239 | 176.3668 | 1.4273 | 0.030991 | 0.96901
| 199 | 0.00411 | -2.3862 | -2.3765 | -0.009662 | 9.3353e-05 | 99.06% | 0.090296 | 243.309 | 1.4255 | 0.030781 | 0.96922
| 200 | 0.00324 | -2.4895 | -2.4834 | -0.006055 | 3.6663e-05 | 99.26% | 0.071889 | 308.642 | 1.4238 | 0.030578 | 0.96942
| 201 | 0.00264 | -2.5784 | -2.5761 | -0.0023239 | 5.4003e-06 | 99.39% | 0.058978 | 378.7879 | 1.4222 | 0.030382 | 0.96962
| 202 | 0.0023 | -2.6383 | -2.6545 | 0.016241 | 0.00026376 | 99.47% | 0.051582 | 434.7826 | 1.4206 | 0.030193 | 0.96981
| 203 | 0.00192 | -2.7167 | -2.7187 | 0.0020232 | 4.0934e-06 | 99.56% | 0.043247 | 520.8333 | 1.4191 | 0.03001 | 0.96999 | 0.029922
| 204 | 0.00168 | -2.7747 | -2.7687 | -0.0059912 | 3.5894e-05 | 99.61% | 0.037945 | 595.2381 | 1.4176 | 0.029833 | 0.97017
| 205 | 0.00156 | -2.8069 | -2.8044 | -0.0024299 | 5.9044e-06 | 99.64% | 0.035283 | 641.0256 | 1.4161 | 0.029662 | 0.97034
| 206 | 0.0015 | -2.8239 | -2.826 | 0.0020512 | 4.2073e-06 | 99.66% | 0.033949 | 666.6667 | 1.4147 | 0.029496 | 0.9705
| 207 | 0.0014 | -2.8539 | -2.8456 | -0.0082493 | 6.805e-05 | 99.68% | 0.031722 | 714.2857 | 1.4133 | 0.029335 | 0.97067
| 208 | 0.00134 | -2.8729 | -2.8638 | -0.0091231 | 8.323e-05 | 99.69% | 0.030383 | 746.2687 | 1.412 | 0.029179 | 0.97082 | 0.029103
| 209 | 0.00125 | -2.9031 | -2.8817 | -0.021378 | 0.00045702 | 99.71% | 0.028372 | 800 | 1.4107 | 0.029027 | 0.97097
| 210 | 0.00126 | -2.8996 | -2.8994 | -0.00018555 | 3.4431e-08 | 99.71% | 0.028596 | 793.6508 | 1.4095 | 0.02888 | 0.97112
| 211 | 0.00117 | -2.9318 | -2.917 | -0.014844 | 0.00022036 | 99.73% | 0.026581 | 854.7009 | 1.4082 | 0.028737 | 0.97126
| 212 | 0.00114 | -2.9431 | -2.9343 | -0.008804 | 7.751e-05 | 99.74% | 0.025908 | 877.193 | 1.4071 | 0.028598 | 0.9714
| 213 | 0.0011 | -2.9586 | -2.9514 | -0.0071973 | 5.1801e-05 | 99.75% | 0.02501 | 909.0909 | 1.4059 | 0.028462 | 0.97154 | 0.028396
| 214 | 0.00106 | -2.9747 | -2.9683 | -0.0063661 | 4.0528e-05 | 99.76% | 0.024112 | 943.3962 | 1.4048 | 0.028331 | 0.97167
| 215 | 0.00101 | -2.9957 | -2.985 | -0.010632 | 0.00011303 | 99.77% | 0.022988 | 990.099 | 1.4037 | 0.028202 | 0.9718
| 216 | 0.000988 | -3.0052 | -3.0016 | -0.0036746 | 1.3503e-05 | 99.77% | 0.022493 | 1012.1457 | 1.4026 | 0.028078 | 0.97192
| 217 | 0.000946 | -3.0241 | -3.0179 | -0.0062145 | 3.862e-05 | 99.78% | 0.021547 | 1057.0825 | 1.4015 | 0.027956 | 0.97204
| 218 | 0.000931 | -3.0311 | -3.034 | 0.0029761 | 8.8571e-06 | 99.79% | 0.021209 | 1074.1139 | 1.4005 | 0.027838 | 0.97216 | 0.02778
| 219 | 0.000918 | -3.0372 | -3.05 | 0.012809 | 0.00016407 | 99.79% | 0.020916 | 1089.3246 | 1.3995 | 0.027722 | 0.97228
| 220 | 0.000805 | -3.0942 | -3.0657 | -0.028488 | 0.00081157 | 99.81% | 0.018365 | 1242.236 | 1.3985 | 0.02761 | 0.97239
| 221 | 0.000786 | -3.1046 | -3.0813 | -0.0233 | 0.00054291 | 99.82% | 0.017936 | 1272.2646 | 1.3976 | 0.0275 | 0.9725
| 222 | 0.000737 | -3.1325 | -3.0967 | -0.035881 | 0.0012875 | 99.83% | 0.016827 | 1356.8521 | 1.3967 | 0.027393 | 0.97261
| 223 | 0.000714 | -3.1463 | -3.1118 | -0.034461 | 0.0011876 | 99.84% | 0.016306 | 1400.5602 | 1.3958 | 0.027288 | 0.97271 | 0.027237
| 224 | 0.0007 | -3.1549 | -3.1268 | -0.028056 | 0.00078713 | 99.84% | 0.015989 | 1428.5714 | 1.3949 | 0.027186 | 0.97281
| 225 | 0.00065 | -3.1871 | -3.1417 | -0.045416 | 0.0020626 | 99.85% | 0.014855 | 1538.4615 | 1.394 | 0.027086 | 0.97291
| 226 | 0.000669 | -3.1746 | -3.1563 | -0.018259 | 0.0003334 | 99.85% | 0.015286 | 1494.7683 | 1.3932 | 0.026989 | 0.97301
| 227 | 0.000643 | -3.1918 | -3.1708 | -0.021008 | 0.00044134 | 99.85% | 0.014697 | 1555.21 | 1.3923 | 0.026893 | 0.97311
| 228 | 0.000615 | -3.2111 | -3.1851 | -0.026054 | 0.00067881 | 99.86% | 0.014061 | 1626.0163 | 1.3915 | 0.0268 | 0.9732 | 0.026755
| 229 | 0.0006 | -3.2218 | -3.1992 | -0.022662 | 0.00051359 | 99.86% | 0.013721 | 1666.6667 | 1.3907 | 0.026709 | 0.97329
| 230 | 0.000595 | -3.2255 | -3.2131 | -0.012354 | 0.00015262 | 99.86% | 0.013607 | 1680.6723 | 1.3899 | 0.02662 | 0.97338
| 231 | 0.0005778 | -3.2382 | -3.2269 | -0.011322 | 0.00012819 | 99.87% | 0.013216 | 1730.7027 | 1.3892 | 0.026533 | 0.97347
| 232 | 0.000542 | -3.266 | -3.2405 | -0.025498 | 0.00065015 | 99.88% | 0.012402 | 1845.0185 | 1.3884 | 0.026448 | 0.97355
| 233 | 0.000534 | -3.2725 | -3.2539 | -0.018522 | 0.00034305 | 99.88% | 0.012221 | 1872.6592 | 1.3877 | 0.026364 | 0.97364 | 0.026323
| 234 | 0.000535 | -3.2716 | -3.2672 | -0.0044402 | 1.9715e-05 | 99.88% | 0.012243 | 1869.1589 | 1.387 | 0.026283 | 0.97372
| 235 | 0.000542 | -3.266 | -3.2803 | 0.01431 | 0.00020478 | 99.88% | 0.012402 | 1845.0185 | 1.3863 | 0.026203 | 0.9738
| 236 | 0.00051 | -3.2924 | -3.2933 | 0.00082348 | 6.7812e-07 | 99.88% | 0.011675 | 1960.7843 | 1.3856 | 0.026124 | 0.97388
| 237 | 0.000493 | -3.3072 | -3.306 | -0.0011179 | 1.2496e-06 | 99.89% | 0.011288 | 2028.3976 | 1.3849 | 0.026047 | 0.97395
| 238 | 0.000452 | -3.3449 | -3.3187 | -0.026203 | 0.00068661 | 99.9% | 0.010354 | 2212.3894 | 1.3842 | 0.025972 | 0.97403 | 0.025935
| 239 | 0.000471 | -3.327 | -3.3311 | 0.0041453 | 1.7184e-05 | 99.89% | 0.010787 | 2123.1423 | 1.3836 | 0.025899 | 0.9741
| 240 | 0.000483 | -3.3161 | -3.3434 | 0.027382 | 0.00074979 | 99.89% | 0.01106 | 2070.3934 | 1.383 | 0.025826 | 0.97417
| 241 | 0.000468 | -3.3298 | -3.3556 | 0.025838 | 0.00066762 | 99.89% | 0.010718 | 2136.7521 | 1.3823 | 0.025755 | 0.97424
| 242 | 0.000439 | -3.3575 | -3.3676 | 0.010062 | 0.00010125 | 99.9% | 0.010057 | 2277.9043 | 1.3817 | 0.025686 | 0.97431
| 243 | 0.000419 | -3.3778 | -3.3795 | 0.0016673 | 2.78e-06 | 99.9% | 0.0096014 | 2386.6348 | 1.3811 | 0.025618 | 0.97438 | 0.025584
| 244 | 0.000405 | -3.3925 | -3.3912 | -0.0013845 | 1.9169e-06 | 99.91% | 0.0092821 | 2469.1358 | 1.3805 | 0.025551 | 0.97445
| 245 | 0.000392 | -3.4067 | -3.4027 | -0.0039929 | 1.5943e-05 | 99.91% | 0.0089855 | 2551.0204 | 1.3799 | 0.025485 | 0.97451
| 246 | 0.000399 | -3.399 | -3.4141 | 0.01511 | 0.00022831 | 99.91% | 0.0091452 | 2506.2657 | 1.3794 | 0.025421 | 0.97458
| 247 | 0.000401 | -3.3969 | -3.4254 | 0.028554 | 0.00081534 | 99.91% | 0.0091909 | 2493.7656 | 1.3788 | 0.025358 | 0.97464
| 248 | 0.000392 | -3.4067 | -3.4365 | 0.029827 | 0.00088967 | 99.91% | 0.0089855 | 2551.0204 | 1.3783 | 0.025296 | 0.9747 | 0.025265
| 249 | 0.00039 | -3.4089 | -3.4475 | 0.038598 | 0.0014898 | 99.91% | 0.0089399 | 2564.1026 | 1.3777 | 0.025235 | 0.97477
| 250 | 0.000376 | -3.4248 | -3.4584 | 0.033575 | 0.0011273 | 99.91% | 0.0086203 | 2659.5745 | 1.3772 | 0.025175 | 0.97483
| 251 | 0.000336 | -3.4737 | -3.4691 | -0.004555 | 2.0748e-05 | 99.92% | 0.0077068 | 2976.1905 | 1.3767 | 0.025116 | 0.97488
| 252 | 0.000351 | -3.4547 | -3.4797 | 0.024997 | 0.00062484 | 99.92% | 0.0080495 | 2849.0028 | 1.3761 | 0.025059 | 0.97494
| 253 | 0.000359 | -3.4449 | -3.4901 | 0.045236 | 0.0020463 | 99.92% | 0.0082322 | 2785.5153 | 1.3756 | 0.025002 | 0.975 | 0.024982
| 253.7 | 0.00031814 | -3.4974 | -3.4974 | 0 | 1.9722e-31 | 99.93% | 0.0072987 | 3143.2527 | 1.3753 | 0.024963 | 0.97504
| 254 | 0.000348 | -3.4584 | -3.5005 | 0.042041 | 0.0017675 | 99.92% | 0.007981 | 2873.5632 | 1.3751 | 0.024946 | 0.97505
| 255 | 0.000326 | -3.4868 | -3.5107 | 0.023871 | 0.00056983 | 99.92% | 0.0074783 | 3067.4847 | 1.3746 | 0.024891 | 0.97511
| 256 | 0.000323 | -3.4908 | -3.5207 | 0.029921 | 0.00089524 | 99.93% | 0.0074098 | 3095.9752 | 1.3742 | 0.024838 | 0.97516
| 257 | 0.000323 | -3.4908 | -3.5307 | 0.03986 | 0.0015888 | 99.93% | 0.0074098 | 3095.9752 | 1.3737 | 0.024785 | 0.97522 | 0.024759
| 258 | 0.000306 | -3.5143 | -3.5405 | 0.026194 | 0.00068611 | 99.93% | 0.0070211 | 3267.9739 | 1.3732 | 0.024733 | 0.97527
| 259 | 0.000328 | -3.4841 | -3.5502 | 0.066039 | 0.0043612 | 99.92% | 0.007524 | 3048.7805 | 1.3728 | 0.024682 | 0.97532
| 260 | 0.000308 | -3.5114 | -3.5597 | 0.048289 | 0.0023318 | 99.93% | 0.0070669 | 3246.7532 | 1.3723 | 0.024631 | 0.97537
| 261 | 0.000331 | -3.4802 | -3.5692 | 0.089021 | 0.0079247 | 99.92% | 0.0075926 | 3021.148 | 1.3719 | 0.024582 | 0.97542
| 262 | 0.000279 | -3.5544 | -3.5785 | 0.024135 | 0.00058249 | 99.94% | 0.0064036 | 3584.2294 | 1.3714 | 0.024533 | 0.97547 | 0.024509
| 263 | 0.000285 | -3.5452 | -3.5878 | 0.042598 | 0.0018146 | 99.93% | 0.0065409 | 3508.7719 | 1.371 | 0.024485 | 0.97551
| 264 | 0.000273 | -3.5638 | -3.5969 | 0.033026 | 0.0010907 | 99.94% | 0.0062663 | 3663.0037 | 1.3706 | 0.024438 | 0.97556
| 265 | 0.000251 | -3.6003 | -3.6059 | 0.0055347 | 3.0633e-05 | 99.94% | 0.0057628 | 3984.0637 | 1.3702 | 0.024392 | 0.97561
| 266 | 0.000264 | -3.5784 | -3.6147 | 0.036353 | 0.0013216 | 99.94% | 0.0060604 | 3787.8788 | 1.3698 | 0.024346 | 0.97565
| 267 | 0.000236 | -3.6271 | -3.6235 | -0.0035583 | 1.2661e-05 | 99.95% | 0.0054194 | 4237.2881 | 1.3694 | 0.024301 | 0.9757 | 0.024279
| 268 | 0.000233 | -3.6326 | -3.6322 | -0.00044022 | 1.9379e-07 | 99.95% | 0.0053507 | 4291.8455 | 1.369 | 0.024257 | 0.97574
| 269 | 0.000239 | -3.6216 | -3.6408 | 0.019171 | 0.00036754 | 99.94% | 0.0054881 | 4184.1004 | 1.3686 | 0.024214 | 0.97579
| 270 | 0.000236 | -3.6271 | -3.6492 | 0.022153 | 0.00049073 | 99.95% | 0.0054194 | 4237.2881 | 1.3682 | 0.024171 | 0.97583
| 271 | 0.000234 | -3.6308 | -3.6576 | 0.026822 | 0.00071944 | 99.95% | 0.0053736 | 4273.5043 | 1.3678 | 0.024129 | 0.97587
| 272 | 0.000196 | -3.7077 | -3.6659 | -0.041871 | 0.0017532 | 99.95% | 0.0045029 | 5102.0408 | 1.3674 | 0.024087 | 0.97591 | 0.024067
| 273 | 0.000214 | -3.6696 | -3.674 | 0.0044564 | 1.986e-05 | 99.95% | 0.0049154 | 4672.8972 | 1.3671 | 0.024046 | 0.97595
| 274 | 0.000225 | -3.6478 | -3.6821 | 0.034299 | 0.0011764 | 99.95% | 0.0051674 | 4444.4444 | 1.3667 | 0.024006 | 0.97599
| 275 | 0.000216 | -3.6655 | -3.6901 | 0.02455 | 0.00060269 | 99.95% | 0.0049612 | 4629.6296 | 1.3663 | 0.023966 | 0.97603
| 276 | 0.000198 | -3.7033 | -3.698 | -0.0053514 | 2.8637e-05 | 99.95% | 0.0045487 | 5050.5051 | 1.366 | 0.023927 | 0.97607
| 277 | 0.000215 | -3.6676 | -3.7058 | 0.038219 | 0.0014607 | 99.95% | 0.0049383 | 4651.1628 | 1.3656 | 0.023889 | 0.97611 | 0.02387
| 278 | 0.000208 | -3.6819 | -3.7135 | 0.031552 | 0.00099554 | 99.95% | 0.0047779 | 4807.6923 | 1.3653 | 0.023851 | 0.97615
| 279 | 0.000209 | -3.6799 | -3.7211 | 0.041256 | 0.0017021 | 99.95% | 0.0048008 | 4784.689 | 1.3649 | 0.023813 | 0.97619
| 280 | 0.000222 | -3.6536 | -3.7286 | 0.074999 | 0.0056249 | 99.95% | 0.0050987 | 4504.5045 | 1.3646 | 0.023776 | 0.97622
| 281 | 0.000213 | -3.6716 | -3.7361 | 0.064479 | 0.0041575 | 99.95% | 0.0048925 | 4694.8357 | 1.3643 | 0.02374 | 0.97626
| 282 | 0.00018 | -3.7447 | -3.7435 | -0.0012572 | 1.5806e-06 | 99.96% | 0.0041361 | 5555.5556 | 1.364 | 0.023704 | 0.9763 | 0.023687
| 283 | 0.000177 | -3.752 | -3.7508 | -0.0012653 | 1.601e-06 | 99.96% | 0.0040673 | 5649.7175 | 1.3636 | 0.023669 | 0.97633
| 284 | 0.000185 | -3.7328 | -3.758 | 0.025146 | 0.00063233 | 99.96% | 0.0042507 | 5405.4054 | 1.3633 | 0.023634 | 0.97637
| 285 | 0.000178 | -3.7496 | -3.7651 | 0.015531 | 0.00024121 | 99.96% | 0.0040902 | 5617.9775 | 1.363 | 0.0236 | 0.9764
| 286 | 0.000175 | -3.757 | -3.7722 | 0.015211 | 0.00023137 | 99.96% | 0.0040214 | 5714.2857 | 1.3627 | 0.023566 | 0.97643
| 287 | 0.000164 | -3.7852 | -3.7792 | -0.0059943 | 3.5932e-05 | 99.96% | 0.0037691 | 6097.561 | 1.3624 | 0.023532 | 0.97647 | 0.023516
| 288 | 0.000168 | -3.7747 | -3.7861 | 0.011389 | 0.00012971 | 99.96% | 0.0038609 | 5952.381 | 1.3621 | 0.0235 | 0.9765
| 289 | 0.000159 | -3.7986 | -3.7929 | -0.0056749 | 3.2204e-05 | 99.96% | 0.0036544 | 6289.3082 | 1.3618 | 0.023467 | 0.97653
| 290 | 0.000163 | -3.7878 | -3.7997 | 0.011896 | 0.00014152 | 99.96% | 0.0037462 | 6134.9693 | 1.3615 | 0.023435 | 0.97657
| 291 | 0.00016 | -3.7959 | -3.8064 | 0.010543 | 0.00011117 | 99.96% | 0.0036774 | 6250 | 1.3612 | 0.023403 | 0.9766
| 292 | 0.000157 | -3.8041 | -3.8131 | 0.0089738 | 8.0529e-05 | 99.96% | 0.0036085 | 6369.4268 | 1.3609 | 0.023372 | 0.97663 | 0.023357
| 293 | 0.00016 | -3.7959 | -3.8197 | 0.023782 | 0.0005656 | 99.96% | 0.0036774 | 6250 | 1.3607 | 0.023341 | 0.97666
| 294 | 0.000151 | -3.821 | -3.8262 | 0.0051669 | 2.6697e-05 | 99.97% | 0.0034709 | 6622.5166 | 1.3604 | 0.023311 | 0.97669
| 295 | 0.000145 | -3.8386 | -3.8327 | -0.0059733 | 3.568e-05 | 99.97% | 0.0033332 | 6896.5517 | 1.3601 | 0.023281 | 0.97672
| 296 | 0.000144 | -3.8416 | -3.8391 | -0.0025673 | 6.5909e-06 | 99.97% | 0.0033102 | 6944.4444 | 1.3598 | 0.023251 | 0.97675
| 297 | 0.000123 | -3.9101 | -3.8454 | -0.064669 | 0.004182 | 99.97% | 0.0028282 | 8130.0813 | 1.3596 | 0.023222 | 0.97678 | 0.023208
| 298 | 0.000114 | -3.9431 | -3.8517 | -0.091366 | 0.0083478 | 99.97% | 0.0026215 | 8771.9298 | 1.3593 | 0.023193 | 0.97681
| 299 | 0.000124 | -3.9066 | -3.858 | -0.048599 | 0.0023618 | 99.97% | 0.0028511 | 8064.5161 | 1.359 | 0.023165 | 0.97684
| 300 | 0.000124 | -3.9066 | -3.8642 | -0.042398 | 0.0017976 | 99.97% | 0.0028511 | 8064.5161 | 1.3588 | 0.023137 | 0.97686
| 301 | 0.000146 | -3.8356 | -3.8703 | 0.034685 | 0.001203 | 99.97% | 0.0033561 | 6849.3151 | 1.3585 | 0.023109 | 0.97689
| 302 | 0.000108 | -3.9666 | -3.8764 | -0.090139 | 0.008125 | 99.98% | 0.0024837 | 9259.2593 | 1.3583 | 0.023082 | 0.97692 | 0.023068
| 303 | 0.000138 | -3.8601 | -3.8825 | 0.022377 | 0.00050074 | 99.97% | 0.0031725 | 7246.3768 | 1.358 | 0.023055 | 0.97695
| 304 | 0.000127 | -3.8962 | -3.8885 | -0.0076807 | 5.8992e-05 | 99.97% | 0.00292 | 7874.0157 | 1.3578 | 0.023028 | 0.97697
| 305 | 0.000112 | -3.9508 | -3.8945 | -0.05629 | 0.0031686 | 99.97% | 0.0025756 | 8928.5714 | 1.3576 | 0.023002 | 0.977
| 306 | 0.000122 | -3.9136 | -3.9004 | -0.013212 | 0.00017456 | 99.97% | 0.0028052 | 8196.7213 | 1.3573 | 0.022975 | 0.97702
| 307 | 0.000116 | -3.9355 | -3.9063 | -0.029215 | 0.00085353 | 99.97% | 0.0026674 | 8620.6897 | 1.3571 | 0.02295 | 0.97705 | 0.022937
| 308 | 0.000104 | -3.983 | -3.9122 | -0.070777 | 0.0050094 | 99.98% | 0.0023918 | 9615.3846 | 1.3568 | 0.022924 | 0.97708
| 309 | 0.000125 | -3.9031 | -3.918 | 0.014927 | 0.00022283 | 99.97% | 0.0028741 | 8000 | 1.3566 | 0.022899 | 0.9771
| 310 | 0.000112 | -3.9508 | -3.9238 | -0.026969 | 0.00072733 | 99.97% | 0.0025756 | 8928.5714 | 1.3564 | 0.022874 | 0.97713
| 311 | 0.00013 | -3.8861 | -3.9296 | 0.043521 | 0.0018941 | 99.97% | 0.0029889 | 7692.3077 | 1.3562 | 0.02285 | 0.97715
| 312 | 0.000119 | -3.9245 | -3.9353 | 0.01086 | 0.00011794 | 99.97% | 0.0027363 | 8403.3613 | 1.3559 | 0.022826 | 0.97717 | 0.022814
| 313 | 0.000113 | -3.9469 | -3.941 | -0.0059003 | 3.4813e-05 | 99.97% | 0.0025985 | 8849.5575 | 1.3557 | 0.022802 | 0.9772
| 314 | 0.000101 | -3.9957 | -3.9467 | -0.048975 | 0.0023985 | 99.98% | 0.0023229 | 9900.9901 | 1.3555 | 0.022778 | 0.97722
| 315 | 0.000105 | -3.979 | -3.9524 | -0.026637 | 0.00070955 | 99.98% | 0.0024148 | 9523.8095 | 1.3553 | 0.022755 | 0.97725
| 316 | 9.6e-05 | -4.0177 | -3.958 | -0.059729 | 0.0035676 | 99.98% | 0.002208 | 10416.6667 | 1.3551 | 0.022732 | 0.97727
| 317 | 0.000103 | -3.9872 | -3.9636 | -0.023547 | 0.00055447 | 99.98% | 0.0023689 | 9708.7379 | 1.3549 | 0.022709 | 0.97729 | 0.022698
| 318 | 9.3e-05 | -4.0315 | -3.9692 | -0.062304 | 0.0038818 | 99.98% | 0.0021391 | 10752.6882 | 1.3547 | 0.022686 | 0.97731
| 319 | 0.000108 | -3.9666 | -3.9748 | 0.008218 | 6.7535e-05 | 99.98% | 0.0024837 | 9259.2593 | 1.3545 | 0.022664 | 0.97734
| 320 | 0.0001 | -4 | -3.9804 | -0.01964 | 0.00038573 | 99.98% | 0.0022999 | 10000 | 1.3542 | 0.022642 | 0.97736
| 325 | 8.0562e-05 | -4.0939 | -4.0831 | -0.010803 | 0.00011671 | 99.98% | 0.0018533 | 12412.8576 | 1.3533 | 0.022535 | 0.97746
| 330 | 6.9487e-05 | -4.1581 | -4.1538 | -0.004259 | 1.8139e-05 | 99.98% | 0.0015987 | 14391.1568 | 1.3523 | 0.022435 | 0.97757
| 335 | 5.9933e-05 | -4.2223 | -4.2225 | 0.00011601 | 1.3458e-08 | 99.99% | 0.001379 | 16685.3992 | 1.3515 | 0.022339 | 0.97766
| 340 | 5.1681e-05 | -4.2867 | -4.2882 | 0.0015481 | 2.3965e-06 | 99.99% | 0.0011893 | 19349.4546 | 1.3506 | 0.022249 | 0.97775
| 345 | 4.5167e-05 | -4.3452 | -4.3504 | 0.0052475 | 2.7537e-05 | 99.99% | 0.0010395 | 22140.2413 | 1.3498 | 0.022163 | 0.97784
| 350 | 3.9521e-05 | -4.4032 | -4.4084 | 0.0052182 | 2.723e-05 | 99.99% | 0.00090959 | 25303.1329 | 1.3491 | 0.022082 | 0.97792
| 355 | 3.4744e-05 | -4.4591 | -4.4614 | 0.0022796 | 5.1964e-06 | 99.99% | 0.00079968 | 28782.3137 | 1.3484 | 0.022004 | 0.978
| 360 | 3.0835e-05 | -4.511 | -4.5088 | -0.0021885 | 4.7897e-06 | 99.99% | 0.00070975 | 32430.776 | 1.3477 | 0.021931 | 0.97807
| 365 | 2.8229e-05 | -4.5493 | -4.5498 | 0.00048141 | 2.3175e-07 | 99.99% | 0.00064979 | 35424.386 | 1.347 | 0.02186 | 0.97814
| 370 | 2.6058e-05 | -4.5841 | -4.5838 | -0.00031374 | 9.8432e-08 | 99.99% | 0.00059982 | 38376.4182 | 1.3464 | 0.021793 | 0.97821
| 375 | 2.4538e-05 | -4.6102 | -4.61 | -0.00019693 | 3.8781e-08 | 99.99% | 0.00056484 | 40753.7185 | 1.3458 | 0.021729 | 0.97827
| 380 | 2.3452e-05 | -4.6298 | -4.6277 | -0.0020787 | 4.3211e-06 | 99.99% | 0.00053985 | 42640.4647 | 1.3452 | 0.021668 | 0.97833
| 385 | 2.3018e-05 | -4.6379 | -4.6364 | -0.0015698 | 2.4644e-06 | 99.99% | 0.00052986 | 43445.0018 | 1.3447 | 0.021609 | 0.97839
| 390 | 2.3452e-05 | -4.6298 | -4.6352 | 0.0053294 | 2.8402e-05 | 99.99% | 0.00053985 | 42640.4647 | 1.3441 | 0.021553 | 0.97845
| 395 | 2.3669e-05 | -4.6258 | -4.6234 | -0.0024314 | 5.9118e-06 | 99.99% | 0.00054485 | 42249.2678 | 1.3436 | 0.021499 | 0.9785
| 400 | 2.5189e-05 | -4.5988 | -4.6004 | 0.0015923 | 2.5354e-06 | 99.99% | 0.00057983 | 39699.743 | 1.3432 | 0.021448 | 0.97855

## Weighting
IR Radiometer Reading |  | 0.14319 | mW/cm2 | from fluence calculations sheet
d sample depth |  | 0.40601 | cm | from fluence calculations sheet
fp radiometer lamp correction factor |  | 3.1372 | unitless | =(sumRLE)/(sumRLE*fR)
Incident Intensity |  | 0.42241 | mW/cm2 | =fp*IR*DF*PF*RF
Average Intensity |  | 0.15796 | mW/cm2 | =sum Iavg
Weighted Average Intensity |  | 0.03253 | mW/cm2 | =sum Iavg*Rweight
Rλnormal |  | 6.1493e-06 | unitless | from calibration data below
COLUMN SUM | 2268.7255 | 1 | 0.0011842 | 192.5739 | 723.1785 | 0.42241 | 19.8989 | 0.15796 | 0.045 | 0 | 0 | 87.1904 | 0 |  | 0.03253
λ | RLE | L(λ) | R | fR | RLE*fR | I(λ) | a | Iavg | ADNA | R DNA | Aphotolyase | R photolyase | R 254 | R etc | Iavg*Rweight
Wavelength (nm) | SpectralOutput, relative lamp energy, Spectra tab | RLE/ Total RLE | S. 2021 Radiometer Calibration Factor | R / Rλpeak |  | IncidentIntensity*L(λ) | Sample Absorbance (cm-1) | Beer Law using abs over sample depth | DNA abs from 2001 LP/MP sheet | DNAabs / λDNAabs (choose wavelength) | Photolyase absorbance from Payne and Sancar 1990 | Absorbance / λAbsorbance (choose wavelength) | weighting rel 254 for germicidal |  | select weight and apply formula
300 | 0.04039 | 1.7803e-05 | 9.7416e-06 | 1.5842 | 0.063984 | 7.5202e-06 | 0.17729 | 6.93e-06 | 0.045 |  |  | 1.0022 |  |  | 6.9451e-06
301 | 0.0546 | 2.4066e-05 | 9.471e-06 | 1.5402 | 0.084093 | 1.0166e-05 | 0.17505 | 9.3777e-06 |  |  |  | 0.80027 |  |  | 7.5047e-06
302 | 0.021953 | 9.6763e-06 | 9.2378e-06 | 1.5022 | 0.032979 | 4.0874e-06 | 0.17276 | 3.7744e-06 |  |  |  | 0.60976 |  |  | 2.3015e-06
303 | 0.01612 | 7.1052e-06 | 9.0158e-06 | 1.4661 | 0.023634 | 3.0013e-06 | 0.17076 | 2.774e-06 |  |  |  | 0.51529 |  |  | 1.4294e-06
304 | 0.016088 | 7.0912e-06 | 8.801e-06 | 1.4312 | 0.023025 | 2.9954e-06 | 0.16882 | 2.771e-06 |  |  |  | 0.45584 |  |  | 1.2632e-06
305 | 0.029619 | 1.3055e-05 | 8.6111e-06 | 1.4003 | 0.041476 | 5.5147e-06 | 0.16691 | 5.106e-06 |  |  |  | 0.39311 |  |  | 2.0072e-06
306 | 0.039393 | 1.7363e-05 | 8.4457e-06 | 1.3734 | 0.054103 | 7.3345e-06 | 0.16504 | 6.7967e-06 |  |  |  | 0.35139 |  |  | 2.3883e-06
307 | 0.062761 | 2.7664e-05 | 8.2961e-06 | 1.3491 | 0.084671 | 1.1685e-05 | 0.16346 | 1.0836e-05 |  |  |  | 0.30495 |  |  | 3.3046e-06
308 | 0.061021 | 2.6897e-05 | 8.1553e-06 | 1.3262 | 0.080927 | 1.1361e-05 | 0.16181 | 1.0544e-05 |  |  |  | 0.27567 |  |  | 2.9067e-06
309 | 0.02974 | 1.3109e-05 | 8.0215e-06 | 1.3045 | 0.038795 | 5.5373e-06 | 0.16026 | 5.1425e-06 |  |  |  | 0.26646 |  |  | 1.3703e-06
310 | 0.01299 | 5.7256e-06 | 7.8939e-06 | 1.2837 | 0.016675 | 2.4186e-06 | 0.15888 | 2.2475e-06 |  |  |  | 0.26223 |  |  | 5.8937e-07
311 | 0.017269 | 7.6119e-06 | 7.7766e-06 | 1.2646 | 0.021839 | 3.2154e-06 | 0.15734 | 2.9901e-06 |  |  |  | 0.26079 |  |  | 7.798e-07
312 | 0.025851 | 1.1394e-05 | 7.6703e-06 | 1.2473 | 0.032245 | 4.8131e-06 | 0.15596 | 4.4787e-06 |  |  |  | 0.26238 |  |  | 1.1751e-06
313 | 0.04439 | 1.9566e-05 | 7.5662e-06 | 1.2304 | 0.054619 | 8.265e-06 | 0.15468 | 7.6952e-06 |  |  |  | 0.26623 |  |  | 2.0487e-06
314 | 0.020515 | 9.0425e-06 | 7.4689e-06 | 1.2146 | 0.024917 | 3.8197e-06 | 0.1534 | 3.5584e-06 |  |  |  | 0.27123 |  |  | 9.6514e-07
315 | 0.017064 | 7.5213e-06 | 7.3775e-06 | 1.1997 | 0.020472 | 3.1771e-06 | 0.15217 | 2.9614e-06 |  |  |  | 0.27753 |  |  | 8.219e-07
316 | 0.019072 | 8.4065e-06 | 7.3057e-06 | 1.188 | 0.022658 | 3.551e-06 | 0.15099 | 3.3118e-06 |  |  |  | 0.28536 |  |  | 9.4503e-07
317 | 0.02648 | 1.1672e-05 | 7.2485e-06 | 1.1787 | 0.031213 | 4.9302e-06 | 0.14977 | 4.6007e-06 |  |  |  | 0.29482 |  |  | 1.3564e-06
318 | 0.026386 | 1.163e-05 | 7.1932e-06 | 1.1698 | 0.030865 | 4.9128e-06 | 0.14851 | 4.587e-06 |  |  |  | 0.30435 |  |  | 1.396e-06
319 | 0.011448 | 5.046e-06 | 7.1367e-06 | 1.1606 | 0.013286 | 2.1315e-06 | 0.1475 | 1.991e-06 |  |  |  | 0.31492 |  |  | 6.2702e-07
320 | 0.017166 | 7.5665e-06 | 7.081e-06 | 1.1515 | 0.019767 | 3.1962e-06 | 0.14638 | 2.9872e-06 |  |  |  | 0.32649 |  |  | 9.7527e-07
321 | 0.0051338 | 2.2628e-06 | 7.0134e-06 | 1.1405 | 0.0058552 | 9.5586e-07 | 0.1452 | 8.9382e-07 |  |  |  | 0.33741 |  |  | 3.0159e-07
322 | 0.036643 | 1.6151e-05 | 6.9527e-06 | 1.1306 | 0.04143 | 6.8226e-06 | 0.14422 | 6.3826e-06 |  |  |  | 0.34832 |  |  | 2.2232e-06
323 | 0.03743 | 1.6498e-05 | 6.9004e-06 | 1.1221 | 0.042002 | 6.969e-06 | 0.14309 | 6.523e-06 |  |  |  | 0.35912 |  |  | 2.3425e-06
324 | 0.0045206 | 1.9926e-06 | 6.854e-06 | 1.1146 | 0.0050386 | 8.4168e-07 | 0.14219 | 7.8814e-07 |  |  |  | 0.36999 |  |  | 2.916e-07
325 | 0.031366 | 1.3825e-05 | 6.822e-06 | 1.1094 | 0.034797 | 5.84e-06 | 0.14098 | 5.4715e-06 |  |  |  | 0.38108 |  |  | 2.0851e-06
326 | 0.04756 | 2.0963e-05 | 6.7948e-06 | 1.105 | 0.052553 | 8.8552e-06 | 0.14023 | 8.2993e-06 |  |  |  | 0.39219 |  |  | 3.2549e-06
327 | 0.036486 | 1.6082e-05 | 6.7753e-06 | 1.1018 | 0.040201 | 6.7934e-06 | 0.13913 | 6.3701e-06 |  |  |  | 0.40173 |  |  | 2.5591e-06
328 | 0.017268 | 7.6111e-06 | 6.7567e-06 | 1.0988 | 0.018973 | 3.215e-06 | 0.1381 | 3.0161e-06 |  |  |  | 0.41178 |  |  | 1.242e-06
329 | 0.031399 | 1.384e-05 | 6.7209e-06 | 1.093 | 0.034318 | 5.8462e-06 | 0.1372 | 5.4868e-06 |  |  |  | 0.42205 |  |  | 2.3157e-06
330 | 0.042079 | 1.8548e-05 | 6.6951e-06 | 1.0888 | 0.045814 | 7.8347e-06 | 0.13616 | 7.3566e-06 |  |  |  | 0.43202 |  |  | 3.1782e-06
331 | 0.03027 | 1.3342e-05 | 6.6708e-06 | 1.0848 | 0.032837 | 5.636e-06 | 0.13548 | 5.2936e-06 |  |  |  | 0.44171 |  |  | 2.3383e-06
332 | 0.012913 | 5.6916e-06 | 6.6465e-06 | 1.0809 | 0.013957 | 2.4042e-06 | 0.13448 | 2.2592e-06 |  |  |  | 0.45155 |  |  | 1.0202e-06
333 | 0.030017 | 1.3231e-05 | 6.638e-06 | 1.0795 | 0.032403 | 5.5889e-06 | 0.13361 | 5.2539e-06 |  |  |  | 0.4614 |  |  | 2.4241e-06
334 | 0.049925 | 2.2006e-05 | 6.6221e-06 | 1.0769 | 0.053764 | 9.2955e-06 | 0.13267 | 8.7422e-06 |  |  |  | 0.47209 |  |  | 4.1271e-06
335 | 0.017359 | 7.6513e-06 | 6.6068e-06 | 1.0744 | 0.01865 | 3.232e-06 | 0.13179 | 3.0409e-06 |  |  |  | 0.48238 |  |  | 1.4668e-06
336 | 0.03727 | 1.6428e-05 | 6.5968e-06 | 1.0728 | 0.039983 | 6.9393e-06 | 0.13101 | 6.5312e-06 |  |  |  | 0.49256 |  |  | 3.217e-06
337 | 0.022561 | 9.9444e-06 | 6.5881e-06 | 1.0714 | 0.024171 | 4.2007e-06 | 0.13005 | 3.9553e-06 |  |  |  | 0.50504 |  |  | 1.9976e-06
338 | -0.00010997 | -4.8472e-08 | 6.5785e-06 | 1.0698 | -0.00011765 | -2.0475e-08 | 0.12947 | -1.9285e-08 |  |  |  | 0.51679 |  |  | -9.9661e-09
339 | 0.012431 | 5.4794e-06 | 6.5716e-06 | 1.0687 | 0.013285 | 2.3146e-06 | 0.12846 | 2.181e-06 |  |  |  | 0.5283 |  |  | 1.1522e-06
340 | 0.032322 | 1.4247e-05 | 6.5694e-06 | 1.0683 | 0.03453 | 6.018e-06 | 0.12752 | 5.6731e-06 |  |  |  | 0.54046 |  |  | 3.0661e-06
341 | 0.050305 | 2.2173e-05 | 6.5746e-06 | 1.0692 | 0.053784 | 9.3662e-06 | 0.12698 | 8.8317e-06 |  |  |  | 0.5526 |  |  | 4.8804e-06
342 | 0.04183 | 1.8437e-05 | 6.5888e-06 | 1.0715 | 0.044819 | 7.7882e-06 | 0.12577 | 7.3478e-06 |  |  |  | 0.56497 |  |  | 4.1513e-06
343 | 0.036653 | 1.6156e-05 | 6.5979e-06 | 1.073 | 0.039327 | 6.8245e-06 | 0.12538 | 6.4397e-06 |  |  |  | 0.57759 |  |  | 3.7195e-06
344 | 0.028496 | 1.256e-05 | 6.5999e-06 | 1.0733 | 0.030584 | 5.3056e-06 | 0.12448 | 5.0085e-06 |  |  |  | 0.58993 |  |  | 2.9547e-06
345 | 0.027469 | 1.2108e-05 | 6.6e-06 | 1.0733 | 0.029482 | 5.1144e-06 | 0.12359 | 4.83e-06 |  |  |  | 0.6025 |  |  | 2.9101e-06
346 | 0.047918 | 2.1121e-05 | 6.5955e-06 | 1.0726 | 0.051395 | 8.9218e-06 | 0.12321 | 8.4272e-06 |  |  |  | 0.61531 |  |  | 5.1853e-06
347 | 0.029554 | 1.3027e-05 | 6.5982e-06 | 1.073 | 0.031711 | 5.5026e-06 | 0.12191 | 5.2007e-06 |  |  |  | 0.62802 |  |  | 3.2661e-06
348 | 0.051236 | 2.2584e-05 | 6.6098e-06 | 1.0749 | 0.055073 | 9.5397e-06 | 0.12184 | 9.0164e-06 |  |  |  | 0.63979 |  |  | 5.7686e-06
349 | 0.072871 | 3.212e-05 | 6.6215e-06 | 1.0768 | 0.078467 | 1.3568e-05 | 0.12046 | 1.2832e-05 |  |  |  | 0.65182 |  |  | 8.364e-06
350 | 0.096584 | 4.2572e-05 | 6.6365e-06 | 1.0792 | 0.10424 | 1.7983e-05 | 0.11961 | 1.7014e-05 |  |  |  | 0.66398 |  |  | 1.1297e-05
351 | 0.078438 | 3.4574e-05 | 6.6646e-06 | 1.0838 | 0.085011 | 1.4604e-05 | 0.1189 | 1.3822e-05 |  |  |  | 0.67577 |  |  | 9.3404e-06
352 | 0.073867 | 3.2559e-05 | 6.6963e-06 | 1.0889 | 0.080438 | 1.3753e-05 | 0.11825 | 1.302e-05 |  |  |  | 0.68616 |  |  | 8.9341e-06
353 | 0.090164 | 3.9742e-05 | 6.733e-06 | 1.0949 | 0.098722 | 1.6788e-05 | 0.11751 | 1.5898e-05 |  |  |  | 0.69765 |  |  | 1.1091e-05
354 | 0.13674 | 6.027e-05 | 6.7653e-06 | 1.1002 | 0.15043 | 2.5459e-05 | 0.11683 | 2.4118e-05 |  |  |  | 0.71087 |  |  | 1.7145e-05
355 | 0.15754 | 6.9439e-05 | 6.7864e-06 | 1.1036 | 0.17386 | 2.9332e-05 | 0.11617 | 2.7795e-05 |  |  |  | 0.72365 |  |  | 2.0114e-05
356 | 0.16741 | 7.3789e-05 | 6.8075e-06 | 1.107 | 0.18533 | 3.117e-05 | 0.11556 | 2.9545e-05 |  |  |  | 0.73948 |  |  | 2.1848e-05
357 | 0.17534 | 7.7284e-05 | 6.8202e-06 | 1.1091 | 0.19446 | 3.2646e-05 | 0.11489 | 3.0954e-05 |  |  |  | 0.75485 |  |  | 2.3365e-05
358 | 0.18327 | 8.0779e-05 | 6.8324e-06 | 1.1111 | 0.20362 | 3.4122e-05 | 0.11417 | 3.2364e-05 |  |  |  | 0.77023 |  |  | 2.4928e-05
359 | 0.22684 | 9.9987e-05 | 6.8379e-06 | 1.112 | 0.25224 | 4.2236e-05 | 0.1136 | 4.0071e-05 |  |  |  | 0.78694 |  |  | 3.1533e-05
360 | 0.24978 | 0.0001101 | 6.8249e-06 | 1.1099 | 0.27722 | 4.6506e-05 | 0.11291 | 4.4136e-05 |  |  |  | 0.80322 |  |  | 3.5451e-05
361 | 0.26822 | 0.00011823 | 6.8229e-06 | 1.1095 | 0.2976 | 4.994e-05 | 0.11226 | 4.7409e-05 |  |  |  | 0.8194 |  |  | 3.8847e-05
362 | 0.31529 | 0.00013897 | 6.8217e-06 | 1.1094 | 0.34977 | 5.8704e-05 | 0.11163 | 5.5745e-05 |  |  |  | 0.8308 |  |  | 4.6313e-05
363 | 0.44094 | 0.00019436 | 6.8158e-06 | 1.1084 | 0.48873 | 8.2098e-05 | 0.11106 | 7.798e-05 |  |  |  | 0.8445 |  |  | 6.5854e-05
364 | 4.3027 | 0.0018965 | 6.8146e-06 | 1.1082 | 4.7682 | 0.00080112 | 0.11045 | 0.00076115 |  |  |  | 0.85883 |  |  | 0.0006537
365 | 8.6517 | 0.0038134 | 6.8053e-06 | 1.1067 | 9.5746 | 0.0016108 | 0.10982 | 0.0015309 |  |  |  | 0.87219 |  |  | 0.0013352
366 | 3.4135 | 0.0015046 | 6.7819e-06 | 1.1029 | 3.7647 | 0.00063556 | 0.10913 | 0.00060422 |  |  |  | 0.88557 |  |  | 0.00053508
367 | 0.9051 | 0.00039894 | 6.7476e-06 | 1.0973 | 0.99316 | 0.00016852 | 0.1086 | 0.00016025 |  |  |  | 0.89948 |  |  | 0.00014414
368 | 0.47032 | 0.00020731 | 6.7221e-06 | 1.0932 | 0.51414 | 8.757e-05 | 0.10799 | 8.3294e-05 |  |  |  | 0.91256 |  |  | 7.6011e-05
369 | 0.48951 | 0.00021576 | 6.7126e-06 | 1.0916 | 0.53435 | 9.1141e-05 | 0.10734 | 8.6718e-05 |  |  |  | 0.92042 |  |  | 7.9816e-05
370 | 0.52584 | 0.00023178 | 6.7137e-06 | 1.0918 | 0.5741 | 9.7906e-05 | 0.10671 | 9.3181e-05 |  |  |  | 0.93063 |  |  | 8.6717e-05
371 | 0.54979 | 0.00024234 | 6.7218e-06 | 1.0931 | 0.60098 | 0.00010237 | 0.10622 | 9.7447e-05 |  |  |  | 0.93644 |  |  | 9.1253e-05
372 | 0.60796 | 0.00026798 | 6.7122e-06 | 1.0915 | 0.66362 | 0.0001132 | 0.10564 | 0.00010779 |  |  |  | 0.94274 |  |  | 0.00010161
373 | 0.65381 | 0.00028819 | 6.6689e-06 | 1.0845 | 0.70906 | 0.00012173 | 0.10503 | 0.00011595 |  |  |  | 0.94955 |  |  | 0.0001101
374 | 0.66501 | 0.00029312 | 6.5933e-06 | 1.0722 | 0.71302 | 0.00012382 | 0.10449 | 0.00011796 |  |  |  | 0.95649 |  |  | 0.00011283
375 | 0.69039 | 0.00030431 | 6.5251e-06 | 1.0611 | 0.73258 | 0.00012854 | 0.10391 | 0.0001225 |  |  |  | 0.96334 |  |  | 0.00011801
376 | 0.72434 | 0.00031927 | 6.4776e-06 | 1.0534 | 0.76301 | 0.00013486 | 0.10341 | 0.00012855 |  |  |  | 0.97055 |  |  | 0.00012476
377 | 0.72978 | 0.00032167 | 6.4352e-06 | 1.0465 | 0.76371 | 0.00013588 | 0.10289 | 0.00012955 |  |  |  | 0.97654 |  |  | 0.00012651
378 | 0.76906 | 0.00033898 | 6.3929e-06 | 1.0396 | 0.79953 | 0.00014319 | 0.10235 | 0.00013655 |  |  |  | 0.98145 |  |  | 0.00013402
379 | 0.82792 | 0.00036493 | 6.3479e-06 | 1.0323 | 0.85466 | 0.00015415 | 0.10175 | 0.00014705 |  |  |  | 0.9872 |  |  | 0.00014516
380 | 0.88189 | 0.00038872 | 6.3001e-06 | 1.0245 | 0.90352 | 0.0001642 | 0.10121 | 0.00015667 |  |  |  | 0.99216 |  |  | 0.00015544
381 | 0.94297 | 0.00041564 | 6.2482e-06 | 1.0161 | 0.95813 | 0.00017557 | 0.10073 | 0.00016756 |  |  |  | 0.99578 |  |  | 0.00016685
382 | 0.9816 | 0.00043267 | 6.1982e-06 | 1.0079 | 0.9894 | 0.00018276 | 0.10025 | 0.00017446 |  |  |  | 0.99846 |  |  | 0.00017419
383 | 1 | 0.00044078 | 6.1493e-06 | 1 | 1 | 0.00018619 | 0.099713 | 0.00017777 |  |  |  | 1 |  |  | 0.00017777
384 | 1.0637 | 0.00046887 | 6.1006e-06 | 0.99208 | 1.0553 | 0.00019806 | 0.099182 | 0.00018915 |  |  |  | 0.99998 |  |  | 0.00018915
385 | 1.1365 | 0.00050096 | 6.0549e-06 | 0.98465 | 1.1191 | 0.00021161 | 0.098622 | 0.00020215 |  |  |  | 0.99786 |  |  | 0.00020172
386 | 1.1817 | 0.00052087 | 6.0072e-06 | 0.97689 | 1.1544 | 0.00022002 | 0.098199 | 0.00021022 |  |  |  | 0.99535 |  |  | 0.00020925
387 | 1.2631 | 0.00055675 | 5.9639e-06 | 0.96984 | 1.225 | 0.00023518 | 0.097638 | 0.00022476 |  |  |  | 0.99227 |  |  | 0.00022303
388 | 1.3055 | 0.00057544 | 5.9217e-06 | 0.963 | 1.2572 | 0.00024307 | 0.097198 | 0.00023236 |  |  |  | 0.98892 |  |  | 0.00022978
389 | 1.3766 | 0.00060679 | 5.8816e-06 | 0.95647 | 1.3167 | 0.00025632 | 0.096662 | 0.00024508 |  |  |  | 0.984 |  |  | 0.00024116
390 | 1.5097 | 0.00066546 | 5.8402e-06 | 0.94973 | 1.4338 | 0.0002811 | 0.096133 | 0.00026884 |  |  |  | 0.97647 |  |  | 0.00026251
391 | 1.5501 | 0.00068325 | 5.8007e-06 | 0.94332 | 1.4622 | 0.00028861 | 0.095691 | 0.00027608 |  |  |  | 0.96842 |  |  | 0.00026736
392 | 1.5646 | 0.00068963 | 5.7597e-06 | 0.93664 | 1.4655 | 0.00029131 | 0.095238 | 0.00027872 |  |  |  | 0.96031 |  |  | 0.00026765
393 | 1.657 | 0.00073037 | 5.7194e-06 | 0.93009 | 1.5412 | 0.00030852 | 0.094768 | 0.00029525 |  |  |  | 0.95206 |  |  | 0.00028109
394 | 1.7567 | 0.0007743 | 5.681e-06 | 0.92385 | 1.6229 | 0.00032707 | 0.094268 | 0.00031308 |  |  |  | 0.93956 |  |  | 0.00029415
395 | 1.824 | 0.00080397 | 5.6419e-06 | 0.91749 | 1.6735 | 0.00033961 | 0.093759 | 0.00032515 |  |  |  | 0.92722 |  |  | 0.00030148
396 | 1.9121 | 0.00084279 | 5.6037e-06 | 0.91128 | 1.7424 | 0.00035601 | 0.093358 | 0.00034091 |  |  |  | 0.91316 |  |  | 0.00031131
397 | 2.0098 | 0.00088585 | 5.567e-06 | 0.90531 | 1.8195 | 0.0003742 | 0.092894 | 0.00035841 |  |  |  | 0.89669 |  |  | 0.00032138
398 | 2.1079 | 0.00092909 | 5.5335e-06 | 0.89987 | 1.8968 | 0.00039246 | 0.092411 | 0.00037599 |  |  |  | 0.87913 |  |  | 0.00033054
399 | 2.192 | 0.0009662 | 5.5024e-06 | 0.8948 | 1.9614 | 0.00040814 | 0.091996 | 0.00039108 |  |  |  | 0.86206 |  |  | 0.00033713
400 | 2.3145 | 0.0010202 | 5.4705e-06 | 0.88961 | 2.059 | 0.00043094 | 0.091572 | 0.00041301 |  |  |  | 0.84425 |  |  | 0.00034869
401 | 2.4324 | 0.0010721 | 5.4361e-06 | 0.88402 | 2.1503 | 0.00045288 | 0.091145 | 0.00043413 |  |  |  | 0.82472 |  |  | 0.00035803
402 | 2.5824 | 0.0011383 | 5.4017e-06 | 0.87842 | 2.2684 | 0.00048082 | 0.090714 | 0.00046099 |  |  |  | 0.80452 |  |  | 0.00037088
403 | 3.9915 | 0.0017594 | 5.3687e-06 | 0.87306 | 3.4849 | 0.00074318 | 0.090274 | 0.00071269 |  |  |  | 0.78548 |  |  | 0.0005598
404 | 21.5045 | 0.0094787 | 5.3362e-06 | 0.86777 | 18.6609 | 0.0040039 | 0.089852 | 0.0038404 |  |  |  | 0.76903 |  |  | 0.0029534
405 | 20.7542 | 0.0091479 | 5.3014e-06 | 0.86212 | 17.8926 | 0.0038642 | 0.089445 | 0.0037071 |  |  |  | 0.75524 |  |  | 0.0027997
406 | 3.8006 | 0.0016752 | 5.2656e-06 | 0.85628 | 3.2544 | 0.00070763 | 0.088985 | 0.00067899 |  |  |  | 0.73764 |  |  | 0.00050085
407 | 4.2481 | 0.0018725 | 5.228e-06 | 0.85018 | 3.6116 | 0.00079095 | 0.088556 | 0.00075909 |  |  |  | 0.72019 |  |  | 0.00054669
408 | 4.6278 | 0.0020398 | 5.1934e-06 | 0.84456 | 3.9085 | 0.00086165 | 0.088228 | 0.00082708 |  |  |  | 0.70253 |  |  | 0.00058105
409 | 3.4364 | 0.0015147 | 5.1585e-06 | 0.83887 | 2.8827 | 0.00063982 | 0.087762 | 0.00061428 |  |  |  | 0.68475 |  |  | 0.00042063
410 | 3.4779 | 0.001533 | 5.1231e-06 | 0.83311 | 2.8975 | 0.00064756 | 0.087359 | 0.00062182 |  |  |  | 0.66773 |  |  | 0.00041521
411 | 3.6094 | 0.0015909 | 5.0904e-06 | 0.8278 | 2.9879 | 0.00067203 | 0.086948 | 0.00064545 |  |  |  | 0.6503 |  |  | 0.00041973
412 | 3.6729 | 0.0016189 | 5.0602e-06 | 0.82289 | 3.0224 | 0.00068385 | 0.086527 | 0.00065693 |  |  |  | 0.63261 |  |  | 0.00041558
413 | 3.7906 | 0.0016708 | 5.031e-06 | 0.81814 | 3.1013 | 0.00070578 | 0.08615 | 0.0006781 |  |  |  | 0.61418 |  |  | 0.00041648
414 | 3.9125 | 0.0017246 | 5.0016e-06 | 0.81337 | 3.1823 | 0.00072847 | 0.085677 | 0.00070006 |  |  |  | 0.59567 |  |  | 0.00041701
415 | 3.9986 | 0.0017625 | 4.971e-06 | 0.80839 | 3.2324 | 0.0007445 | 0.085378 | 0.00071556 |  |  |  | 0.57701 |  |  | 0.00041289
416 | 4.0376 | 0.0017797 | 4.9406e-06 | 0.80344 | 3.244 | 0.00075177 | 0.084934 | 0.00072269 |  |  |  | 0.55806 |  |  | 0.00040331
417 | 4.1364 | 0.0018232 | 4.9111e-06 | 0.79864 | 3.3035 | 0.00077015 | 0.084562 | 0.00074049 |  |  |  | 0.53764 |  |  | 0.00039812
418 | 4.2492 | 0.0018729 | 4.8839e-06 | 0.79422 | 3.3748 | 0.00079115 | 0.084164 | 0.00076083 |  |  |  | 0.51709 |  |  | 0.00039342
419 | 4.3473 | 0.0019162 | 4.8589e-06 | 0.79015 | 3.435 | 0.00080942 | 0.083772 | 0.00077854 |  |  |  | 0.49378 |  |  | 0.00038443
420 | 4.482 | 0.0019756 | 4.8365e-06 | 0.78651 | 3.5251 | 0.0008345 | 0.083397 | 0.0008028 |  |  |  | 0.47091 |  |  | 0.00037805
421 | 4.5786 | 0.0020181 | 4.8165e-06 | 0.78325 | 3.5862 | 0.00085248 | 0.08302 | 0.00082024 |  |  |  | 0.44854 |  |  | 0.00036791
422 | 4.6227 | 0.0020376 | 4.7958e-06 | 0.77989 | 3.6052 | 0.00086069 | 0.082681 | 0.00082827 |  |  |  | 0.4242 |  |  | 0.00035135
423 | 4.7509 | 0.0020941 | 4.7844e-06 | 0.77805 | 3.6964 | 0.00088457 | 0.082278 | 0.0008514 |  |  |  | 0.39819 |  |  | 0.00033902
424 | 4.8803 | 0.0021511 | 4.8096e-06 | 0.78213 | 3.817 | 0.00090865 | 0.081917 | 0.00087473 |  |  |  | 0.36925 |  |  | 0.000323
425 | 4.9813 | 0.0021957 | 4.8951e-06 | 0.79605 | 3.9654 | 0.00092747 | 0.081515 | 0.00089301 |  |  |  | 0.34567 |  |  | 0.00030869
426 | 5.0425 | 0.0022226 | 5.0166e-06 | 0.8158 | 4.1137 | 0.00093887 | 0.081147 | 0.00090414 |  |  |  | 0.32064 |  |  | 0.0002899
427 | 5.0486 | 0.0022253 | 5.1004e-06 | 0.82943 | 4.1874 | 0.00093999 | 0.080748 | 0.00090538 |  |  |  | 0.29521 |  |  | 0.00026728
428 | 5.2384 | 0.002309 | 5.0907e-06 | 0.82785 | 4.3367 | 0.00097534 | 0.080396 | 0.00093959 |  |  |  | 0.27136 |  |  | 0.00025497
429 | 5.3508 | 0.0023585 | 4.9762e-06 | 0.80923 | 4.33 | 0.00099626 | 0.079985 | 0.00095992 |  |  |  | 0.24729 |  |  | 0.00023738
430 | 5.5467 | 0.0024448 | 4.8484e-06 | 0.78844 | 4.3732 | 0.0010327 | 0.079663 | 0.00099521 |  |  |  | 0.22566 |  |  | 0.00022458
431 | 5.7518 | 0.0025353 | 4.7886e-06 | 0.77872 | 4.479 | 0.0010709 | 0.079352 | 0.0010322 |  |  |  | 0.20598 |  |  | 0.0002126
432 | 5.7688 | 0.0025427 | 4.7913e-06 | 0.77916 | 4.4948 | 0.0010741 | 0.078977 | 0.0010354 |  |  |  | 0.18685 |  |  | 0.00019346
433 | 6.04 | 0.0026623 | 4.846e-06 | 0.78806 | 4.7599 | 0.0011246 | 0.078573 | 0.0010843 |  |  |  | 0.16982 |  |  | 0.00018413
434 | 7.9831 | 0.0035188 | 4.9043e-06 | 0.79754 | 6.3668 | 0.0014864 | 0.078229 | 0.0014333 |  |  |  | 0.15458 |  |  | 0.00022157
435 | 43.0803 | 0.018989 | 4.9423e-06 | 0.80372 | 34.6245 | 0.0080211 | 0.077891 | 0.007736 |  |  |  | 0.14148 |  |  | 0.0010945
436 | 66.6687 | 0.029386 | 4.9698e-06 | 0.80819 | 53.8809 | 0.012413 | 0.077549 | 0.011974 |  |  |  | 0.12877 |  |  | 0.0015418
437 | 13.9619 | 0.0061541 | 4.9761e-06 | 0.80922 | 11.2982 | 0.0025996 | 0.077187 | 0.002508 |  |  |  | 0.11796 |  |  | 0.00029585
438 | 6.5376 | 0.0028816 | 4.964e-06 | 0.80725 | 5.2774 | 0.0012172 | 0.076871 | 0.0011745 |  |  |  | 0.10703 |  |  | 0.00012571
439 | 6.7072 | 0.0029564 | 4.9442e-06 | 0.80403 | 5.3928 | 0.0012488 | 0.076525 | 0.0012052 |  |  |  | 0.096564 |  |  | 0.00011638
440 | 6.8748 | 0.0030303 | 4.9226e-06 | 0.80052 | 5.5034 | 0.00128 | 0.076182 | 0.0012355 |  |  |  | 0.08659 |  |  | 0.00010698
441 | 6.9169 | 0.0030488 | 4.9003e-06 | 0.79689 | 5.512 | 0.0012879 | 0.075871 | 0.0012432 |  |  |  | 0.078607 |  |  | 9.7728e-05
442 | 6.9482 | 0.0030626 | 4.8799e-06 | 0.79357 | 5.5139 | 0.0012937 | 0.075587 | 0.001249 |  |  |  | 0.073701 |  |  | 9.2054e-05
443 | 7.0385 | 0.0031024 | 4.8635e-06 | 0.79091 | 5.5668 | 0.0013105 | 0.075215 | 0.0012655 |  |  |  | 0.069356 |  |  | 8.7768e-05
444 | 7.1262 | 0.0031411 | 4.8511e-06 | 0.78889 | 5.6218 | 0.0013268 | 0.074925 | 0.0012814 |  |  |  | 0.065135 |  |  | 8.3465e-05
445 | 7.1681 | 0.0031595 | 4.8401e-06 | 0.78709 | 5.6419 | 0.0013346 | 0.074543 | 0.0012892 |  |  |  | 0.06118 |  |  | 7.8872e-05
446 | 7.1732 | 0.0031618 | 4.8293e-06 | 0.78533 | 5.6334 | 0.0013356 | 0.074286 | 0.0012903 |  |  |  | 0.058316 |  |  | 7.5243e-05
447 | 7.1518 | 0.0031523 | 4.8162e-06 | 0.78322 | 5.6014 | 0.0013316 | 0.073921 | 0.0012866 |  |  |  | 0.055344 |  |  | 7.1206e-05
448 | 7.2107 | 0.0031783 | 4.8044e-06 | 0.78129 | 5.6337 | 0.0013426 | 0.07364 | 0.0012974 |  |  |  | 0.052288 |  |  | 6.7839e-05
449 | 7.2802 | 0.0032089 | 4.7937e-06 | 0.77956 | 5.6753 | 0.0013555 | 0.073311 | 0.0013101 |  |  |  | 0.050195 |  |  | 6.576e-05
450 | 7.3308 | 0.0032312 | 4.7853e-06 | 0.77818 | 5.7047 | 0.0013649 | 0.072975 | 0.0013194 |  |  |  | 0.047894 |  |  | 6.3191e-05
451 | 7.4026 | 0.0032629 | 4.7792e-06 | 0.77719 | 5.7532 | 0.0013783 | 0.072726 | 0.0013325 |  |  |  | 0.04506 |  |  | 6.0042e-05
452 | 7.472 | 0.0032935 | 4.7757e-06 | 0.77662 | 5.8029 | 0.0013912 | 0.072458 | 0.0013451 |  |  |  | 0.042275 |  |  | 5.6866e-05
453 | 7.5153 | 0.0033126 | 4.7721e-06 | 0.77604 | 5.8321 | 0.0013993 | 0.072119 | 0.0013531 |  |  |  | 0.040219 |  |  | 5.4422e-05
454 | 7.5635 | 0.0033338 | 4.7687e-06 | 0.77548 | 5.8654 | 0.0014082 | 0.071744 | 0.0013621 |  |  |  | 0.039258 |  |  | 5.3472e-05
455 | 7.6369 | 0.0033662 | 4.7654e-06 | 0.77495 | 5.9182 | 0.0014219 | 0.071493 | 0.0013754 |  |  |  | 0.037635 |  |  | 5.1765e-05
456 | 7.692 | 0.0033904 | 4.763e-06 | 0.77456 | 5.9579 | 0.0014322 | 0.071182 | 0.0013856 |  |  |  | 0.035934 |  |  | 4.9788e-05
457 | 7.7987 | 0.0034375 | 4.7621e-06 | 0.77441 | 6.0394 | 0.001452 | 0.070886 | 0.001405 |  |  |  | 0.034232 |  |  | 4.8095e-05
458 | 7.8379 | 0.0034548 | 4.7621e-06 | 0.77441 | 6.0698 | 0.0014593 | 0.070561 | 0.0014123 |  |  |  | 0.033874 |  |  | 4.7838e-05
459 | 7.878 | 0.0034724 | 4.7647e-06 | 0.77484 | 6.1042 | 0.0014668 | 0.070321 | 0.0014196 |  |  |  | 0.033082 |  |  | 4.6963e-05
460 | 7.9355 | 0.0034978 | 4.7692e-06 | 0.77557 | 6.1545 | 0.0014775 | 0.07004 | 0.0014302 |  |  |  | 0.032155 |  |  | 4.5987e-05
461 | 8.0057 | 0.0035287 | 4.774e-06 | 0.77635 | 6.2152 | 0.0014906 | 0.069682 | 0.0014431 |  |  |  | 0.031227 |  |  | 4.5063e-05
462 | 7.915 | 0.0034888 | 4.782e-06 | 0.77765 | 6.1551 | 0.0014737 | 0.069395 | 0.0014269 |  |  |  | 0.030619 |  |  | 4.3691e-05
463 | 8.0619 | 0.0035535 | 4.7908e-06 | 0.77908 | 6.2809 | 0.001501 | 0.069167 | 0.0014535 |  |  |  | 0.029915 |  |  | 4.3483e-05
464 | 8.1319 | 0.0035843 | 4.8009e-06 | 0.78072 | 6.3488 | 0.0015141 | 0.068877 | 0.0014664 |  |  |  | 0.029167 |  |  | 4.2769e-05
465 | 8.192 | 0.0036108 | 4.8107e-06 | 0.78231 | 6.4087 | 0.0015253 | 0.068561 | 0.0014774 |  |  |  | 0.028418 |  |  | 4.1985e-05
466 | 8.2498 | 0.0036363 | 4.8175e-06 | 0.78342 | 6.4631 | 0.001536 | 0.068321 | 0.001488 |  |  |  | 0.027669 |  |  | 4.1172e-05
467 | 8.2908 | 0.0036544 | 4.8227e-06 | 0.78427 | 6.5022 | 0.0015437 | 0.068022 | 0.0014956 |  |  |  | 0.026823 |  |  | 4.0116e-05
468 | 8.3207 | 0.0036676 | 4.8287e-06 | 0.78525 | 6.5338 | 0.0015492 | 0.06776 | 0.0015012 |  |  |  | 0.025897 |  |  | 3.8876e-05
469 | 8.3688 | 0.0036888 | 4.8371e-06 | 0.78661 | 6.583 | 0.0015582 | 0.067443 | 0.0015101 |  |  |  | 0.025091 |  |  | 3.7889e-05
470 | 8.3926 | 0.0036992 | 4.8483e-06 | 0.78844 | 6.617 | 0.0015626 | 0.067181 | 0.0015145 |  |  |  | 0.024294 |  |  | 3.6794e-05
471 | 8.3942 | 0.0036999 | 4.8612e-06 | 0.79052 | 6.6358 | 0.0015629 | 0.066955 | 0.001515 |  |  |  | 0.023496 |  |  | 3.5597e-05
472 | 8.4395 | 0.0037199 | 4.8735e-06 | 0.79253 | 6.6885 | 0.0015713 | 0.066687 | 0.0015234 |  |  |  | 0.022634 |  |  | 3.448e-05
473 | 8.4984 | 0.0037459 | 4.8858e-06 | 0.79453 | 6.7523 | 0.0015823 | 0.066395 | 0.0015342 |  |  |  | 0.021683 |  |  | 3.3267e-05
474 | 8.517 | 0.0037541 | 4.8985e-06 | 0.7966 | 6.7846 | 0.0015858 | 0.066132 | 0.0015377 |  |  |  | 0.020793 |  |  | 3.1975e-05
475 | 8.5188 | 0.0037549 | 4.9121e-06 | 0.79881 | 6.8049 | 0.0015861 | 0.065874 | 0.0015383 |  |  |  | 0.019908 |  |  | 3.0624e-05
476 | 8.5107 | 0.0037513 | 4.9264e-06 | 0.80112 | 6.8181 | 0.0015846 | 0.065586 | 0.001537 |  |  |  | 0.019023 |  |  | 2.9239e-05
477 | 8.534 | 0.0037616 | 4.9413e-06 | 0.80356 | 6.8576 | 0.0015889 | 0.065347 | 0.0015414 |  |  |  | 0.018085 |  |  | 2.7876e-05
478 | 8.5621 | 0.003774 | 4.9561e-06 | 0.80597 | 6.9007 | 0.0015942 | 0.065074 | 0.0015466 |  |  |  | 0.017071 |  |  | 2.6404e-05
479 | 8.5555 | 0.0037711 | 4.9712e-06 | 0.80842 | 6.9165 | 0.001593 | 0.064852 | 0.0015456 |  |  |  | 0.016134 |  |  | 2.4936e-05
480 | 8.5524 | 0.0037697 | 4.9873e-06 | 0.81103 | 6.9363 | 0.0015924 | 0.064547 | 0.0015453 |  |  |  | 0.015204 |  |  | 2.3494e-05
481 | 8.5884 | 0.0037856 | 5.0025e-06 | 0.81351 | 6.9868 | 0.0015991 | 0.064315 | 0.0015519 |  |  |  | 0.014273 |  |  | 2.2152e-05
482 | 8.624 | 0.0038012 | 5.018e-06 | 0.81603 | 7.0374 | 0.0016057 | 0.064094 | 0.0015585 |  |  |  | 0.01245 |  |  | 1.9404e-05
483 | 8.6481 | 0.0038119 | 5.0341e-06 | 0.81864 | 7.0797 | 0.0016102 | 0.063809 | 0.0015631 |  |  |  | 0.01094 |  |  | 1.71e-05
484 | 8.6784 | 0.0038252 | 5.0512e-06 | 0.82142 | 7.1286 | 0.0016158 | 0.063579 | 0.0015687 |  |  |  | 0.00965 |  |  | 1.5138e-05
485 | 8.738 | 0.0038515 | 5.068e-06 | 0.82415 | 7.2014 | 0.0016269 | 0.063332 | 0.0015797 |  |  |  | 0.0083599 |  |  | 1.3206e-05
486 | 8.8048 | 0.003881 | 5.0847e-06 | 0.82687 | 7.2804 | 0.0016394 | 0.063119 | 0.0015919 |  |  |  | 0.0072087 |  |  | 1.1476e-05
487 | 8.8056 | 0.0038813 | 5.0999e-06 | 0.82935 | 7.3029 | 0.0016395 | 0.062885 | 0.0015923 |  |  |  | 0.0059105 |  |  | 9.411e-06
488 | 8.7688 | 0.0038651 | 5.1144e-06 | 0.8317 | 7.2931 | 0.0016327 | 0.062589 | 0.0015858 |  |  |  | 0.0045334 |  |  | 7.1892e-06
489 | 8.7484 | 0.0038561 | 5.1301e-06 | 0.83426 | 7.2984 | 0.0016289 | 0.06241 | 0.0015822 |  |  |  | 0.0031563 |  |  | 4.9941e-06
490 | 8.72 | 0.0038436 | 5.1456e-06 | 0.83678 | 7.2968 | 0.0016236 | 0.06214 | 0.0015773 |  |  |  | 0.0033199 |  |  | 5.2365e-06
491 | 8.8329 | 0.0038933 | 5.1597e-06 | 0.83908 | 7.4115 | 0.0016446 | 0.061972 | 0.0015979 |  |  |  | 0.0033022 |  |  | 5.2765e-06
492 | 8.8306 | 0.0038923 | 5.1736e-06 | 0.84134 | 7.4295 | 0.0016442 | 0.061767 | 0.0015976 |  |  |  | 0.0027881 |  |  | 4.4542e-06
493 | 8.5973 | 0.0037895 | 5.1881e-06 | 0.84369 | 7.2534 | 0.0016007 | 0.061562 | 0.0015555 |  |  |  | 0.002274 |  |  | 3.5372e-06
494 | 8.549 | 0.0037682 | 5.2033e-06 | 0.84616 | 7.2338 | 0.0015917 | 0.061305 | 0.001547 |  |  |  | 0.0018737 |  |  | 2.8985e-06
495 | 8.4653 | 0.0037313 | 5.2182e-06 | 0.84858 | 7.1834 | 0.0015761 | 0.061042 | 0.001532 |  |  |  | 0.0022014 |  |  | 3.3726e-06
496 | 8.4525 | 0.0037256 | 5.232e-06 | 0.85082 | 7.1915 | 0.0015738 | 0.06082 | 0.0015299 |  |  |  | 0.0022545 |  |  | 3.4491e-06
497 | 8.4161 | 0.0037096 | 5.2452e-06 | 0.85297 | 7.1787 | 0.001567 | 0.060621 | 0.0015234 |  |  |  | 0.0022434 |  |  | 3.4177e-06
498 | 8.3569 | 0.0036835 | 5.2589e-06 | 0.8552 | 7.1468 | 0.001556 | 0.060364 | 0.0015129 |  |  |  | 0.0022324 |  |  | 3.3774e-06
499 | 8.2894 | 0.0036538 | 5.2715e-06 | 0.85725 | 7.106 | 0.0015434 | 0.060144 | 0.0015008 |  |  |  | 0.0022213 |  |  | 3.3338e-06
500 | 8.2458 | 0.0036345 | 5.2843e-06 | 0.85934 | 7.0859 | 0.0015353 | 0.059853 | 0.0014931 |  |  |  | 0.0022103 |  |  | 3.3002e-06
501 | 8.1871 | 0.0036087 |  | 0 | 0 | 0.0015243 | [8 empty] | 0
502 | 8.1326 | 0.0035847 |  | 0 | 0 | 0.0015142 | [8 empty] | 0
503 | 8.0927 | 0.0035671 |  | 0 | 0 | 0.0015068 | [8 empty] | 0
504 | 8.0574 | 0.0035515 |  | 0 | 0 | 0.0015002 | [8 empty] | 0
505 | 8.0417 | 0.0035446 |  | 0 | 0 | 0.0014973 | [8 empty] | 0
506 | 7.9901 | 0.0035218 |  | 0 | 0 | 0.0014877 | [8 empty] | 0
507 | 7.9523 | 0.0035052 |  | 0 | 0 | 0.0014806 | [8 empty] | 0
508 | 7.8873 | 0.0034765 |  | 0 | 0 | 0.0014685 | [8 empty] | 0
509 | 7.8808 | 0.0034737 |  | 0 | 0 | 0.0014673 | [8 empty] | 0
510 | 7.8551 | 0.0034623 |  | 0 | 0 | 0.0014625 | [8 empty] | 0
511 | 7.7977 | 0.003437 |  | 0 | 0 | 0.0014518 | [8 empty] | 0
512 | 7.784 | 0.003431 |  | 0 | 0 | 0.0014493 | [8 empty] | 0
513 | 7.7392 | 0.0034113 |  | 0 | 0 | 0.001441 | [8 empty] | 0
514 | 7.7049 | 0.0033961 |  | 0 | 0 | 0.0014346 | [8 empty] | 0
515 | 7.6898 | 0.0033895 |  | 0 | 0 | 0.0014318 | [8 empty] | 0
516 | 7.6308 | 0.0033635 |  | 0 | 0 | 0.0014208 | [8 empty] | 0
517 | 7.6043 | 0.0033518 |  | 0 | 0 | 0.0014158 | [8 empty] | 0
518 | 7.5897 | 0.0033454 |  | 0 | 0 | 0.0014131 | [8 empty] | 0
519 | 7.5851 | 0.0033433 |  | 0 | 0 | 0.0014123 | [8 empty] | 0
520 | 7.552 | 0.0033287 |  | 0 | 0 | 0.0014061 | [8 empty] | 0
521 | 7.572 | 0.0033376 |  | 0 | 0 | 0.0014098 | [8 empty] | 0
522 | 7.551 | 0.0033283 |  | 0 | 0 | 0.0014059 | [8 empty] | 0
523 | 7.5131 | 0.0033116 |  | 0 | 0 | 0.0013989 | [8 empty] | 0
524 | 7.5492 | 0.0033275 |  | 0 | 0 | 0.0014056 | [8 empty] | 0
525 | 7.5261 | 0.0033173 |  | 0 | 0 | 0.0014013 | [8 empty] | 0
526 | 7.5291 | 0.0033187 |  | 0 | 0 | 0.0014018 | [8 empty] | 0
527 | 7.5244 | 0.0033166 |  | 0 | 0 | 0.001401 | [8 empty] | 0
528 | 7.5574 | 0.0033311 |  | 0 | 0 | 0.0014071 | [8 empty] | 0
529 | 7.5736 | 0.0033383 |  | 0 | 0 | 0.0014101 | [8 empty] | 0
530 | 7.5899 | 0.0033454 |  | 0 | 0 | 0.0014132 | [8 empty] | 0
531 | 7.5791 | 0.0033407 |  | 0 | 0 | 0.0014112 | [8 empty] | 0
532 | 7.6157 | 0.0033568 |  | 0 | 0 | 0.001418 | [8 empty] | 0
533 | 7.693 | 0.0033909 |  | 0 | 0 | 0.0014324 | [8 empty] | 0
534 | 7.7875 | 0.0034326 |  | 0 | 0 | 0.00145 | [8 empty] | 0
535 | 7.8003 | 0.0034382 |  | 0 | 0 | 0.0014523 | [8 empty] | 0
536 | 7.8218 | 0.0034476 |  | 0 | 0 | 0.0014563 | [8 empty] | 0
537 | 7.9424 | 0.0035008 |  | 0 | 0 | 0.0014788 | [8 empty] | 0
538 | 8.0716 | 0.0035578 |  | 0 | 0 | 0.0015028 | [8 empty] | 0
539 | 8.227 | 0.0036262 |  | 0 | 0 | 0.0015318 | [8 empty] | 0
540 | 8.486 | 0.0037404 |  | 0 | 0 | 0.00158 | [8 empty] | 0
541 | 8.9436 | 0.0039421 |  | 0 | 0 | 0.0016652 | [8 empty] | 0
542 | 9.5313 | 0.0042012 |  | 0 | 0 | 0.0017746 | [8 empty] | 0
543 | 9.6644 | 0.0042598 |  | 0 | 0 | 0.0017994 | [8 empty] | 0
544 | 9.8092 | 0.0043236 |  | 0 | 0 | 0.0018264 | [8 empty] | 0
545 | 17.5769 | 0.0077475 |  | 0 | 0 | 0.0032726 | [8 empty] | 0
546 | 44.8352 | 0.019762 |  | 0 | 0 | 0.0083478 | [8 empty] | 0
547 | 24.742 | 0.010906 |  | 0 | 0 | 0.0046067 | [8 empty] | 0
548 | 9.5513 | 0.00421 |  | 0 | 0 | 0.0017783 | [8 empty] | 0
549 | 9.4723 | 0.0041752 |  | 0 | 0 | 0.0017636 | [8 empty] | 0
550 | 9.5156 | 0.0041942 |  | 0 | 0 | 0.0017717 | [8 empty] | 0
551 | 9.5296 | 0.0042004 |  | 0 | 0 | 0.0017743 | [8 empty] | 0
552 | 9.5544 | 0.0042113 |  | 0 | 0 | 0.0017789 | [8 empty] | 0
553 | 9.6217 | 0.004241 |  | 0 | 0 | 0.0017915 | [8 empty] | 0
554 | 9.6542 | 0.0042553 |  | 0 | 0 | 0.0017975 | [8 empty] | 0
555 | 9.6653 | 0.0042602 |  | 0 | 0 | 0.0017996 | [8 empty] | 0
556 | 9.7056 | 0.004278 |  | 0 | 0 | 0.0018071 | [8 empty] | 0
557 | 9.6971 | 0.0042743 |  | 0 | 0 | 0.0018055 | [8 empty] | 0
558 | 9.7418 | 0.0042939 |  | 0 | 0 | 0.0018138 | [8 empty] | 0
559 | 9.8529 | 0.0043429 |  | 0 | 0 | 0.0018345 | [8 empty] | 0
560 | 9.929 | 0.0043765 |  | 0 | 0 | 0.0018487 | [8 empty] | 0
561 | 10.0173 | 0.0044154 |  | 0 | 0 | 0.0018651 | [8 empty] | 0
562 | 10.061 | 0.0044347 |  | 0 | 0 | 0.0018733 | [8 empty] | 0
563 | 10.109 | 0.0044558 |  | 0 | 0 | 0.0018822 | [8 empty] | 0
564 | 10.1627 | 0.0044795 |  | 0 | 0 | 0.0018922 | [8 empty] | 0
565 | 10.2495 | 0.0045177 |  | 0 | 0 | 0.0019084 | [8 empty] | 0
566 | 10.3214 | 0.0045494 |  | 0 | 0 | 0.0019217 | [8 empty] | 0
567 | 10.3661 | 0.0045691 |  | 0 | 0 | 0.0019301 | [8 empty] | 0
568 | 10.4368 | 0.0046003 |  | 0 | 0 | 0.0019432 | [8 empty] | 0
569 | 10.4735 | 0.0046165 |  | 0 | 0 | 0.0019501 | [8 empty] | 0
570 | 10.4996 | 0.004628 |  | 0 | 0 | 0.0019549 | [8 empty] | 0
571 | 10.5451 | 0.004648 |  | 0 | 0 | 0.0019634 | [8 empty] | 0
572 | 10.59 | 0.0046678 |  | 0 | 0 | 0.0019717 | [8 empty] | 0
573 | 10.5706 | 0.0046593 |  | 0 | 0 | 0.0019681 | [8 empty] | 0
574 | 10.6134 | 0.0046781 |  | 0 | 0 | 0.0019761 | [8 empty] | 0
575 | 10.6314 | 0.0046861 |  | 0 | 0 | 0.0019795 | [8 empty] | 0
576 | 12.2979 | 0.0054206 |  | 0 | 0 | 0.0022897 | [8 empty] | 0
577 | 16.6147 | 0.0073233 |  | 0 | 0 | 0.0030935 | [8 empty] | 0
578 | 13.6231 | 0.0060048 |  | 0 | 0 | 0.0025365 | [8 empty] | 0
579 | 16.0458 | 0.0070726 |  | 0 | 0 | 0.0029876 | [8 empty] | 0
580 | 13.0105 | 0.0057347 |  | 0 | 0 | 0.0024224 | [8 empty] | 0
581 | 10.7676 | 0.0047461 |  | 0 | 0 | 0.0020048 | [8 empty] | 0
582 | 10.673 | 0.0047044 |  | 0 | 0 | 0.0019872 | [8 empty] | 0
583 | 10.7035 | 0.0047178 |  | 0 | 0 | 0.0019929 | [8 empty] | 0
584 | 10.663 | 0.0047 |  | 0 | 0 | 0.0019853 | [8 empty] | 0
585 | 10.554 | 0.004652 |  | 0 | 0 | 0.001965 | [8 empty] | 0
586 | 10.4805 | 0.0046195 |  | 0 | 0 | 0.0019514 | [8 empty] | 0
587 | 10.5955 | 0.0046702 |  | 0 | 0 | 0.0019728 | [8 empty] | 0
588 | 10.4215 | 0.0045936 |  | 0 | 0 | 0.0019404 | [8 empty] | 0
589 | 10.0762 | 0.0044413 |  | 0 | 0 | 0.0018761 | [8 empty] | 0
590 | 10.0028 | 0.004409 |  | 0 | 0 | 0.0018624 | [8 empty] | 0
591 | 9.9235 | 0.0043741 |  | 0 | 0 | 0.0018477 | [8 empty] | 0
592 | 9.9124 | 0.0043692 |  | 0 | 0 | 0.0018456 | [8 empty] | 0
593 | 9.9037 | 0.0043653 |  | 0 | 0 | 0.001844 | [8 empty] | 0
594 | 9.7053 | 0.0042779 |  | 0 | 0 | 0.001807 | [8 empty] | 0
595 | 9.4137 | 0.0041493 |  | 0 | 0 | 0.0017527 | [8 empty] | 0
596 | 9.2361 | 0.0040711 |  | 0 | 0 | 0.0017197 | [8 empty] | 0
597 | 9.0827 | 0.0040034 |  | 0 | 0 | 0.0016911 | [8 empty] | 0
598 | 9.0092 | 0.003971 |  | 0 | 0 | 0.0016774 | [8 empty] | 0
599 | 9.0101 | 0.0039715 |  | 0 | 0 | 0.0016776 | [8 empty] | 0
600 | 8.864 | 0.003907 |  | 0 | 0 | 0.0016504 | [8 empty] | 0
601 | 8.6277 | 0.0038029 |  | 0 | 0 | 0.0016064 | [8 empty] | 0
602 | 8.3896 | 0.0036979 |  | 0 | 0 | 0.0015621 | [8 empty] | 0
603 | 8.2823 | 0.0036506 |  | 0 | 0 | 0.0015421 | [8 empty] | 0
604 | 8.1465 | 0.0035908 |  | 0 | 0 | 0.0015168 | [8 empty] | 0
605 | 8.0212 | 0.0035355 |  | 0 | 0 | 0.0014935 | [8 empty] | 0
606 | 7.8844 | 0.0034752 |  | 0 | 0 | 0.001468 | [8 empty] | 0
607 | 7.7718 | 0.0034256 |  | 0 | 0 | 0.001447 | [8 empty] | 0
608 | 7.7428 | 0.0034128 |  | 0 | 0 | 0.0014416 | [8 empty] | 0
609 | 7.7846 | 0.0034313 |  | 0 | 0 | 0.0014494 | [8 empty] | 0
610 | 8.4893 | 0.0037419 |  | 0 | 0 | 0.0015806 | [8 empty] | 0
611 | 11.0803 | 0.0048839 |  | 0 | 0 | 0.002063 | [8 empty] | 0
612 | 10.6007 | 0.0046725 |  | 0 | 0 | 0.0019737 | [8 empty] | 0
613 | 8.4163 | 0.0037097 |  | 0 | 0 | 0.001567 | [8 empty] | 0
614 | 7.8753 | 0.0034712 |  | 0 | 0 | 0.0014663 | [8 empty] | 0
615 | 7.4822 | 0.003298 |  | 0 | 0 | 0.0013931 | [8 empty] | 0
616 | 6.9181 | 0.0030493 |  | 0 | 0 | 0.0012881 | [8 empty] | 0
617 | 6.5824 | 0.0029013 |  | 0 | 0 | 0.0012256 | [8 empty] | 0
618 | 6.2966 | 0.0027754 |  | 0 | 0 | 0.0011724 | [8 empty] | 0
619 | 6.14 | 0.0027064 |  | 0 | 0 | 0.0011432 | [8 empty] | 0
620 | 6.0917 | 0.0026851 |  | 0 | 0 | 0.0011342 | [8 empty] | 0
621 | 5.9409 | 0.0026186 |  | 0 | 0 | 0.0011061 | [8 empty] | 0
622 | 5.7928 | 0.0025533 |  | 0 | 0 | 0.0010786 | [8 empty] | 0
623 | 5.6967 | 0.002511 |  | 0 | 0 | 0.0010607 | [8 empty] | 0
624 | 5.5686 | 0.0024545 |  | 0 | 0 | 0.0010368 | [8 empty] | 0
625 | 5.3709 | 0.0023673 |  | 0 | 0 | 0.001 | [8 empty] | 0
626 | 5.2691 | 0.0023225 |  | 0 | 0 | 0.00098105 | [8 empty] | 0
627 | 5.1626 | 0.0022755 |  | 0 | 0 | 0.00096122 | [8 empty] | 0
628 | 5.0066 | 0.0022068 |  | 0 | 0 | 0.00093217 | [8 empty] | 0
629 | 4.8922 | 0.0021564 |  | 0 | 0 | 0.00091087 | [8 empty] | 0
630 | 4.9148 | 0.0021663 |  | 0 | 0 | 0.00091509 | [8 empty] | 0
631 | 4.9921 | 0.0022004 |  | 0 | 0 | 0.00092948 | [8 empty] | 0
632 | 4.6227 | 0.0020376 |  | 0 | 0 | 0.0008607 | [8 empty] | 0
633 | 4.2818 | 0.0018873 |  | 0 | 0 | 0.00079722 | [8 empty] | 0
634 | 4.1469 | 0.0018278 |  | 0 | 0 | 0.00077211 | [8 empty] | 0
635 | 4.0077 | 0.0017665 |  | 0 | 0 | 0.00074619 | [8 empty] | 0
636 | 3.898 | 0.0017182 |  | 0 | 0 | 0.00072577 | [8 empty] | 0
637 | 3.8168 | 0.0016823 |  | 0 | 0 | 0.00071065 | [8 empty] | 0
638 | 3.7135 | 0.0016368 |  | 0 | 0 | 0.00069141 | [8 empty] | 0
639 | 3.6151 | 0.0015935 |  | 0 | 0 | 0.0006731 | [8 empty] | 0
640 | 3.4905 | 0.0015385 |  | 0 | 0 | 0.0006499 | [8 empty] | 0
641 | 3.3909 | 0.0014946 |  | 0 | 0 | 0.00063135 | [8 empty] | 0
642 | 3.3359 | 0.0014704 |  | 0 | 0 | 0.00062112 | [8 empty] | 0
643 | 3.2558 | 0.0014351 |  | 0 | 0 | 0.0006062 | [8 empty] | 0
644 | 3.1642 | 0.0013947 |  | 0 | 0 | 0.00058914 | [8 empty] | 0
645 | 3.0727 | 0.0013544 |  | 0 | 0 | 0.00057211 | [8 empty] | 0
646 | 3.0105 | 0.0013269 |  | 0 | 0 | 0.00056051 | [8 empty] | 0
647 | 2.9015 | 0.0012789 |  | 0 | 0 | 0.00054023 | [8 empty] | 0
648 | 2.8413 | 0.0012524 |  | 0 | 0 | 0.00052901 | [8 empty] | 0
649 | 2.8179 | 0.0012421 |  | 0 | 0 | 0.00052467 | [8 empty] | 0
650 | 2.8101 | 0.0012386 |  | 0 | 0 | 0.00052321 | [8 empty] | 0
651 | 2.7463 | 0.0012105 |  | 0 | 0 | 0.00051133 | [8 empty] | 0
652 | 2.6332 | 0.0011606 |  | 0 | 0 | 0.00049027 | [8 empty] | 0
653 | 2.5648 | 0.0011305 |  | 0 | 0 | 0.00047753 | [8 empty] | 0
654 | 2.4517 | 0.0010806 |  | 0 | 0 | 0.00045647 | [8 empty] | 0
655 | 2.3422 | 0.0010324 |  | 0 | 0 | 0.00043609 | [8 empty] | 0
656 | 2.2742 | 0.0010024 |  | 0 | 0 | 0.00042343 | [8 empty] | 0
657 | 2.2623 | 0.00099718 |  | 0 | 0 | 0.00042122 | [8 empty] | 0
658 | 2.2099 | 0.00097405 |  | 0 | 0 | 0.00041145 | [8 empty] | 0
659 | 2.152 | 0.00094855 |  | 0 | 0 | 0.00040068 | [8 empty] | 0
660 | 2.0719 | 0.00091325 |  | 0 | 0 | 0.00038577 | [8 empty] | 0
661 | 1.999 | 0.00088113 |  | 0 | 0 | 0.0003722 | [8 empty] | 0
662 | 2.0264 | 0.00089318 |  | 0 | 0 | 0.00037729 | [8 empty] | 0
663 | 1.9839 | 0.00087446 |  | 0 | 0 | 0.00036939 | [8 empty] | 0
664 | 1.8956 | 0.00083555 |  | 0 | 0 | 0.00035295 | [8 empty] | 0
665 | 1.833 | 0.00080793 |  | 0 | 0 | 0.00034128 | [8 empty] | 0
666 | 1.7587 | 0.00077518 |  | 0 | 0 | 0.00032745 | [8 empty] | 0
667 | 1.7507 | 0.00077167 |  | 0 | 0 | 0.00032596 | [8 empty] | 0
668 | 1.6966 | 0.00074783 |  | 0 | 0 | 0.00031589 | [8 empty] | 0
669 | 1.624 | 0.00071583 |  | 0 | 0 | 0.00030238 | [8 empty] | 0
670 | 1.591 | 0.00070129 |  | 0 | 0 | 0.00029624 | [8 empty] | 0
671 | 1.5827 | 0.00069762 |  | 0 | 0 | 0.00029468 | [8 empty] | 0
672 | 1.5287 | 0.0006738 |  | 0 | 0 | 0.00028462 | [8 empty] | 0
673 | 1.481 | 0.0006528 |  | 0 | 0 | 0.00027575 | [8 empty] | 0
674 | 1.4462 | 0.00063747 |  | 0 | 0 | 0.00026928 | [8 empty] | 0
675 | 1.4308 | 0.00063067 |  | 0 | 0 | 0.00026641 | [8 empty] | 0
676 | 1.4117 | 0.00062225 |  | 0 | 0 | 0.00026285 | [8 empty] | 0
677 | 1.357 | 0.00059812 |  | 0 | 0 | 0.00025265 | [8 empty] | 0
678 | 1.2917 | 0.00056934 |  | 0 | 0 | 0.0002405 | [8 empty] | 0
679 | 1.2805 | 0.00056443 |  | 0 | 0 | 0.00023842 | [8 empty] | 0
680 | 1.2687 | 0.0005592 |  | 0 | 0 | 0.00023621 | [8 empty] | 0
681 | 1.2278 | 0.00054118 |  | 0 | 0 | 0.0002286 | [8 empty] | 0
682 | 1.1978 | 0.00052798 |  | 0 | 0 | 0.00022303 | [8 empty] | 0
683 | 1.1558 | 0.00050943 |  | 0 | 0 | 0.00021519 | [8 empty] | 0
684 | 1.1139 | 0.00049098 |  | 0 | 0 | 0.00020739 | [8 empty] | 0
685 | 1.0882 | 0.00047965 |  | 0 | 0 | 0.00020261 | [8 empty] | 0
686 | 1.0607 | 0.00046752 |  | 0 | 0 | 0.00019749 | [8 empty] | 0
687 | 1.1287 | 0.00049749 |  | 0 | 0 | 0.00021015 | [8 empty] | 0
688 | 1.1402 | 0.00050256 |  | 0 | 0 | 0.00021229 | [8 empty] | 0
689 | 1.0723 | 0.00047266 |  | 0 | 0 | 0.00019966 | [8 empty] | 0
690 | 1.112 | 0.00049013 |  | 0 | 0 | 0.00020704 | [8 empty] | 0
691 | 1.1677 | 0.00051468 |  | 0 | 0 | 0.00021741 | [8 empty] | 0
692 | 0.9383 | 0.00041358 |  | 0 | 0 | 0.0001747 | [8 empty] | 0
693 | 0.96348 | 0.00042468 |  | 0 | 0 | 0.00017939 | [8 empty] | 0
694 | 0.95738 | 0.00042199 |  | 0 | 0 | 0.00017825 | [8 empty] | 0
695 | 0.88894 | 0.00039182 |  | 0 | 0 | 0.00016551 | [8 empty] | 0
696 | 0.94608 | 0.00041701 |  | 0 | 0 | 0.00017615 | [8 empty] | 0
697 | 0.94087 | 0.00041471 |  | 0 | 0 | 0.00017518 | [8 empty] | 0
698 | 0.84114 | 0.00037075 |  | 0 | 0 | 0.00015661 | [8 empty] | 0
699 | 0.80407 | 0.00035441 |  | 0 | 0 | 0.00014971 | [8 empty] | 0
700 | 0.79349 | 0.00034975 |  | 0 | 0 | 0.00014774 | [8 empty] | 0
701 | 0.77626 | 0.00034216 |  | 0 | 0 | 0.00014453 | [8 empty] | 0
702 | 0.78581 | 0.00034637 |  | 0 | 0 | 0.00014631 | [8 empty] | 0
703 | 0.76828 | 0.00033864 |  | 0 | 0 | 0.00014305 | [8 empty] | 0
704 | 0.76246 | 0.00033607 |  | 0 | 0 | 0.00014196 | [8 empty] | 0
705 | 0.79872 | 0.00035206 |  | 0 | 0 | 0.00014871 | [8 empty] | 0
706 | 0.97295 | 0.00042885 |  | 0 | 0 | 0.00018115 | [8 empty] | 0
707 | 1.0761 | 0.00047431 |  | 0 | 0 | 0.00020035 | [8 empty] | 0
708 | 0.94913 | 0.00041835 |  | 0 | 0 | 0.00017672 | [8 empty] | 0
709 | 0.99064 | 0.00043665 |  | 0 | 0 | 0.00018445 | [8 empty] | 0
710 | 0.82788 | 0.00036491 |  | 0 | 0 | 0.00015414 | [8 empty] | 0
711 | 1.0128 | 0.00044642 |  | 0 | 0 | 0.00018857 | [8 empty] | 0
712 | 1.1585 | 0.00051064 |  | 0 | 0 | 0.0002157 | [8 empty] | 0
713 | 0.80351 | 0.00035417 |  | 0 | 0 | 0.00014961 | [8 empty] | 0
714 | 0.67476 | 0.00029742 |  | 0 | 0 | 0.00012563 | [8 empty] | 0
715 | 0.60059 | 0.00026473 |  | 0 | 0 | 0.00011182 | [8 empty] | 0
716 | 0.58488 | 0.0002578 |  | 0 | 0 | 0.0001089 | [8 empty] | 0
717 | 0.59943 | 0.00026422 |  | 0 | 0 | 0.00011161 | [8 empty] | 0
718 | 0.53457 | 0.00023562 |  | 0 | 0 | 9.9531e-05 | [8 empty] | 0
719 | 0.54869 | 0.00024185 |  | 0 | 0 | 0.00010216 | [8 empty] | 0
720 | 0.58376 | 0.00025731 |  | 0 | 0 | 0.00010869 | [8 empty] | 0
721 | 0.5597 | 0.0002467 |  | 0 | 0 | 0.00010421 | [8 empty] | 0
722 | 0.53402 | 0.00023539 |  | 0 | 0 | 9.943e-05 | [8 empty] | 0
723 | 0.46113 | 0.00020325 |  | 0 | 0 | 8.5857e-05 | [8 empty] | 0
724 | 0.48533 | 0.00021392 |  | 0 | 0 | 9.0363e-05 | [8 empty] | 0
725 | 0.49112 | 0.00021647 |  | 0 | 0 | 9.1441e-05 | [8 empty] | 0
726 | 0.47754 | 0.00021049 |  | 0 | 0 | 8.8913e-05 | [8 empty] | 0
727 | 0.47413 | 0.00020899 |  | 0 | 0 | 8.8278e-05 | [8 empty] | 0
728 | 0.48118 | 0.00021209 |  | 0 | 0 | 8.959e-05 | [8 empty] | 0
729 | 0.48457 | 0.00021359 |  | 0 | 0 | 9.0222e-05 | [8 empty] | 0
730 | 0.41849 | 0.00018446 |  | 0 | 0 | 7.7918e-05 | [8 empty] | 0
731 | 0.42787 | 0.00018859 |  | 0 | 0 | 7.9665e-05 | [8 empty] | 0
732 | 0.44421 | 0.0001958 |  | 0 | 0 | 8.2707e-05 | [8 empty] | 0
733 | 0.43748 | 0.00019283 |  | 0 | 0 | 8.1455e-05 | [8 empty] | 0
734 | 0.40017 | 0.00017639 |  | 0 | 0 | 7.4508e-05 | [8 empty] | 0
735 | 0.40155 | 0.00017699 |  | 0 | 0 | 7.4765e-05 | [8 empty] | 0
736 | 0.43666 | 0.00019247 |  | 0 | 0 | 8.1301e-05 | [8 empty] | 0
737 | 0.41471 | 0.00018279 |  | 0 | 0 | 7.7214e-05 | [8 empty] | 0
738 | 0.55044 | 0.00024262 |  | 0 | 0 | 0.00010249 | [8 empty] | 0
739 | 0.55566 | 0.00024492 |  | 0 | 0 | 0.00010346 | [8 empty] | 0
740 | 0.38943 | 0.00017165 |  | 0 | 0 | 7.2508e-05 | [8 empty] | 0
741 | 0.39824 | 0.00017554 |  | 0 | 0 | 7.4149e-05 | [8 empty] | 0
742 | 0.39651 | 0.00017477 |  | 0 | 0 | 7.3827e-05 | [8 empty] | 0
743 | 0.38917 | 0.00017154 |  | 0 | 0 | 7.2459e-05 | [8 empty] | 0
744 | 0.3706 | 0.00016335 |  | 0 | 0 | 6.9002e-05 | [8 empty] | 0
745 | 0.35877 | 0.00015814 |  | 0 | 0 | 6.68e-05 | [8 empty] | 0
746 | 0.39575 | 0.00017444 |  | 0 | 0 | 7.3684e-05 | [8 empty] | 0
747 | 0.3727 | 0.00016428 |  | 0 | 0 | 6.9392e-05 | [8 empty] | 0
748 | 0.33818 | 0.00014906 |  | 0 | 0 | 6.2965e-05 | [8 empty] | 0
749 | 0.32956 | 0.00014526 |  | 0 | 0 | 6.136e-05 | [8 empty] | 0
750 | 0.43141 | 0.00019015 |  | 0 | 0 | 8.0323e-05 | [8 empty] | 0
751 | 0.60595 | 0.00026709 |  | 0 | 0 | 0.00011282 | [8 empty] | 0
752 | 0.50743 | 0.00022366 |  | 0 | 0 | 9.4478e-05 | [8 empty] | 0
753 | 0.35952 | 0.00015847 |  | 0 | 0 | 6.694e-05 | [8 empty] | 0
754 | 0.3534 | 0.00015577 |  | 0 | 0 | 6.5799e-05 | [8 empty] | 0
755 | 0.35075 | 0.0001546 |  | 0 | 0 | 6.5306e-05 | [8 empty] | 0
756 | 0.30549 | 0.00013465 |  | 0 | 0 | 5.6879e-05 | [8 empty] | 0
757 | 0.28937 | 0.00012755 |  | 0 | 0 | 5.3878e-05 | [8 empty] | 0
758 | 0.27662 | 0.00012193 |  | 0 | 0 | 5.1503e-05 | [8 empty] | 0
759 | 0.29278 | 0.00012905 |  | 0 | 0 | 5.4512e-05 | [8 empty] | 0
760 | 0.2715 | 0.00011967 |  | 0 | 0 | 5.055e-05 | [8 empty] | 0
761 | 0.25923 | 0.00011426 |  | 0 | 0 | 4.8266e-05 | [8 empty] | 0
762 | 0.29787 | 0.0001313 |  | 0 | 0 | 5.5461e-05 | [8 empty] | 0
763 | 0.64219 | 0.00028306 |  | 0 | 0 | 0.00011957 | [8 empty] | 0
764 | 0.81354 | 0.00035859 |  | 0 | 0 | 0.00015147 | [8 empty] | 0
765 | 0.32276 | 0.00014226 |  | 0 | 0 | 6.0095e-05 | [8 empty] | 0
766 | 0.26569 | 0.00011711 |  | 0 | 0 | 4.9469e-05 | [8 empty] | 0
767 | 0.24149 | 0.00010644 |  | 0 | 0 | 4.4963e-05 | [8 empty] | 0
768 | 0.25071 | 0.00011051 |  | 0 | 0 | 4.6679e-05 | [8 empty] | 0
769 | 0.22976 | 0.00010127 |  | 0 | 0 | 4.2779e-05 | [8 empty] | 0
770 | 0.24368 | 0.00010741 |  | 0 | 0 | 4.5371e-05 | [8 empty] | 0
771 | 0.23896 | 0.00010533 |  | 0 | 0 | 4.4491e-05 | [8 empty] | 0
772 | 0.43188 | 0.00019036 |  | 0 | 0 | 8.0411e-05 | [8 empty] | 0
773 | 0.4883 | 0.00021523 |  | 0 | 0 | 9.0916e-05 | [8 empty] | 0
774 | 0.24273 | 0.00010699 |  | 0 | 0 | 4.5193e-05 | [8 empty] | 0
775 | 0.19287 | 8.5013e-05 |  | 0 | 0 | 3.591e-05 | [8 empty] | 0
776 | 0.22426 | 9.885e-05 |  | 0 | 0 | 4.1755e-05 | [8 empty] | 0
777 | 0.21639 | 9.538e-05 |  | 0 | 0 | 4.029e-05 | [8 empty] | 0
778 | 0.17349 | 7.6472e-05 |  | 0 | 0 | 3.2303e-05 | [8 empty] | 0
779 | 0.2029 | 8.9434e-05 |  | 0 | 0 | 3.7778e-05 | [8 empty] | 0
780 | 0.18713 | 8.2481e-05 |  | 0 | 0 | 3.4841e-05 | [8 empty] | 0
781 | 0.18617 | 8.206e-05 |  | 0 | 0 | 3.4663e-05 | [8 empty] | 0
782 | 0.19347 | 8.5278e-05 |  | 0 | 0 | 3.6023e-05 | [8 empty] | 0
783 | 0.1755 | 7.7356e-05 |  | 0 | 0 | 3.2676e-05 | [8 empty] | 0
784 | 0.20802 | 9.1692e-05 |  | 0 | 0 | 3.8732e-05 | [8 empty] | 0
785 | 0.18007 | 7.9369e-05 |  | 0 | 0 | 3.3526e-05 | [8 empty] | 0
786 | 0.18703 | 8.2437e-05 |  | 0 | 0 | 3.4823e-05 | [8 empty] | 0
787 | 0.20351 | 8.9703e-05 |  | 0 | 0 | 3.7892e-05 | [8 empty] | 0
788 | 0.18842 | 8.3053e-05 |  | 0 | 0 | 3.5083e-05 | [8 empty] | 0
789 | 0.073101 | 3.2221e-05 |  | 0 | 0 | 1.3611e-05 | [8 empty] | 0
790 | 0.12546 | 5.53e-05 |  | 0 | 0 | 2.3359e-05 | [8 empty] | 0
791 | 0.14252 | 6.2819e-05 |  | 0 | 0 | 2.6535e-05 | [8 empty] | 0
792 | 0.16779 | 7.3958e-05 |  | 0 | 0 | 3.1241e-05 | [8 empty] | 0
793 | 0.1276 | 5.6244e-05 |  | 0 | 0 | 2.3758e-05 | [8 empty] | 0
794 | 0.12755 | 5.622e-05 |  | 0 | 0 | 2.3748e-05 | [8 empty] | 0
795 | 0.29253 | 0.00012894 |  | 0 | 0 | 5.4467e-05 | [8 empty] | 0
796 | 0.23524 | 0.00010369 |  | 0 | 0 | 4.38e-05 | [8 empty] | 0
797 | 0.19104 | 8.4206e-05 |  | 0 | 0 | 3.557e-05 | [8 empty] | 0
798 | 0.17212 | 7.5865e-05 |  | 0 | 0 | 3.2046e-05 | [8 empty] | 0
799 | 0.17729 | 7.8145e-05 |  | 0 | 0 | 3.301e-05 | [8 empty] | 0
800 | 0.21443 | 9.4516e-05 |  | 0 | 0 | 3.9925e-05 | [8 empty] | 0
199 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
198 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
197 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
196 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
195 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
194 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
193 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
192 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
191 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
190 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
189 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
188 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
187 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
186 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
185 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
184 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
183 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
182 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
181 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
180 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
179 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
178 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A
177 | #N/A | #N/A |  | 0 | #N/A | #N/A | #N/A | #N/A | [6 empty] | #N/A

## Lamp spectra
Date |  | 2021-10-01
Spectrometer |  | Ocean HDX
Operator(s) |  | DPM
Lamp(s) |  | GE LED lamp
λavg (300-500) |  | 447.7086
λ peak |  | 436
Plotted output |  | GE LED lamp NEW INTERPOLATION
| 13.1765 | 5899.2118
wavelength | NEW INTERPOLATION | NEW*lambda
178 | 0 | 0
179 | 0 | 0
180 | 0 | 0
181 | 0 | 0
182 | 0 | 0
183 | 0 | 0
184 | 0 | 0
185 | 0 | 0
186 | 0 | 0
187 | 0 | 0
188 | 0 | 0
189 | 0 | 0
190 | 0 | 0
191 | 0 | 0
192 | 0 | 0
193 | 0 | 0
194 | 0 | 0
195 | 0 | 0
196 | 0 | 0
197 | 0 | 0
198 | 0 | 0
199 | 0 | 0
200 | 0.00090114 | 0.18023
201 | 0.0015609 | 0.31375
202 | 0.0011137 | 0.22497
203 | 0.00053034 | 0.10766
204 | 0.00079572 | 0.16233
205 | 0.00034981 | 0.07171
206 | 0 | 0
207 | 0.0006128 | 0.12685
208 | 0.0013151 | 0.27355
209 | 0.001292 | 0.27003
210 | 0.0011154 | 0.23424
211 | 0.0016479 | 0.3477
212 | 0.0011269 | 0.23891
213 | 0.00083724 | 0.17833
214 | 0.00032788 | 0.070167
215 | 0.0013596 | 0.29231
216 | 0.0015133 | 0.32687
217 | 0.00081219 | 0.17625
218 | 0.0007436 | 0.1621
219 | 0.00099525 | 0.21796
220 | 0.0011937 | 0.26261
221 | 0.00075975 | 0.16791
222 | 0.00058683 | 0.13028
223 | 0.0010999 | 0.24529
224 | 0.0010388 | 0.23268
225 | 0.0010279 | 0.23127
226 | 0.0010991 | 0.2484
227 | 0.0011239 | 0.25512
228 | 0.00094443 | 0.21533
229 | 0.0012171 | 0.27872
230 | 0.0012828 | 0.29505
231 | 0.00072785 | 0.16813
232 | 0.00038864 | 0.090165
233 | 0.00064925 | 0.15127
234 | 0.00028445 | 0.066561
235 | 0.00080077 | 0.18818
236 | 0.0011421 | 0.26953
237 | 0.0015981 | 0.37875
238 | 0.00098994 | 0.2356
239 | 0.0010092 | 0.24119
240 | 0.0013246 | 0.3179
241 | 0.0011631 | 0.28031
242 | 0.0012163 | 0.29433
243 | 0.00075223 | 0.18279
244 | 0.00040296 | 0.098322
245 | 0.00043579 | 0.10677
246 | 0.00078757 | 0.19374
247 | 0.00053034 | 0.13099
248 | 0.00036627 | 0.090836
249 | 0.0001491 | 0.037127
250 | 0.000648 | 0.162
251 | 0.0006426 | 0.16129
252 | 0.00075197 | 0.1895
253 | 0.00056862 | 0.14386
254 | 0.00074804 | 0.19
255 | 0.00034649 | 0.088355
256 | 0 | 0
257 | 0.00069982 | 0.17985
258 | 0.0013355 | 0.34455
259 | 0.0011711 | 0.30332
260 | 0.0010382 | 0.26994
261 | 0.0015536 | 0.4055
262 | 0.0010532 | 0.27593
263 | 0.00064357 | 0.16926
264 | 0.00059642 | 0.15745
265 | 0.0016058 | 0.42553
266 | 0.0011264 | 0.29963
267 | 0.00058666 | 0.15664
268 | 0.00097463 | 0.2612
269 | 0.0011444 | 0.30786
270 | 0.0011188 | 0.30208
271 | 0.00034585 | 0.093726
272 | 0.00085706 | 0.23312
273 | 0.0012335 | 0.33673
274 | 0.001314 | 0.36004
275 | 0.0011462 | 0.3152
276 | 0.0011093 | 0.30618
277 | 0.00083221 | 0.23052
278 | 0.0010269 | 0.28549
279 | 0.0013834 | 0.38598
280 | 0.0009096 | 0.25469
281 | 0.00040696 | 0.11436
282 | 0.00071497 | 0.20162
283 | 0.00039564 | 0.11197
284 | 0.0008023 | 0.22785
285 | 0.0011062 | 0.31528
286 | 0.0015862 | 0.45366
287 | 0.001052 | 0.30193
288 | 0.0010358 | 0.29831
289 | 0.0013408 | 0.38749
290 | 0.001185 | 0.34365
291 | 0.0012582 | 0.36614
292 | 0.00070053 | 0.20455
293 | 0.00045289 | 0.1327
294 | 0.00066459 | 0.19539
295 | 0.00084223 | 0.24846
296 | 0.00043213 | 0.12791
297 | 0.00023265 | 0.069096
298 | 0.00025386 | 0.075649
299 | 0.00075975 | 0.22717
300 | 0.00060583 | 0.18175
301 | 0.00081897 | 0.24651
302 | 0.00032928 | 0.099443
303 | 0.00024179 | 0.073262
304 | 0.00024131 | 0.073359
305 | 0.00044427 | 0.1355
306 | 0.00059087 | 0.18081
307 | 0.00094139 | 0.28901
308 | 0.00091528 | 0.28191
309 | 0.00044609 | 0.13784
310 | 0.00019484 | 0.060401
311 | 0.00025903 | 0.080559
312 | 0.00038775 | 0.12098
313 | 0.00066584 | 0.20841
314 | 0.00030771 | 0.096622
315 | 0.00025595 | 0.080623
316 | 0.00028607 | 0.090399
317 | 0.00039718 | 0.12591
318 | 0.00039578 | 0.12586
319 | 0.00017171 | 0.054776
320 | 0.00025749 | 0.082396
321 | 7.7004e-05 | 0.024718
322 | 0.00054963 | 0.17698
323 | 0.00056143 | 0.18134
324 | 6.7807e-05 | 0.021969
325 | 0.00047048 | 0.1529
326 | 0.00071338 | 0.23256
327 | 0.00054728 | 0.17896
328 | 0.000259 | 0.084954
329 | 0.00047098 | 0.15495
330 | 0.00063117 | 0.20829
331 | 0.00045404 | 0.15029
332 | 0.00019368 | 0.064303
333 | 0.00045024 | 0.14993
334 | 0.00074885 | 0.25012
335 | 0.00026037 | 0.087225
336 | 0.00055904 | 0.18784
337 | 0.00033841 | 0.11404
338 | -1.6495e-06 | -0.00055753
339 | 0.00018646 | 0.063211
340 | 0.00048481 | 0.16484
341 | 0.00075455 | 0.2573
342 | 0.00062742 | 0.21458
343 | 0.00054978 | 0.18858
344 | 0.00042742 | 0.14703
345 | 0.00041202 | 0.14215
346 | 0.00071875 | 0.24869
347 | 0.0004433 | 0.15382
348 | 0.00076852 | 0.26745
349 | 0.001093 | 0.38147
350 | 0.0014487 | 0.50705
351 | 0.0011765 | 0.41296
352 | 0.001108 | 0.39001
353 | 0.0013524 | 0.4774
354 | 0.002051 | 0.72605
355 | 0.002363 | 0.83886
356 | 0.002511 | 0.89393
357 | 0.00263 | 0.9389
358 | 0.0027489 | 0.98411
359 | 0.0034025 | 1.2215
360 | 0.0037466 | 1.3488
361 | 0.0040232 | 1.4524
362 | 0.0047293 | 1.712
363 | 0.0066139 | 2.4008
364 | 0.064539 | 23.4922
365 | 0.12977 | 47.3664
366 | 0.051201 | 18.7397
367 | 0.013576 | 4.9824
368 | 0.0070547 | 2.5961
369 | 0.0073424 | 2.7093
370 | 0.0078873 | 2.9183
371 | 0.0082466 | 3.0595
372 | 0.0091192 | 3.3923
373 | 0.0098069 | 3.658
374 | 0.0099748 | 3.7306
375 | 0.010356 | 3.8833
376 | 0.010865 | 4.0851
377 | 0.010946 | 4.1268
378 | 0.011536 | 4.3605
379 | 0.012418 | 4.7066
380 | 0.013228 | 5.0266
381 | 0.014144 | 5.3889
382 | 0.014724 | 5.6244
383 | 0.015 | 5.7448
384 | 0.015955 | 6.1269
385 | 0.017047 | 6.5633
386 | 0.017725 | 6.8419
387 | 0.018946 | 7.3321
388 | 0.019582 | 7.5979
389 | 0.020649 | 8.0325
390 | 0.022645 | 8.8317
391 | 0.023251 | 9.0911
392 | 0.023468 | 9.1994
393 | 0.024854 | 9.7677
394 | 0.026349 | 10.3816
395 | 0.027359 | 10.8067
396 | 0.02868 | 11.3574
397 | 0.030145 | 11.9677
398 | 0.031617 | 12.5835
399 | 0.03288 | 13.119
400 | 0.034717 | 13.8868
401 | 0.036485 | 14.6303
402 | 0.038735 | 15.5714
403 | 0.059871 | 24.1282
404 | 0.32256 | 130.3135
405 | 0.3113 | 126.0778
406 | 0.057007 | 23.1447
407 | 0.063719 | 25.9338
408 | 0.069415 | 28.3215
409 | 0.051544 | 21.0816
410 | 0.052168 | 21.3887
411 | 0.054139 | 22.2513
412 | 0.055092 | 22.6978
413 | 0.056858 | 23.4823
414 | 0.058686 | 24.2961
415 | 0.059977 | 24.8906
416 | 0.060563 | 25.1941
417 | 0.062044 | 25.8722
418 | 0.063736 | 26.6415
419 | 0.065207 | 27.3219
420 | 0.067228 | 28.2357
421 | 0.068677 | 28.9128
422 | 0.069338 | 29.2606
423 | 0.071261 | 30.1435
424 | 0.073202 | 31.0375
425 | 0.074718 | 31.755
426 | 0.075636 | 32.2209
427 | 0.075726 | 32.335
428 | 0.078574 | 33.6297
429 | 0.080259 | 34.4311
430 | 0.083197 | 35.7749
431 | 0.086274 | 37.1843
432 | 0.086529 | 37.3806
433 | 0.090597 | 39.2286
434 | 0.11974 | 51.9685
435 | 0.64619 | 281.0905
436 | 1 | 436
437 | 0.20942 | 91.5175
438 | 0.098061 | 42.9505
439 | 0.1006 | 44.1653
440 | 0.10312 | 45.3724
441 | 0.10375 | 45.754
442 | 0.10422 | 46.065
443 | 0.10557 | 46.7691
444 | 0.10689 | 47.4591
445 | 0.10752 | 47.8454
446 | 0.10759 | 47.9871
447 | 0.10727 | 47.9512
448 | 0.10816 | 48.4548
449 | 0.1092 | 49.0305
450 | 0.10996 | 49.4815
451 | 0.11104 | 50.0773
452 | 0.11208 | 50.6585
453 | 0.11273 | 51.0648
454 | 0.11345 | 51.506
455 | 0.11455 | 52.1201
456 | 0.11538 | 52.6116
457 | 0.11698 | 53.4582
458 | 0.11757 | 53.845
459 | 0.11817 | 54.2383
460 | 0.11903 | 54.7536
461 | 0.12008 | 55.3575
462 | 0.11872 | 54.8496
463 | 0.12093 | 55.9884
464 | 0.12197 | 56.5963
465 | 0.12288 | 57.1375
466 | 0.12374 | 57.6644
467 | 0.12436 | 58.0753
468 | 0.12481 | 58.4094
469 | 0.12553 | 58.8729
470 | 0.12588 | 59.1658
471 | 0.12591 | 59.303
472 | 0.12659 | 59.7496
473 | 0.12747 | 60.2947
474 | 0.12775 | 60.554
475 | 0.12778 | 60.6946
476 | 0.12766 | 60.7646
477 | 0.12801 | 61.0591
478 | 0.12843 | 61.3882
479 | 0.12833 | 61.4697
480 | 0.12828 | 61.5753
481 | 0.12882 | 61.9633
482 | 0.12936 | 62.3496
483 | 0.12972 | 62.6532
484 | 0.13017 | 63.0031
485 | 0.13107 | 63.5667
486 | 0.13207 | 64.1854
487 | 0.13208 | 64.3231
488 | 0.13153 | 64.186
489 | 0.13122 | 64.1674
490 | 0.1308 | 64.0902
491 | 0.13249 | 65.0525
492 | 0.13245 | 65.1675
493 | 0.12896 | 63.575
494 | 0.12823 | 63.3464
495 | 0.12698 | 62.8528
496 | 0.12678 | 62.8843
497 | 0.12624 | 62.7402
498 | 0.12535 | 62.4244
499 | 0.12434 | 62.044
500 | 0.12368 | 61.8412
501 | 0.1228 | 61.524
502 | 0.12199 | 61.237
503 | 0.12139 | 61.0577
504 | 0.12086 | 60.9121
505 | 0.12062 | 60.914
506 | 0.11985 | 60.6429
507 | 0.11928 | 60.4758
508 | 0.11831 | 60.0991
509 | 0.11821 | 60.1682
510 | 0.11782 | 60.0895
511 | 0.11696 | 59.7673
512 | 0.11676 | 59.7797
513 | 0.11608 | 59.5513
514 | 0.11557 | 59.403
515 | 0.11534 | 59.4019
516 | 0.11446 | 59.0604
517 | 0.11406 | 58.9696
518 | 0.11384 | 58.9701
519 | 0.11377 | 59.0478
520 | 0.11328 | 58.9036
521 | 0.11358 | 59.1735
522 | 0.11326 | 59.1222
523 | 0.11269 | 58.9384
524 | 0.11323 | 59.3345
525 | 0.11289 | 59.2661
526 | 0.11293 | 59.4031
527 | 0.11286 | 59.4789
528 | 0.11336 | 59.8529
529 | 0.1136 | 60.0951
530 | 0.11385 | 60.3379
531 | 0.11368 | 60.366
532 | 0.11423 | 60.7716
533 | 0.11539 | 61.5037
534 | 0.11681 | 62.3763
535 | 0.117 | 62.5952
536 | 0.11732 | 62.8851
537 | 0.11913 | 63.9737
538 | 0.12107 | 65.1357
539 | 0.1234 | 66.5129
540 | 0.12729 | 68.7342
541 | 0.13415 | 72.5752
542 | 0.14296 | 77.487
543 | 0.14496 | 78.714
544 | 0.14713 | 80.0403
545 | 0.26365 | 143.6868
546 | 0.67251 | 367.1888
547 | 0.37112 | 203.0022
548 | 0.14326 | 78.5091
549 | 0.14208 | 78.0019
550 | 0.14273 | 78.5009
551 | 0.14294 | 78.7601
552 | 0.14331 | 79.108
553 | 0.14432 | 79.8097
554 | 0.14481 | 80.2239
555 | 0.14498 | 80.4614
556 | 0.14558 | 80.9419
557 | 0.14545 | 81.017
558 | 0.14612 | 81.5361
559 | 0.14779 | 82.614
560 | 0.14893 | 83.4009
561 | 0.15026 | 84.2933
562 | 0.15091 | 84.8119
563 | 0.15163 | 85.3677
564 | 0.15244 | 85.9737
565 | 0.15374 | 86.862
566 | 0.15482 | 87.6257
567 | 0.15549 | 88.1609
568 | 0.15655 | 88.9187
569 | 0.1571 | 89.3888
570 | 0.15749 | 89.769
571 | 0.15817 | 90.3161
572 | 0.15884 | 90.8591
573 | 0.15855 | 90.8514
574 | 0.1592 | 91.3783
575 | 0.15947 | 91.6928
576 | 0.18446 | 106.2509
577 | 0.24921 | 143.7956
578 | 0.20434 | 118.109
579 | 0.24068 | 139.3533
580 | 0.19515 | 113.1881
581 | 0.16151 | 93.8365
582 | 0.16009 | 93.1724
583 | 0.16055 | 93.599
584 | 0.15994 | 93.4053
585 | 0.15831 | 92.6087
586 | 0.1572 | 92.1206
587 | 0.15893 | 93.2902
588 | 0.15632 | 91.9153
589 | 0.15114 | 89.0204
590 | 0.15004 | 88.5225
591 | 0.14885 | 87.9694
592 | 0.14868 | 88.0196
593 | 0.14855 | 88.0903
594 | 0.14557 | 86.4713
595 | 0.1412 | 84.0144
596 | 0.13854 | 82.5682
597 | 0.13624 | 81.3327
598 | 0.13513 | 80.8101
599 | 0.13515 | 80.9536
600 | 0.13296 | 79.7733
601 | 0.12941 | 77.7761
602 | 0.12584 | 75.7556
603 | 0.12423 | 74.9108
604 | 0.12219 | 73.805
605 | 0.12031 | 72.7899
606 | 0.11826 | 71.6667
607 | 0.11657 | 70.76
608 | 0.11614 | 70.6118
609 | 0.11677 | 71.1104
610 | 0.12734 | 77.6751
611 | 0.1662 | 101.5479
612 | 0.15901 | 97.3115
613 | 0.12624 | 77.3854
614 | 0.11813 | 72.5292
615 | 0.11223 | 69.0208
616 | 0.10377 | 63.9212
617 | 0.098732 | 60.9179
618 | 0.094447 | 58.3682
619 | 0.092098 | 57.0085
620 | 0.091373 | 56.6515
621 | 0.08911 | 55.3375
622 | 0.08689 | 54.0453
623 | 0.085448 | 53.2342
624 | 0.083527 | 52.1206
625 | 0.080561 | 50.3503
626 | 0.079034 | 49.4752
627 | 0.077436 | 48.5524
628 | 0.075096 | 47.1605
629 | 0.07338 | 46.1563
630 | 0.07372 | 46.4435
631 | 0.074879 | 47.2488
632 | 0.069339 | 43.8221
633 | 0.064224 | 40.654
634 | 0.062201 | 39.4356
635 | 0.060114 | 38.1722
636 | 0.058469 | 37.1861
637 | 0.05725 | 36.4683
638 | 0.0557 | 35.5366
639 | 0.054225 | 34.6499
640 | 0.052356 | 33.5079
641 | 0.050862 | 32.6026
642 | 0.050038 | 32.1241
643 | 0.048836 | 31.4014
644 | 0.047462 | 30.5652
645 | 0.04609 | 29.7278
646 | 0.045155 | 29.1704
647 | 0.043521 | 28.1583
648 | 0.042617 | 27.6161
649 | 0.042268 | 27.4316
650 | 0.04215 | 27.3974
651 | 0.041193 | 26.8169
652 | 0.039497 | 25.7518
653 | 0.03847 | 25.1211
654 | 0.036774 | 24.05
655 | 0.035131 | 23.0111
656 | 0.034112 | 22.3775
657 | 0.033934 | 22.2945
658 | 0.033147 | 21.8107
659 | 0.032279 | 21.2719
660 | 0.031078 | 20.5113
661 | 0.029985 | 19.8199
662 | 0.030395 | 20.1214
663 | 0.029758 | 19.7295
664 | 0.028434 | 18.8799
665 | 0.027494 | 18.2834
666 | 0.026379 | 17.5686
667 | 0.02626 | 17.5153
668 | 0.025448 | 16.9995
669 | 0.02436 | 16.2967
670 | 0.023865 | 15.9895
671 | 0.02374 | 15.9295
672 | 0.022929 | 15.4084
673 | 0.022215 | 14.9505
674 | 0.021693 | 14.6211
675 | 0.021462 | 14.4867
676 | 0.021175 | 14.3144
677 | 0.020354 | 13.7796
678 | 0.019375 | 13.136
679 | 0.019208 | 13.0419
680 | 0.01903 | 12.9401
681 | 0.018416 | 12.5415
682 | 0.017967 | 12.2536
683 | 0.017336 | 11.8404
684 | 0.016708 | 11.4282
685 | 0.016322 | 11.1809
686 | 0.01591 | 10.914
687 | 0.016929 | 11.6305
688 | 0.017102 | 11.7662
689 | 0.016085 | 11.0824
690 | 0.016679 | 11.5086
691 | 0.017515 | 12.1026
692 | 0.014074 | 9.7392
693 | 0.014452 | 10.015
694 | 0.01436 | 9.966
695 | 0.013334 | 9.2669
696 | 0.014191 | 9.8767
697 | 0.014113 | 9.8365
698 | 0.012617 | 8.8064
699 | 0.012061 | 8.4304
700 | 0.011902 | 8.3314
701 | 0.011644 | 8.1621
702 | 0.011787 | 8.2744
703 | 0.011524 | 8.1013
704 | 0.011436 | 8.0513
705 | 0.01198 | 8.4462
706 | 0.014594 | 10.3033
707 | 0.016141 | 11.4114
708 | 0.014237 | 10.0795
709 | 0.014859 | 10.5351
710 | 0.012418 | 8.8167
711 | 0.015192 | 10.8013
712 | 0.017377 | 12.3723
713 | 0.012052 | 8.5933
714 | 0.010121 | 7.2264
715 | 0.0090086 | 6.4411
716 | 0.008773 | 6.2815
717 | 0.0089912 | 6.4467
718 | 0.0080183 | 5.7571
719 | 0.00823 | 5.9174
720 | 0.0087561 | 6.3044
721 | 0.0083953 | 6.053
722 | 0.0080101 | 5.7833
723 | 0.0069167 | 5.0008
724 | 0.0072797 | 5.2705
725 | 0.0073666 | 5.3408
726 | 0.0071629 | 5.2003
727 | 0.0071117 | 5.1702
728 | 0.0072174 | 5.2543
729 | 0.0072683 | 5.2986
730 | 0.0062771 | 4.5823
731 | 0.0064178 | 4.6914
732 | 0.0066629 | 4.8773
733 | 0.0065621 | 4.81
734 | 0.0060024 | 4.4058
735 | 0.0060231 | 4.427
736 | 0.0065497 | 4.8205
737 | 0.0062204 | 4.5844
738 | 0.0082563 | 6.0932
739 | 0.0083346 | 6.1593
740 | 0.0058413 | 4.3226
741 | 0.0059735 | 4.4264
742 | 0.0059475 | 4.4131
743 | 0.0058373 | 4.3371
744 | 0.0055588 | 4.1358
745 | 0.0053814 | 4.0092
746 | 0.005936 | 4.4283
747 | 0.0055903 | 4.1759
748 | 0.0050725 | 3.7942
749 | 0.0049432 | 3.7025
750 | 0.0064709 | 4.8532
751 | 0.0090889 | 6.8258
752 | 0.0076112 | 5.7236
753 | 0.0053927 | 4.0607
754 | 0.0053008 | 3.9968
755 | 0.0052611 | 3.9721
756 | 0.0045822 | 3.4642
757 | 0.0043404 | 3.2857
758 | 0.0041491 | 3.145
759 | 0.0043915 | 3.3332
760 | 0.0040723 | 3.095
761 | 0.0038883 | 2.959
762 | 0.004468 | 3.4046
763 | 0.0096326 | 7.3497
764 | 0.012203 | 9.3229
765 | 0.0048413 | 3.7036
766 | 0.0039853 | 3.0527
767 | 0.0036222 | 2.7782
768 | 0.0037605 | 2.8881
769 | 0.0034463 | 2.6502
770 | 0.0036551 | 2.8144
771 | 0.0035842 | 2.7635
772 | 0.006478 | 5.001
773 | 0.0073242 | 5.6616
774 | 0.0036408 | 2.818
775 | 0.002893 | 2.2421
776 | 0.0033638 | 2.6103
777 | 0.0032458 | 2.522
778 | 0.0026023 | 2.0246
779 | 0.0030434 | 2.3708
780 | 0.0028068 | 2.1893
781 | 0.0027925 | 2.1809
782 | 0.002902 | 2.2694
783 | 0.0026324 | 2.0612
784 | 0.0031203 | 2.4463
785 | 0.0027009 | 2.1202
786 | 0.0028053 | 2.205
787 | 0.0030526 | 2.4024
788 | 0.0028263 | 2.2271
789 | 0.0010965 | 0.86513
790 | 0.0018818 | 1.4867
791 | 0.0021377 | 1.6909
792 | 0.0025168 | 1.9933
793 | 0.001914 | 1.5178
794 | 0.0019131 | 1.519
795 | 0.0043879 | 3.4884
796 | 0.0035286 | 2.8087
797 | 0.0028655 | 2.2838
798 | 0.0025817 | 2.0602
799 | 0.0026593 | 2.1248
800 | 0.0032164 | 2.5731
801 | 0.0045746
802 | 0.0059868
803 | 0.0027027
804 | 0.0023169
805 | 0.002356
806 | 0.0022233
807 | 0.0012349

## Raw Sample absorbance
UV VIS Wavelength | Absorbance | paste absorbance measurements here
500 | 0.059853
499 | 0.060144
498 | 0.060364
497 | 0.060621
496 | 0.06082
495 | 0.061042
494 | 0.061305
493 | 0.061562
492 | 0.061767
491 | 0.061972
490 | 0.06214
489 | 0.06241
488 | 0.062589
487 | 0.062885
486 | 0.063119
485 | 0.063332
484 | 0.063579
483 | 0.063809
482 | 0.064094
481 | 0.064315
480 | 0.064547
479 | 0.064852
478 | 0.065074
477 | 0.065347
476 | 0.065586
475 | 0.065874
474 | 0.066132
473 | 0.066395
472 | 0.066687
471 | 0.066955
470 | 0.067181
469 | 0.067443
468 | 0.06776
467 | 0.068022
466 | 0.068321
465 | 0.068561
464 | 0.068877
463 | 0.069167
462 | 0.069395
461 | 0.069682
460 | 0.07004
459 | 0.070321
458 | 0.070561
457 | 0.070886
456 | 0.071182
455 | 0.071493
454 | 0.071744
453 | 0.072119
452 | 0.072458
451 | 0.072726
450 | 0.072975
449 | 0.073311
448 | 0.07364
447 | 0.073921
446 | 0.074286
445 | 0.074543
444 | 0.074925
443 | 0.075215
442 | 0.075587
441 | 0.075871
440 | 0.076182
439 | 0.076525
438 | 0.076871
437 | 0.077187
436 | 0.077549
435 | 0.077891
434 | 0.078229
433 | 0.078573
432 | 0.078977
431 | 0.079352
430 | 0.079663
429 | 0.079985
428 | 0.080396
427 | 0.080748
426 | 0.081147
425 | 0.081515
424 | 0.081917
423 | 0.082278
422 | 0.082681
421 | 0.08302
420 | 0.083397
419 | 0.083772
418 | 0.084164
417 | 0.084562
416 | 0.084934
415 | 0.085378
414 | 0.085677
413 | 0.08615
412 | 0.086527
411 | 0.086948
410 | 0.087359
409 | 0.087762
408 | 0.088228
407 | 0.088556
406 | 0.088985
405 | 0.089445
404 | 0.089852
403 | 0.090274
402 | 0.090714
401 | 0.091145
400 | 0.091572
399 | 0.091996
398 | 0.092411
397 | 0.092894
396 | 0.093358
395 | 0.093759
394 | 0.094268
393 | 0.094768
392 | 0.095238
391 | 0.095691
390 | 0.096133
389 | 0.096662
388 | 0.097198
387 | 0.097638
386 | 0.098199
385 | 0.098622
384 | 0.099182
383 | 0.099713
382 | 0.10025
381 | 0.10073
380 | 0.10121
379 | 0.10175
378 | 0.10235
377 | 0.10289
376 | 0.10341
375 | 0.10391
374 | 0.10449
373 | 0.10503
372 | 0.10564
371 | 0.10622
370 | 0.10671
369 | 0.10734
368 | 0.10799
367 | 0.1086
366 | 0.10913
365 | 0.10982
364 | 0.11045
363 | 0.11106
362 | 0.11163
361 | 0.11226
360 | 0.11291
359 | 0.1136
358 | 0.11417
357 | 0.11489
356 | 0.11556
355 | 0.11617
354 | 0.11683
353 | 0.11751
352 | 0.11825
351 | 0.1189
350 | 0.11961
349 | 0.12046
348 | 0.12184
347 | 0.12191
346 | 0.12321
345 | 0.12359
344 | 0.12448
343 | 0.12538
342 | 0.12577
341 | 0.12698
340 | 0.12752
339 | 0.12846
338 | 0.12947
337 | 0.13005
336 | 0.13101
335 | 0.13179
334 | 0.13267
333 | 0.13361
332 | 0.13448
331 | 0.13548
330 | 0.13616
329 | 0.1372
328 | 0.1381
327 | 0.13913
326 | 0.14023
325 | 0.14098
324 | 0.14219
323 | 0.14309
322 | 0.14422
321 | 0.1452
320 | 0.14638
319 | 0.1475
318 | 0.14851
317 | 0.14977
316 | 0.15099
315 | 0.15217
314 | 0.1534
313 | 0.15468
312 | 0.15596
311 | 0.15734
310 | 0.15888
309 | 0.16026
308 | 0.16181
307 | 0.16346
306 | 0.16504
305 | 0.16691
304 | 0.16882
303 | 0.17076
302 | 0.17276
301 | 0.17505
300 | 0.17729
299 | 0.17961
298 | 0.18216
297 | 0.18479
296 | 0.18747
295 | 0.19019
294 | 0.19297
293 | 0.19573
292 | 0.1983
291 | 0.20092
290 | 0.20334
289 | 0.20582
288 | 0.2083
287 | 0.21081
286 | 0.21325
285 | 0.21543
284 | 0.21765
283 | 0.21973
282 | 0.22163
281 | 0.22355
280 | 0.22529
279 | 0.22712
278 | 0.22865
277 | 0.23018
276 | 0.23166
275 | 0.23299
274 | 0.23422
273 | 0.23535
272 | 0.23631
271 | 0.23731
270 | 0.23818
269 | 0.23894
268 | 0.23951
267 | 0.24007
266 | 0.24053
265 | 0.24097
264 | 0.24116
263 | 0.24133
262 | 0.24138
261 | 0.24142
260 | 0.24125
259 | 0.24104
258 | 0.24084
257 | 0.24052
256 | 0.24032
255 | 0.24009
254 | 0.23997
253 | 0.2397
252 | 0.23951
251 | 0.23949
250 | 0.23953
249 | 0.23977
248 | 0.24018
247 | 0.24077
246 | 0.24159
245 | 0.24271
244 | 0.24422
243 | 0.24599
242 | 0.24809
241 | 0.25064
240 | 0.25351
239 | 0.25695
238 | 0.26066
237 | 0.26465
236 | 0.26905
235 | 0.27368
234 | 0.2785
233 | 0.28338
232 | 0.2884
231 | 0.29322
230 | 0.29806
229 | 0.30281
228 | 0.30735
227 | 0.31201
226 | 0.31652
225 | 0.32121
224 | 0.32567
223 | 0.33036
222 | 0.33496
221 | 0.3395
220 | 0.34412
219 | 0.3488
218 | 0.35365
217 | 0.35878
216 | 0.36439
215 | 0.37099
214 | 0.3786
213 | 0.38794
212 | 0.40027
211 | 0.41669
210 | 0.43939
209 | 0.47133
208 | 0.51697
207 | 0.58286
206 | 0.67836
205 | 0.81652
204 | 1.0143
203 | 1.2903
202 | 1.6678
201 | 2.1831
200 | 2.8706

## irradiance interpolation
wavelength | raw blanked output (averaged) | new wavelength | forecast over 1, 1/2 nm increments | new wavelength | forecast over 1, 0.2 nm increments | new wavelength | forecast over 1, 0.5 nm increments | new wavelength | forecast over 1, 1 nm increments | 1 | RLE is shown in column k | Irradiance
195.423 | 0 | 199.8 | 0 | 199.8 | 0 | 199.5 | 0 | 200 | -0.86354 | -0.86354 |  | 141.2575
195.796 | 0 | 199.9 | -0.75746 | 200 | -1.0545 | 200 | -0.83795 | 201 | -1.5017 | -1.5017
196.169 | 0 | 200 | -1.0545 | 200.2 | -1.4593 | 200.5 | -1.7527 | 202 | -1.1861 | -1.1861
196.542 | 0 | 200.1 | -1.3515 | 200.4 | -1.6046 | 201 | -1.5425 | 203 | -1.3303 | -1.3303
196.915 | 0 | 200.2 | -1.6486 | 200.6 | -1.8792 | 201.5 | -1.2099 | 204 | -1.332 | -1.332
197.287 | 0 | 200.3 | -1.3779 | 200.8 | -1.6562 | 202 | -1.1654 | 205 | -1.1876 | -1.1876
197.66 | 0 | 200.4 | -1.6046 | 201 | -1.5117 | 202.5 | -1.1831 | 206 | -1.1743 | -1.1743
198.032 | 0 | 200.5 | -1.8313 | 201.2 | -1.4596 | 203 | -1.3671 | 207 | -1.1368 | -1.1368
198.404 | 0 | 200.6 | -2.0581 | 201.4 | -1.2991 | 203.5 | -1.4407 | 208 | -1.0248 | -1.0248
198.777 | 0 | 200.7 | -1.7481 | 201.6 | -1.1232 | 204 | -1.2896 | 209 | -0.99538 | -0.99538
199.149 | 0 | 200.8 | -1.6562 | 201.8 | -1.16 | 204.5 | -1.2656 | 210 | -1.0323 | -1.0323
199.521 | 0 | 200.9 | -1.5642 | 202 | -1.2034 | 205 | -1.1649 | 211 | -0.96965 | -0.96965
199.893 | 0 | 201 | -1.4723 | 202.2 | -1.1328 | 205.5 | -1.1324 | 212 | -0.97433 | -0.97433
200.265 | -2.21 | 201.1 | -1.4986 | 202.4 | -1.1433 | 206 | -1.1791 | 213 | -0.78046 | -0.78046
200.636 | -1.6833 | 201.2 | -1.4596 | 202.6 | -1.213 | 206.5 | -1.2114 | 214 | -0.64941 | -0.64941
201.008 | -1.5267 | 201.3 | -1.4206 | 202.8 | -1.2439 | 207 | -1.0923 | 215 | -0.59996 | -0.59996
201.38 | -1.3933 | 201.4 | -1.2766 | 203 | -1.3821 | 207.5 | -1.1068 | 216 | -0.66533 | -0.66533
201.751 | -0.95667 | 201.5 | -1.1999 | 203.2 | -1.4752 | 208 | -1.0137 | 217 | -0.6548 | -0.6548
202.123 | -1.3267 | 201.6 | -1.1232 | 203.4 | -1.4619 | 208.5 | -0.95385 | 218 | -0.58568 | -0.58568
202.494 | -1 | 201.7 | -1.0465 | 203.6 | -1.4237 | 209 | -0.89044 | 219 | -0.50295 | -0.50295
202.865 | -1.3167 | 201.8 | -1.2212 | 203.8 | -1.2323 | 209.5 | -1.1419 | 220 | -0.49265 | -0.49265
203.236 | -1.5367 | 201.9 | -1.2123 | 204 | -1.2848 | 210 | -1.0599 | 221 | -0.47265 | -0.47265
203.608 | -1.4467 | 202 | -1.2034 | 204.2 | -1.3516 | 210.5 | -0.89518 | 222 | -0.45931 | -0.45931
203.979 | -1.0733 | 202.1 | -1.1944 | 204.4 | -1.2771 | 211 | -0.92328 | 223 | -0.43373 | -0.43373
204.349 | -1.4933 | 202.2 | -1.099 | 204.6 | -1.2468 | 211.5 | -1.0905 | 224 | -0.40943 | -0.40943
204.72 | -1.1233 | 202.3 | -1.1049 | 204.8 | -1.217 | 212 | -0.96825 | 225 | -0.34121 | -0.34121
205.091 | -1.1567 | 202.4 | -1.1107 | 205 | -1.1463 | 212.5 | -0.86426 | 226 | -0.31575 | -0.31575
205.462 | -1.1167 | 202.5 | -1.2144 | 205.2 | -1.1312 | 213 | -0.76961 | 227 | -0.26467 | -0.26467
205.832 | -1.1367 | 202.6 | -1.213 | 205.4 | -1.1318 | 213.5 | -0.7075 | 228 | -0.21653 | -0.21653
206.203 | -1.2233 | 202.7 | -1.2117 | 205.6 | -1.1329 | 214 | -0.66069 | 229 | -0.18053 | -0.18053
206.573 | -1.27 | 202.8 | -1.2103 | 205.8 | -1.1421 | 214.5 | -0.58005 | 230 | -0.17088 | -0.17088
206.944 | -1.0333 | 202.9 | -1.3098 | 206 | -1.183 | 215 | -0.61976 | 231 | -0.15898 | -0.15898
207.314 | -1.08 | 203 | -1.3821 | 206.2 | -1.2123 | 215.5 | -0.60006 | 232 | -0.13334 | -0.13334
207.684 | -1.1967 | 203.1 | -1.4544 | 206.4 | -1.2455 | 216 | -0.67849 | 233 | -0.1102 | -0.1102
208.054 | -0.85333 | 203.2 | -1.5267 | 206.6 | -1.1918 | 216.5 | -0.71745 | 234 | -0.10785 | -0.10785
208.424 | -1.12 | 203.3 | -1.4445 | 206.8 | -1.1174 | 217 | -0.61999 | 235 | -0.10216 | -0.10216
208.794 | -0.81333 | 203.4 | -1.4619 | 207 | -1.0976 | 217.5 | -0.62697 | 236 | -0.085219 | -0.085219
209.164 | -0.79667 | 203.5 | -1.4794 | 207.2 | -1.062 | 218 | -0.59653 | 237 | -0.087287 | -0.087287
209.533 | -1.3433 | 203.6 | -1.4969 | 207.4 | -1.101 | 218.5 | -0.53354 | 238 | -0.086889 | -0.086889
209.903 | -1.1033 | 203.7 | -1.2947 | 207.6 | -1.1164 | 219 | -0.50334 | 239 | -0.082122 | -0.082122
210.273 | -0.81667 | 203.8 | -1.2323 | 207.8 | -1.0078 | 219.5 | -0.47196 | 240 | -0.072187 | -0.072187
210.642 | -0.92333 | 203.9 | -1.17 | 208 | -0.99187 | 220 | -0.49001 | 241 | -0.069213 | -0.069213
211.012 | -0.87667 | 204 | -1.3391 | 208.2 | -1.0415 | 220.5 | -0.51599 | 242 | -0.06783 | -0.06783
211.381 | -1.0267 | 204.1 | -1.3454 | 208.4 | -0.99226 | 221 | -0.47857 | 243 | -0.064163 | -0.064163
211.75 | -1.21 | 204.2 | -1.3516 | 208.6 | -0.91938 | 221.5 | -0.42338 | 244 | -0.057777 | -0.057777
212.119 | -0.77333 | 204.3 | -1.3579 | 208.8 | -0.89501 | 222 | -0.46679 | 245 | -0.061637 | -0.061637
212.488 | -1.0433 | 204.4 | -1.2334 | 209 | -0.81999 | 222.5 | -0.48776 | 246 | -0.061556 | -0.061556
212.857 | -0.69 | 204.5 | -1.2401 | 209.2 | -0.95632 | 223 | -0.39533 | 247 | -0.061827 | -0.061827
213.226 | -0.76667 | 204.6 | -1.2468 | 209.4 | -1.1539 | 223.5 | -0.41811 | 248 | -0.060425 | -0.060425
213.595 | -0.68 | 204.7 | -1.2535 | 209.6 | -1.1615 | 224 | -0.42448 | 249 | -0.062085 | -0.062085
213.964 | -0.68 | 204.8 | -1.2215 | 209.8 | -1.1916 | 224.5 | -0.38569 | 250 | -0.060514 | -0.060514
214.332 | -0.61667 | 204.9 | -1.1761 | 210 | -1.0665 | 225 | -0.31475 | 251 | -0.058499 | -0.058499
214.701 | -0.53667 | 205 | -1.1307 | 210.2 | -0.92169 | 225.5 | -0.32318 | 252 | -0.049771 | -0.049771
215.069 | -0.72667 | 205.1 | -1.1321 | 210.4 | -0.91673 | 226 | -0.32279 | 253 | -0.048531 | -0.048531
215.437 | -0.54333 | 205.2 | -1.1312 | 210.6 | -0.87907 | 226.5 | -0.30127 | 254 | -0.05356 | -0.05356
215.806 | -0.6 | 205.3 | -1.1303 | 210.8 | -0.88501 | 227 | -0.26028 | 255 | -0.057469 | -0.057469
216.174 | -0.79 | 205.4 | -1.1294 | 211 | -0.9163 | 227.5 | -0.23247 | 256 | -0.055284 | -0.055284
216.542 | -0.70667 | 205.5 | -1.1356 | 211.2 | -0.96853 | 228 | -0.22345 | 257 | -0.055749 | -0.055749
216.91 | -0.64 | 205.6 | -1.1329 | 211.4 | -1.0401 | 228.5 | -0.19367 | 258 | -0.056659 | -0.056659
217.278 | -0.53667 | 205.7 | -1.1302 | 211.6 | -1.1367 | 229 | -0.17409 | 259 | -0.04562 | -0.04562
217.646 | -0.69667 | 205.8 | -1.1275 | 211.8 | -1.04 | 229.5 | -0.17384 | 260 | -0.05204 | -0.05204
218.014 | -0.54667 | 205.9 | -1.1686 | 212 | -0.91752 | 230 | -0.17455 | 261 | -0.056899 | -0.056899
218.381 | -0.60667 | 206 | -1.183 | 212.2 | -0.94727 | 230.5 | -0.16425 | 262 | -0.051547 | -0.051547
218.749 | -0.46667 | 206.1 | -1.1974 | 212.4 | -0.91588 | 231 | -0.15577 | 263 | -0.057487 | -0.057487
219.116 | -0.54 | 206.2 | -1.2118 | 212.6 | -0.82291 | 231.5 | -0.15691 | 264 | -0.058383 | -0.058383
219.484 | -0.40333 | 206.3 | -1.2275 | 212.8 | -0.80972 | 232 | -0.13219 | 265 | -0.055849 | -0.055849
219.851 | -0.50667 | 206.4 | -1.2455 | 213 | -0.77972 | 232.5 | -0.11091 | 266 | -0.060852 | -0.060852
220.218 | -0.49667 | 206.5 | -1.2635 | 213.2 | -0.7194 | 233 | -0.10707 | 267 | -0.058238 | -0.058238
220.586 | -0.53333 | 206.6 | -1.1687 | 213.4 | -0.70986 | 233.5 | -0.11261 | 268 | -0.055255 | -0.055255
220.953 | -0.50333 | 206.7 | -1.1431 | 213.6 | -0.70446 | 234 | -0.10606 | 269 | -0.057286 | -0.057286
221.32 | -0.41 | 206.8 | -1.1174 | 213.8 | -0.68481 | 234.5 | -0.10489 | 270 | -0.056831 | -0.056831
221.687 | -0.41333 | 206.9 | -1.0917 | 214 | -0.65867 | 235 | -0.10085 | 271 | -0.058967 | -0.058967
222.053 | -0.46667 | 207 | -1.1133 | 214.2 | -0.63859 | 235.5 | -0.10075 | 272 | -0.059899 | -0.059899
222.42 | -0.54333 | 207.1 | -1.0877 | 214.4 | -0.60215 | 236 | -0.079944 | 273 | -0.064962 | -0.064962
222.787 | -0.43333 | 207.2 | -1.062 | 214.6 | -0.55905 | 236.5 | -0.074962 | 274 | -0.057922 | -0.057922
223.153 | -0.32 | 207.3 | -1.0363 | 214.8 | -0.61249 | 237 | -0.089446 | 275 | -0.056897 | -0.056897
223.52 | -0.47 | 207.4 | -1.1223 | 215 | -0.64339 | 237.5 | -0.097453 | 276 | -0.06155 | -0.06155
223.886 | -0.41667 | 207.5 | -1.1444 | 215.2 | -0.60341 | 238 | -0.083871 | 277 | -0.062689 | -0.062689
224.253 | -0.42 | 207.6 | -1.1665 | 215.4 | -0.60737 | 238.5 | -0.079343 | 278 | -0.060666 | -0.060666
224.619 | -0.37667 | 207.7 | -1.0384 | 215.6 | -0.5954 | 239 | -0.089868 | 279 | -0.052484 | -0.052484
224.985 | -0.29667 | 207.8 | -1.0078 | 215.8 | -0.6051 | 239.5 | -0.077154 | 280 | -0.051949 | -0.051949
225.351 | -0.29 | 207.9 | -0.97717 | 216 | -0.70947 | 240 | -0.071492 | 281 | -0.05367 | -0.05367
225.717 | -0.35667 | 208 | -0.94654 | 216.2 | -0.72091 | 240.5 | -0.067914 | 282 | -0.052341 | -0.052341
226.083 | -0.30667 | 208.1 | -1.0519 | 216.4 | -0.73164 | 241 | -0.069971 | 283 | -0.056092 | -0.056092
226.449 | -0.31 | 208.2 | -1.0415 | 216.6 | -0.70885 | 241.5 | -0.069754 | 284 | -0.06066 | -0.06066
226.814 | -0.28667 | 208.3 | -1.0312 | 216.8 | -0.65964 | 242 | -0.071218 | 285 | -0.055731 | -0.055731
227.18 | -0.22667 | 208.4 | -1.0208 | 217 | -0.61005 | 242.5 | -0.062519 | 286 | -0.052388 | -0.052388
227.546 | -0.23667 | 208.5 | -0.92478 | 217.2 | -0.59027 | 243 | -0.065551 | 287 | -0.049113 | -0.049113
227.911 | -0.22667 | 208.6 | -0.91938 | 217.4 | -0.63384 | 243.5 | -0.06442 | 288 | -0.046902 | -0.046902
228.276 | -0.21667 | 208.7 | -0.91397 | 217.6 | -0.62828 | 244 | -0.048941 | 289 | -0.03915 | -0.03915
228.642 | -0.17667 | 208.8 | -0.90738 | 217.8 | -0.59543 | 244.5 | -0.05997 | 290 | -0.041628 | -0.041628
229.007 | -0.18 | 208.9 | -0.86368 | 218 | -0.60034 | 245 | -0.060366 | 291 | -0.041996 | -0.041996
229.372 | -0.15667 | 209 | -0.81999 | 218.2 | -0.59383 | 245.5 | -0.064574 | 292 | -0.036983 | -0.036983
229.737 | -0.18667 | 209.1 | -0.7763 | 218.4 | -0.54887 | 246 | -0.05993 | 293 | -0.038859 | -0.038859
230.102 | -0.17333 | 209.2 | -1.0105 | 218.6 | -0.51617 | 246.5 | -0.060165 | 294 | -0.037955 | -0.037955
230.467 | -0.15333 | 209.3 | -1.0822 | 218.8 | -0.52081 | 247 | -0.055905 | 295 | -0.033483 | -0.033483
230.832 | -0.17 | 209.4 | -1.1539 | 219 | -0.51495 | 247.5 | -0.069412 | 296 | -0.030776 | -0.030776
231.196 | -0.13667 | 209.5 | -1.2255 | 219.2 | -0.47426 | 248 | -0.055367 | 297 | -0.028452 | -0.028452
231.561 | -0.18 | 209.6 | -1.1087 | 219.4 | -0.46075 | 248.5 | -0.056497 | 298 | -0.0271 | -0.0271
231.926 | -0.12333 | 209.7 | -1.1502 | 219.6 | -0.47804 | 249 | -0.067185 | 299 | -0.026352 | -0.026352
232.29 | -0.12 | 209.8 | -1.1916 | 219.8 | -0.47252 | 249.5 | -0.062574 | 300 | -0.025519 | -0.025519
232.654 | -0.10333 | 209.9 | -1.2331 | 220 | -0.48784 | 250 | -0.059367 | 301 | -0.029777 | -0.029777
233.019 | -0.10667 | 210 | -1.0187 | 220.2 | -0.50967 | 250.5 | -0.059602 | 302 | -0.032269 | -0.032269
233.383 | -0.11 | 210.1 | -0.94757 | 220.4 | -0.51882 | 251 | -0.063699 | 303 | -0.031945 | -0.031945
233.747 | -0.11667 | 210.2 | -0.8764 | 220.6 | -0.51528 | 251.5 | -0.052195 | 304 | -0.030832 | -0.030832
234.111 | -0.096667 | 210.3 | -0.94111 | 220.8 | -0.51306 | 252 | -0.048946 | 305 | -0.028338 | -0.028338
234.475 | -0.11333 | 210.4 | -0.91673 | 221 | -0.48194 | 252.5 | -0.048171 | 306 | -0.026998 | -0.026998
234.839 | -0.1 | 210.5 | -0.89235 | 221.2 | -0.44072 | 253 | -0.046187 | 307 | -0.025695 | -0.025695
235.203 | -0.096667 | 210.6 | -0.86797 | 221.4 | -0.42549 | 253.5 | -0.051234 | 308 | -0.023593 | -0.023593
235.566 | -0.11333 | 210.7 | -0.8769 | 221.6 | -0.41969 | 254 | -0.051605 | 309 | -0.023762 | -0.023762
235.93 | -0.073333 | 210.8 | -0.88501 | 221.8 | -0.43876 | 254.5 | -0.05784 | 310 | -0.02145 | -0.02145
236.293 | -0.07 | 210.9 | -0.89312 | 222 | -0.46114 | 255 | -0.062057 | 311 | -0.022987 | -0.022987
236.657 | -0.076667 | 211 | -0.90123 | 222.2 | -0.50046 | 255.5 | -0.05251 | 312 | -0.022819 | -0.022819
237.02 | -0.086667 | 211.1 | -0.95456 | 222.4 | -0.51054 | 256 | -0.055478 | 313 | -0.01958 | -0.01958
237.384 | -0.11 | 211.2 | -0.96853 | 222.6 | -0.47294 | 256.5 | -0.057865 | 314 | -0.021098 | -0.021098
237.747 | -0.09 | 211.3 | -0.9825 | 222.8 | -0.43142 | 257 | -0.049003 | 315 | -0.021362 | -0.021362
238.11 | -0.08 | 211.4 | -1.0464 | 223 | -0.36722 | 257.5 | -0.06038 | 316 | -0.024316 | -0.024316
238.473 | -0.07 | 211.5 | -1.0915 | 223.2 | -0.38734 | 258 | -0.058518 | 317 | -0.02364 | -0.02364
238.836 | -0.09 | 211.6 | -1.1367 | 223.4 | -0.42016 | 258.5 | -0.051078 | 318 | -0.020397 | -0.020397
239.199 | -0.10333 | 211.7 | -1.1819 | 223.6 | -0.42134 | 259 | -0.04357 | 319 | -0.022725 | -0.022725
239.561 | -0.06 | 211.8 | -0.98617 | 223.8 | -0.43329 | 259.5 | -0.042212 | 320 | -0.02114 | -0.02114
239.924 | -0.08 | 211.9 | -0.95184 | 224 | -0.42781 | 260 | -0.0537 | 321 | -0.019593 | -0.019593
240.287 | -0.066667 | 212 | -0.91752 | 224.2 | -0.41234 | 260.5 | -0.060207 | 322 | -0.019867 | -0.019867
240.649 | -0.066667 | 212.1 | -0.88319 | 224.4 | -0.39641 | 261 | -0.056442 | 323 | -0.021431 | -0.021431
241.012 | -0.07 | 212.2 | -0.9906 | 224.6 | -0.37575 | 261.5 | -0.054049 | 324 | -0.024705 | -0.024705
241.374 | -0.073333 | 212.3 | -0.96801 | 224.8 | -0.33395 | 262 | -0.048655 | 325 | -0.022849 | -0.022849
241.736 | -0.066667 | 212.4 | -0.94543 | 225 | -0.31464 | 262.5 | -0.051937 | 326 | -0.022728 | -0.022728
242.098 | -0.086667 | 212.5 | -0.8342 | 225.2 | -0.29566 | 263 | -0.05978 | 327 | -0.02226 | -0.02226
242.46 | -0.046667 | 212.6 | -0.82291 | 225.4 | -0.30964 | 263.5 | -0.060745 | 328 | -0.021257 | -0.021257
242.823 | -0.066667 | 212.7 | -0.81162 | 225.6 | -0.33485 | 264 | -0.058122 | 329 | -0.018031 | -0.018031
243.184 | -0.07 | 212.8 | -0.80033 | 225.8 | -0.32822 | 264.5 | -0.056281 | 330 | -0.019489 | -0.019489
243.546 | -0.07 | 212.9 | -0.81721 | 226 | -0.32318 | 265 | -0.055139 | 331 | -0.021891 | -0.021891
243.908 | -0.04 | 213 | -0.77972 | 226.2 | -0.31699 | 265.5 | -0.056126 | 332 | -0.020513 | -0.020513
244.27 | -0.05 | 213.1 | -0.74224 | 226.4 | -0.30485 | 266 | -0.062111 | 333 | -0.013164 | -0.013164
244.631 | -0.07 | 213.2 | -0.70475 | 226.6 | -0.29697 | 266.5 | -0.064317 | 334 | -0.0078947 | -0.0078947
244.993 | -0.05 | 213.3 | -0.71122 | 226.8 | -0.28347 | 267 | -0.055708 | 335 | -0.019286 | -0.019286
245.354 | -0.07 | 213.4 | -0.70986 | 227 | -0.25327 | 267.5 | -0.054687 | 336 | -0.020286 | -0.020286
245.716 | -0.063333 | 213.5 | -0.70851 | 227.2 | -0.2441 | 268 | -0.05412 | 337 | -0.019201 | -0.019201
246.077 | -0.053333 | 213.6 | -0.7083 | 227.4 | -0.23497 | 268.5 | -0.056958 | 338 | -0.01632 | -0.01632
246.438 | -0.063333 | 213.7 | -0.69656 | 227.6 | -0.22938 | 269 | -0.056913 | 339 | -0.015497 | -0.015497
246.799 | -0.06 | 213.8 | -0.68481 | 227.8 | -0.23 | 269.5 | -0.057988 | 340 | -0.016576 | -0.016576
247.16 | -0.043333 | 213.9 | -0.67307 | 228 | -0.22524 | 270 | -0.052947 | 341 | -0.012836 | -0.012836
247.521 | -0.09 | 214 | -0.65577 | 228.2 | -0.21509 | 270.5 | -0.059557 | 342 | -0.01217 | -0.01217
247.882 | -0.05 | 214.1 | -0.64718 | 228.4 | -0.19821 | 271 | -0.062199 | 343 | -0.011538 | -0.011538
248.243 | -0.046667 | 214.2 | -0.63859 | 228.6 | -0.18802 | 271.5 | -0.055146 | 344 | -0.010447 | -0.010447
248.603 | -0.06 | 214.3 | -0.63 | 228.8 | -0.18316 | 272 | -0.057446 | 345 | -0.01038 | -0.01038
248.964 | -0.073333 | 214.4 | -0.59795 | 229 | -0.17328 | 272.5 | -0.067106 | 346 | -0.009092 | -0.009092
249.324 | -0.063333 | 214.5 | -0.5785 | 229.2 | -0.16582 | 273 | -0.064378 | 347 | -0.0082108 | -0.0082108
249.685 | -0.06 | 214.6 | -0.55905 | 229.4 | -0.17113 | 273.5 | -0.063402 | 348 | -0.0032475 | -0.0032475
250.045 | -0.056667 | 214.7 | -0.5396 | 229.6 | -0.17653 | 274 | -0.058143 | 349 | -0.0020708 | -0.0020708
250.405 | -0.063333 | 214.8 | -0.64148 | 229.8 | -0.17568 | 274.5 | -0.05222 | 350 | -0.0027911 | -0.0027911
250.765 | -0.056667 | 214.9 | -0.65638 | 230 | -0.17823 | 275 | -0.053829 | 351 | -0.0018138 | -0.0018138
251.126 | -0.073333 | 215 | -0.67129 | 230.2 | -0.16974 | 275.5 | -0.064641 | 352 | 0.0014841 | 0.0014841
251.486 | -0.043333 | 215.1 | -0.6025 | 230.4 | -0.16166 | 276 | -0.058528 | 353 | 0.0072206 | 0.0072206
251.845 | -0.046667 | 215.2 | -0.60341 | 230.6 | -0.16495 | 276.5 | -0.06148 | 354 | 0.0098967 | 0.0098967
252.205 | -0.053333 | 215.3 | -0.60431 | 230.8 | -0.1601 | 277 | -0.061755 | 355 | 0.014876 | 0.014876
252.565 | -0.046667 | 215.4 | -0.60522 | 231 | -0.14949 | 277.5 | -0.06483 | 356 | 0.016394 | 0.016394
252.925 | -0.043333 | 215.5 | -0.61257 | 231.2 | -0.15771 | 278 | -0.057642 | 357 | 0.018334 | 0.018334
253.284 | -0.05 | 215.6 | -0.5954 | 231.4 | -0.16503 | 278.5 | -0.059525 | 358 | 0.023922 | 0.023922
253.644 | -0.053333 | 215.7 | -0.57823 | 231.6 | -0.15216 | 279 | -0.04926 | 359 | 0.028005 | 0.028005
254.003 | -0.05 | 215.8 | -0.56106 | 231.8 | -0.1423 | 279.5 | -0.048667 | 360 | 0.030307 | 0.030307
254.362 | -0.053333 | 215.9 | -0.67601 | 232 | -0.13407 | 280 | -0.052909 | 361 | 0.038046 | 0.038046
254.722 | -0.063333 | 216 | -0.70947 | 232.2 | -0.12019 | 280.5 | -0.05427 | 362 | 0.046735 | 0.046735
255.081 | -0.07 | 216.1 | -0.74293 | 232.4 | -0.11253 | 281 | -0.053673 | 363 | 0.067896 | 0.067896
255.44 | -0.046667 | 216.2 | -0.70266 | 232.6 | -0.10866 | 281.5 | -0.053065 | 364 | 0.77471 | 0.77471
255.799 | -0.05 | 216.3 | -0.71715 | 232.8 | -0.10734 | 282 | -0.054152 | 365 | 1.6376 | 1.6376
256.158 | -0.06 | 216.4 | -0.73164 | 233 | -0.10553 | 282.5 | -0.049807 | 366 | 0.67839 | 0.67839
256.517 | -0.063333 | 216.5 | -0.74614 | 233.2 | -0.10832 | 283 | -0.056736 | 367 | 0.1721 | 0.1721
256.875 | -0.04 | 216.6 | -0.7004 | 233.4 | -0.1111 | 283.5 | -0.061733 | 368 | 0.080078 | 0.080078
257.234 | -0.05 | 216.7 | -0.68002 | 233.6 | -0.11409 | 284 | -0.060559 | 369 | 0.084736 | 0.084736
257.593 | -0.07 | 216.8 | -0.65964 | 233.8 | -0.10908 | 284.5 | -0.059689 | 370 | 0.092869 | 0.092869
257.951 | -0.06 | 216.9 | -0.63926 | 234 | -0.10314 | 285 | -0.0551 | 371 | 0.10209 | 0.10209
258.31 | -0.05 | 217 | -0.60699 | 234.2 | -0.10594 | 285.5 | -0.052405 | 372 | 0.10923 | 0.10923
258.668 | -0.053333 | 217.1 | -0.58389 | 234.4 | -0.10635 | 286 | -0.054049 | 373 | 0.11469 | 0.11469
259.026 | -0.036667 | 217.2 | -0.56079 | 234.6 | -0.10391 | 286.5 | -0.050711 | 374 | 0.12284 | 0.12284
259.384 | -0.046667 | 217.3 | -0.62614 | 234.8 | -0.10371 | 287 | -0.048512 | 375 | 0.13047 | 0.13047
259.742 | -0.04 | 217.4 | -0.63384 | 235 | -0.099647 | 287.5 | -0.048116 | 376 | 0.13811 | 0.13811
260.1 | -0.06 | 217.5 | -0.64154 | 235.2 | -0.099181 | 288 | -0.049364 | 377 | 0.14469 | 0.14469
260.458 | -0.066667 | 217.6 | -0.64924 | 235.4 | -0.10695 | 288.5 | -0.043226 | 378 | 0.15522 | 0.15522
260.816 | -0.053333 | 217.7 | -0.59407 | 235.6 | -0.097432 | 289 | -0.038834 | 379 | 0.16623 | 0.16623
261.174 | -0.056667 | 217.8 | -0.59543 | 235.8 | -0.086936 | 289.5 | -0.035388 | 380 | 0.17367 | 0.17367
261.532 | -0.056667 | 217.9 | -0.59678 | 236 | -0.080161 | 290 | -0.043033 | 381 | 0.18572 | 0.18572
261.889 | -0.043333 | 218 | -0.59814 | 236.2 | -0.072733 | 290.5 | -0.046463 | 382 | 0.19919 | 0.19919
262.247 | -0.05 | 218.1 | -0.60608 | 236.4 | -0.073823 | 291 | -0.039364 | 383 | 0.20656 | 0.20656
262.604 | -0.053333 | 218.2 | -0.59383 | 236.6 | -0.075932 | 291.5 | -0.040162 | 384 | 0.2195 | 0.2195
262.961 | -0.053333 | 218.3 | -0.58157 | 236.8 | -0.081063 | 292 | -0.033905 | 385 | 0.23063 | 0.23063
263.319 | -0.073333 | 218.4 | -0.53797 | 237 | -0.087923 | 292.5 | -0.036882 | 386 | 0.24345 | 0.24345
263.676 | -0.05 | 218.5 | -0.52707 | 237.2 | -0.09935 | 293 | -0.040107 | 387 | 0.25409 | 0.25409
264.033 | -0.063333 | 218.6 | -0.51617 | 237.4 | -0.098553 | 293.5 | -0.039588 | 388 | 0.27101 | 0.27101
264.39 | -0.053333 | 218.7 | -0.50528 | 237.6 | -0.096553 | 294 | -0.038703 | 389 | 0.28729 | 0.28729
264.747 | -0.056667 | 218.8 | -0.53311 | 237.8 | -0.091723 | 294.5 | -0.035575 | 390 | 0.3107 | 0.3107
265.104 | -0.053333 | 218.9 | -0.52403 | 238 | -0.082879 | 295 | -0.032749 | 391 | 0.32642 | 0.32642
265.46 | -0.06 | 219 | -0.51495 | 238.2 | -0.077011 | 295.5 | -0.032125 | 392 | 0.33618 | 0.33618
265.817 | -0.053333 | 219.1 | -0.50586 | 238.4 | -0.075716 | 296 | -0.029395 | 393 | 0.35045 | 0.35045
266.174 | -0.07 | 219.2 | -0.46278 | 238.6 | -0.081749 | 296.5 | -0.030808 | 394 | 0.36876 | 0.36876
266.53 | -0.066667 | 219.3 | -0.45415 | 238.8 | -0.086116 | 297 | -0.027788 | 395 | 0.38786 | 0.38786
266.886 | -0.05 | 219.4 | -0.44552 | 239 | -0.095308 | 297.5 | -0.026761 | 396 | 0.40831 | 0.40831
267.243 | -0.056667 | 219.5 | -0.48259 | 239.2 | -0.088181 | 298 | -0.026677 | 397 | 0.43045 | 0.43045
267.599 | -0.056667 | 219.6 | -0.47804 | 239.4 | -0.076121 | 298.5 | -0.027863 | 398 | 0.45252 | 0.45252
267.955 | -0.046667 | 219.7 | -0.47349 | 239.6 | -0.076169 | 299 | -0.027219 | 399 | 0.47609 | 0.47609
268.311 | -0.063333 | 219.8 | -0.46894 | 239.8 | -0.073439 | 299.5 | -0.023973 | 400 | 0.49949 | 0.49949
268.667 | -0.053333 | 219.9 | -0.47512 | 240 | -0.070106 | 300 | -0.02497 | 401 | 0.53082 | 0.53082
269.023 | -0.056667 | 220 | -0.48784 | 240.2 | -0.070931 | 300.5 | -0.027612 | 402 | 0.56754 | 0.56754
269.379 | -0.06 | 220.1 | -0.50055 | 240.4 | -0.069026 | 301 | -0.029046 | 403 | 0.86902 | 0.86902
269.735 | -0.056667 | 220.2 | -0.51327 | 240.6 | -0.066848 | 301.5 | -0.032674 | 404 | 4.77 | 4.77
270.091 | -0.046667 | 220.3 | -0.51519 | 240.8 | -0.068471 | 302 | -0.033472 | 405 | 4.7699 | 4.7699
270.446 | -0.056667 | 220.4 | -0.51882 | 241 | -0.069711 | 302.5 | -0.03066 | 406 | 0.86813 | 0.86813
270.802 | -0.07 | 220.5 | -0.52245 | 241.2 | -0.071732 | 303 | -0.031595 | 407 | 0.93285 | 0.93285
271.157 | -0.056667 | 220.6 | -0.51124 | 241.4 | -0.070651 | 303.5 | -0.033579 | 408 | 1.0256 | 1.0256
271.512 | -0.056667 | 220.7 | -0.51215 | 241.6 | -0.068959 | 304 | -0.030819 | 409 | 0.77009 | 0.77009
271.868 | -0.05 | 220.8 | -0.51306 | 241.8 | -0.074603 | 304.5 | -0.028096 | 410 | 0.77669 | 0.77669
272.223 | -0.063333 | 220.9 | -0.51398 | 242 | -0.075202 | 305 | -0.028416 | 411 | 0.80097 | 0.80097
272.578 | -0.07 | 221 | -0.47432 | 242.2 | -0.063849 | 305.5 | -0.028501 | 412 | 0.82186 | 0.82186
272.933 | -0.07 | 221.1 | -0.45752 | 242.4 | -0.061662 | 306 | -0.025816 | 413 | 0.84593 | 0.84593
273.288 | -0.05 | 221.2 | -0.44072 | 242.6 | -0.062819 | 306.5 | -0.026677 | 414 | 0.86668 | 0.86668
273.643 | -0.073333 | 221.3 | -0.42392 | 242.8 | -0.06033 | 307 | -0.026614 | 415 | 0.88431 | 0.88431
273.997 | -0.05 | 221.4 | -0.43241 | 243 | -0.066841 | 307.5 | -0.023793 | 416 | 0.90693 | 0.90693
274.352 | -0.06 | 221.5 | -0.42015 | 243.2 | -0.069483 | 308 | -0.022385 | 417 | 0.92814 | 0.92814
274.707 | -0.046667 | 221.6 | -0.40789 | 243.4 | -0.069883 | 308.5 | -0.024599 | 418 | 0.95219 | 0.95219
275.061 | -0.056667 | 221.7 | -0.43103 | 243.6 | -0.060575 | 309 | -0.023784 | 419 | 0.97882 | 0.97882
275.416 | -0.056667 | 221.8 | -0.43876 | 243.8 | -0.049475 | 309.5 | -0.022902 | 420 | 1.0028 | 1.0028
275.77 | -0.076667 | 221.9 | -0.44649 | 244 | -0.048051 | 310 | -0.021261 | 421 | 1.0185 | 1.0185
276.124 | -0.043333 | 222 | -0.45421 | 244.2 | -0.049296 | 310.5 | -0.020187 | 422 | 1.0381 | 1.0381
276.478 | -0.07 | 222.1 | -0.48272 | 244.4 | -0.058741 | 311 | -0.021825 | 423 | 1.068 | 1.068
276.833 | -0.063333 | 222.2 | -0.50046 | 244.6 | -0.062197 | 311.5 | -0.026949 | 424 | 1.1046 | 1.1046
277.187 | -0.053333 | 222.3 | -0.5182 | 244.8 | -0.056662 | 312 | -0.021837 | 425 | 1.125 | 1.125
277.541 | -0.076667 | 222.4 | -0.53593 | 245 | -0.061108 | 312.5 | -0.019672 | 426 | 1.1297 | 1.1297
277.894 | -0.05 | 222.5 | -0.47748 | 245.2 | -0.063328 | 313 | -0.01715 | 427 | 1.1382 | 1.1382
278.248 | -0.056667 | 222.6 | -0.47294 | 245.4 | -0.063024 | 313.5 | -0.021919 | 428 | 1.1745 | 1.1745
278.602 | -0.063333 | 222.7 | -0.4684 | 245.6 | -0.065637 | 314 | -0.021184 | 429 | 1.2044 | 1.2044
278.955 | -0.05 | 222.8 | -0.42816 | 245.8 | -0.061911 | 314.5 | -0.02019 | 430 | 1.2466 | 1.2466
279.309 | -0.04 | 222.9 | -0.39769 | 246 | -0.057881 | 315 | -0.023194 | 431 | 1.2929 | 1.2929
279.662 | -0.056667 | 223 | -0.36722 | 246.2 | -0.06 | 315.5 | -0.020702 | 432 | 1.2979 | 1.2979
280.016 | -0.05 | 223.1 | -0.33676 | 246.4 | -0.05982 | 316 | -0.02471 | 433 | 1.3537 | 1.3537
280.369 | -0.056667 | 223.2 | -0.41012 | 246.6 | -0.060385 | 316.5 | -0.027536 | 434 | 1.7557 | 1.7557
280.722 | -0.053333 | 223.3 | -0.41514 | 246.8 | -0.056531 | 317 | -0.023427 | 435 | 7.9077 | 7.9077
281.075 | -0.05 | 223.4 | -0.42016 | 247 | -0.049988 | 317.5 | -0.019956 | 436 | 11.4582 | 11.4582
281.428 | -0.06 | 223.5 | -0.42518 | 247.2 | -0.061195 | 318 | -0.017689 | 437 | 3.0056 | 3.0056
281.781 | -0.046667 | 223.6 | -0.41283 | 247.4 | -0.074417 | 318.5 | -0.023546 | 438 | 1.4781 | 1.4781
282.134 | -0.063333 | 223.7 | -0.42603 | 247.6 | -0.067725 | 319 | -0.023183 | 439 | 1.5105 | 1.5105
282.487 | -0.033333 | 223.8 | -0.43923 | 247.8 | -0.062531 | 319.5 | -0.021445 | 440 | 1.5382 | 1.5382
282.84 | -0.06 | 223.9 | -0.43462 | 248 | -0.05514 | 320 | -0.020611 | 441 | 1.5526 | 1.5526
283.192 | -0.06 | 224 | -0.42781 | 248.2 | -0.04843 | 320.5 | -0.021364 | 442 | 1.5619 | 1.5619
283.545 | -0.063333 | 224.1 | -0.42099 | 248.4 | -0.054403 | 321 | -0.019218 | 443 | 1.5875 | 1.5875
283.897 | -0.06 | 224.2 | -0.41417 | 248.6 | -0.058846 | 321.5 | -0.018198 | 444 | 1.6078 | 1.6078
284.25 | -0.06 | 224.3 | -0.40186 | 248.8 | -0.067274 | 322 | -0.021428 | 445 | 1.6114 | 1.6114
284.602 | -0.06 | 224.4 | -0.39641 | 249 | -0.067628 | 322.5 | -0.019974 | 446 | 1.6134 | 1.6134
284.954 | -0.056667 | 224.5 | -0.39095 | 249.2 | -0.066652 | 323 | -0.020354 | 447 | 1.6123 | 1.6123
285.306 | -0.05 | 224.6 | -0.3855 | 249.4 | -0.064527 | 323.5 | -0.023966 | 448 | 1.6251 | 1.6251
285.658 | -0.053333 | 224.7 | -0.3508 | 249.6 | -0.060875 | 324 | -0.026334 | 449 | 1.6409 | 1.6409
286.01 | -0.056667 | 224.8 | -0.33395 | 249.8 | -0.058934 | 324.5 | -0.023814 | 450 | 1.654 | 1.654
286.362 | -0.05 | 224.9 | -0.3171 | 250 | -0.058449 | 325 | -0.023369 | 451 | 1.6739 | 1.6739
286.714 | -0.05 | 225 | -0.31934 | 250.2 | -0.060718 | 325.5 | -0.021366 | 452 | 1.6848 | 1.6848
287.066 | -0.046667 | 225.1 | -0.3075 | 250.4 | -0.060571 | 326 | -0.025085 | 453 | 1.6945 | 1.6945
287.417 | -0.05 | 225.2 | -0.29566 | 250.6 | -0.058889 | 326.5 | -0.021735 | 454 | 1.7102 | 1.7102
287.769 | -0.046667 | 225.3 | -0.28382 | 250.8 | -0.063376 | 327 | -0.020071 | 455 | 1.7253 | 1.7253
288.12 | -0.053333 | 225.4 | -0.31846 | 251 | -0.067703 | 327.5 | -0.024973 | 456 | 1.7377 | 1.7377
288.472 | -0.04 | 225.5 | -0.32666 | 251.2 | -0.060019 | 328 | -0.022141 | 457 | 1.7545 | 1.7545
288.823 | -0.04 | 225.6 | -0.33485 | 251.4 | -0.053728 | 328.5 | -0.016657 | 458 | 1.7674 | 1.7674
289.174 | -0.04 | 225.7 | -0.34305 | 251.6 | -0.050202 | 329 | -0.018767 | 459 | 1.7771 | 1.7771
289.526 | -0.03 | 225.8 | -0.31967 | 251.8 | -0.045936 | 329.5 | -0.018669 | 460 | 1.7919 | 1.7919
289.877 | -0.043333 | 225.9 | -0.32194 | 252 | -0.049929 | 330 | -0.019798 | 461 | 1.8002 | 1.8002
290.228 | -0.046667 | 226 | -0.32422 | 252.2 | -0.050973 | 330.5 | -0.02 | 462 | 1.8107 | 1.8107
290.579 | -0.05 | 226.1 | -0.32336 | 252.4 | -0.048889 | 331 | -0.024044 | 463 | 1.8245 | 1.8245
290.929 | -0.036667 | 226.2 | -0.31699 | 252.6 | -0.047361 | 331.5 | -0.02163 | 464 | 1.8357 | 1.8357
291.28 | -0.036667 | 226.3 | -0.31061 | 252.8 | -0.044514 | 332 | -0.019909 | 465 | 1.8469 | 1.8469
291.631 | -0.043333 | 226.4 | -0.30423 | 253 | -0.045873 | 332.5 | -0.02 | 466 | 1.8567 | 1.8567
291.981 | -0.036667 | 226.5 | -0.29971 | 253.2 | -0.048175 | 333 | -0.014885 | 467 | 1.8676 | 1.8676
292.332 | -0.023333 | 226.6 | -0.29697 | 253.4 | -0.050497 | 333.5 | -0.0046056 | 468 | 1.879 | 1.879
292.682 | -0.05 | 226.7 | -0.29424 | 253.6 | -0.052093 | 334 | -0.0036379 | 469 | 1.8835 | 1.8835
293.033 | -0.036667 | 226.8 | -0.2915 | 253.8 | -0.051112 | 334.5 | -0.015441 | 470 | 1.8871 | 1.8871
293.383 | -0.04 | 226.9 | -0.26468 | 254 | -0.051482 | 335 | -0.021541 | 471 | 1.8923 | 1.8923
293.733 | -0.04 | 227 | -0.25327 | 254.2 | -0.052222 | 335.5 | -0.020876 | 472 | 1.9032 | 1.9032
294.083 | -0.036667 | 227.1 | -0.24187 | 254.4 | -0.055529 | 336 | -0.019982 | 473 | 1.9067 | 1.9067
294.433 | -0.04 | 227.2 | -0.24863 | 254.6 | -0.059964 | 336.5 | -0.02 | 474 | 1.9093 | 1.9093
294.783 | -0.03 | 227.3 | -0.2418 | 254.8 | -0.064071 | 337 | -0.018595 | 475 | 1.9179 | 1.9179
295.133 | -0.033333 | 227.4 | -0.23497 | 255 | -0.064863 | 337.5 | -0.019007 | 476 | 1.9212 | 1.9212
295.483 | -0.033333 | 227.5 | -0.22814 | 255.2 | -0.057238 | 338 | -0.016806 | 477 | 1.9226 | 1.9226
295.832 | -0.03 | 227.6 | -0.23 | 255.4 | -0.053799 | 338.5 | -0.013147 | 478 | 1.9284 | 1.9284
296.182 | -0.026667 | 227.7 | -0.23 | 255.6 | -0.051099 | 339 | -0.017983 | 479 | 1.9314 | 1.9314
296.531 | -0.033333 | 227.8 | -0.23 | 255.8 | -0.051551 | 339.5 | -0.015362 | 480 | 1.9304 | 1.9304
296.881 | -0.03 | 227.9 | -0.23 | 256 | -0.055955 | 340 | -0.018036 | 481 | 1.9352 | 1.9352
297.23 | -0.023333 | 228 | -0.22423 | 256.2 | -0.058928 | 340.5 | -0.016332 | 482 | 1.9395 | 1.9395
297.579 | -0.03 | 228.1 | -0.22149 | 256.4 | -0.062272 | 341 | -0.0094505 | 483 | 1.9478 | 1.9478
297.929 | -0.023333 | 228.2 | -0.21875 | 256.6 | -0.055195 | 341.5 | -0.012726 | 484 | 1.9566 | 1.9566
298.278 | -0.03 | 228.3 | -0.20505 | 256.8 | -0.048844 | 342 | -0.010392 | 485 | 1.968 | 1.968
298.627 | -0.026667 | 228.4 | -0.19821 | 257 | -0.048796 | 342.5 | -0.01339 | 486 | 1.9803 | 1.9803
298.976 | -0.03 | 228.5 | -0.19136 | 257.2 | -0.04937 | 343 | -0.011328 | 487 | 1.9825 | 1.9825
299.325 | -0.023333 | 228.6 | -0.18452 | 257.4 | -0.060269 | 343.5 | -0.0098947 | 488 | 1.9761 | 1.9761
299.673 | -0.023333 | 228.7 | -0.18818 | 257.6 | -0.062016 | 344 | -0.01 | 489 | 1.9683 | 1.9683
300.022 | -0.023333 | 228.8 | -0.18316 | 257.8 | -0.062896 | 344.5 | -0.011447 | 490 | 1.9655 | 1.9655
300.371 | -0.03 | 228.9 | -0.17815 | 258 | -0.059596 | 345 | -0.0098814 | 491 | 1.984 | 1.984
300.719 | -0.026667 | 229 | -0.17313 | 258.2 | -0.053064 | 345.5 | -0.0098124 | 492 | 1.9851 | 1.9851
301.068 | -0.03 | 229.1 | -0.16856 | 258.4 | -0.052184 | 346 | -0.0090452 | 493 | 1.9364 | 1.9364
301.416 | -0.03 | 229.2 | -0.16582 | 258.6 | -0.050162 | 346.5 | -0.0084184 | 494 | 1.9244 | 1.9244
301.764 | -0.036667 | 229.3 | -0.16308 | 258.8 | -0.044209 | 347 | -0.010755 | 495 | 1.9132 | 1.9132
302.113 | -0.033333 | 229.4 | -0.1747 | 259 | -0.042566 | 347.5 | -0.0054588 | 496 | 1.9029 | 1.9029
302.461 | -0.026667 | 229.5 | -0.17561 | 259.2 | -0.043935 | 348 | -0.0025516 | 497 | 1.8934 | 1.8934
302.809 | -0.033333 | 229.6 | -0.17653 | 259.4 | -0.041947 | 348.5 | -0.0017319 | 498 | 1.8829 | 1.8829
303.157 | -0.03 | 229.7 | -0.17744 | 259.6 | -0.042117 | 349 | -0.0010413 | 499 | 1.8684 | 1.8684
303.505 | -0.036667 | 229.8 | -0.17366 | 259.8 | -0.048127 | 349.5 | -0.003439 | 500 | 1.8592 | 1.8592
303.853 | -0.03 | 229.9 | -0.17594 | 260 | -0.053693 | 350 | -0.00080078 | 501 | 1.8472 | 1.8472
304.2 | -0.03 | 230 | -0.17823 | 260.2 | -0.05928 | 350.5 | -0.0041334 | 502 | 1.8329 | 1.8329
304.548 | -0.03 | 230.1 | -0.18051 | 260.4 | -0.063114 | 351 | -0.0018644 | 503 | 1.8265 | 1.8265
304.895 | -0.02 | 230.2 | -0.16664 | 260.6 | -0.058678 | 351.5 | 0.00055652 | 504 | 1.8164 | 1.8164
305.243 | -0.036667 | 230.3 | -0.16207 | 260.8 | -0.057426 | 352 | -3.1505e-05 | 505 | 1.8048 | 1.8048
305.59 | -0.023333 | 230.4 | -0.1575 | 261 | -0.056319 | 352.5 | 0.0039272 | 506 | 1.7952 | 1.7952
305.938 | -0.03 | 230.5 | -0.1654 | 261.2 | -0.05558 | 353 | 0.0090853 | 507 | 1.7853 | 1.7853
306.285 | -0.02 | 230.6 | -0.16495 | 261.4 | -0.056608 | 353.5 | 0.0086494 | 508 | 1.7818 | 1.7818
306.632 | -0.03 | 230.7 | -0.16449 | 261.6 | -0.052369 | 354 | 0.0097761 | 509 | 1.7759 | 1.7759
306.979 | -0.03 | 230.8 | -0.16404 | 261.8 | -0.048735 | 354.5 | 0.011265 | 510 | 1.7663 | 1.7663
307.326 | -0.02 | 230.9 | -0.15177 | 262 | -0.04897 | 355 | 0.01709 | 511 | 1.7551 | 1.7551
307.673 | -0.026667 | 231 | -0.14949 | 262.2 | -0.04826 | 355.5 | 0.016274 | 512 | 1.7461 | 1.7461
308.02 | -0.02 | 231.1 | -0.14721 | 262.4 | -0.051034 | 356 | 0.016678 | 513 | 1.7404 | 1.7404
308.367 | -0.023333 | 231.2 | -0.16227 | 262.6 | -0.052978 | 356.5 | 0.016231 | 514 | 1.7298 | 1.7298
308.713 | -0.026667 | 231.3 | -0.16365 | 262.8 | -0.053137 | 357 | 0.018219 | 515 | 1.725 | 1.725
309.06 | -0.02 | 231.4 | -0.16503 | 263 | -0.059522 | 357.5 | 0.020552 | 516 | 1.7222 | 1.7222
309.407 | -0.026667 | 231.5 | -0.1664 | 263.2 | -0.066679 | 358 | 0.023507 | 517 | 1.7155 | 1.7155
309.753 | -0.02 | 231.6 | -0.14595 | 263.4 | -0.062013 | 358.5 | 0.027707 | 518 | 1.7098 | 1.7098
310.099 | -0.02 | 231.7 | -0.14413 | 263.6 | -0.059173 | 359 | 0.029734 | 519 | 1.7062 | 1.7062
310.446 | -0.023333 | 231.8 | -0.1423 | 263.8 | -0.060486 | 359.5 | 0.026576 | 520 | 1.7051 | 1.7051
310.792 | -0.016667 | 231.9 | -0.14047 | 264 | -0.057546 | 360 | 0.027192 | 521 | 1.6963 | 1.6963
311.138 | -0.023333 | 232 | -0.13499 | 264.2 | -0.056335 | 360.5 | 0.037154 | 522 | 1.6958 | 1.6958
311.484 | -0.03 | 232.1 | -0.12676 | 264.4 | -0.057079 | 361 | 0.036022 | 523 | 1.6894 | 1.6894
311.83 | -0.026667 | 232.2 | -0.11852 | 264.6 | -0.055817 | 361.5 | 0.040962 | 524 | 1.6912 | 1.6912
312.176 | -0.013333 | 232.3 | -0.11528 | 264.8 | -0.054591 | 362 | 0.049219 | 525 | 1.695 | 1.695
312.522 | -0.026667 | 232.4 | -0.11253 | 265 | -0.054444 | 362.5 | 0.050025 | 526 | 1.6969 | 1.6969
312.867 | -0.01 | 232.5 | -0.10979 | 265.2 | -0.056381 | 363 | 0.05727 | 527 | 1.6966 | 1.6966
313.213 | -0.02 | 232.6 | -0.10704 | 265.4 | -0.057063 | 363.5 | 0.096394 | 528 | 1.6993 | 1.6993
313.558 | -0.023333 | 232.7 | -0.10917 | 265.6 | -0.055554 | 364 | 0.51282 | 529 | 1.704 | 1.704
313.904 | -0.023333 | 232.8 | -0.10734 | 265.8 | -0.057793 | 364.5 | 1.7149 | 530 | 1.7093 | 1.7093
314.249 | -0.016667 | 232.9 | -0.10551 | 266 | -0.063674 | 365 | 2.0728 | 531 | 1.7136 | 1.7136
314.595 | -0.02 | 233 | -0.10368 | 266.2 | -0.064866 | 365.5 | 1.1252 | 532 | 1.7204 | 1.7204
314.94 | -0.026667 | 233.1 | -0.10741 | 266.4 | -0.067569 | 366 | 0.58496 | 533 | 1.7311 | 1.7311
315.285 | -0.02 | 233.2 | -0.10832 | 266.6 | -0.062381 | 366.5 | 0.325 | 534 | 1.7414 | 1.7414
315.63 | -0.02 | 233.3 | -0.10924 | 266.8 | -0.056557 | 367 | 0.11366 | 535 | 1.7472 | 1.7472
315.975 | -0.023333 | 233.4 | -0.11134 | 267 | -0.056185 | 367.5 | 0.077625 | 536 | 1.7632 | 1.7632
316.32 | -0.03 | 233.5 | -0.11272 | 267.2 | -0.054383 | 368 | 0.079268 | 537 | 1.7885 | 1.7885
316.665 | -0.026667 | 233.6 | -0.11409 | 267.4 | -0.055916 | 368.5 | 0.083342 | 538 | 1.8224 | 1.8224
317.01 | -0.023333 | 233.7 | -0.11547 | 267.6 | -0.054029 | 369 | 0.083886 | 539 | 1.8577 | 1.8577
317.354 | -0.02 | 233.8 | -0.10681 | 267.8 | -0.05051 | 369.5 | 0.08698 | 540 | 1.9096 | 1.9096
317.699 | -0.02 | 233.9 | -0.10498 | 268 | -0.053999 | 370 | 0.094701 | 541 | 2.0143 | 2.0143
318.043 | -0.013333 | 234 | -0.10314 | 268.2 | -0.05785 | 370.5 | 0.096927 | 542 | 2.147 | 2.147
318.388 | -0.023333 | 234.1 | -0.10131 | 268.4 | -0.056759 | 371 | 0.1018 | 543 | 2.1825 | 2.1825
318.732 | -0.026667 | 234.2 | -0.10848 | 268.6 | -0.056944 | 371.5 | 0.10753 | 544 | 2.2216 | 2.2216
319.076 | -0.02 | 234.3 | -0.10802 | 268.8 | -0.056532 | 372 | 0.11017 | 545 | 3.9614 | 3.9614
319.42 | -0.023333 | 234.4 | -0.10757 | 269 | -0.055881 | 372.5 | 0.11 | 546 | 10.4534 | 10.4534
319.765 | -0.02 | 234.5 | -0.10345 | 269.2 | -0.058324 | 373 | 0.1146 | 547 | 5.9744 | 5.9744
320.109 | -0.02 | 234.6 | -0.10391 | 269.4 | -0.058272 | 373.5 | 0.11947 | 548 | 2.1649 | 2.1649
320.453 | -0.023333 | 234.7 | -0.10436 | 269.6 | -0.057778 | 374 | 0.12339 | 549 | 2.1356 | 2.1356
320.796 | -0.02 | 234.8 | -0.10482 | 269.8 | -0.05412 | 374.5 | 0.12566 | 550 | 2.1458 | 2.1458
321.14 | -0.02 | 234.9 | -0.10194 | 270 | -0.05139 | 375 | 0.13066 | 551 | 2.156 | 2.156
321.484 | -0.01 | 235 | -0.099647 | 270.2 | -0.053332 | 375.5 | 0.13511 | 552 | 2.1737 | 2.1737
321.828 | -0.026667 | 235.1 | -0.097358 | 270.4 | -0.0554 | 376 | 0.13753 | 553 | 2.1837 | 2.1837
322.171 | -0.02 | 235.2 | -0.095069 | 270.6 | -0.062821 | 376.5 | 0.1417 | 554 | 2.1903 | 2.1903
322.515 | -0.02 | 235.3 | -0.10512 | 270.8 | -0.065534 | 377 | 0.14427 | 555 | 2.1926 | 2.1926
322.858 | -0.02 | 235.4 | -0.10695 | 271 | -0.061115 | 377.5 | 0.14809 | 556 | 2.183 | 2.183
323.201 | -0.02 | 235.5 | -0.10878 | 271.2 | -0.059949 | 378 | 0.15541 | 557 | 2.1875 | 2.1875
323.544 | -0.026667 | 235.6 | -0.093363 | 271.4 | -0.056548 | 378.5 | 0.16216 | 558 | 2.2101 | 2.2101
323.888 | -0.023333 | 235.7 | -0.09015 | 271.6 | -0.053659 | 379 | 0.16656 | 559 | 2.222 | 2.222
324.231 | -0.03 | 235.8 | -0.086936 | 271.8 | -0.0538 | 379.5 | 0.16998 | 560 | 2.2362 | 2.2362
324.574 | -0.02 | 235.9 | -0.083723 | 272 | -0.057906 | 380 | 0.17341 | 561 | 2.2569 | 2.2569
324.917 | -0.023333 | 236 | -0.081362 | 272.2 | -0.060634 | 380.5 | 0.17763 | 562 | 2.2768 | 2.2768
325.259 | -0.026667 | 236.1 | -0.075399 | 272.4 | -0.066097 | 381 | 0.18521 | 563 | 2.2875 | 2.2875
325.602 | -0.016667 | 236.2 | -0.069436 | 272.6 | -0.068607 | 381.5 | 0.19432 | 564 | 2.3029 | 2.3029
325.945 | -0.03 | 236.3 | -0.073364 | 272.8 | -0.069862 | 382 | 0.20152 | 565 | 2.317 | 2.317
326.287 | -0.023333 | 236.4 | -0.073823 | 273 | -0.063625 | 382.5 | 0.20173 | 566 | 2.3281 | 2.3281
326.63 | -0.02 | 236.5 | -0.074282 | 273.2 | -0.059647 | 383 | 0.20388 | 567 | 2.3418 | 2.3418
326.972 | -0.02 | 236.6 | -0.074741 | 273.4 | -0.06497 | 383.5 | 0.21407 | 568 | 2.353 | 2.353
327.315 | -0.02 | 236.7 | -0.078771 | 273.6 | -0.063043 | 384 | 0.21972 | 569 | 2.3587 | 2.3587
327.657 | -0.03 | 236.8 | -0.081063 | 273.8 | -0.057783 | 384.5 | 0.22471 | 570 | 2.3688 | 2.3688
327.999 | -0.026667 | 236.9 | -0.083356 | 274 | -0.059343 | 385 | 0.22953 | 571 | 2.3771 | 2.3771
328.341 | -0.01 | 237 | -0.085648 | 274.2 | -0.057304 | 385.5 | 0.23765 | 572 | 2.3829 | 2.3829
328.683 | -0.02 | 237.1 | -0.094765 | 274.4 | -0.052983 | 386 | 0.24372 | 573 | 2.3912 | 2.3912
329.025 | -0.02 | 237.2 | -0.09935 | 274.6 | -0.051058 | 386.5 | 0.24896 | 574 | 2.3926 | 2.3926
329.367 | -0.016667 | 237.3 | -0.10394 | 274.8 | -0.052708 | 387 | 0.25306 | 575 | 2.3982 | 2.3982
329.709 | -0.02 | 237.4 | -0.095631 | 275 | -0.053491 | 387.5 | 0.26024 | 576 | 2.7556 | 2.7556
330.05 | -0.02 | 237.5 | -0.096092 | 275.2 | -0.055288 | 388 | 0.27239 | 577 | 3.7379 | 3.7379
330.392 | -0.02 | 237.6 | -0.096553 | 275.4 | -0.060172 | 388.5 | 0.28038 | 578 | 3.0972 | 3.0972
330.734 | -0.02 | 237.7 | -0.097015 | 275.6 | -0.068531 | 389 | 0.28647 | 579 | 3.6404 | 3.6404
331.075 | -0.03 | 237.8 | -0.091143 | 275.8 | -0.062038 | 389.5 | 0.29501 | 580 | 2.9739 | 2.9739
331.417 | -0.02 | 237.9 | -0.087011 | 276 | -0.054557 | 390 | 0.31339 | 581 | 2.4313 | 2.4313
331.758 | -0.02 | 238 | -0.082879 | 276.2 | -0.058989 | 390.5 | 0.32369 | 582 | 2.4091 | 2.4091
332.099 | -0.02 | 238.1 | -0.078747 | 276.4 | -0.060637 | 391 | 0.32704 | 583 | 2.4107 | 2.4107
332.44 | -0.02 | 238.2 | -0.077521 | 276.6 | -0.062318 | 391.5 | 0.32854 | 584 | 2.4006 | 2.4006
332.781 | -0.02 | 238.3 | -0.074766 | 276.8 | -0.064577 | 392 | 0.33734 | 585 | 2.3781 | 2.3781
333.122 | -0.013333 | 238.4 | -0.072011 | 277 | -0.058289 | 392.5 | 0.34268 | 586 | 2.3612 | 2.3612
333.463 | -0.0033333 | 238.5 | -0.080372 | 277.2 | -0.0624 | 393 | 0.3502 | 587 | 2.3897 | 2.3897
333.804 | 0 | 238.6 | -0.081749 | 277.4 | -0.068456 | 393.5 | 0.35846 | 588 | 2.363 | 2.363
334.145 | -0.0033333 | 238.7 | -0.083127 | 277.6 | -0.063105 | 394 | 0.3681 | 589 | 2.2827 | 2.2827
334.485 | -0.02 | 238.8 | -0.084504 | 277.8 | -0.059664 | 394.5 | 0.37973 | 590 | 2.2536 | 2.2536
334.826 | -0.02 | 238.9 | -0.090716 | 278 | -0.058124 | 395 | 0.38718 | 591 | 2.23 | 2.23
335.167 | -0.023333 | 239 | -0.095308 | 278.2 | -0.055138 | 395.5 | 0.39667 | 592 | 2.2204 | 2.2204
335.507 | -0.02 | 239.1 | -0.099899 | 278.4 | -0.059529 | 396 | 0.40708 | 593 | 2.2308 | 2.2308
335.847 | -0.02 | 239.2 | -0.084389 | 278.6 | -0.06015 | 396.5 | 0.42119 | 594 | 2.1782 | 2.1782
336.188 | -0.02 | 239.3 | -0.080255 | 278.8 | -0.054799 | 397 | 0.43177 | 595 | 2.1133 | 2.1133
336.528 | -0.02 | 239.4 | -0.076121 | 279 | -0.049944 | 397.5 | 0.4384 | 596 | 2.0789 | 2.0789
336.868 | -0.02 | 239.5 | -0.071986 | 279.2 | -0.043037 | 398 | 0.45272 | 597 | 2.0443 | 2.0443
337.208 | -0.016667 | 239.6 | -0.079868 | 279.4 | -0.046725 | 398.5 | 0.46644 | 598 | 2.0222 | 2.0222
337.548 | -0.02 | 239.7 | -0.076654 | 279.6 | -0.050581 | 399 | 0.47632 | 599 | 2.0161 | 2.0161
337.888 | -0.02 | 239.8 | -0.073439 | 279.8 | -0.050834 | 399.5 | 0.4855 | 600 | 1.9877 | 1.9877
338.228 | -0.013333 | 239.9 | -0.070225 | 280 | -0.05345 | 400 | 0.49934 | 601 | 1.9378 | 1.9378
338.567 | -0.01 | 240 | -0.069587 | 280.2 | -0.054443 | 400.5 | 0.51364 | 602 | 1.8888 | 1.8888
338.907 | -0.02 | 240.1 | -0.070505 | 280.4 | -0.053958 | 401 | 0.5309 | 603 | 1.8541 | 1.8541
339.246 | -0.02 | 240.2 | -0.071423 | 280.6 | -0.054424 | 401.5 | 0.54792 | 604 | 1.8241 | 1.8241
339.586 | -0.01 | 240.3 | -0.070866 | 280.8 | -0.053048 | 402 | 0.56318 | 605 | 1.7959 | 1.7959
339.925 | -0.02 | 240.4 | -0.069026 | 281 | -0.052347 | 402.5 | 0.5915 | 606 | 1.7727 | 1.7727
340.265 | -0.02 | 240.5 | -0.067186 | 281.2 | -0.055625 | 403 | 0.66296 | 607 | 1.7459 | 1.7459
340.604 | -0.016667 | 240.6 | -0.065346 | 281.4 | -0.055322 | 403.5 | 1.3526 | 608 | 1.7268 | 1.7268
340.943 | 0 | 240.7 | -0.068011 | 281.6 | -0.05141 | 404 | 4.6233 | 609 | 1.7361 | 1.7361
341.282 | -0.02 | 240.8 | -0.068471 | 281.8 | -0.054974 | 404.5 | 8.3342 | 610 | 1.8941 | 1.8941
341.621 | -0.01 | 240.9 | -0.068931 | 282 | -0.057701 | 405 | 4.7756 | 611 | 2.4701 | 2.4701
341.96 | -0.01 | 241 | -0.069391 | 282.2 | -0.049782 | 405.5 | 1.2 | 612 | 2.3942 | 2.3942
342.299 | -0.01 | 241.1 | -0.070812 | 282.4 | -0.046519 | 406 | 0.68844 | 613 | 1.8961 | 1.8961
342.638 | -0.016667 | 241.2 | -0.071732 | 282.6 | -0.051689 | 406.5 | 0.71599 | 614 | 1.766 | 1.766
342.976 | -0.01 | 241.3 | -0.072651 | 282.8 | -0.051785 | 407 | 0.88954 | 615 | 1.6763 | 1.6763
343.315 | -0.01 | 241.4 | -0.06988 | 283 | -0.057179 | 407.5 | 1.193 | 616 | 1.5523 | 1.5523
343.654 | -0.01 | 241.5 | -0.06942 | 283.2 | -0.061244 | 408 | 1.0913 | 617 | 1.4699 | 1.4699
343.992 | -0.01 | 241.6 | -0.068959 | 283.4 | -0.062093 | 408.5 | 0.79261 | 618 | 1.4217 | 1.4217
344.33 | -0.01 | 241.7 | -0.068499 | 283.6 | -0.061597 | 409 | 0.75438 | 619 | 1.3908 | 1.3908
344.669 | -0.013333 | 241.8 | -0.076734 | 283.8 | -0.061108 | 409.5 | 0.76326 | 620 | 1.3622 | 1.3622
345.007 | -0.01 | 241.9 | -0.078576 | 284 | -0.060626 | 410 | 0.7763 | 621 | 1.3309 | 1.3309
345.345 | -0.0066667 | 242 | -0.080417 | 284.2 | -0.059945 | 410.5 | 0.79052 | 622 | 1.2923 | 1.2923
345.683 | -0.013333 | 242.1 | -0.066611 | 284.4 | -0.06 | 411 | 0.80285 | 623 | 1.2662 | 1.2662
346.021 | -0.0066667 | 242.2 | -0.063849 | 284.6 | -0.059475 | 411.5 | 0.80955 | 624 | 1.2404 | 1.2404
346.359 | -0.01 | 242.3 | -0.061087 | 284.8 | -0.057951 | 412 | 0.82151 | 625 | 1.2086 | 1.2086
346.697 | -0.0066667 | 242.4 | -0.058324 | 285 | -0.055287 | 412.5 | 0.83451 | 626 | 1.1729 | 1.1729
347.035 | -0.013333 | 242.5 | -0.065574 | 285.2 | -0.052061 | 413 | 0.84484 | 627 | 1.1452 | 1.1452
347.372 | -0.01 | 242.6 | -0.062819 | 285.4 | -0.051981 | 413.5 | 0.85845 | 628 | 1.1118 | 1.1118
347.71 | 0 | 242.7 | -0.060064 | 285.6 | -0.052696 | 414 | 0.86821 | 629 | 1.0902 | 1.0902
348.047 | -0.0033333 | 242.8 | -0.057309 | 285.8 | -0.054678 | 414.5 | 0.87338 | 630 | 1.0867 | 1.0867
348.385 | -0.0033333 | 242.9 | -0.063616 | 286 | -0.055035 | 415 | 0.88261 | 631 | 1.1123 | 1.1123
348.722 | 0 | 243 | -0.066841 | 286.2 | -0.052434 | 415.5 | 0.89694 | 632 | 1.0339 | 1.0339
349.059 | 0.0033333 | 243.1 | -0.070066 | 286.4 | -0.051579 | 416 | 0.90569 | 633 | 0.94902 | 0.94902
349.397 | -0.01 | 243.2 | -0.068961 | 286.6 | -0.049968 | 416.5 | 0.91816 | 634 | 0.91727 | 0.91727
349.734 | 0 | 243.3 | -0.069422 | 286.8 | -0.048504 | 417 | 0.92874 | 635 | 0.89751 | 0.89751
350.071 | 0 | 243.4 | -0.069883 | 287 | -0.048144 | 417.5 | 0.93752 | 636 | 0.87122 | 0.87122
350.408 | 0 | 243.5 | -0.070344 | 287.2 | -0.048888 | 418 | 0.95429 | 637 | 0.84468 | 0.84468
350.744 | -0.01 | 243.6 | -0.057762 | 287.4 | -0.048518 | 418.5 | 0.96475 | 638 | 0.81712 | 0.81712
351.081 | 0.0033333 | 243.7 | -0.053619 | 287.6 | -0.047777 | 419 | 0.97835 | 639 | 0.79712 | 0.79712
351.418 | 0 | 243.8 | -0.049475 | 287.8 | -0.049516 | 419.5 | 0.99336 | 640 | 0.7747 | 0.7747
351.755 | 0 | 243.9 | -0.045331 | 288 | -0.051095 | 420 | 1.005 | 641 | 0.75386 | 0.75386
352.091 | 0 | 244 | -0.050792 | 288.2 | -0.04748 | 420.5 | 1.01 | 642 | 0.73473 | 0.73473
352.428 | 0 | 244.1 | -0.048029 | 288.4 | -0.044293 | 421 | 1.0173 | 643 | 0.7171 | 0.7171
352.764 | 0.01 | 244.2 | -0.045267 | 288.6 | -0.042009 | 421.5 | 1.0281 | 644 | 0.69571 | 0.69571
353.1 | 0.01 | 244.3 | -0.054592 | 288.8 | -0.039442 | 422 | 1.0352 | 645 | 0.67901 | 0.67901
353.437 | 0.0066667 | 244.4 | -0.058741 | 289 | -0.04 | 422.5 | 1.0512 | 646 | 0.65932 | 0.65932
353.773 | 0.01 | 244.5 | -0.062889 | 289.2 | -0.03706 | 423 | 1.0661 | 647 | 0.64273 | 0.64273
354.109 | 0.01 | 244.6 | -0.067038 | 289.4 | -0.033455 | 423.5 | 1.0866 | 648 | 0.62512 | 0.62512
354.445 | 0.01 | 244.7 | -0.056665 | 289.6 | -0.036254 | 424 | 1.104 | 649 | 0.6089 | 0.6089
354.781 | 0.013333 | 244.8 | -0.056662 | 289.8 | -0.039407 | 424.5 | 1.1232 | 650 | 0.61311 | 0.61311
355.117 | 0.02 | 244.9 | -0.05666 | 290 | -0.04292 | 425 | 1.1256 | 651 | 0.60389 | 0.60389
355.453 | 0.02 | 245 | -0.063333 | 290.2 | -0.046771 | 425.5 | 1.1261 | 652 | 0.57259 | 0.57259
355.788 | 0.01 | 245.1 | -0.063331 | 290.4 | -0.0483 | 426 | 1.1347 | 653 | 0.55885 | 0.55885
356.124 | 0.02 | 245.2 | -0.063328 | 290.6 | -0.045368 | 426.5 | 1.1281 | 654 | 0.53675 | 0.53675
356.459 | 0.02 | 245.3 | -0.063325 | 290.8 | -0.04129 | 427 | 1.1292 | 655 | 0.51659 | 0.51659
356.795 | 0.01 | 245.4 | -0.061953 | 291 | -0.039166 | 427.5 | 1.1573 | 656 | 0.50204 | 0.50204
357.13 | 0.023333 | 245.5 | -0.063795 | 291.2 | -0.037637 | 428 | 1.1774 | 657 | 0.48734 | 0.48734
357.466 | 0.02 | 245.6 | -0.065637 | 291.4 | -0.040028 | 428.5 | 1.1887 | 658 | 0.48174 | 0.48174
357.801 | 0.02 | 245.7 | -0.06748 | 291.6 | -0.040598 | 429 | 1.2033 | 659 | 0.46658 | 0.46658
358.136 | 0.026667 | 245.8 | -0.060278 | 291.8 | -0.03889 | 429.5 | 1.2211 | 660 | 0.45089 | 0.45089
358.471 | 0.026667 | 245.9 | -0.057973 | 292 | -0.03462 | 430 | 1.2449 | 661 | 0.44144 | 0.44144
358.806 | 0.03 | 246 | -0.055668 | 292.2 | -0.028205 | 430.5 | 1.2738 | 662 | 0.44305 | 0.44305
359.141 | 0.03 | 246.1 | -0.06 | 292.4 | -0.034393 | 431 | 1.2991 | 663 | 0.43032 | 0.43032
359.476 | 0.03 | 246.2 | -0.06 | 292.6 | -0.039543 | 431.5 | 1.3058 | 664 | 0.40686 | 0.40686
359.811 | 0.02 | 246.3 | -0.06 | 292.8 | -0.038902 | 432 | 1.2902 | 665 | 0.39076 | 0.39076
360.145 | 0.03 | 246.4 | -0.06 | 293 | -0.041587 | 432.5 | 1.2978 | 666 | 0.38177 | 0.38177
360.48 | 0.04 | 246.5 | -0.059461 | 293.2 | -0.039833 | 433 | 1.3426 | 667 | 0.37644 | 0.37644
360.814 | 0.04 | 246.6 | -0.060385 | 293.4 | -0.03894 | 433.5 | 1.4206 | 668 | 0.36395 | 0.36395
361.149 | 0.03 | 246.7 | -0.061308 | 293.6 | -0.039922 | 434 | 1.5684 | 669 | 0.35503 | 0.35503
361.483 | 0.04 | 246.8 | -0.055528 | 293.8 | -0.039021 | 434.5 | 2.2781 | 670 | 0.34568 | 0.34568
361.817 | 0.05 | 246.9 | -0.052758 | 294 | -0.0382 | 435 | 7.1494 | 671 | 0.33645 | 0.33645
362.152 | 0.05 | 247 | -0.049988 | 294.2 | -0.038889 | 435.5 | 14.2957 | 672 | 0.3285 | 0.3285
362.486 | 0.05 | 247.1 | -0.047218 | 294.4 | -0.037565 | 436 | 14.363 | 673 | 0.31801 | 0.31801
362.82 | 0.05 | 247.2 | -0.066106 | 294.6 | -0.033965 | 436.5 | 5.7158 | 674 | 0.30822 | 0.30822
363.154 | 0.06 | 247.3 | -0.070262 | 294.8 | -0.033542 | 437 | 1.8362 | 675 | 0.29914 | 0.29914
363.488 | 0.08 | 247.4 | -0.074417 | 295 | -0.032378 | 437.5 | 1.4649 | 676 | 0.2867 | 0.2867
363.822 | 0.14667 | 247.5 | -0.078572 | 295.2 | -0.032328 | 438 | 1.475 | 677 | 0.27972 | 0.27972
364.155 | 0.63 | 247.6 | -0.061841 | 295.4 | -0.032884 | 438.5 | 1.4944 | 678 | 0.27285 | 0.27285
364.489 | 1.8067 | 247.7 | -0.062764 | 295.6 | -0.031663 | 439 | 1.5127 | 679 | 0.26586 | 0.26586
364.823 | 2.54 | 247.8 | -0.063687 | 295.8 | -0.030417 | 439.5 | 1.5243 | 680 | 0.26297 | 0.26297
365.156 | 1.95 | 247.9 | -0.061142 | 296 | -0.028401 | 440 | 1.5398 | 681 | 0.25615 | 0.25615
365.489 | 0.93 | 248 | -0.05514 | 296.2 | -0.029366 | 440.5 | 1.5504 | 682 | 0.24866 | 0.24866
365.823 | 0.62333 | 248.1 | -0.049138 | 296.4 | -0.03104 | 441 | 1.5545 | 683 | 0.2451 | 0.2451
366.156 | 0.56 | 248.2 | -0.043136 | 296.6 | -0.030882 | 441.5 | 1.5528 | 684 | 0.24107 | 0.24107
366.489 | 0.34 | 248.3 | -0.053017 | 296.8 | -0.030232 | 442 | 1.5599 | 685 | 0.23424 | 0.23424
366.822 | 0.11 | 248.4 | -0.054403 | 297 | -0.027182 | 442.5 | 1.5729 | 686 | 0.229 | 0.229
367.156 | 0.08 | 248.5 | -0.055789 | 297.2 | -0.02595 | 443 | 1.5883 | 687 | 0.23584 | 0.23584
367.488 | 0.073333 | 248.6 | -0.057174 | 297.4 | -0.027778 | 443.5 | 1.6014 | 688 | 0.23244 | 0.23244
367.821 | 0.08 | 248.7 | -0.063575 | 297.6 | -0.026296 | 444 | 1.6119 | 689 | 0.21428 | 0.21428
368.154 | 0.08 | 248.8 | -0.067274 | 297.8 | -0.025554 | 444.5 | 1.6101 | 690 | 0.22871 | 0.22871
368.487 | 0.08 | 248.9 | -0.070972 | 298 | -0.027035 | 445 | 1.6116 | 691 | 0.24519 | 0.24519
368.82 | 0.09 | 249 | -0.065724 | 298.2 | -0.027441 | 445.5 | 1.6125 | 692 | 0.19472 | 0.19472
369.152 | 0.08 | 249.1 | -0.066188 | 298.4 | -0.027249 | 446 | 1.6158 | 693 | 0.19513 | 0.19513
369.485 | 0.08 | 249.2 | -0.066652 | 298.6 | -0.028273 | 446.5 | 1.612 | 694 | 0.19683 | 0.19683
369.817 | 0.1 | 249.3 | -0.067116 | 298.8 | -0.028889 | 447 | 1.6099 | 695 | 0.18086 | 0.18086
370.149 | 0.093333 | 249.4 | -0.064157 | 299 | -0.027172 | 447.5 | 1.6148 | 696 | 0.18158 | 0.18158
370.482 | 0.093333 | 249.5 | -0.062308 | 299.2 | -0.025597 | 448 | 1.6274 | 697 | 0.17623 | 0.17623
370.814 | 0.10333 | 249.6 | -0.060459 | 299.4 | -0.024611 | 448.5 | 1.633 | 698 | 0.1609 | 0.1609
371.146 | 0.1 | 249.7 | -0.059858 | 299.6 | -0.023377 | 449 | 1.6428 | 699 | 0.15554 | 0.15554
371.478 | 0.11 | 249.8 | -0.058934 | 299.8 | -0.023333 | 449.5 | 1.6468 | 700 | 0.1584 | 0.1584
371.81 | 0.11 | 249.9 | -0.058009 | 300 | -0.024322 | 450 | 1.6512 | 701 | 0.15444 | 0.15444
372.142 | 0.11 | 250 | -0.057084 | 300.2 | -0.027256 | 450.5 | 1.664 | 702 | 0.15086 | 0.15086
372.474 | 0.11 | 250.1 | -0.060255 | 300.4 | -0.027435 | 451 | 1.6731 | 703 | 0.15238 | 0.15238
372.805 | 0.11 | 250.2 | -0.060718 | 300.6 | -0.027765 | 451.5 | 1.6845 | 704 | 0.14792 | 0.14792
373.137 | 0.12 | 250.3 | -0.061181 | 300.8 | -0.028674 | 452 | 1.6833 | 705 | 0.14955 | 0.14955
373.469 | 0.11333 | 250.4 | -0.061644 | 301 | -0.028941 | 452.5 | 1.6865 | 706 | 0.17476 | 0.17476
373.8 | 0.12667 | 250.5 | -0.058889 | 301.2 | -0.029522 | 453 | 1.6936 | 707 | 0.2007 | 0.2007
374.132 | 0.12333 | 250.6 | -0.058889 | 301.4 | -0.031169 | 453.5 | 1.7034 | 708 | 0.18328 | 0.18328
374.463 | 0.12 | 250.7 | -0.058889 | 301.6 | -0.033985 | 454 | 1.7099 | 709 | 0.19708 | 0.19708
374.794 | 0.13333 | 250.8 | -0.064926 | 301.8 | -0.034142 | 454.5 | 1.7172 | 710 | 0.15547 | 0.15547
375.125 | 0.13 | 250.9 | -0.066314 | 302 | -0.034459 | 455 | 1.7285 | 711 | 0.13637 | 0.13637
375.456 | 0.13333 | 251 | -0.067703 | 302.2 | -0.031814 | 455.5 | 1.73 | 712 | 0.17173 | 0.17173
375.787 | 0.14 | 251.1 | -0.069091 | 302.4 | -0.029582 | 456 | 1.736 | 713 | 0.15635 | 0.15635
376.118 | 0.13667 | 251.2 | -0.056405 | 302.6 | -0.031111 | 456.5 | 1.7472 | 714 | 0.11865 | 0.11865
376.449 | 0.13667 | 251.3 | -0.054559 | 302.8 | -0.030886 | 457 | 1.7561 | 715 | 0.10858 | 0.10858
376.78 | 0.15 | 251.4 | -0.052713 | 303 | -0.030915 | 457.5 | 1.76 | 716 | 0.10298 | 0.10298
377.111 | 0.14 | 251.5 | -0.053913 | 303.2 | -0.032984 | 458 | 1.7717 | 717 | 0.10484 | 0.10484
377.441 | 0.15 | 251.6 | -0.050202 | 303.4 | -0.034497 | 458.5 | 1.7703 | 718 | 0.10313 | 0.10313
377.772 | 0.15 | 251.7 | -0.046491 | 303.6 | -0.03314 | 459 | 1.7751 | 719 | 0.10074 | 0.10074
378.102 | 0.15667 | 251.8 | -0.04278 | 303.8 | -0.032071 | 459.5 | 1.786 | 720 | 0.094184 | 0.094184
378.433 | 0.16667 | 251.9 | -0.048538 | 304 | -0.030808 | 460 | 1.7956 | 721 | 0.092785 | 0.092785
378.763 | 0.16 | 252 | -0.049929 | 304.2 | -0.029579 | 460.5 | 1.7942 | 722 | 0.087065 | 0.087065
379.093 | 0.17 | 252.1 | -0.05132 | 304.4 | -0.03 | 461 | 1.8003 | 723 | 0.083923 | 0.083923
379.424 | 0.17 | 252.2 | -0.052711 | 304.6 | -0.026797 | 461.5 | 1.8062 | 724 | 0.085614 | 0.085614
379.754 | 0.17 | 252.3 | -0.048889 | 304.8 | -0.025482 | 462 | 1.8097 | 725 | 0.084125 | 0.084125
380.084 | 0.17333 | 252.4 | -0.048889 | 305 | -0.029895 | 462.5 | 1.8163 | 726 | 0.07838 | 0.07838
380.414 | 0.18 | 252.5 | -0.048889 | 305.2 | -0.029872 | 463 | 1.8263 | 727 | 0.077811 | 0.077811
380.744 | 0.17667 | 252.6 | -0.047292 | 305.4 | -0.027425 | 463.5 | 1.8307 | 728 | 0.076313 | 0.076313
381.073 | 0.19 | 252.7 | -0.045903 | 305.6 | -0.028921 | 464 | 1.8363 | 729 | 0.073306 | 0.073306
381.403 | 0.19 | 252.8 | -0.044514 | 305.8 | -0.027992 | 464.5 | 1.8401 | 730 | 0.07205 | 0.07205
381.733 | 0.2 | 252.9 | -0.043125 | 306 | -0.024949 | 465 | 1.847 | 731 | 0.069855 | 0.069855
382.062 | 0.20333 | 253 | -0.047015 | 306.2 | -0.024508 | 465.5 | 1.8537 | 732 | 0.068234 | 0.068234
382.392 | 0.20333 | 253.1 | -0.047478 | 306.4 | -0.026667 | 466 | 1.8565 | 733 | 0.064263 | 0.064263
382.721 | 0.2 | 253.2 | -0.047941 | 306.6 | -0.026993 | 466.5 | 1.8598 | 734 | 0.063254 | 0.063254
383.051 | 0.20333 | 253.3 | -0.049107 | 306.8 | -0.029087 | 467 | 1.867 | 735 | 0.058873 | 0.058873
383.38 | 0.21 | 253.4 | -0.050497 | 307 | -0.027272 | 467.5 | 1.8762 | 736 | 0.061839 | 0.061839
383.709 | 0.22 | 253.5 | -0.051888 | 307.2 | -0.023482 | 468 | 1.8792 | 737 | 0.06318 | 0.06318
384.038 | 0.22 | 253.6 | -0.053279 | 307.4 | -0.023987 | 468.5 | 1.8816 | 738 | 0.078772 | 0.078772
384.367 | 0.22 | 253.7 | -0.051111 | 307.6 | -0.023727 | 469 | 1.885 | 739 | 0.074995 | 0.074995
384.696 | 0.23 | 253.8 | -0.051112 | 307.8 | -0.022222 | 469.5 | 1.8839 | 740 | 0.052591 | 0.052591
385.025 | 0.23 | 253.9 | -0.051112 | 308 | -0.022465 | 470 | 1.8897 | 741 | 0.05633 | 0.05633
385.354 | 0.23 | 254 | -0.051113 | 308.2 | -0.022469 | 470.5 | 1.8876 | 742 | 0.056339 | 0.056339
385.682 | 0.24667 | 254.1 | -0.052222 | 308.4 | -0.023419 | 471 | 1.892 | 743 | 0.057467 | 0.057467
386.011 | 0.24 | 254.2 | -0.052222 | 308.6 | -0.025578 | 471.5 | 1.8973 | 744 | 0.048485 | 0.048485
386.34 | 0.25 | 254.3 | -0.052222 | 308.8 | -0.023963 | 472 | 1.9029 | 745 | 0.045517 | 0.045517
386.668 | 0.25 | 254.4 | -0.056254 | 309 | -0.022944 | 472.5 | 1.9096 | 746 | 0.046895 | 0.046895
386.997 | 0.25 | 254.5 | -0.058109 | 309.2 | -0.024444 | 473 | 1.9065 | 747 | 0.046696 | 0.046696
387.325 | 0.26 | 254.6 | -0.059964 | 309.4 | -0.023704 | 473.5 | 1.9039 | 748 | 0.044264 | 0.044264
387.653 | 0.26 | 254.7 | -0.061819 | 309.6 | -0.022224 | 474 | 1.91 | 749 | 0.041074 | 0.041074
387.981 | 0.27333 | 254.8 | -0.064038 | 309.8 | -0.0216 | 474.5 | 1.9139 | 750 | 0.061098 | 0.061098
388.309 | 0.28 | 254.9 | -0.066356 | 310 | -0.020588 | 475 | 1.9164 | 751 | 0.074378 | 0.074378
388.637 | 0.28 | 255 | -0.068675 | 310.2 | -0.021596 | 475.5 | 1.9235 | 752 | 0.056193 | 0.056193
388.965 | 0.28667 | 255.1 | -0.059559 | 310.4 | -0.021458 | 476 | 1.921 | 753 | 0.033474 | 0.033474
389.293 | 0.29 | 255.2 | -0.057238 | 310.6 | -0.019259 | 476.5 | 1.9192 | 754 | 0.027735 | 0.027735
389.621 | 0.29667 | 255.3 | -0.054916 | 310.8 | -0.020334 | 477 | 1.9237 | 755 | 0.028337 | 0.028337
389.949 | 0.31 | 255.4 | -0.052595 | 311 | -0.021111 | 477.5 | 1.9251 | 756 | 0.026989 | 0.026989
390.276 | 0.33 | 255.5 | -0.053884 | 311.2 | -0.024031 | 478 | 1.929 | 757 | 0.026616 | 0.026616
390.604 | 0.32 | 255.6 | -0.051099 | 311.4 | -0.027193 | 478.5 | 1.9311 | 758 | 0.022609 | 0.022609
390.932 | 0.33 | 255.7 | -0.048313 | 311.6 | -0.027225 | 479 | 1.9326 | 759 | 0.018918 | 0.018918
391.259 | 0.32667 | 255.8 | -0.052241 | 311.8 | -0.025848 | 479.5 | 1.9304 | 760 | 0.023203 | 0.023203
391.586 | 0.32667 | 255.9 | -0.054098 | 312 | -0.019239 | 480 | 1.9266 | 761 | 0.02056 | 0.02056
391.914 | 0.33667 | 256 | -0.055955 | 312.2 | -0.020425 | 480.5 | 1.9342 | 762 | 0.017471 | 0.017471
392.241 | 0.34333 | 256.1 | -0.057812 | 312.4 | -0.022222 | 481 | 1.9364 | 763 | 0.061319 | 0.061319
392.568 | 0.34 | 256.2 | -0.058558 | 312.6 | -0.018108 | 481.5 | 1.9349 | 764 | 0.089232 | 0.089232
392.895 | 0.35 | 256.3 | -0.060415 | 312.8 | -0.016571 | 482 | 1.9403 | 765 | 0.023745 | 0.023745
393.222 | 0.35333 | 256.4 | -0.062272 | 313 | -0.017611 | 482.5 | 1.9433 | 766 | 0.016221 | 0.016221
393.549 | 0.36 | 256.5 | -0.064129 | 313.2 | -0.017266 | 483 | 1.9464 | 767 | 0.015876 | 0.015876
393.876 | 0.36333 | 256.6 | -0.052121 | 313.4 | -0.021393 | 483.5 | 1.9538 | 768 | 0.01666 | 0.01666
394.202 | 0.37333 | 256.7 | -0.049334 | 313.6 | -0.022884 | 484 | 1.9567 | 769 | 0.014753 | 0.014753
394.529 | 0.38333 | 256.8 | -0.046546 | 313.8 | -0.023387 | 484.5 | 1.9593 | 770 | 0.012445 | 0.012445
394.855 | 0.38 | 256.9 | -0.050653 | 314 | -0.02109 | 485 | 1.9667 | 771 | 0.0052528 | 0.0052528
395.182 | 0.39333 | 257 | -0.048796 | 314.2 | -0.019076 | 485.5 | 1.978 | 772 | 0.026575 | 0.026575
395.508 | 0.39667 | 257.1 | -0.046938 | 314.4 | -0.019274 | 486 | 1.9814 | 773 | 0.039248 | 0.039248
395.835 | 0.4 | 257.2 | -0.045081 | 314.6 | -0.020872 | 486.5 | 1.9814 | 774 | 0.011634 | 0.011634
396.161 | 0.41333 | 257.3 | -0.056091 | 314.8 | -0.024082 | 487 | 1.9852 | 775 | 0.0050946 | 0.0050946
396.487 | 0.42 | 257.4 | -0.060269 | 315 | -0.023325 | 487.5 | 1.9808 | 776 | 0.0036419 | 0.0036419
396.813 | 0.43 | 257.5 | -0.064448 | 315.2 | -0.022174 | 488 | 1.9791 | 777 | 0.0037901 | 0.0037901
397.139 | 0.43333 | 257.6 | -0.060102 | 315.4 | -0.021111 | 488.5 | 1.9684 | 778 | 0.005959 | 0.005959
397.465 | 0.44 | 257.7 | -0.061499 | 315.6 | -0.020258 | 489 | 1.9665 | 779 | 0.0033049 | 0.0033049
397.791 | 0.44 | 257.8 | -0.062896 | 315.8 | -0.021932 | 489.5 | 1.97 | 780 | -0.0020174 | -0.0020174
398.117 | 0.46 | 257.9 | -0.064292 | 316 | -0.024493 | 490 | 1.9646 | 781 | -0.00067018 | -0.00067018
398.443 | 0.46667 | 258 | -0.058642 | 316.2 | -0.027705 | 490.5 | 1.9619 | 782 | 0.005164 | 0.005164
398.768 | 0.47 | 258.1 | -0.055853 | 316.4 | -0.027915 | 491 | 1.9842 | 783 | -0.0024539 | -0.0024539
399.094 | 0.48 | 258.2 | -0.053064 | 316.6 | -0.027295 | 491.5 | 2.006 | 784 | -0.0031916 | -0.0031916
399.42 | 0.48333 | 258.3 | -0.050274 | 316.8 | -0.025362 | 492 | 1.9997 | 785 | -0.0066376 | -0.0066376
399.745 | 0.49 | 258.4 | -0.053604 | 317 | -0.023428 | 492.5 | 1.9496 | 786 | -0.013123 | -0.013123
400.07 | 0.50333 | 258.5 | -0.052673 | 317.2 | -0.021492 | 493 | 1.9298 | 787 | -0.012979 | -0.012979
400.396 | 0.51 | 258.6 | -0.051742 | 317.4 | -0.020607 | 493.5 | 1.9299 | 788 | -0.01157 | -0.01157
400.721 | 0.52 | 258.7 | -0.046071 | 317.6 | -0.019365 | 494 | 1.9262 | 789 | -0.014559 | -0.014559
401.046 | 0.53 | 258.8 | -0.044209 | 317.8 | -0.016798 | 494.5 | 1.9173 | 790 | -0.011981 | -0.011981
401.371 | 0.55 | 258.9 | -0.042346 | 318 | -0.016619 | 495 | 1.9095 | 791 | -0.014142 | -0.014142
401.696 | 0.55 | 259 | -0.040484 | 318.2 | -0.019649 | 495.5 | 1.9129 | 792 | -0.010666 | -0.010666
402.021 | 0.56 | 259.1 | -0.044867 | 318.4 | -0.02159 | 496 | 1.8995 | 793 | -0.012342 | -0.012342
402.346 | 0.58333 | 259.2 | -0.043935 | 318.6 | -0.025221 | 496.5 | 1.8962 | 794 | -0.015841 | -0.015841
402.67 | 0.6 | 259.3 | -0.043004 | 318.8 | -0.024227 | 497 | 1.8947 | 795 | 0.0043499 | 0.0043499
402.995 | 0.63 | 259.4 | -0.041186 | 319 | -0.02259 | 497.5 | 1.8894 | 796 | -0.0048898 | -0.0048898
403.32 | 0.76333 | 259.5 | -0.041651 | 319.2 | -0.022733 | 498 | 1.8843 | 797 | -0.021285 | -0.021285
403.644 | 1.53 | 259.6 | -0.042117 | 319.4 | -0.021707 | 498.5 | 1.875 | 798 | -0.017508 | -0.017508
403.969 | 3.8633 | 259.7 | -0.042582 | 319.6 | -0.02111 | 499 | 1.8679 | 799 | -0.020065 | -0.020065
404.293 | 7.84 | 259.8 | -0.049969 | 319.8 | -0.020835 | 499.5 | 1.8622 | 800 | -0.021413 | -0.021413
404.617 | 9.6967 | 259.9 | -0.051831 | 320 | -0.019972 | 500 | 1.8606 | 801 | -4.2477e-05 | -4.2477e-05
404.942 | 5.3167 | 260 | -0.053693 | 320.2 | -0.021026 | 500.5 | 1.8547 | 802 | 0.020115 | 0.020115
405.266 | 1.29 | 260.1 | -0.055556 | 320.4 | -0.02189 | 501 | 1.8511 | 803 | -0.0033264 | -0.0033264
405.59 | 0.74 | 260.2 | -0.05928 | 320.6 | -0.021112 | 501.5 | 1.8357 | 804 | -0.0042143 | -0.0042143
405.914 | 0.69 | 260.3 | -0.063004 | 320.8 | -0.020938 | 502 | 1.8329 | 805 | -0.0021897 | -0.0021897
406.238 | 0.68 | 260.4 | -0.066729 | 321 | -0.020123 | 502.5 | 1.8299 | 806 | -0.010157 | -0.010157
406.562 | 0.71 | 260.5 | -0.059609 | 321.2 | -0.016591 | 503 | 1.8262 | 807 | -0.032328 | -0.032328
406.885 | 0.79667 | 260.6 | -0.058678 | 321.4 | -0.015424 | 503.5 | 1.8234
407.209 | 0.99333 | 260.7 | -0.057747 | 321.6 | -0.020013 | 504 | 1.8164
407.533 | 1.28 | 260.8 | -0.056816 | 321.8 | -0.020959 | 504.5 | 1.8094
407.856 | 1.2767 | 260.9 | -0.057716 | 322 | -0.0214 | 505 | 1.8041
408.18 | 0.92333 | 261 | -0.056319 | 322.2 | -0.021925 | 505.5 | 1.8008
408.503 | 0.76 | 261.1 | -0.054922 | 322.4 | -0.020004 | 506 | 1.7964
408.826 | 0.75 | 261.2 | -0.055677 | 322.6 | -0.019678 | 506.5 | 1.7883
409.15 | 0.76 | 261.3 | -0.056142 | 322.8 | -0.02 | 507 | 1.7858
409.473 | 0.76 | 261.4 | -0.056608 | 323 | -0.02 | 507.5 | 1.7818
409.796 | 0.77 | 261.5 | -0.057073 | 323.2 | -0.021061 | 508 | 1.7805
410.119 | 0.78 | 261.6 | -0.050949 | 323.4 | -0.024156 | 508.5 | 1.783
410.442 | 0.79 | 261.7 | -0.049085 | 323.6 | -0.024273 | 509 | 1.7728
410.765 | 0.79667 | 261.8 | -0.047221 | 323.8 | -0.025129 | 509.5 | 1.7718
411.087 | 0.80667 | 261.9 | -0.049901 | 324 | -0.027211 | 510 | 1.7666
411.41 | 0.81 | 262 | -0.04897 | 324.2 | -0.026661 | 510.5 | 1.7604
411.733 | 0.81 | 262.1 | -0.048038 | 324.4 | -0.023623 | 511 | 1.7583
412.055 | 0.82667 | 262.2 | -0.047107 | 324.6 | -0.023516 | 511.5 | 1.7465
412.378 | 0.83 | 262.3 | -0.049635 | 324.8 | -0.022248 | 512 | 1.7458
412.7 | 0.84 | 262.4 | -0.051034 | 325 | -0.023513 | 512.5 | 1.7461
413.022 | 0.84 | 262.5 | -0.052433 | 325.2 | -0.024345 | 513 | 1.7433
413.345 | 0.86 | 262.6 | -0.053831 | 325.4 | -0.020852 | 513.5 | 1.7318
413.667 | 0.86 | 262.7 | -0.05267 | 325.6 | -0.021234 | 514 | 1.7322
413.989 | 0.86667 | 262.8 | -0.053137 | 325.8 | -0.025407 | 514.5 | 1.7253
414.311 | 0.87667 | 262.9 | -0.053604 | 326 | -0.024871 | 515 | 1.7254
414.633 | 0.87 | 263 | -0.061082 | 326.2 | -0.024976 | 515.5 | 1.7243
414.955 | 0.88 | 263.1 | -0.063881 | 326.4 | -0.0228 | 516 | 1.7213
415.277 | 0.89333 | 263.2 | -0.066679 | 326.6 | -0.020663 | 516.5 | 1.721
415.598 | 0.9 | 263.3 | -0.069478 | 326.8 | -0.020282 | 517 | 1.7152
415.92 | 0.90333 | 263.4 | -0.058512 | 327 | -0.019932 | 517.5 | 1.7103
416.242 | 0.91 | 263.5 | -0.058049 | 327.2 | -0.02 | 518 | 1.71
416.563 | 0.92 | 263.6 | -0.057585 | 327.4 | -0.023539 | 518.5 | 1.7092
416.885 | 0.93 | 263.7 | -0.061886 | 327.6 | -0.026503 | 519 | 1.7026
417.206 | 0.93 | 263.8 | -0.060486 | 327.8 | -0.026949 | 519.5 | 1.7067
417.527 | 0.93667 | 263.9 | -0.059085 | 328 | -0.023129 | 520 | 1.7056
417.848 | 0.95 | 264 | -0.057684 | 328.2 | -0.016345 | 520.5 | 1.7029
418.17 | 0.96 | 264.1 | -0.055868 | 328.4 | -0.016358 | 521 | 1.6938
418.491 | 0.97 | 264.2 | -0.056335 | 328.6 | -0.016873 | 521.5 | 1.6923
418.812 | 0.96333 | 264.3 | -0.056802 | 328.8 | -0.018377 | 522 | 1.7004
419.132 | 0.99 | 264.4 | -0.057684 | 329 | -0.019888 | 522.5 | 1.6947
419.453 | 0.99 | 264.5 | -0.056751 | 329.2 | -0.018036 | 523 | 1.6864
419.774 | 1 | 264.6 | -0.055817 | 329.4 | -0.018442 | 523.5 | 1.687
420.095 | 1.01 | 264.7 | -0.054883 | 329.6 | -0.018889 | 524 | 1.6922
420.415 | 1.01 | 264.8 | -0.054444 | 329.8 | -0.019349 | 524.5 | 1.6942
420.736 | 1.01 | 264.9 | -0.054444 | 330 | -0.020045 | 525 | 1.6933
421.056 | 1.0167 | 265 | -0.054444 | 330.2 | -0.02 | 525.5 | 1.6974
421.377 | 1.03 | 265.1 | -0.054444 | 330.4 | -0.02 | 526 | 1.6978
421.697 | 1.03 | 265.2 | -0.057116 | 330.6 | -0.02 | 526.5 | 1.6954
422.017 | 1.03 | 265.3 | -0.057583 | 330.8 | -0.023357 | 527 | 1.6946
422.337 | 1.05 | 265.4 | -0.05805 | 331 | -0.025444 | 527.5 | 1.6998
422.657 | 1.0533 | 265.5 | -0.055555 | 331.2 | -0.023332 | 528 | 1.6971
422.977 | 1.07 | 265.6 | -0.055554 | 331.4 | -0.022924 | 528.5 | 1.7011
423.297 | 1.0667 | 265.7 | -0.055553 | 331.6 | -0.020648 | 529 | 1.7078
423.617 | 1.1 | 265.8 | -0.055553 | 331.8 | -0.019728 | 529.5 | 1.7031
423.937 | 1.0933 | 265.9 | -0.062274 | 332 | -0.02 | 530 | 1.7114
424.257 | 1.12 | 266 | -0.063674 | 332.2 | -0.02 | 530.5 | 1.7135
424.576 | 1.13 | 266.1 | -0.065075 | 332.4 | -0.02 | 531 | 1.7138
424.896 | 1.12 | 266.2 | -0.063826 | 332.6 | -0.02 | 531.5 | 1.7135
425.215 | 1.13 | 266.3 | -0.065697 | 332.8 | -0.018069 | 532 | 1.7237
425.535 | 1.12 | 266.4 | -0.067569 | 333 | -0.015637 | 532.5 | 1.724
425.854 | 1.1367 | 266.5 | -0.06944 | 333.2 | -0.010949 | 533 | 1.73
426.173 | 1.14 | 266.6 | -0.060256 | 333.4 | -0.0060443 | 533.5 | 1.7394
426.493 | 1.1267 | 266.7 | -0.057447 | 333.6 | -0.0028772 | 534 | 1.7407
426.812 | 1.12 | 266.8 | -0.054638 | 333.8 | -0.00070381 | 534.5 | 1.7441
427.131 | 1.13 | 266.9 | -0.057586 | 334 | -0.0022222 | 535 | 1.7453
427.45 | 1.16 | 267 | -0.056185 | 334.2 | -0.0079876 | 535.5 | 1.7521
427.769 | 1.17 | 267.1 | -0.054784 | 334.4 | -0.014139 | 536 | 1.7636
428.087 | 1.18 | 267.2 | -0.053383 | 334.6 | -0.017249 | 536.5 | 1.774
428.406 | 1.19 | 267.3 | -0.054981 | 334.8 | -0.021103 | 537 | 1.7861
428.725 | 1.19 | 267.4 | -0.055916 | 335 | -0.021962 | 537.5 | 1.8055
429.043 | 1.21 | 267.5 | -0.056852 | 335.2 | -0.021558 | 538 | 1.8204
429.362 | 1.21 | 267.6 | -0.053319 | 335.4 | -0.021112 | 538.5 | 1.8412
429.68 | 1.23 | 267.7 | -0.051915 | 335.6 | -0.020644 | 539 | 1.8557
429.999 | 1.25 | 267.8 | -0.05051 | 335.8 | -0.019947 | 539.5 | 1.8763
430.317 | 1.2533 | 267.9 | -0.049106 | 336 | -0.02 | 540 | 1.9076
430.635 | 1.29 | 268 | -0.055977 | 336.2 | -0.02 | 540.5 | 1.9448
430.953 | 1.2967 | 268.1 | -0.056913 | 336.4 | -0.02 | 541 | 2.0058
431.271 | 1.31 | 268.2 | -0.05785 | 336.6 | -0.02 | 541.5 | 2.0922
431.589 | 1.31 | 268.3 | -0.058786 | 336.8 | -0.019577 | 542 | 2.167
431.907 | 1.29 | 268.4 | -0.055278 | 337 | -0.018242 | 542.5 | 2.182
432.225 | 1.28 | 268.5 | -0.056214 | 337.2 | -0.017967 | 543 | 2.1783
432.543 | 1.3 | 268.6 | -0.05715 | 337.4 | -0.018889 | 543.5 | 2.1871
432.861 | 1.32 | 268.7 | -0.057469 | 337.6 | -0.019222 | 544 | 2.215
433.178 | 1.3633 | 268.8 | -0.056532 | 337.8 | -0.019139 | 544.5 | 2.2626
433.496 | 1.41 | 268.9 | -0.055596 | 338 | -0.01668 | 545 | 2.9614
433.813 | 1.4867 | 269 | -0.05466 | 338.2 | -0.014599 | 545.5 | 6.6602
434.131 | 1.6 | 269.1 | -0.057388 | 338.4 | -0.011906 | 546 | 12.6729
434.448 | 1.8267 | 269.2 | -0.058324 | 338.6 | -0.013649 | 546.5 | 12.027
434.765 | 3.14 | 269.3 | -0.05926 | 338.8 | -0.016731 | 547 | 3.6875
435.082 | 7.74 | 269.4 | -0.057778 | 339 | -0.018424 | 547.5 | 2.2087
435.399 | 14.01 | 269.5 | -0.057778 | 339.2 | -0.018793 | 548 | 2.1492
435.717 | 16.6233 | 269.6 | -0.057778 | 339.4 | -0.014402 | 548.5 | 2.1368
436.033 | 16.6133 | 269.7 | -0.057778 | 339.6 | -0.01542 | 549 | 2.1344
436.35 | 7.55 | 269.8 | -0.053227 | 339.8 | -0.016664 | 549.5 | 2.1355
436.667 | 2.0333 | 269.9 | -0.051355 | 340 | -0.017889 | 550 | 2.1501
436.984 | 1.5333 | 270 | -0.049482 | 340.2 | -0.019554 | 550.5 | 2.1518
437.301 | 1.47 | 270.1 | -0.053333 | 340.4 | -0.018225 | 551 | 2.154
437.617 | 1.4733 | 270.2 | -0.053332 | 340.6 | -0.014789 | 551.5 | 2.1622
437.934 | 1.47 | 270.3 | -0.053331 | 340.8 | -0.0064405 | 552 | 2.175
438.25 | 1.48 | 270.4 | -0.053329 | 341 | -0.0096624 | 552.5 | 2.184
438.566 | 1.5 | 270.5 | -0.059539 | 341.2 | -0.012248 | 553 | 2.1805
438.883 | 1.51 | 270.6 | -0.062821 | 341.4 | -0.01174 | 553.5 | 2.1867
439.199 | 1.52 | 270.7 | -0.066103 | 341.6 | -0.013358 | 554 | 2.1935
439.515 | 1.52 | 270.8 | -0.069385 | 341.8 | -0.010693 | 554.5 | 2.1908
439.831 | 1.54 | 270.9 | -0.061113 | 342 | -0.0097394 | 555 | 2.1966
440.147 | 1.54 | 271 | -0.061115 | 342.2 | -0.010744 | 555.5 | 2.1902
440.463 | 1.5567 | 271.1 | -0.061116 | 342.4 | -0.013215 | 556 | 2.1788
440.779 | 1.55 | 271.2 | -0.060304 | 342.6 | -0.013868 | 556.5 | 2.1799
441.095 | 1.56 | 271.3 | -0.058426 | 342.8 | -0.012224 | 557 | 2.1836
441.41 | 1.55 | 271.4 | -0.056548 | 343 | -0.01174 | 557.5 | 2.199
441.726 | 1.5533 | 271.5 | -0.05467 | 343.2 | -0.010021 | 558 | 2.2089
442.042 | 1.56 | 271.6 | -0.053622 | 343.4 | -0.0096788 | 558.5 | 2.2223
442.357 | 1.57 | 271.7 | -0.052684 | 343.6 | -0.01 | 559 | 2.2192
442.672 | 1.5767 | 271.8 | -0.051746 | 343.8 | -0.01 | 559.5 | 2.2244
442.988 | 1.5867 | 271.9 | -0.056969 | 344 | -0.01 | 560 | 2.2393
443.303 | 1.6 | 272 | -0.057906 | 344.2 | -0.01 | 560.5 | 2.2448
443.618 | 1.6 | 272.1 | -0.058842 | 344.4 | -0.011134 | 561 | 2.2563
443.933 | 1.62 | 272.2 | -0.059778 | 344.6 | -0.011833 | 561.5 | 2.2698
444.248 | 1.6067 | 272.3 | -0.06328 | 344.8 | -0.011112 | 562 | 2.2769
444.563 | 1.61 | 272.4 | -0.066097 | 345 | -0.010436 | 562.5 | 2.2838
444.878 | 1.6133 | 272.5 | -0.068914 | 345.2 | -0.0080966 | 563 | 2.2836
445.193 | 1.61 | 272.6 | -0.067984 | 345.4 | -0.009382 | 563.5 | 2.295
445.508 | 1.6133 | 272.7 | -0.068923 | 345.6 | -0.010304 | 564 | 2.3042
445.822 | 1.6133 | 272.8 | -0.069862 | 345.8 | -0.0088889 | 564.5 | 2.3096
446.137 | 1.62 | 272.9 | -0.070801 | 346 | -0.0091294 | 565 | 2.3167
446.452 | 1.61 | 273 | -0.061446 | 346.2 | -0.0091174 | 565.5 | 2.3247
446.766 | 1.61 | 273.1 | -0.058629 | 346.4 | -0.0080599 | 566 | 2.3243
447.08 | 1.61 | 273.2 | -0.055812 | 346.6 | -0.0085234 | 566.5 | 2.3352
447.395 | 1.61 | 273.3 | -0.064501 | 346.8 | -0.010508 | 567 | 2.3445
447.709 | 1.62 | 273.4 | -0.06497 | 347 | -0.010939 | 567.5 | 2.3458
448.023 | 1.63 | 273.5 | -0.06544 | 347.2 | -0.010818 | 568 | 2.3561
448.337 | 1.6333 | 273.6 | -0.065909 | 347.4 | -0.0079329 | 568.5 | 2.3572
448.651 | 1.63 | 273.7 | -0.05778 | 347.6 | -0.0032796 | 569 | 2.3588
448.965 | 1.65 | 273.8 | -0.057783 | 347.8 | -0.0024726 | 569.5 | 2.3601
449.279 | 1.64 | 273.9 | -0.057786 | 348 | -0.0022064 | 570 | 2.3678
449.593 | 1.65 | 274 | -0.061061 | 348.2 | -0.0029758 | 570.5 | 2.3784
449.906 | 1.65 | 274.1 | -0.059183 | 348.4 | -0.0024229 | 571 | 2.3783
450.22 | 1.65 | 274.2 | -0.057304 | 348.6 | -0.0011594 | 571.5 | 2.3744
450.534 | 1.67 | 274.3 | -0.055426 | 348.8 | 0.00062212 | 572 | 2.3828
450.847 | 1.67 | 274.4 | -0.051997 | 349 | 0.00056173 | 572.5 | 2.3914
451.161 | 1.6733 | 274.5 | -0.051527 | 349.2 | -0.0043079 | 573 | 2.3932
451.474 | 1.6867 | 274.6 | -0.051058 | 349.4 | -0.0035879 | 573.5 | 2.3891
451.787 | 1.69 | 274.7 | -0.050588 | 349.6 | -0.0032298 | 574 | 2.389
452.1 | 1.6767 | 274.8 | -0.054004 | 349.8 | -0.0023166 | 574.5 | 2.3996
452.413 | 1.69 | 274.9 | -0.053533 | 350 | -8.5724e-05 | 575 | 2.3908
452.727 | 1.6867 | 275 | -0.053061 | 350.2 | 0 | 575.5 | 2.4041
453.04 | 1.6967 | 275.1 | -0.053878 | 350.4 | -0.0015682 | 576 | 2.561
453.352 | 1.6967 | 275.2 | -0.055288 | 350.6 | -0.0061898 | 576.5 | 3.3018
453.665 | 1.71 | 275.3 | -0.056698 | 350.8 | -0.0036897 | 577 | 4.3208
453.978 | 1.71 | 275.4 | -0.058108 | 351 | -0.0014468 | 577.5 | 3.5911
454.291 | 1.71 | 275.5 | -0.065711 | 351.2 | -0.00045664 | 578 | 2.647
454.603 | 1.72 | 275.6 | -0.068531 | 351.4 | 0.0014144 | 578.5 | 3.0534
454.916 | 1.73 | 275.7 | -0.07135 | 351.6 | 0.00021101 | 579 | 4.0516
455.228 | 1.73 | 275.8 | -0.058324 | 351.8 | -9.4516e-05 | 579.5 | 3.8161
455.541 | 1.73 | 275.9 | -0.056441 | 352 | 0 | 580 | 2.6525
455.853 | 1.73 | 276 | -0.054557 | 352.2 | 0 | 580.5 | 2.4532
456.165 | 1.74 | 276.1 | -0.052674 | 352.4 | 0.0014692 | 581 | 2.4279
456.477 | 1.75 | 276.2 | -0.062618 | 352.6 | 0.0058927 | 581.5 | 2.4128
456.79 | 1.75 | 276.3 | -0.061676 | 352.8 | 0.0077569 | 582 | 2.403
457.102 | 1.76 | 276.4 | -0.060734 | 353 | 0.010179 | 582.5 | 2.4115
457.414 | 1.76 | 276.5 | -0.0595 | 353.2 | 0.0093204 | 583 | 2.4109
457.725 | 1.76 | 276.6 | -0.062318 | 353.4 | 0.008064 | 583.5 | 2.4098
458.037 | 1.78 | 276.7 | -0.065137 | 353.6 | 0.0088881 | 584 | 2.4007
458.349 | 1.7733 | 276.8 | -0.067956 | 353.8 | 0.0091431 | 584.5 | 2.3914
458.661 | 1.7633 | 276.9 | -0.06064 | 354 | 0.010015 | 585 | 2.3759
458.972 | 1.78 | 277 | -0.058289 | 354.2 | 0.01017 | 585.5 | 2.367
459.284 | 1.7733 | 277.1 | -0.055939 | 354.4 | 0.010461 | 586 | 2.356
459.595 | 1.7933 | 277.2 | -0.064689 | 354.6 | 0.01188 | 586.5 | 2.3605
459.907 | 1.7933 | 277.3 | -0.066573 | 354.8 | 0.014439 | 587 | 2.3917
460.218 | 1.8 | 277.4 | -0.068456 | 355 | 0.017703 | 587.5 | 2.4168
460.529 | 1.79 | 277.5 | -0.070339 | 355.2 | 0.019129 | 588 | 2.3663
460.84 | 1.7967 | 277.6 | -0.059722 | 355.4 | 0.018713 | 588.5 | 2.306
461.151 | 1.8067 | 277.7 | -0.059254 | 355.6 | 0.014472 | 589 | 2.276
461.462 | 1.8 | 277.8 | -0.058786 | 355.8 | 0.015439 | 589.5 | 2.266
461.773 | 1.8133 | 277.9 | -0.060951 | 356 | 0.01667 | 590 | 2.2541
462.084 | 1.8033 | 278 | -0.058124 | 356.2 | 0.017924 | 590.5 | 2.2408
462.395 | 1.82 | 278.1 | -0.055298 | 356.4 | 0.018714 | 591 | 2.2327
462.706 | 1.8167 | 278.2 | -0.052471 | 356.6 | 0.014569 | 591.5 | 2.2166
463.016 | 1.83 | 278.3 | -0.057646 | 356.8 | 0.016394 | 592 | 2.2198
463.327 | 1.83 | 278.4 | -0.059529 | 357 | 0.018794 | 592.5 | 2.2249
463.638 | 1.83 | 278.5 | -0.061412 | 357.2 | 0.01947 | 593 | 2.2401
463.948 | 1.8367 | 278.6 | -0.063296 | 357.4 | 0.021013 | 593.5 | 2.2272
464.258 | 1.84 | 278.7 | -0.055741 | 357.6 | 0.020443 | 594 | 2.174
464.569 | 1.8367 | 278.8 | -0.054799 | 357.8 | 0.020868 | 594.5 | 2.1335
464.879 | 1.85 | 278.9 | -0.053858 | 358 | 0.024202 | 595 | 2.1111
465.189 | 1.8467 | 279 | -0.049637 | 358.2 | 0.025452 | 595.5 | 2.0953
465.499 | 1.8533 | 279.1 | -0.046337 | 358.4 | 0.027023 | 596 | 2.0788
465.809 | 1.86 | 279.2 | -0.043037 | 358.6 | 0.02842 | 596.5 | 2.0626
466.119 | 1.85 | 279.3 | -0.039737 | 358.8 | 0.029229 | 597 | 2.0441
466.429 | 1.87 | 279.4 | -0.049748 | 359 | 0.029854 | 597.5 | 2.0263
466.739 | 1.8533 | 279.5 | -0.05069 | 359.2 | 0.030117 | 598 | 2.0224
467.048 | 1.87 | 279.6 | -0.051631 | 359.4 | 0.028769 | 598.5 | 2.018
467.358 | 1.8767 | 279.7 | -0.049421 | 359.6 | 0.024816 | 599 | 2.0149
467.668 | 1.8767 | 279.8 | -0.050834 | 359.8 | 0.02394 | 599.5 | 2.0153
467.977 | 1.88 | 279.9 | -0.052247 | 360 | 0.026664 | 600 | 1.9839
468.286 | 1.88 | 280 | -0.05366 | 360.2 | 0.030973 | 600.5 | 1.9638
468.596 | 1.88 | 280.1 | -0.054444 | 360.4 | 0.036403 | 601 | 1.9367
468.905 | 1.89 | 280.2 | -0.054443 | 360.6 | 0.038466 | 601.5 | 1.913
469.214 | 1.88 | 280.3 | -0.054442 | 360.8 | 0.038935 | 602 | 1.8847
469.523 | 1.8833 | 280.4 | -0.05348 | 361 | 0.03389 | 602.5 | 1.8686
469.832 | 1.89 | 280.5 | -0.053952 | 361.2 | 0.035242 | 603 | 1.8529
470.141 | 1.89 | 280.6 | -0.054424 | 361.4 | 0.037945 | 603.5 | 1.8407
470.45 | 1.8933 | 280.7 | -0.054896 | 361.6 | 0.043503 | 604 | 1.8229
470.759 | 1.88 | 280.8 | -0.052597 | 361.8 | 0.047963 | 604.5 | 1.8088
471.068 | 1.9 | 280.9 | -0.051653 | 362 | 0.049396 | 605 | 1.7974
471.377 | 1.8933 | 281 | -0.050708 | 362.2 | 0.050297 | 605.5 | 1.7815
471.685 | 1.9 | 281.1 | -0.054681 | 362.4 | 0.05 | 606 | 1.7748
471.994 | 1.9 | 281.2 | -0.055625 | 362.6 | 0.05 | 606.5 | 1.762
472.302 | 1.91 | 281.3 | -0.056569 | 362.8 | 0.05151 | 607 | 1.7436
472.611 | 1.91 | 281.4 | -0.057513 | 363 | 0.056028 | 607.5 | 1.7323
472.919 | 1.91 | 281.5 | -0.051882 | 363.2 | 0.064271 | 608 | 1.727
473.227 | 1.9 | 281.6 | -0.05141 | 363.4 | 0.080461 | 608.5 | 1.7211
473.536 | 1.9033 | 281.7 | -0.050938 | 363.6 | 0.11009 | 609 | 1.7275
473.844 | 1.91 | 281.8 | -0.056756 | 363.8 | 0.20307 | 609.5 | 1.7597
474.152 | 1.91 | 281.9 | -0.057229 | 364 | 0.43255 | 610 | 1.844
474.46 | 1.9167 | 282 | -0.057701 | 364.2 | 0.90283 | 610.5 | 2.0786
474.768 | 1.9133 | 282.1 | -0.058173 | 364.4 | 1.4606 | 611 | 2.5397
475.075 | 1.9167 | 282.2 | -0.046531 | 364.6 | 1.9763 | 611.5 | 2.7919
475.383 | 1.92 | 282.3 | -0.044643 | 364.8 | 2.3086 | 612 | 2.3802
475.691 | 1.93 | 282.4 | -0.042754 | 365 | 2.1372 | 612.5 | 2.0107
475.999 | 1.92 | 282.5 | -0.052161 | 365.2 | 1.7725 | 613 | 1.8715
476.306 | 1.9167 | 282.6 | -0.051689 | 365.4 | 1.274 | 613.5 | 1.8061
476.614 | 1.92 | 282.7 | -0.051217 | 365.6 | 0.94772 | 614 | 1.759
476.921 | 1.9267 | 282.8 | -0.050744 | 365.8 | 0.65348 | 614.5 | 1.7328
477.228 | 1.92 | 282.9 | -0.053394 | 366 | 0.60604 | 615 | 1.6847
477.536 | 1.93 | 283 | -0.057179 | 366.2 | 0.49538 | 615.5 | 1.6113
477.843 | 1.92 | 283.1 | -0.060963 | 366.4 | 0.39324 | 616 | 1.5436
478.15 | 1.9367 | 283.2 | -0.061147 | 366.6 | 0.26167 | 616.5 | 1.502
478.457 | 1.93 | 283.3 | -0.06162 | 366.8 | 0.15568 | 617 | 1.472
478.764 | 1.93 | 283.4 | -0.062093 | 367 | 0.10744 | 617.5 | 1.4358
479.071 | 1.9367 | 283.5 | -0.062566 | 367.2 | 0.077866 | 618 | 1.4227
479.378 | 1.93 | 283.6 | -0.061111 | 367.4 | 0.077295 | 618.5 | 1.4065
479.684 | 1.93 | 283.7 | -0.061112 | 367.6 | 0.077779 | 619 | 1.3883
479.991 | 1.9233 | 283.8 | -0.061112 | 367.8 | 0.078043 | 619.5 | 1.3776
480.298 | 1.9267 | 283.9 | -0.061099 | 368 | 0.07957 | 620 | 1.3606
480.604 | 1.94 | 284 | -0.060626 | 368.2 | 0.08019 | 620.5 | 1.3484
480.911 | 1.9333 | 284.1 | -0.060153 | 368.4 | 0.081176 | 621 | 1.3318
481.217 | 1.94 | 284.2 | -0.059681 | 368.6 | 0.08503 | 621.5 | 1.3126
481.523 | 1.93 | 284.3 | -0.06 | 368.8 | 0.085966 | 622 | 1.2908
481.83 | 1.94 | 284.4 | -0.06 | 369 | 0.083336 | 622.5 | 1.2734
482.136 | 1.9433 | 284.5 | -0.06 | 369.2 | 0.082356 | 623 | 1.2622
482.442 | 1.94 | 284.6 | -0.06 | 369.4 | 0.082617 | 623.5 | 1.263
482.748 | 1.9467 | 284.7 | -0.058425 | 369.6 | 0.090134 | 624 | 1.2341
483.054 | 1.9433 | 284.8 | -0.057951 | 369.8 | 0.094021 | 624.5 | 1.2243
483.36 | 1.9533 | 284.9 | -0.057478 | 370 | 0.094786 | 625 | 1.2106
483.666 | 1.9567 | 285 | -0.054902 | 370.2 | 0.095296 | 625.5 | 1.1908
483.971 | 1.9567 | 285.1 | -0.053482 | 370.4 | 0.094677 | 626 | 1.1656
484.277 | 1.9567 | 285.2 | -0.052061 | 370.6 | 0.098445 | 626.5 | 1.1621
484.583 | 1.96 | 285.3 | -0.050641 | 370.8 | 0.10038 | 627 | 1.1458
484.888 | 1.9633 | 285.4 | -0.052888 | 371 | 0.10076 | 627.5 | 1.1278
485.194 | 1.97 | 285.5 | -0.052415 | 371.2 | 0.10425 | 628 | 1.1104
485.499 | 1.98 | 285.6 | -0.051941 | 371.4 | 0.10666 | 628.5 | 1.097
485.805 | 1.9833 | 285.7 | -0.053731 | 371.6 | 0.1085 | 629 | 1.0935
486.11 | 1.98 | 285.8 | -0.054678 | 371.8 | 0.11051 | 629.5 | 1.08
486.415 | 1.98 | 285.9 | -0.055625 | 372 | 0.11 | 630 | 1.0769
486.72 | 1.9833 | 286 | -0.056572 | 372.2 | 0.11 | 630.5 | 1.1034
487.025 | 1.99 | 286.1 | -0.052907 | 372.4 | 0.11 | 631 | 1.1305
487.33 | 1.98 | 286.2 | -0.052434 | 372.6 | 0.11 | 631.5 | 1.103
487.635 | 1.98 | 286.3 | -0.05196 | 372.8 | 0.11159 | 632 | 1.028
487.94 | 1.98 | 286.4 | -0.051862 | 373 | 0.11627 | 632.5 | 0.97075
488.245 | 1.98 | 286.5 | -0.050915 | 373.2 | 0.11593 | 633 | 0.94788
488.549 | 1.96 | 286.6 | -0.049968 | 373.4 | 0.11711 | 633.5 | 0.92843
488.854 | 1.97 | 286.7 | -0.049021 | 373.6 | 0.12132 | 634 | 0.91848
489.159 | 1.9633 | 286.8 | -0.048482 | 373.8 | 0.12275 | 634.5 | 0.9049
489.463 | 1.9733 | 286.9 | -0.048008 | 374 | 0.12412 | 635 | 0.89689
489.767 | 1.97 | 287 | -0.047535 | 374.2 | 0.1233 | 635.5 | 0.89075
490.072 | 1.9667 | 287.1 | -0.048889 | 374.4 | 0.1228 | 636 | 0.86477
490.376 | 1.95 | 287.2 | -0.048888 | 374.6 | 0.12763 | 636.5 | 0.85813
490.68 | 1.9667 | 287.3 | -0.048888 | 374.8 | 0.12879 | 637 | 0.84661
490.984 | 1.99 | 287.4 | -0.048887 | 375 | 0.13089 | 637.5 | 0.82929
491.288 | 1.99 | 287.5 | -0.047777 | 375.2 | 0.13228 | 638 | 0.81246
491.592 | 2.01 | 287.6 | -0.047777 | 375.4 | 0.13318 | 638.5 | 0.80961
491.896 | 2.0333 | 287.7 | -0.047777 | 375.6 | 0.13662 | 639 | 0.79594
492.2 | 1.9733 | 287.8 | -0.050148 | 375.8 | 0.13737 | 639.5 | 0.78581
492.504 | 1.9433 | 287.9 | -0.050622 | 376 | 0.13774 | 640 | 0.77342
492.808 | 1.94 | 288 | -0.051095 | 376.2 | 0.13749 | 640.5 | 0.76487
493.111 | 1.92 | 288.1 | -0.051569 | 376.4 | 0.13845 | 641 | 0.7534
493.415 | 1.93 | 288.2 | -0.04591 | 376.6 | 0.14415 | 641.5 | 0.74331
493.718 | 1.9333 | 288.3 | -0.04496 | 376.8 | 0.14377 | 642 | 0.73622
494.022 | 1.9267 | 288.4 | -0.044011 | 377 | 0.14333 | 642.5 | 0.72464
494.325 | 1.9167 | 288.5 | -0.043907 | 377.2 | 0.14572 | 643 | 0.71869
494.628 | 1.92 | 288.6 | -0.042009 | 377.4 | 0.14696 | 643.5 | 0.70795
494.931 | 1.9 | 288.7 | -0.040112 | 377.6 | 0.14907 | 644 | 0.69145
495.235 | 1.9167 | 288.8 | -0.038214 | 377.8 | 0.1522 | 644.5 | 0.68773
495.538 | 1.9167 | 288.9 | -0.04 | 378 | 0.15452 | 645 | 0.68055
495.841 | 1.9 | 289 | -0.04 | 378.2 | 0.15951 | 645.5 | 0.66875
496.144 | 1.8933 | 289.1 | -0.04 | 378.4 | 0.16317 | 646 | 0.65556
496.446 | 1.9033 | 289.2 | -0.036301 | 378.6 | 0.16196 | 646.5 | 0.65367
496.749 | 1.89 | 289.3 | -0.034878 | 378.8 | 0.16482 | 647 | 0.64299
497.052 | 1.8933 | 289.4 | -0.033455 | 379 | 0.16659 | 647.5 | 0.63154
497.355 | 1.9 | 289.5 | -0.032032 | 379.2 | 0.16828 | 648 | 0.62682
497.657 | 1.88 | 289.6 | -0.038129 | 379.4 | 0.17037 | 648.5 | 0.617
497.96 | 1.8833 | 289.7 | -0.038602 | 379.6 | 0.17 | 649 | 0.60636
498.262 | 1.89 | 289.8 | -0.039074 | 379.8 | 0.17106 | 649.5 | 0.60334
498.564 | 1.8667 | 289.9 | -0.040546 | 380 | 0.17296 | 650 | 0.61368
498.867 | 1.87 | 290 | -0.04292 | 380.2 | 0.1762 | 650.5 | 0.62231
499.169 | 1.87 | 290.1 | -0.045294 | 380.4 | 0.17802 | 651 | 0.60669
499.471 | 1.8533 | 290.2 | -0.047669 | 380.6 | 0.17761 | 651.5 | 0.58266
499.773 | 1.8667 | 290.3 | -0.04735 | 380.8 | 0.18193 | 652 | 0.56676
500.075 | 1.86 | 290.4 | -0.0483 | 381 | 0.1856 | 652.5 | 0.56835
500.377 | 1.8567 | 290.5 | -0.04925 | 381.2 | 0.18812 | 653 | 0.56001
500.679 | 1.85 | 290.6 | -0.04414 | 381.4 | 0.19237 | 653.5 | 0.54818
500.981 | 1.86 | 290.7 | -0.042715 | 381.6 | 0.19632 | 654 | 0.53621
501.282 | 1.84 | 290.8 | -0.04129 | 381.8 | 0.19938 | 654.5 | 0.52587
501.584 | 1.83 | 290.9 | -0.039864 | 382 | 0.20225 | 655 | 0.51696
501.886 | 1.8333 | 291 | -0.039768 | 382.2 | 0.20292 | 655.5 | 0.50694
502.187 | 1.8333 | 291.1 | -0.037867 | 382.4 | 0.20243 | 656 | 0.50337
502.489 | 1.8333 | 291.2 | -0.035965 | 382.6 | 0.20117 | 656.5 | 0.49579
502.79 | 1.8233 | 291.3 | -0.039079 | 382.8 | 0.2017 | 657 | 0.48291
503.091 | 1.83 | 291.4 | -0.040028 | 383 | 0.20321 | 657.5 | 0.48331
503.392 | 1.82 | 291.5 | -0.040978 | 383.2 | 0.20671 | 658 | 0.48528
503.694 | 1.8267 | 291.6 | -0.041928 | 383.4 | 0.21133 | 658.5 | 0.47662
503.995 | 1.81 | 291.7 | -0.03889 | 383.6 | 0.21668 | 659 | 0.46528
504.296 | 1.82 | 291.8 | -0.03889 | 383.8 | 0.21895 | 659.5 | 0.45784
504.597 | 1.8 | 291.9 | -0.038891 | 384 | 0.22022 | 660 | 0.44973
504.898 | 1.8133 | 292 | -0.033912 | 384.2 | 0.22 | 660.5 | 0.44509
505.198 | 1.7967 | 292.1 | -0.031058 | 384.4 | 0.22306 | 661 | 0.44008
505.499 | 1.8 | 292.2 | -0.028205 | 384.6 | 0.22632 | 661.5 | 0.43915
505.8 | 1.8033 | 292.3 | -0.025351 | 384.8 | 0.22825 | 662 | 0.44662
506.1 | 1.79 | 292.4 | -0.037965 | 385 | 0.23035 | 662.5 | 0.44336
506.401 | 1.7967 | 292.5 | -0.039864 | 385.2 | 0.23 | 663 | 0.43031
506.701 | 1.78 | 292.6 | -0.041763 | 385.4 | 0.23533 | 663.5 | 0.4173
507.002 | 1.79 | 292.7 | -0.037002 | 385.6 | 0.24007 | 664 | 0.40615
507.302 | 1.7833 | 292.8 | -0.038902 | 385.8 | 0.24068 | 664.5 | 0.39713
507.602 | 1.78 | 292.9 | -0.040801 | 386 | 0.24397 | 665 | 0.39042
507.903 | 1.78 | 293 | -0.0427 | 386.2 | 0.24651 | 665.5 | 0.38472
508.203 | 1.78 | 293.1 | -0.041261 | 386.4 | 0.2479 | 666 | 0.37977
508.503 | 1.79 | 293.2 | -0.039833 | 386.6 | 0.24991 | 666.5 | 0.38083
508.803 | 1.77 | 293.3 | -0.038406 | 386.8 | 0.25 | 667 | 0.37929
509.103 | 1.7733 | 293.4 | -0.03897 | 387 | 0.25276 | 667.5 | 0.36919
509.402 | 1.77 | 293.5 | -0.039446 | 387.2 | 0.25643 | 668 | 0.3634
509.702 | 1.7733 | 293.6 | -0.039922 | 387.4 | 0.25836 | 668.5 | 0.35925
510.002 | 1.7667 | 293.7 | -0.040398 | 387.6 | 0.26186 | 669 | 0.35434
510.301 | 1.76 | 293.8 | -0.03857 | 387.8 | 0.26743 | 669.5 | 0.35149
510.601 | 1.76 | 293.9 | -0.038094 | 388 | 0.27196 | 670 | 0.34464
510.9 | 1.7633 | 294 | -0.037617 | 388.2 | 0.27779 | 670.5 | 0.3409
511.2 | 1.7533 | 294.1 | -0.038889 | 388.4 | 0.27975 | 671 | 0.3362
511.499 | 1.74 | 294.2 | -0.038889 | 388.6 | 0.28111 | 671.5 | 0.33225
511.798 | 1.7467 | 294.3 | -0.038889 | 388.8 | 0.28388 | 672 | 0.33208
512.098 | 1.75 | 294.4 | -0.038889 | 389 | 0.2862 | 672.5 | 0.32119
512.397 | 1.74 | 294.5 | -0.034917 | 389.2 | 0.28932 | 673 | 0.31884
512.696 | 1.75 | 294.6 | -0.033965 | 389.4 | 0.29274 | 673.5 | 0.31399
512.995 | 1.75 | 294.7 | -0.033013 | 389.6 | 0.29712 | 674 | 0.30697
513.294 | 1.73 | 294.8 | -0.034283 | 389.8 | 0.30435 | 674.5 | 0.30369
513.593 | 1.7267 | 294.9 | -0.03333 | 390 | 0.31405 | 675 | 0.30031
513.891 | 1.74 | 295 | -0.032378 | 390.2 | 0.32177 | 675.5 | 0.29342
514.19 | 1.73 | 295.1 | -0.031425 | 390.4 | 0.32189 | 676 | 0.28677
514.489 | 1.72 | 295.2 | -0.032541 | 390.6 | 0.325 | 676.5 | 0.27992
514.787 | 1.7267 | 295.3 | -0.033017 | 390.8 | 0.32667 | 677 | 0.28113
515.086 | 1.7233 | 295.4 | -0.033494 | 391 | 0.32673 | 677.5 | 0.27811
515.384 | 1.73 | 295.5 | -0.03214 | 391.2 | 0.32771 | 678 | 0.2748
515.683 | 1.72 | 295.6 | -0.031663 | 391.4 | 0.32706 | 678.5 | 0.26563
515.981 | 1.72 | 295.7 | -0.031186 | 391.6 | 0.3295 | 679 | 0.26485
516.279 | 1.7233 | 295.8 | -0.03071 | 391.8 | 0.33326 | 679.5 | 0.26709
516.577 | 1.72 | 295.9 | -0.029355 | 392 | 0.33761 | 680 | 0.26055
516.875 | 1.72 | 296 | -0.028401 | 392.2 | 0.34115 | 680.5 | 0.26128
517.173 | 1.71 | 296.1 | -0.027447 | 392.4 | 0.34081 | 681 | 0.25812
517.471 | 1.71 | 296.2 | -0.030087 | 392.6 | 0.34396 | 681.5 | 0.24905
517.769 | 1.71 | 296.3 | -0.030563 | 392.8 | 0.34683 | 682 | 0.25258
518.067 | 1.71 | 296.4 | -0.03104 | 393 | 0.34992 | 682.5 | 0.24436
518.365 | 1.71 | 296.5 | -0.031516 | 393.2 | 0.35386 | 683 | 0.24463
518.662 | 1.71 | 296.6 | -0.030327 | 393.4 | 0.35717 | 683.5 | 0.24632
518.96 | 1.7 | 296.7 | -0.030803 | 393.6 | 0.35985 | 684 | 0.23886
519.258 | 1.7 | 296.8 | -0.031279 | 393.8 | 0.36333 | 684.5 | 0.23804
519.555 | 1.71 | 296.9 | -0.028612 | 394 | 0.36809 | 685 | 0.23266
519.852 | 1.71 | 297 | -0.027182 | 394.2 | 0.37288 | 685.5 | 0.23201
520.15 | 1.7 | 297.1 | -0.025752 | 394.4 | 0.37939 | 686 | 0.22535
520.447 | 1.7067 | 297.2 | -0.024321 | 394.6 | 0.3809 | 686.5 | 0.22963
520.744 | 1.7 | 297.3 | -0.027778 | 394.8 | 0.38285 | 687 | 0.23783
521.041 | 1.69 | 297.4 | -0.027778 | 395 | 0.38777 | 687.5 | 0.24006
521.338 | 1.69 | 297.5 | -0.027778 | 395.2 | 0.39093 | 688 | 0.23795
521.635 | 1.6933 | 297.6 | -0.025555 | 395.4 | 0.39557 | 688.5 | 0.21929
521.932 | 1.7 | 297.7 | -0.025554 | 395.6 | 0.39812 | 689 | 0.2159
522.229 | 1.7067 | 297.8 | -0.025554 | 395.8 | 0.40109 | 689.5 | 0.20764
522.526 | 1.69 | 297.9 | -0.025553 | 396 | 0.40755 | 690 | 0.2084
522.823 | 1.69 | 298 | -0.027777 | 396.2 | 0.4126 | 690.5 | 0.2701
523.119 | 1.68 | 298.1 | -0.027776 | 396.4 | 0.41842 | 691 | 0.26546
523.416 | 1.6967 | 298.2 | -0.027775 | 396.6 | 0.424 | 691.5 | 0.20002
523.712 | 1.68 | 298.3 | -0.026772 | 396.8 | 0.42841 | 692 | 0.19533
524.009 | 1.7 | 298.4 | -0.027249 | 397 | 0.4316 | 692.5 | 0.18881
524.305 | 1.69 | 298.5 | -0.027727 | 397.2 | 0.43531 | 693 | 0.19392
524.601 | 1.6967 | 298.6 | -0.028204 | 397.4 | 0.43783 | 693.5 | 0.20266
524.898 | 1.6933 | 298.7 | -0.028889 | 397.6 | 0.43916 | 694 | 0.20166
525.194 | 1.69 | 298.8 | -0.028889 | 397.8 | 0.44571 | 694.5 | 0.18617
525.49 | 1.7 | 298.9 | -0.028889 | 398 | 0.45308 | 695 | 0.17977
525.786 | 1.7 | 299 | -0.026552 | 398.2 | 0.45938 | 695.5 | 0.17663
526.082 | 1.7 | 299.1 | -0.026074 | 398.4 | 0.46554 | 696 | 0.1752
526.378 | 1.69 | 299.2 | -0.025597 | 398.6 | 0.46797 | 696.5 | 0.19292
526.673 | 1.7 | 299.3 | -0.025119 | 398.8 | 0.47243 | 697 | 0.17473
526.969 | 1.69 | 299.4 | -0.024835 | 399 | 0.4766 | 697.5 | 0.16104
527.265 | 1.6967 | 299.5 | -0.023878 | 399.2 | 0.47995 | 698 | 0.16694
527.56 | 1.7067 | 299.6 | -0.022921 | 399.4 | 0.4839 | 698.5 | 0.15471
527.856 | 1.69 | 299.7 | -0.023333 | 399.6 | 0.48721 | 699 | 0.15347
528.151 | 1.7 | 299.8 | -0.023333 | 399.8 | 0.49322 | 699.5 | 0.15845
528.447 | 1.7 | 299.9 | -0.023333 | 400 | 0.49969 | 700 | 0.15789
528.742 | 1.7033 | 300 | -0.023333 | 400.2 | 0.50509 | 700.5 | 0.15886
529.037 | 1.7167 | 300.1 | -0.026301 | 400.4 | 0.51106 | 701 | 0.15348
529.333 | 1.7 | 300.2 | -0.027256 | 400.6 | 0.51634 | 701.5 | 0.15098
529.628 | 1.7 | 300.3 | -0.028211 | 400.8 | 0.52228 | 702 | 0.14863
529.923 | 1.72 | 300.4 | -0.026807 | 401 | 0.52997 | 702.5 | 0.15297
530.218 | 1.7033 | 300.5 | -0.027286 | 401.2 | 0.54044 | 703 | 0.15476
530.513 | 1.72 | 300.6 | -0.027765 | 401.4 | 0.54553 | 703.5 | 0.14939
530.808 | 1.71 | 300.7 | -0.028244 | 401.6 | 0.55036 | 704 | 0.14645
531.102 | 1.7167 | 300.8 | -0.028889 | 401.8 | 0.55493 | 704.5 | 0.14791
531.397 | 1.71 | 300.9 | -0.02889 | 402 | 0.56099 | 705 | 0.14909
531.692 | 1.7133 | 301 | -0.02889 | 402.2 | 0.57362 | 705.5 | 0.15166
531.986 | 1.7267 | 301.1 | -0.029044 | 402.4 | 0.58461 | 706 | 0.17043
532.281 | 1.73 | 301.2 | -0.029522 | 402.6 | 0.598 | 706.5 | 0.2022
532.575 | 1.72 | 301.3 | -0.030001 | 402.8 | 0.61377 | 707 | 0.21061
532.87 | 1.7233 | 301.4 | -0.030479 | 403 | 0.65916 | 707.5 | 0.18929
533.164 | 1.74 | 301.5 | -0.033027 | 403.2 | 0.71596 | 708 | 0.18057
533.458 | 1.73 | 301.6 | -0.033985 | 403.4 | 1.0171 | 708.5 | 0.17997
533.752 | 1.75 | 301.7 | -0.034943 | 403.6 | 1.6352 | 709 | 0.22003
534.046 | 1.7367 | 301.8 | -0.033504 | 403.8 | 2.796 | 709.5 | 0.19123
534.34 | 1.74 | 301.9 | -0.033981 | 404 | 4.5591 | 710 | 0.14562
534.634 | 1.75 | 302 | -0.034459 | 404.2 | 6.5147 | 710.5 | 0.12957
534.928 | 1.74 | 302.1 | -0.034936 | 404.4 | 8.0966 | 711 | 0.13551
535.222 | 1.75 | 302.2 | -0.030969 | 404.6 | 8.7299 | 711.5 | 0.14403
535.516 | 1.75 | 302.3 | -0.029535 | 404.8 | 6.9067 | 712 | 0.18089
535.809 | 1.76 | 302.4 | -0.0281 | 405 | 4.8599 | 712.5 | 0.19027
536.103 | 1.7667 | 302.5 | -0.031111 | 405.2 | 2.5601 | 713 | 0.15526
536.397 | 1.77 | 302.6 | -0.031111 | 405.4 | 1.5025 | 713.5 | 0.12352
536.69 | 1.78 | 302.7 | -0.031111 | 405.6 | 0.83281 | 714 | 0.11752
536.983 | 1.78 | 302.8 | -0.031111 | 405.8 | 0.71222 | 714.5 | 0.11492
537.277 | 1.8 | 302.9 | -0.030436 | 406 | 0.66704 | 715 | 0.10799
537.57 | 1.8067 | 303 | -0.030915 | 406.2 | 0.68607 | 715.5 | 0.10282
537.863 | 1.82 | 303.1 | -0.031394 | 406.4 | 0.69833 | 716 | 0.10232
538.156 | 1.82 | 303.2 | -0.033539 | 406.6 | 0.73035 | 716.5 | 0.1038
538.449 | 1.84 | 303.3 | -0.034018 | 406.8 | 0.78848 | 717 | 0.10801
538.742 | 1.8533 | 303.4 | -0.034497 | 407 | 0.88356 | 717.5 | 0.1027
539.035 | 1.85 | 303.5 | -0.034976 | 407.2 | 0.99658 | 718 | 0.10227
539.328 | 1.87 | 303.6 | -0.032222 | 407.4 | 1.1658 | 718.5 | 0.10442
539.621 | 1.88 | 303.7 | -0.032222 | 407.6 | 1.2366 | 719 | 0.10173
539.914 | 1.9033 | 303.8 | -0.032222 | 407.8 | 1.231 | 719.5 | 0.096086
540.206 | 1.92 | 303.9 | -0.031768 | 408 | 1.0808 | 720 | 0.095009
540.499 | 1.9433 | 304 | -0.030808 | 408.2 | 0.9622 | 720.5 | 0.091455
540.792 | 1.97 | 304.1 | -0.029849 | 408.4 | 0.81068 | 721 | 0.093527
541.084 | 2.02 | 304.2 | -0.028889 | 408.6 | 0.75805 | 721.5 | 0.093373
541.376 | 2.06 | 304.3 | -0.03 | 408.8 | 0.74878 | 722 | 0.084819
541.669 | 2.13 | 304.4 | -0.03 | 409 | 0.75667 | 722.5 | 0.083002
541.961 | 2.18 | 304.5 | -0.03 | 409.2 | 0.7577 | 723 | 0.082101
542.253 | 2.18 | 304.6 | -0.025914 | 409.4 | 0.76109 | 723.5 | 0.086667
542.545 | 2.18 | 304.7 | -0.024476 | 409.6 | 0.7653 | 724 | 0.083883
542.837 | 2.19 | 304.8 | -0.023038 | 409.8 | 0.77006 | 724.5 | 0.086294
543.129 | 2.17 | 304.9 | -0.028934 | 410 | 0.77632 | 725 | 0.088049
543.421 | 2.17 | 305 | -0.029895 | 410.2 | 0.78251 | 725.5 | 0.078031
543.713 | 2.21 | 305.1 | -0.030856 | 410.4 | 0.78823 | 726 | 0.079171
544.005 | 2.21 | 305.2 | -0.031817 | 410.6 | 0.79297 | 726.5 | 0.077937
544.296 | 2.23 | 305.3 | -0.026943 | 410.8 | 0.7985 | 727 | 0.076349
544.588 | 2.2533 | 305.4 | -0.027425 | 411 | 0.80328 | 727.5 | 0.079147
544.88 | 2.37 | 305.5 | -0.027906 | 411.2 | 0.80677 | 728 | 0.074658
545.171 | 3.03 | 305.6 | -0.029907 | 411.4 | 0.8097 | 728.5 | 0.075133
545.463 | 5.86 | 305.7 | -0.02895 | 411.6 | 0.80987 | 729 | 0.077172
545.754 | 9.72 | 305.8 | -0.027992 | 411.8 | 0.81585 | 729.5 | 0.067612
546.045 | 12.6867 | 305.9 | -0.027034 | 412 | 0.82198 | 730 | 0.076161
546.337 | 17.7067 | 306 | -0.024146 | 412.2 | 0.82671 | 730.5 | 0.072376
546.628 | 8.87 | 306.1 | -0.023668 | 412.4 | 0.83241 | 731 | 0.069511
546.919 | 3.09 | 306.2 | -0.023189 | 412.6 | 0.83682 | 731.5 | 0.067677
547.21 | 2.3633 | 306.3 | -0.026667 | 412.8 | 0.83896 | 732 | 0.071703
547.501 | 2.21 | 306.4 | -0.026667 | 413 | 0.84339 | 732.5 | 0.065321
547.792 | 2.16 | 306.5 | -0.026667 | 413.2 | 0.85218 | 733 | 0.062109
548.083 | 2.13 | 306.6 | -0.026667 | 413.4 | 0.85616 | 733.5 | 0.065358
548.373 | 2.14 | 306.7 | -0.027646 | 413.6 | 0.86066 | 734 | 0.063396
548.664 | 2.1367 | 306.8 | -0.029087 | 413.8 | 0.8636 | 734.5 | 0.061007
548.955 | 2.13 | 306.9 | -0.030528 | 414 | 0.86778 | 735 | 0.057464
549.245 | 2.14 | 307 | -0.026364 | 414.2 | 0.87324 | 735.5 | 0.058148
549.536 | 2.13 | 307.1 | -0.024923 | 414.4 | 0.87316 | 736 | 0.064978
549.826 | 2.1433 | 307.2 | -0.023482 | 414.6 | 0.87353 | 736.5 | 0.06239
550.116 | 2.16 | 307.3 | -0.022041 | 414.8 | 0.87642 | 737 | 0.065404
550.407 | 2.15 | 307.4 | -0.0252 | 415 | 0.88201 | 737.5 | 0.061745
550.697 | 2.15 | 307.5 | -0.02472 | 415.2 | 0.8894 | 738 | 0.074311
550.987 | 2.15 | 307.6 | -0.02424 | 415.4 | 0.89495 | 738.5 | 0.10026
551.277 | 2.1633 | 307.7 | -0.022222 | 415.6 | 0.89915 | 739 | 0.067207
551.567 | 2.16 | 307.8 | -0.022222 | 415.8 | 0.90202 | 739.5 | 0.057518
551.857 | 2.17 | 307.9 | -0.022222 | 416 | 0.9055 | 740 | 0.048411
552.147 | 2.18 | 308 | -0.022222 | 416.2 | 0.90955 | 740.5 | 0.051843
552.437 | 2.19 | 308.1 | -0.022949 | 416.4 | 0.91521 | 741 | 0.060068
552.727 | 2.18 | 308.2 | -0.022469 | 416.6 | 0.92107 | 741.5 | 0.057078
553.016 | 2.18 | 308.3 | -0.021988 | 416.8 | 0.92617 | 742 | 0.053785
553.306 | 2.18 | 308.4 | -0.023654 | 417 | 0.92846 | 742.5 | 0.058153
553.595 | 2.19 | 308.5 | -0.024616 | 417.2 | 0.9316 | 743 | 0.057507
553.885 | 2.1967 | 308.6 | -0.025578 | 417.4 | 0.93424 | 743.5 | 0.056742
554.174 | 2.1933 | 308.7 | -0.02654 | 417.6 | 0.94024 | 744 | 0.043327
554.464 | 2.19 | 308.8 | -0.022916 | 417.8 | 0.94748 | 744.5 | 0.045387
554.753 | 2.19 | 308.9 | -0.022434 | 418 | 0.95439 | 745 | 0.045211
555.042 | 2.2 | 309 | -0.021952 | 418.2 | 0.96101 | 745.5 | 0.045952
555.331 | 2.2 | 309.1 | -0.024444 | 418.4 | 0.96524 | 746 | 0.048903
555.62 | 2.1833 | 309.2 | -0.024444 | 418.6 | 0.96501 | 746.5 | 0.045831
555.909 | 2.1767 | 309.3 | -0.024444 | 418.8 | 0.96959 | 747 | 0.047488
556.198 | 2.18 | 309.4 | -0.024444 | 419 | 0.98032 | 747.5 | 0.04677
556.487 | 2.18 | 309.5 | -0.022223 | 419.2 | 0.98515 | 748 | 0.048364
556.776 | 2.18 | 309.6 | -0.022224 | 419.4 | 0.99146 | 748.5 | 0.03766
557.065 | 2.1833 | 309.7 | -0.022225 | 419.6 | 0.99562 | 749 | 0.040281
557.353 | 2.19 | 309.8 | -0.021769 | 419.8 | 1.0006 | 749.5 | 0.045281
557.642 | 2.21 | 309.9 | -0.020806 | 420 | 1.0059 | 750 | 0.051301
557.931 | 2.2 | 310 | -0.019843 | 420.2 | 1.0083 | 750.5 | 0.086711
558.219 | 2.22 | 310.1 | -0.021114 | 420.4 | 1.0104 | 751 | 0.061801
558.507 | 2.22 | 310.2 | -0.021596 | 420.6 | 1.01 | 751.5 | 0.07462
558.796 | 2.23 | 310.3 | -0.022077 | 420.8 | 1.0123 | 752 | 0.055282
559.084 | 2.21 | 310.4 | -0.022558 | 421 | 1.0164 | 752.5 | 0.038676
559.372 | 2.22 | 310.5 | -0.019739 | 421.2 | 1.0234 | 753 | 0.033777
559.66 | 2.23 | 310.6 | -0.019259 | 421.4 | 1.0269 | 753.5 | 0.02797
559.948 | 2.2367 | 310.7 | -0.018778 | 421.6 | 1.0294 | 754 | 0.026814
560.236 | 2.25 | 310.8 | -0.021111 | 421.8 | 1.03 | 754.5 | 0.028419
560.524 | 2.24 | 310.9 | -0.021111 | 422 | 1.0331 | 755 | 0.027916
560.812 | 2.25 | 311 | -0.021111 | 422.2 | 1.0424 | 755.5 | 0.028676
561.1 | 2.26 | 311.1 | -0.021111 | 422.4 | 1.0475 | 756 | 0.026953
561.388 | 2.27 | 311.2 | -0.024528 | 422.6 | 1.0545 | 756.5 | 0.025339
561.675 | 2.2733 | 311.3 | -0.026455 | 422.8 | 1.0622 | 757 | 0.028733
561.963 | 2.27 | 311.4 | -0.028382 | 423 | 1.065 | 757.5 | 0.025775
562.251 | 2.29 | 311.5 | -0.026744 | 423.2 | 1.071 | 758 | 0.023334
562.538 | 2.28 | 311.6 | -0.027225 | 423.4 | 1.0837 | 758.5 | 0.018717
562.825 | 2.2867 | 311.7 | -0.027707 | 423.6 | 1.0905 | 759 | 0.018945
563.113 | 2.28 | 311.8 | -0.028189 | 423.8 | 1.0943 | 759.5 | 0.019092
563.4 | 2.29 | 311.9 | -0.021647 | 424 | 1.1048 | 760 | 0.025656
563.687 | 2.3067 | 312 | -0.019239 | 424.2 | 1.113 | 760.5 | 0.024862
563.974 | 2.3 | 312.1 | -0.01683 | 424.4 | 1.1227 | 761 | 0.018577
564.261 | 2.31 | 312.2 | -0.022222 | 424.6 | 1.125 | 761.5 | 0.018241
564.548 | 2.31 | 312.3 | -0.022222 | 424.8 | 1.1244 | 762 | 0.017448
564.835 | 2.31 | 312.4 | -0.022222 | 425 | 1.1267 | 762.5 | 0.016725
565.122 | 2.32 | 312.5 | -0.022222 | 425.2 | 1.1256 | 763 | 0.030473
565.409 | 2.33 | 312.6 | -0.01629 | 425.4 | 1.1233 | 763.5 | 0.13676
565.696 | 2.32 | 312.7 | -0.01581 | 425.6 | 1.1278 | 764 | 0.09796
565.983 | 2.3267 | 312.8 | -0.01533 | 425.8 | 1.132 | 764.5 | 0.032977
566.269 | 2.32 | 312.9 | -0.018574 | 426 | 1.1368 | 765 | 0.024216
566.556 | 2.3433 | 313 | -0.017611 | 426.2 | 1.1355 | 765.5 | 0.014041
566.842 | 2.34 | 313.1 | -0.016648 | 426.4 | 1.1307 | 766 | 0.018175
567.129 | 2.35 | 313.2 | -0.015686 | 426.6 | 1.1255 | 766.5 | 0.016446
567.415 | 2.34 | 313.3 | -0.019463 | 426.8 | 1.1226 | 767 | 0.015386
567.701 | 2.35 | 313.4 | -0.021393 | 427 | 1.1265 | 767.5 | 0.015795
567.987 | 2.36 | 313.5 | -0.023323 | 427.2 | 1.1384 | 768 | 0.01679
568.273 | 2.36 | 313.6 | -0.022423 | 427.4 | 1.1524 | 768.5 | 0.017396
568.56 | 2.35 | 313.7 | -0.022905 | 427.6 | 1.1627 | 769 | 0.011284
568.846 | 2.37 | 313.8 | -0.023387 | 427.8 | 1.1714 | 769.5 | 0.01558
569.132 | 2.35 | 313.9 | -0.02387 | 428 | 1.1773 | 770 | 0.014638
569.417 | 2.3633 | 314 | -0.020182 | 428.2 | 1.1835 | 770.5 | 0.0071187
569.703 | 2.36 | 314.1 | -0.019218 | 428.4 | 1.1882 | 771 | 0.0050756
569.989 | 2.37 | 314.2 | -0.018254 | 428.6 | 1.1897 | 771.5 | 0.003564
570.275 | 2.37 | 314.3 | -0.019756 | 428.8 | 1.1975 | 772 | 0.017136
570.56 | 2.38 | 314.4 | -0.019274 | 429 | 1.2042 | 772.5 | 0.059025
570.846 | 2.39 | 314.5 | -0.018793 | 429.2 | 1.2082 | 773 | 0.040301
571.131 | 2.37 | 314.6 | -0.021188 | 429.4 | 1.2168 | 773.5 | 0.018419
571.417 | 2.37 | 314.7 | -0.022635 | 429.6 | 1.2255 | 774 | 0.010831
571.702 | 2.38 | 314.8 | -0.024082 | 429.8 | 1.2375 | 774.5 | 0.005653
571.987 | 2.38 | 314.9 | -0.025529 | 430 | 1.2455 | 775 | 0.0031149
572.272 | 2.39 | 315 | -0.022222 | 430.2 | 1.2518 | 775.5 | 0.0065158
572.558 | 2.39 | 315.1 | -0.022222 | 430.4 | 1.267 | 776 | 0.0025613
572.843 | 2.4 | 315.2 | -0.022222 | 430.6 | 1.2809 | 776.5 | 0.0018485
573.128 | 2.39 | 315.3 | -0.022077 | 430.8 | 1.2912 | 777 | 0.0053592
573.413 | 2.3867 | 315.4 | -0.021111 | 431 | 1.3006 | 777.5 | 0.0041627
573.698 | 2.3933 | 315.5 | -0.020145 | 431.2 | 1.3054 | 778 | 0.0091068
573.982 | 2.38 | 315.6 | -0.019179 | 431.4 | 1.3083 | 778.5 | 0.0046074
574.267 | 2.3967 | 315.7 | -0.021449 | 431.6 | 1.3044 | 779 | 0.0024048
574.552 | 2.41 | 315.8 | -0.021932 | 431.8 | 1.2967 | 779.5 | 0.0029026
574.836 | 2.3833 | 315.9 | -0.022415 | 432 | 1.2889 | 780 | -0.00030544
575.121 | 2.39 | 316 | -0.024807 | 432.2 | 1.285 | 780.5 | -0.0086495
575.405 | 2.4 | 316.1 | -0.026256 | 432.4 | 1.2928 | 781 | 0.0021509
575.69 | 2.4067 | 316.2 | -0.027705 | 432.6 | 1.3026 | 781.5 | 0.0044881
575.974 | 2.45 | 316.3 | -0.029155 | 432.8 | 1.3192 | 782 | 0.010598
576.259 | 2.7233 | 316.4 | -0.027053 | 433 | 1.3417 | 782.5 | 0.00040561
576.543 | 3.43 | 316.5 | -0.027536 | 433.2 | 1.3669 | 783 | -0.0065238
576.827 | 3.9467 | 316.6 | -0.028019 | 433.4 | 1.3995 | 783.5 | -0.0012434
577.111 | 4.8267 | 316.7 | -0.026329 | 433.6 | 1.4403 | 784 | -0.0057022
577.395 | 4.22 | 316.8 | -0.025362 | 433.8 | 1.4879 | 784.5 | -0.0026293
577.679 | 2.71 | 316.9 | -0.024396 | 434 | 1.5547 | 785 | -0.0081622
577.963 | 2.5033 | 317 | -0.02343 | 434.2 | 1.6627 | 785.5 | -0.0091213
578.247 | 2.57 | 317.1 | -0.022459 | 434.4 | 1.9419 | 786 | -0.013584
578.53 | 3.07 | 317.2 | -0.021492 | 434.6 | 2.5581 | 786.5 | -0.016664
578.814 | 3.6933 | 317.3 | -0.020524 | 434.8 | 4.2859 | 787 | -0.009606
579.098 | 4.2967 | 317.4 | -0.02089 | 435 | 6.8425 | 787.5 | -0.012667
579.381 | 4.7367 | 317.5 | -0.020407 | 435.2 | 10.3198 | 788 | -0.010958
579.665 | 3.07 | 317.6 | -0.019923 | 435.4 | 13.0113 | 788.5 | -0.011083
579.948 | 2.5367 | 317.7 | -0.017765 | 435.6 | 15.5977 | 789 | -0.015759
580.232 | 2.4567 | 317.8 | -0.016798 | 435.8 | 16.531 | 789.5 | -0.016835
580.515 | 2.46 | 317.9 | -0.015831 | 436 | 15.3527 | 790 | -0.0079638
580.798 | 2.4333 | 318 | -0.014864 | 436.2 | 11.2053 | 790.5 | -0.011144
581.081 | 2.4233 | 318.1 | -0.019164 | 436.4 | 7.5454 | 791 | -0.018677
581.364 | 2.4133 | 318.2 | -0.019649 | 436.6 | 3.886 | 791.5 | -0.012606
581.648 | 2.4133 | 318.3 | -0.020134 | 436.8 | 2.4434 | 792 | -0.0094317
581.931 | 2.3933 | 318.4 | -0.02135 | 437 | 1.5783 | 792.5 | -0.0099605
582.213 | 2.41 | 318.5 | -0.023285 | 437.2 | 1.487 | 793 | -0.014157
582.496 | 2.41 | 318.6 | -0.025221 | 437.4 | 1.4514 | 793.5 | -0.012907
582.779 | 2.4167 | 318.7 | -0.027157 | 437.6 | 1.4694 | 794 | -0.013546
583.062 | 2.4033 | 318.8 | -0.023004 | 437.8 | 1.4711 | 794.5 | -0.021069
583.344 | 2.4167 | 318.9 | -0.022519 | 438 | 1.4741 | 795 | 0.019755
583.627 | 2.4067 | 319 | -0.022035 | 438.2 | 1.4797 | 795.5 | 0.014364
583.909 | 2.4033 | 319.1 | -0.023217 | 438.4 | 1.4905 | 796 | -0.0075927
584.192 | 2.3933 | 319.2 | -0.022733 | 438.6 | 1.4988 | 796.5 | -0.021441
584.474 | 2.4 | 319.3 | -0.022248 | 438.8 | 1.5071 | 797 | -0.021688
584.757 | 2.38 | 319.4 | -0.021764 | 439 | 1.5137 | 797.5 | -0.020727
585.039 | 2.3667 | 319.5 | -0.021111 | 439.2 | 1.5173 | 798 | -0.01563
585.321 | 2.38 | 319.6 | -0.02111 | 439.4 | 1.5198 | 798.5 | -0.016168
585.603 | 2.36 | 319.7 | -0.02111 | 439.6 | 1.5278 | 799 | -0.022627
585.885 | 2.3533 | 319.8 | -0.02094 | 439.8 | 1.5346 | 799.5 | -0.0214
586.167 | 2.3567 | 319.9 | -0.020456 | 440 | 1.5387 | 800 | -0.028445
586.449 | 2.3533 | 320 | -0.019972 | 440.2 | 1.5461 | 800.5 | -0.014393
586.731 | 2.37 | 320.1 | -0.019488 | 440.4 | 1.5504 | 801 | 0.0012116
587.013 | 2.39 | 320.2 | -0.021552 | 440.6 | 1.5511 | 801.5 | 0.013054
587.295 | 2.4233 | 320.3 | -0.022036 | 440.8 | 1.5548 | 802 | 0.044065
587.576 | 2.4267 | 320.4 | -0.022521 | 441 | 1.5554 | 802.5 | 0.0032261
587.858 | 2.3933 | 320.5 | -0.021111 | 441.2 | 1.5533 | 803 | -0.0047523
588.139 | 2.34 | 320.6 | -0.021112 | 441.4 | 1.5534 | 803.5 | -0.0084529
588.421 | 2.31 | 320.7 | -0.021112 | 441.6 | 1.5524 | 804 | -0.0037085
588.702 | 2.29 | 320.8 | -0.021093 | 441.8 | 1.5547 | 804.5 | -0.00048137
588.984 | 2.27 | 320.9 | -0.020608 | 442 | 1.5595 | 805 | -0.0025725
589.265 | 2.27 | 321 | -0.020123 | 442.2 | 1.5653 | 805.5 | -0.0035153
589.546 | 2.27 | 321.1 | -0.019638 | 442.4 | 1.5702 | 806 | -0.0040169
589.827 | 2.25 | 321.2 | -0.015795 | 442.6 | 1.5755 | 806.5 | -0.02294
590.108 | 2.2567 | 321.3 | -0.014341 | 442.8 | 1.5812 | 807 | -0.040237
590.389 | 2.2467 | 321.4 | -0.012888 | 443 | 1.588 | 807.5 | -0.033807
590.67 | 2.23 | 321.5 | -0.019044 | 443.2 | 1.5956
590.951 | 2.24 | 321.6 | -0.020013 | 443.4 | 1.5989
591.232 | 2.2267 | 321.7 | -0.020982 | 443.6 | 1.6036
591.513 | 2.2067 | 321.8 | -0.021951 | 443.8 | 1.6124
591.794 | 2.2267 | 321.9 | -0.019943 | 444 | 1.612
592.074 | 2.22 | 322 | -0.0214 | 444.2 | 1.6113
592.355 | 2.2167 | 322.1 | -0.022857 | 444.4 | 1.6098
592.635 | 2.23 | 322.2 | -0.021944 | 444.6 | 1.61
592.916 | 2.24 | 322.3 | -0.020974 | 444.8 | 1.6117
593.196 | 2.25 | 322.4 | -0.020004 | 445 | 1.6111
593.476 | 2.23 | 322.5 | -0.019034 | 445.2 | 1.6119
593.757 | 2.21 | 322.6 | -0.02 | 445.4 | 1.6122
594.037 | 2.1633 | 322.7 | -0.02 | 445.6 | 1.6127
594.317 | 2.14 | 322.8 | -0.02 | 445.8 | 1.6145
594.597 | 2.13 | 322.9 | -0.02 | 446 | 1.6174
594.877 | 2.1167 | 323 | -0.02 | 446.2 | 1.6154
595.157 | 2.1033 | 323.1 | -0.02 | 446.4 | 1.6131
595.437 | 2.0933 | 323.2 | -0.02 | 446.6 | 1.611
595.716 | 2.0933 | 323.3 | -0.023184 | 446.8 | 1.6098
595.996 | 2.0733 | 323.4 | -0.024156 | 447 | 1.61
596.276 | 2.0733 | 323.5 | -0.025128 | 447.2 | 1.61
596.555 | 2.06 | 323.6 | -0.023603 | 447.4 | 1.6128
596.835 | 2.05 | 323.7 | -0.024088 | 447.6 | 1.6166
597.114 | 2.04 | 323.8 | -0.024572 | 447.8 | 1.6224
597.394 | 2.03 | 323.9 | -0.026726 | 448 | 1.6283
597.673 | 2.0167 | 324 | -0.027211 | 448.2 | 1.6315
597.952 | 2.03 | 324.1 | -0.027695 | 448.4 | 1.632
598.232 | 2.0133 | 324.2 | -0.02818 | 448.6 | 1.6338
598.511 | 2.0267 | 324.3 | -0.024109 | 448.8 | 1.6417
598.79 | 2.0033 | 324.4 | -0.023623 | 449 | 1.6424
599.069 | 2.0133 | 324.5 | -0.023137 | 449.2 | 1.6442
599.348 | 2.0267 | 324.6 | -0.024192 | 449.4 | 1.6467
599.627 | 2.0133 | 324.7 | -0.02322 | 449.6 | 1.6473
599.905 | 1.98 | 324.8 | -0.022248 | 449.8 | 1.65
600.184 | 1.9833 | 324.9 | -0.021276 | 450 | 1.6505
600.463 | 1.96 | 325 | -0.024144 | 450.2 | 1.6531
600.742 | 1.9567 | 325.1 | -0.025118 | 450.4 | 1.6624
601.02 | 1.9367 | 325.2 | -0.026091 | 450.6 | 1.6666
601.299 | 1.91 | 325.3 | -0.021826 | 450.8 | 1.6706
601.577 | 1.92 | 325.4 | -0.020852 | 451 | 1.6719
601.855 | 1.89 | 325.5 | -0.019878 | 451.2 | 1.6768
602.134 | 1.87 | 325.6 | -0.018904 | 451.4 | 1.6825
602.412 | 1.8833 | 325.7 | -0.024921 | 451.6 | 1.6867
602.69 | 1.8533 | 325.8 | -0.025407 | 451.8 | 1.6854
602.968 | 1.8533 | 325.9 | -0.025892 | 452 | 1.681
603.246 | 1.85 | 326 | -0.023873 | 452.2 | 1.6835
603.524 | 1.84 | 326.1 | -0.024847 | 452.4 | 1.6856
603.802 | 1.83 | 326.2 | -0.025822 | 452.6 | 1.6874
604.08 | 1.8133 | 326.3 | -0.02426 | 452.8 | 1.6913
604.358 | 1.82 | 326.4 | -0.0228 | 453 | 1.6938
604.636 | 1.8 | 326.5 | -0.02134 | 453.2 | 1.6959
604.913 | 1.8 | 326.6 | -0.019881 | 453.4 | 1.7013
605.191 | 1.7967 | 326.7 | -0.020769 | 453.6 | 1.7057
605.468 | 1.7833 | 326.8 | -0.020282 | 453.8 | 1.7084
605.746 | 1.77 | 326.9 | -0.019795 | 454 | 1.7102
606.023 | 1.78 | 327 | -0.02 | 454.2 | 1.7112
606.301 | 1.77 | 327.1 | -0.02 | 454.4 | 1.7151
606.578 | 1.76 | 327.2 | -0.02 | 454.6 | 1.7194
606.855 | 1.7433 | 327.3 | -0.02 | 454.8 | 1.7263
607.132 | 1.74 | 327.4 | -0.024578 | 455 | 1.729
607.409 | 1.74 | 327.5 | -0.026038 | 455.2 | 1.7303
607.686 | 1.72 | 327.6 | -0.027497 | 455.4 | 1.73
607.963 | 1.73 | 327.7 | -0.025975 | 455.6 | 1.73
608.24 | 1.7267 | 327.8 | -0.026949 | 455.8 | 1.7314
608.517 | 1.72 | 327.9 | -0.027924 | 456 | 1.7357
608.794 | 1.7167 | 328 | -0.022193 | 456.2 | 1.7409
609.071 | 1.73 | 328.1 | -0.019269 | 456.4 | 1.7463
609.347 | 1.74 | 328.2 | -0.016345 | 456.6 | 1.7486
609.624 | 1.7733 | 328.3 | -0.013421 | 456.8 | 1.7529
609.9 | 1.8 | 328.4 | -0.018314 | 457 | 1.7567
610.177 | 1.89 | 328.5 | -0.017339 | 457.2 | 1.7588
610.453 | 2.01 | 328.6 | -0.016365 | 457.4 | 1.7604
610.729 | 2.25 | 328.7 | -0.016915 | 457.6 | 1.76
611.006 | 2.56 | 328.8 | -0.018377 | 457.8 | 1.7671
611.282 | 2.8633 | 328.9 | -0.019839 | 458 | 1.7734
611.558 | 2.8733 | 329 | -0.021301 | 458.2 | 1.7746
611.834 | 2.57 | 329.1 | -0.018523 | 458.4 | 1.7719
612.11 | 2.2267 | 329.2 | -0.018036 | 458.6 | 1.7688
612.386 | 2.03 | 329.3 | -0.017549 | 458.8 | 1.7737
612.662 | 1.9367 | 329.4 | -0.018889 | 459 | 1.7739
612.938 | 1.8767 | 329.5 | -0.018889 | 459.2 | 1.7776
613.213 | 1.8333 | 329.6 | -0.018889 | 459.4 | 1.7847
613.489 | 1.8033 | 329.7 | -0.018889 | 459.6 | 1.7879
613.765 | 1.7833 | 329.8 | -0.019335 | 459.8 | 1.7932
614.04 | 1.7433 | 329.9 | -0.019823 | 460 | 1.7969
614.316 | 1.7467 | 330 | -0.020311 | 460.2 | 1.7968
614.591 | 1.7333 | 330.1 | -0.02 | 460.4 | 1.7935
614.867 | 1.7 | 330.2 | -0.02 | 460.6 | 1.7942
615.142 | 1.6733 | 330.3 | -0.02 | 460.8 | 1.796
615.417 | 1.62 | 330.4 | -0.02 | 461 | 1.8021
615.692 | 1.5833 | 330.5 | -0.02 | 461.2 | 1.8027
615.967 | 1.54 | 330.6 | -0.02 | 461.4 | 1.8038
616.242 | 1.52 | 330.7 | -0.02 | 461.6 | 1.8081
616.517 | 1.5 | 330.8 | -0.024304 | 461.8 | 1.8071
616.792 | 1.4833 | 330.9 | -0.025767 | 462 | 1.8085
617.067 | 1.47 | 331 | -0.027231 | 462.2 | 1.8135
617.342 | 1.45 | 331.1 | -0.023333 | 462.4 | 1.8145
617.617 | 1.42 | 331.2 | -0.023332 | 462.6 | 1.8177
617.891 | 1.43 | 331.3 | -0.02333 | 462.8 | 1.823
618.166 | 1.42 | 331.4 | -0.023329 | 463 | 1.8266
618.441 | 1.4067 | 331.5 | -0.022113 | 463.2 | 1.8295
618.715 | 1.4 | 331.6 | -0.020648 | 463.4 | 1.8305
618.99 | 1.39 | 331.7 | -0.019183 | 463.6 | 1.831
619.264 | 1.3733 | 331.8 | -0.02 | 463.8 | 1.834
619.538 | 1.3867 | 331.9 | -0.02 | 464 | 1.8365
619.812 | 1.36 | 332 | -0.02 | 464.2 | 1.8385
620.087 | 1.36 | 332.1 | -0.02 | 464.4 | 1.8378
620.361 | 1.35 | 332.2 | -0.02 | 464.6 | 1.8416
620.635 | 1.3467 | 332.3 | -0.02 | 464.8 | 1.845
620.909 | 1.3367 | 332.4 | -0.02 | 465 | 1.8464
621.183 | 1.32 | 332.5 | -0.02 | 465.2 | 1.8496
621.456 | 1.32 | 332.6 | -0.02 | 465.4 | 1.8517
621.73 | 1.3 | 332.7 | -0.02 | 465.6 | 1.8555
622.004 | 1.29 | 332.8 | -0.017592 | 465.8 | 1.8571
622.278 | 1.28 | 332.9 | -0.016615 | 466 | 1.8534
622.551 | 1.2733 | 333 | -0.015637 | 466.2 | 1.859
622.825 | 1.26 | 333.1 | -0.014659 | 466.4 | 1.8619
623.098 | 1.26 | 333.2 | -0.010316 | 466.6 | 1.8587
623.372 | 1.2633 | 333.3 | -0.0078723 | 466.8 | 1.8627
623.645 | 1.27 | 333.4 | -0.0054285 | 467 | 1.8658
623.918 | 1.23 | 333.5 | -0.0048322 | 467.2 | 1.8724
624.192 | 1.22 | 333.6 | -0.0028772 | 467.4 | 1.8757
624.465 | 1.23 | 333.7 | -0.00092212 | 467.6 | 1.877
624.738 | 1.22 | 333.8 | 0.0010329 | 467.8 | 1.8785
625.011 | 1.21 | 333.9 | -0.0022222 | 468 | 1.8792
625.284 | 1.2 | 334 | -0.0022222 | 468.2 | 1.8799
625.557 | 1.19 | 334.1 | -0.0022222 | 468.4 | 1.88
625.83 | 1.1733 | 334.2 | -0.0094023 | 468.6 | 1.8828
626.102 | 1.1567 | 334.3 | -0.012338 | 468.8 | 1.8866
626.375 | 1.16 | 334.4 | -0.015274 | 469 | 1.885
626.648 | 1.17 | 334.5 | -0.014803 | 469.2 | 1.8834
626.92 | 1.14 | 334.6 | -0.017249 | 469.4 | 1.8824
627.193 | 1.1433 | 334.7 | -0.019696 | 469.6 | 1.8848
627.465 | 1.1367 | 334.8 | -0.022142 | 469.8 | 1.8882
627.738 | 1.11 | 334.9 | -0.021473 | 470 | 1.8896
628.01 | 1.12 | 335 | -0.021962 | 470.2 | 1.8914
628.282 | 1.09 | 335.1 | -0.02245 | 470.4 | 1.8905
628.555 | 1.1 | 335.2 | -0.021111 | 470.6 | 1.8854
628.827 | 1.0933 | 335.3 | -0.021112 | 470.8 | 1.8893
629.099 | 1.1 | 335.4 | -0.021112 | 471 | 1.8927
629.371 | 1.08 | 335.5 | -0.021113 | 471.2 | 1.894
629.643 | 1.0767 | 335.6 | -0.020655 | 471.4 | 1.8972
629.915 | 1.07 | 335.7 | -0.020165 | 471.6 | 1.8978
630.187 | 1.08 | 335.8 | -0.019675 | 471.8 | 1.899
630.458 | 1.1 | 335.9 | -0.02 | 472 | 1.9029
630.73 | 1.12 | 336 | -0.02 | 472.2 | 1.9067
631.002 | 1.1367 | 336.1 | -0.02 | 472.4 | 1.9088
631.273 | 1.14 | 336.2 | -0.02 | 472.6 | 1.9105
631.545 | 1.1 | 336.3 | -0.02 | 472.8 | 1.91
631.816 | 1.0533 | 336.4 | -0.02 | 473 | 1.9064
632.088 | 1.01 | 336.5 | -0.02 | 473.2 | 1.9032
632.359 | 0.98333 | 336.6 | -0.02 | 473.4 | 1.9026
632.63 | 0.95333 | 336.7 | -0.02 | 473.6 | 1.9047
632.902 | 0.96 | 336.8 | -0.02 | 473.8 | 1.9081
633.173 | 0.93333 | 336.9 | -0.018732 | 474 | 1.9095
633.444 | 0.93 | 337 | -0.018242 | 474.2 | 1.9124
633.715 | 0.92333 | 337.1 | -0.017752 | 474.4 | 1.9141
633.986 | 0.92667 | 337.2 | -0.017261 | 474.6 | 1.9141
634.257 | 0.90333 | 337.3 | -0.018889 | 474.8 | 1.9152
634.528 | 0.90667 | 337.4 | -0.018889 | 475 | 1.916
634.798 | 0.9 | 337.5 | -0.018889 | 475.2 | 1.918
635.069 | 0.89 | 337.6 | -0.019144 | 475.4 | 1.9221
635.34 | 0.9 | 337.7 | -0.019634 | 475.6 | 1.925
635.61 | 0.89 | 337.8 | -0.020124 | 475.8 | 1.9233
635.881 | 0.86 | 337.9 | -0.01766 | 476 | 1.9219
636.151 | 0.86 | 338 | -0.01668 | 476.2 | 1.9179
636.422 | 0.86 | 338.1 | -0.015699 | 476.4 | 1.9178
636.692 | 0.85667 | 338.2 | -0.014719 | 476.6 | 1.9201
636.962 | 0.85 | 338.3 | -0.013379 | 476.8 | 1.9241
637.233 | 0.83667 | 338.4 | -0.011906 | 477 | 1.9234
637.503 | 0.83 | 338.5 | -0.010433 | 477.2 | 1.9235
637.773 | 0.82 | 338.6 | -0.014765 | 477.4 | 1.9265
638.043 | 0.81 | 338.7 | -0.015748 | 477.6 | 1.9246
638.313 | 0.80333 | 338.8 | -0.016731 | 477.8 | 1.9254
638.583 | 0.82 | 338.9 | -0.017714 | 478 | 1.9306
638.853 | 0.8 | 339 | -0.018042 | 478.2 | 1.9309
639.122 | 0.79 | 339.1 | -0.019515 | 478.4 | 1.932
639.392 | 0.78333 | 339.2 | -0.020989 | 478.6 | 1.9307
639.662 | 0.79 | 339.3 | -0.015876 | 478.8 | 1.932
639.931 | 0.76667 | 339.4 | -0.014402 | 479 | 1.9336
640.201 | 0.77667 | 339.5 | -0.012929 | 479.2 | 1.9322
640.47 | 0.76 | 339.6 | -0.016666 | 479.4 | 1.9317
640.74 | 0.76333 | 339.7 | -0.016665 | 479.6 | 1.9294
641.009 | 0.75 | 339.8 | -0.016664 | 479.8 | 1.9265
641.279 | 0.75 | 339.9 | -0.016662 | 480 | 1.926
641.548 | 0.74 | 340 | -0.017766 | 480.2 | 1.9272
641.817 | 0.74 | 340.1 | -0.019238 | 480.4 | 1.9328
642.086 | 0.73333 | 340.2 | -0.02071 | 480.6 | 1.936
642.355 | 0.73333 | 340.3 | -0.018716 | 480.8 | 1.9355
642.624 | 0.71667 | 340.4 | -0.018225 | 481 | 1.9374
642.893 | 0.72 | 340.5 | -0.017734 | 481.2 | 1.9365
643.162 | 0.72 | 340.6 | -0.017243 | 481.4 | 1.9334
643.431 | 0.71 | 340.7 | -0.0093904 | 481.6 | 1.9354
643.699 | 0.7 | 340.8 | -0.0064405 | 481.8 | 1.9375
643.968 | 0.7 | 340.9 | -0.0034907 | 482 | 1.9415
644.237 | 0.67 | 341 | -0.012502 | 482.2 | 1.942
644.505 | 0.69667 | 341.1 | -0.012994 | 482.4 | 1.942
644.774 | 0.68667 | 341.2 | -0.013486 | 482.6 | 1.9442
645.042 | 0.68 | 341.3 | -0.010265 | 482.8 | 1.9442
645.311 | 0.67 | 341.4 | -0.01174 | 483 | 1.9457
645.579 | 0.67 | 341.5 | -0.013215 | 483.2 | 1.9494
645.847 | 0.66 | 341.6 | -0.01469 | 483.4 | 1.9522
646.115 | 0.65 | 341.7 | -0.012168 | 483.6 | 1.9554
646.384 | 0.65 | 341.8 | -0.010693 | 483.8 | 1.9563
646.652 | 0.66 | 341.9 | -0.0092183 | 484 | 1.9567
646.92 | 0.64333 | 342 | -0.01 | 484.2 | 1.9571
647.188 | 0.63667 | 342.1 | -0.01 | 484.4 | 1.9584
647.455 | 0.63 | 342.2 | -0.01 | 484.6 | 1.9602
647.723 | 0.63 | 342.3 | -0.012232 | 484.8 | 1.9628
647.991 | 0.62667 | 342.4 | -0.013215 | 485 | 1.9663
648.259 | 0.62667 | 342.5 | -0.014199 | 485.2 | 1.9711
648.526 | 0.61 | 342.6 | -0.015182 | 485.4 | 1.9762
648.794 | 0.62 | 342.7 | -0.012223 | 485.6 | 1.98
649.061 | 0.59333 | 342.8 | -0.012224 | 485.8 | 1.9825
649.329 | 0.61 | 342.9 | -0.012225 | 486 | 1.9811
649.596 | 0.59667 | 343 | -0.011989 | 486.2 | 1.9806
649.864 | 0.61 | 343.1 | -0.011005 | 486.4 | 1.9804
650.131 | 0.62 | 343.2 | -0.010021 | 486.6 | 1.9821
650.398 | 0.62 | 343.3 | -0.0090365 | 486.8 | 1.9853
650.665 | 0.62667 | 343.4 | -0.01 | 487 | 1.9868
650.932 | 0.61667 | 343.5 | -0.01 | 487.2 | 1.9835
651.199 | 0.58 | 343.6 | -0.01 | 487.4 | 1.9819
651.466 | 0.59667 | 343.7 | -0.01 | 487.6 | 1.9798
651.733 | 0.56667 | 343.8 | -0.01 | 487.8 | 1.98
652 | 0.56333 | 343.9 | -0.01 | 488 | 1.98
652.267 | 0.56667 | 344 | -0.01 | 488.2 | 1.9772
652.534 | 0.57333 | 344.1 | -0.01 | 488.4 | 1.9682
652.8 | 0.56 | 344.2 | -0.01 | 488.6 | 1.9672
653.067 | 0.56333 | 344.3 | -0.01 | 488.8 | 1.966
653.333 | 0.55 | 344.4 | -0.011454 | 489 | 1.9652
653.6 | 0.54667 | 344.5 | -0.011947 | 489.2 | 1.9682
653.866 | 0.54 | 344.6 | -0.01244 | 489.4 | 1.9697
654.133 | 0.53333 | 344.7 | -0.011111 | 489.6 | 1.9704
654.399 | 0.52333 | 344.8 | -0.011112 | 489.8 | 1.9699
654.665 | 0.52667 | 344.9 | -0.011112 | 490 | 1.9658
654.931 | 0.51667 | 345 | -0.011113 | 490.2 | 1.958
655.197 | 0.51333 | 345.1 | -0.0090828 | 490.4 | 1.959
655.463 | 0.51 | 345.2 | -0.0080966 | 490.6 | 1.9641
655.729 | 0.5 | 345.3 | -0.0071105 | 490.8 | 1.9768
655.995 | 0.50333 | 345.4 | -0.010271 | 491 | 1.9843
656.261 | 0.50667 | 345.5 | -0.010764 | 491.2 | 1.9914
656.527 | 0.49667 | 345.6 | -0.011257 | 491.4 | 2.0004
656.793 | 0.48 | 345.7 | -0.0088889 | 491.6 | 2.0114
657.059 | 0.48333 | 345.8 | -0.0088889 | 491.8 | 2.0167
657.324 | 0.48 | 345.9 | -0.0088889 | 492 | 1.9993
657.59 | 0.48333 | 346 | -0.0088889 | 492.2 | 1.983
657.855 | 0.49 | 346.1 | -0.0096105 | 492.4 | 1.9537
658.121 | 0.48667 | 346.2 | -0.0091174 | 492.6 | 1.9425
658.386 | 0.47 | 346.3 | -0.0086243 | 492.8 | 1.9361
658.651 | 0.48 | 346.4 | -0.0077778 | 493 | 1.9271
658.917 | 0.47 | 346.5 | -0.0077778 | 493.2 | 1.9262
659.182 | 0.45333 | 346.6 | -0.0077778 | 493.4 | 1.9273
659.447 | 0.45333 | 346.7 | -0.010015 | 493.6 | 1.9318
659.712 | 0.46333 | 346.8 | -0.010508 | 493.8 | 1.9309
659.977 | 0.44333 | 346.9 | -0.011001 | 494 | 1.927
660.242 | 0.44667 | 347 | -0.011494 | 494.2 | 1.9207
660.507 | 0.44667 | 347.1 | -0.010323 | 494.4 | 1.9191
660.772 | 0.44 | 347.2 | -0.010818 | 494.6 | 1.9158
661.037 | 0.44333 | 347.3 | -0.011312 | 494.8 | 1.9075
661.301 | 0.43333 | 347.4 | -0.0072311 | 495 | 1.9093
661.566 | 0.44 | 347.5 | -0.0052553 | 495.2 | 1.9117
661.831 | 0.44667 | 347.6 | -0.0032796 | 495.4 | 1.9157
662.095 | 0.44667 | 347.7 | -0.0013038 | 495.6 | 1.9115
662.36 | 0.45 | 347.8 | -0.0035514 | 495.8 | 1.9039
662.624 | 0.44 | 347.9 | -0.0025628 | 496 | 1.8972
662.888 | 0.43 | 348 | -0.0015741 | 496.2 | 1.8974
663.153 | 0.43 | 348.1 | -0.0024822 | 496.4 | 1.8984
663.417 | 0.41667 | 348.2 | -0.0029758 | 496.6 | 1.8947
663.681 | 0.41333 | 348.3 | -0.0034693 | 496.8 | 1.894
663.945 | 0.41 | 348.4 | -0.0021465 | 497 | 1.8932
664.209 | 0.39667 | 348.5 | -0.001653 | 497.2 | 1.8969
664.473 | 0.40333 | 348.6 | -0.0011594 | 497.4 | 1.8922
664.737 | 0.39 | 348.7 | -0.00066579 | 497.6 | 1.8867
665.001 | 0.39 | 348.8 | 0.00077151 | 497.8 | 1.8839
665.265 | 0.39 | 348.9 | 0.0017606 | 498 | 1.8843
665.529 | 0.38667 | 349 | 0.0027498 | 498.2 | 1.8847
665.792 | 0.37333 | 349.1 | -0.0028252 | 498.4 | 1.8762
666.056 | 0.38667 | 349.2 | -0.0043079 | 498.6 | 1.873
666.319 | 0.37333 | 349.3 | -0.0057906 | 498.8 | 1.8693
666.583 | 0.38667 | 349.4 | -0.0022387 | 499 | 1.8696
666.846 | 0.38 | 349.5 | -0.0027343 | 499.2 | 1.8649
667.11 | 0.38 | 349.6 | -0.0032298 | 499.4 | 1.8607
667.373 | 0.37 | 349.7 | -0.0037253 | 499.6 | 1.8626
667.636 | 0.36667 | 349.8 | -0.0023541 | 499.8 | 1.8613
667.9 | 0.36333 | 349.9 | -0.00087043 | 500 | 1.8615
668.163 | 0.36333 | 350 | 0.00061325 | 500.2 | 1.859
668.426 | 0.36 | 350.1 | 0 | 500.4 | 1.8554
668.689 | 0.35667 | 350.2 | 0 | 500.6 | 1.8537
668.952 | 0.35667 | 350.3 | 0 | 500.8 | 1.8562
669.215 | 0.35 | 350.4 | 0 | 501 | 1.8515
669.478 | 0.35 | 350.5 | -0.0047046 | 501.2 | 1.8456
669.74 | 0.35333 | 350.6 | -0.0061898 | 501.4 | 1.8375
670.003 | 0.34 | 350.7 | -0.0076749 | 501.6 | 1.8333
670.266 | 0.34333 | 350.8 | -0.0019456 | 501.8 | 1.8325
670.528 | 0.34 | 350.9 | -0.0014485 | 502 | 1.8329
670.791 | 0.34 | 351 | -0.00095153 | 502.2 | 1.8334
671.054 | 0.33667 | 351.1 | -0.0019403 | 502.4 | 1.8322
671.316 | 0.33 | 351.2 | -0.00045664 | 502.6 | 1.8282
671.578 | 0.33 | 351.3 | 0.001027 | 502.8 | 1.8279
671.841 | 0.34 | 351.4 | 0.0025107 | 503 | 1.8268
672.103 | 0.33 | 351.5 | 0.00070557 | 503.2 | 1.8238
672.365 | 0.32 | 351.6 | 0.00021101 | 503.4 | 1.8246
672.627 | 0.32 | 351.7 | -0.00028355 | 503.6 | 1.8227
672.889 | 0.32 | 351.8 | 0 | 503.8 | 1.8171
673.151 | 0.32 | 351.9 | 0 | 504 | 1.8173
673.413 | 0.31 | 352 | 0 | 504.2 | 1.8148
673.675 | 0.31667 | 352.1 | 0 | 504.4 | 1.8083
673.937 | 0.31 | 352.2 | 0 | 504.6 | 1.8092
674.199 | 0.29667 | 352.3 | 0 | 504.8 | 1.8074
674.461 | 0.30333 | 352.4 | 0 | 505 | 1.8028
674.722 | 0.30667 | 352.5 | 0.0044076 | 505.2 | 1.8022
674.984 | 0.29667 | 352.6 | 0.0058927 | 505.4 | 1.8
675.245 | 0.3 | 352.7 | 0.0073779 | 505.6 | 1.8011
675.507 | 0.29333 | 352.8 | 0.0072024 | 505.8 | 1.8006
675.768 | 0.28667 | 352.9 | 0.0086905 | 506 | 1.7944
676.03 | 0.29 | 353 | 0.010179 | 506.2 | 1.7943
676.291 | 0.28 | 353.1 | 0.011667 | 506.4 | 1.7917
676.552 | 0.28 | 353.2 | 0.008395 | 506.6 | 1.7856
676.813 | 0.27667 | 353.3 | 0.0078995 | 506.8 | 1.7861
677.074 | 0.28667 | 353.4 | 0.0074039 | 507 | 1.7858
677.336 | 0.28 | 353.5 | 0.0088886 | 507.2 | 1.7855
677.597 | 0.27333 | 353.6 | 0.0088881 | 507.4 | 1.7834
677.857 | 0.28 | 353.7 | 0.0088876 | 507.6 | 1.7804
678.118 | 0.27333 | 353.8 | 0.0090228 | 507.8 | 1.78
678.379 | 0.26667 | 353.9 | 0.0095188 | 508 | 1.7798
678.64 | 0.26333 | 354 | 0.010015 | 508.2 | 1.7816
678.901 | 0.26 | 354.1 | 0.010511 | 508.4 | 1.7866
679.161 | 0.27 | 354.2 | 0.01 | 508.6 | 1.7811
679.422 | 0.27333 | 354.3 | 0.01 | 508.8 | 1.7756
679.682 | 0.26 | 354.4 | 0.01 | 509 | 1.7723
679.943 | 0.26 | 354.5 | 0.011384 | 509.2 | 1.7706
680.203 | 0.26 | 354.6 | 0.01188 | 509.4 | 1.7715
680.464 | 0.26333 | 354.7 | 0.012376 | 509.6 | 1.7722
680.724 | 0.26 | 354.8 | 0.014727 | 509.8 | 1.7702
680.984 | 0.25667 | 354.9 | 0.016215 | 510 | 1.7672
681.244 | 0.26 | 355 | 0.017703 | 510.2 | 1.7623
681.505 | 0.24333 | 355.1 | 0.019191 | 510.4 | 1.7604
681.765 | 0.25 | 355.2 | 0.018601 | 510.6 | 1.7602
682.025 | 0.26 | 355.3 | 0.019593 | 510.8 | 1.7622
682.285 | 0.24667 | 355.4 | 0.020585 | 511 | 1.7591
682.544 | 0.24333 | 355.5 | 0.015962 | 511.2 | 1.7535
682.804 | 0.24 | 355.6 | 0.014472 | 511.4 | 1.7465
683.064 | 0.24333 | 355.7 | 0.012982 | 511.6 | 1.7455
683.324 | 0.25333 | 355.8 | 0.016667 | 511.8 | 1.7458
683.583 | 0.24333 | 355.9 | 0.016668 | 512 | 1.7472
683.843 | 0.24 | 356 | 0.01667 | 512.2 | 1.7444
684.102 | 0.23333 | 356.1 | 0.016671 | 512.4 | 1.7455
684.362 | 0.25 | 356.2 | 0.017805 | 512.6 | 1.7467
684.621 | 0.23 | 356.3 | 0.019296 | 512.8 | 1.7484
684.881 | 0.23667 | 356.4 | 0.020787 | 513 | 1.7444
685.14 | 0.23 | 356.5 | 0.01606 | 513.2 | 1.7372
685.399 | 0.23 | 356.6 | 0.014569 | 513.4 | 1.7314
685.658 | 0.23667 | 356.7 | 0.013078 | 513.6 | 1.7313
685.917 | 0.22 | 356.8 | 0.017804 | 513.8 | 1.734
686.177 | 0.22667 | 356.9 | 0.018299 | 514 | 1.7328
686.436 | 0.22667 | 357 | 0.018794 | 514.2 | 1.7298
686.694 | 0.23333 | 357.1 | 0.019289 | 514.4 | 1.7249
686.953 | 0.24 | 357.2 | 0.018815 | 514.6 | 1.7249
687.212 | 0.24 | 357.3 | 0.020304 | 514.8 | 1.7239
687.471 | 0.23 | 357.4 | 0.021793 | 515 | 1.7251
687.73 | 0.25 | 357.5 | 0.02094 | 515.2 | 1.7273
687.988 | 0.24 | 357.6 | 0.020443 | 515.4 | 1.7253
688.247 | 0.23 | 357.7 | 0.019946 | 515.6 | 1.7234
688.505 | 0.21667 | 357.8 | 0.019449 | 515.8 | 1.7214
688.764 | 0.21333 | 357.9 | 0.023207 | 516 | 1.7209
689.022 | 0.21667 | 358 | 0.024202 | 516.2 | 1.7217
689.281 | 0.22 | 358.1 | 0.025197 | 516.4 | 1.7211
689.539 | 0.20333 | 358.2 | 0.025081 | 516.6 | 1.7208
689.797 | 0.2 | 358.3 | 0.026076 | 516.8 | 1.7188
690.055 | 0.21 | 358.4 | 0.027071 | 517 | 1.7146
690.314 | 0.21667 | 358.5 | 0.027922 | 517.2 | 1.7123
690.572 | 0.29667 | 358.6 | 0.02842 | 517.4 | 1.7102
690.83 | 0.33333 | 358.7 | 0.028917 | 517.6 | 1.71
691.088 | 0.22667 | 358.8 | 0.029415 | 517.8 | 1.71
691.345 | 0.20333 | 358.9 | 0.029357 | 518 | 1.71
691.603 | 0.19333 | 359 | 0.029854 | 518.2 | 1.71
691.861 | 0.2 | 359.1 | 0.030352 | 518.4 | 1.71
692.119 | 0.19667 | 359.2 | 0.03 | 518.6 | 1.7087
692.376 | 0.18 | 359.3 | 0.03 | 518.8 | 1.7044
692.634 | 0.19333 | 359.4 | 0.03 | 519 | 1.7021
692.891 | 0.19 | 359.5 | 0.026308 | 519.2 | 1.7014
693.149 | 0.19667 | 359.6 | 0.024816 | 519.4 | 1.7057
693.406 | 0.2 | 359.7 | 0.023323 | 519.6 | 1.708
693.664 | 0.20667 | 359.8 | 0.021831 | 519.8 | 1.7086
693.921 | 0.21 | 359.9 | 0.026665 | 520 | 1.7042
694.178 | 0.19 | 360 | 0.026664 | 520.2 | 1.7042
694.435 | 0.19 | 360.1 | 0.026662 | 520.4 | 1.7037
694.693 | 0.18 | 360.2 | 0.031634 | 520.6 | 1.7022
694.95 | 0.17667 | 360.3 | 0.034624 | 520.8 | 1.698
695.207 | 0.18333 | 360.4 | 0.037613 | 521 | 1.6929
695.464 | 0.17667 | 360.5 | 0.036971 | 521.2 | 1.6907
695.721 | 0.17333 | 360.6 | 0.038466 | 521.4 | 1.6908
695.977 | 0.16333 | 360.7 | 0.039962 | 521.6 | 1.6934
696.234 | 0.18667 | 360.8 | 0.041457 | 521.8 | 1.6972
696.491 | 0.20667 | 360.9 | 0.035386 | 522 | 1.7014
696.748 | 0.18667 | 361 | 0.03389 | 522.2 | 1.7025
697.004 | 0.17 | 361.1 | 0.032394 | 522.4 | 1.696
697.261 | 0.16667 | 361.2 | 0.036666 | 522.6 | 1.6928
697.517 | 0.15667 | 361.3 | 0.036664 | 522.8 | 1.688
697.774 | 0.16333 | 361.4 | 0.036663 | 523 | 1.6837
698.03 | 0.17667 | 361.5 | 0.040509 | 523.2 | 1.6876
698.286 | 0.16 | 361.6 | 0.043503 | 523.4 | 1.6895
698.543 | 0.15 | 361.7 | 0.046497 | 523.6 | 1.6856
698.799 | 0.15333 | 361.8 | 0.049491 | 523.8 | 1.6905
699.055 | 0.15667 | 361.9 | 0.047902 | 524 | 1.6929
699.311 | 0.15 | 362 | 0.049396 | 524.2 | 1.6932
699.567 | 0.17 | 362.1 | 0.05089 | 524.4 | 1.6948
699.823 | 0.15 | 362.2 | 0.05 | 524.6 | 1.6941
700.079 | 0.16 | 362.3 | 0.05 | 524.8 | 1.6939
700.335 | 0.16 | 362.4 | 0.05 | 525 | 1.6922
700.591 | 0.16 | 362.5 | 0.05 | 525.2 | 1.6937
700.846 | 0.15333 | 362.6 | 0.05 | 525.4 | 1.6964
701.102 | 0.15 | 362.7 | 0.05 | 525.6 | 1.6985
701.358 | 0.16 | 362.8 | 0.05 | 525.8 | 1.7001
701.613 | 0.14667 | 362.9 | 0.054531 | 526 | 1.6988
701.869 | 0.14 | 363 | 0.056028 | 526.2 | 1.6947
702.124 | 0.15667 | 363.1 | 0.057525 | 526.4 | 1.6954
702.38 | 0.14667 | 363.2 | 0.065399 | 526.6 | 1.6956
702.635 | 0.15667 | 363.3 | 0.06989 | 526.8 | 1.6933
702.89 | 0.15667 | 363.4 | 0.074381 | 527 | 1.6945
703.145 | 0.15333 | 363.5 | 0.097112 | 527.2 | 1.6959
703.401 | 0.15333 | 363.6 | 0.11009 | 527.4 | 1.7016
703.656 | 0.14333 | 363.7 | 0.12306 | 527.6 | 1.6993
703.911 | 0.15 | 363.8 | 0.13603 | 527.8 | 1.6966
704.166 | 0.14 | 363.9 | 0.35012 | 528 | 1.6973
704.421 | 0.15333 | 364 | 0.43255 | 528.2 | 1.6976
704.675 | 0.14333 | 364.1 | 0.51498 | 528.4 | 1.7005
704.93 | 0.15667 | 364.2 | 0.9723 | 528.6 | 1.702
705.185 | 0.14 | 364.3 | 1.2212 | 528.8 | 1.7073
705.44 | 0.15 | 364.4 | 1.4702 | 529 | 1.7105
705.694 | 0.15667 | 364.5 | 1.6903 | 529.2 | 1.7057
705.949 | 0.16667 | 364.6 | 1.9763 | 529.4 | 1.7032
706.203 | 0.18 | 364.7 | 2.2622 | 529.6 | 1.7027
706.458 | 0.19667 | 364.8 | 2.5481 | 529.8 | 1.7125
706.712 | 0.22 | 364.9 | 2.1156 | 530 | 1.711
706.967 | 0.21667 | 365 | 2.1372 | 530.2 | 1.7109
707.221 | 0.2 | 365.1 | 2.1588 | 530.4 | 1.7144
707.475 | 0.19 | 365.2 | 1.7003 | 530.6 | 1.7133
707.729 | 0.18 | 365.3 | 1.4586 | 530.8 | 1.7142
707.984 | 0.18333 | 365.4 | 1.2168 | 531 | 1.7145
708.238 | 0.18 | 365.5 | 1.1466 | 531.2 | 1.7128
708.492 | 0.16 | 365.6 | 0.94772 | 531.4 | 1.7128
708.746 | 0.2 | 365.7 | 0.74887 | 531.6 | 1.7139
709 | 0.23667 | 365.8 | 0.55003 | 531.8 | 1.7197
709.253 | 0.23 | 365.9 | 0.66153 | 532 | 1.7243
709.507 | 0.18667 | 366 | 0.60604 | 532.2 | 1.7271
709.761 | 0.16 | 366.1 | 0.55055 | 532.4 | 1.7242
710.015 | 0.14 | 366.2 | 0.48906 | 532.6 | 1.7234
710.268 | 0.13333 | 366.3 | 0.44652 | 532.8 | 1.7246
710.522 | 0.12667 | 366.4 | 0.40397 | 533 | 1.7322
710.775 | 0.13 | 366.5 | 0.32923 | 533.2 | 1.7333
711.029 | 0.14333 | 366.6 | 0.26167 | 533.4 | 1.7357
711.282 | 0.13 | 366.7 | 0.1941 | 533.6 | 1.7424
711.535 | 0.14 | 366.8 | 0.12653 | 533.8 | 1.7414
711.789 | 0.17 | 366.9 | 0.1464 | 534 | 1.7412
712.042 | 0.18667 | 367 | 0.10744 | 534.2 | 1.7396
712.295 | 0.2 | 367.1 | 0.068472 | 534.4 | 1.7424
712.548 | 0.19 | 367.2 | 0.085317 | 534.6 | 1.7458
712.801 | 0.18 | 367.3 | 0.079808 | 534.8 | 1.7433
713.054 | 0.14333 | 367.4 | 0.074299 | 535 | 1.7456
713.307 | 0.13 | 367.5 | 0.077778 | 535.2 | 1.7471
713.56 | 0.12 | 367.6 | 0.077779 | 535.4 | 1.7497
713.813 | 0.11667 | 367.7 | 0.07778 | 535.6 | 1.7542
714.066 | 0.12333 | 367.8 | 0.077781 | 535.8 | 1.7587
714.318 | 0.11 | 367.9 | 0.078569 | 536 | 1.7643
714.571 | 0.12 | 368 | 0.07957 | 536.2 | 1.7678
714.823 | 0.11 | 368.1 | 0.080571 | 536.4 | 1.7719
715.076 | 0.11 | 368.2 | 0.08 | 536.6 | 1.7761
715.328 | 0.096667 | 368.3 | 0.08 | 536.8 | 1.7785
715.581 | 0.10667 | 368.4 | 0.08 | 537 | 1.786
715.833 | 0.10333 | 368.5 | 0.083529 | 537.2 | 1.7938
716.086 | 0.096667 | 368.6 | 0.08503 | 537.4 | 1.8012
716.338 | 0.10667 | 368.7 | 0.086532 | 537.6 | 1.8096
716.59 | 0.1 | 368.8 | 0.088033 | 537.8 | 1.8155
716.842 | 0.11 | 368.9 | 0.083335 | 538 | 1.8187
717.094 | 0.11 | 369 | 0.083336 | 538.2 | 1.8269
717.346 | 0.10667 | 369.1 | 0.083338 | 538.4 | 1.8358
717.598 | 0.096667 | 369.2 | 0.082617 | 538.6 | 1.8464
717.85 | 0.10333 | 369.3 | 0.081114 | 538.8 | 1.8504
718.102 | 0.10333 | 369.4 | 0.079611 | 539 | 1.8541
718.354 | 0.1 | 369.5 | 0.087128 | 539.2 | 1.8625
718.605 | 0.10667 | 369.6 | 0.090134 | 539.4 | 1.8704
718.857 | 0.11 | 369.7 | 0.09314 | 539.6 | 1.8817
719.109 | 0.096667 | 369.8 | 0.096146 | 539.8 | 1.8946
719.36 | 0.093333 | 369.9 | 0.092778 | 540 | 1.9071
719.612 | 0.1 | 370 | 0.094786 | 540.2 | 1.9211
719.863 | 0.086667 | 370.1 | 0.096794 | 540.4 | 1.9362
720.115 | 0.1 | 370.2 | 0.095048 | 540.6 | 1.9531
720.366 | 0.096667 | 370.3 | 0.094046 | 540.8 | 1.9775
720.617 | 0.086667 | 370.4 | 0.093044 | 541 | 2.0054
720.869 | 0.086667 | 370.5 | 0.096942 | 541.2 | 2.0345
721.12 | 0.1 | 370.6 | 0.098445 | 541.4 | 2.0726
721.371 | 0.096667 | 370.7 | 0.099948 | 541.6 | 2.1117
721.622 | 0.09 | 370.8 | 0.10145 | 541.8 | 2.1503
721.873 | 0.09 | 370.9 | 0.099752 | 542 | 2.1709
722.124 | 0.08 | 371 | 0.10076 | 542.2 | 2.1797
722.375 | 0.08 | 371.1 | 0.10176 | 542.4 | 2.18
722.626 | 0.086667 | 371.2 | 0.10499 | 542.6 | 2.1834
722.876 | 0.08 | 371.3 | 0.10599 | 542.8 | 2.1842
723.127 | 0.08 | 371.4 | 0.10699 | 543 | 2.1772
723.378 | 0.09 | 371.5 | 0.107 | 543.2 | 2.1735
723.628 | 0.086667 | 371.6 | 0.1085 | 543.4 | 2.1756
723.879 | 0.083333 | 371.7 | 0.11001 | 543.6 | 2.1956
724.129 | 0.083333 | 371.8 | 0.11152 | 543.8 | 2.2048
724.38 | 0.083333 | 371.9 | 0.11 | 544 | 2.2152
724.63 | 0.086667 | 372 | 0.11 | 544.2 | 2.2249
724.881 | 0.1 | 372.1 | 0.11 | 544.4 | 2.2388
725.131 | 0.083333 | 372.2 | 0.11 | 544.6 | 2.2816
725.381 | 0.073333 | 372.3 | 0.11 | 544.8 | 2.4082
725.631 | 0.08 | 372.4 | 0.11 | 545 | 2.7114
725.881 | 0.076667 | 372.5 | 0.11 | 545.2 | 3.7645
726.131 | 0.08 | 372.6 | 0.11 | 545.4 | 5.4261
726.381 | 0.083333 | 372.7 | 0.11 | 545.6 | 7.7791
726.631 | 0.073333 | 372.8 | 0.11 | 545.8 | 10.0077
726.881 | 0.076667 | 372.9 | 0.11476 | 546 | 12.5209
727.131 | 0.073333 | 373 | 0.11627 | 546.2 | 15.4902
727.381 | 0.086667 | 373.1 | 0.11778 | 546.4 | 13.8517
727.631 | 0.076667 | 373.2 | 0.11476 | 546.6 | 10.4896
727.88 | 0.07 | 373.3 | 0.11526 | 546.8 | 5.5692
728.13 | 0.076667 | 373.4 | 0.11576 | 547 | 3.2258
728.379 | 0.076667 | 373.5 | 0.12031 | 547.2 | 2.2674
728.629 | 0.073333 | 373.6 | 0.12132 | 547.4 | 2.2672
728.878 | 0.076667 | 373.7 | 0.12232 | 547.6 | 2.1669
729.128 | 0.083333 | 373.8 | 0.12333 | 547.8 | 2.1641
729.377 | 0.063333 | 373.9 | 0.12261 | 548 | 2.1442
729.626 | 0.063333 | 374 | 0.12412 | 548.2 | 2.1393
729.875 | 0.083333 | 374.1 | 0.12563 | 548.4 | 2.1362
730.125 | 0.076667 | 374.2 | 0.12265 | 548.6 | 2.1367
730.374 | 0.07 | 374.3 | 0.12164 | 548.8 | 2.1332
730.623 | 0.073333 | 374.4 | 0.12064 | 549 | 2.1346
730.872 | 0.07 | 374.5 | 0.12611 | 549.2 | 2.1356
731.121 | 0.07 | 374.6 | 0.12763 | 549.4 | 2.1333
731.37 | 0.063333 | 374.7 | 0.12914 | 549.6 | 2.1367
731.618 | 0.07 | 374.8 | 0.12787 | 549.8 | 2.1421
731.867 | 0.07 | 374.9 | 0.12938 | 550 | 2.1534
732.116 | 0.073333 | 375 | 0.13089 | 550.2 | 2.1546
732.365 | 0.073333 | 375.1 | 0.1324 | 550.4 | 2.1531
732.613 | 0.06 | 375.2 | 0.13222 | 550.6 | 2.1506
732.862 | 0.053333 | 375.3 | 0.13222 | 550.8 | 2.15
733.11 | 0.073333 | 375.4 | 0.13222 | 551 | 2.1539
733.359 | 0.05 | 375.5 | 0.13511 | 551.2 | 2.1582
733.607 | 0.076667 | 375.6 | 0.13662 | 551.4 | 2.1599
733.855 | 0.06 | 375.7 | 0.13813 | 551.6 | 2.1641
734.104 | 0.063333 | 375.8 | 0.13673 | 551.8 | 2.1682
734.352 | 0.063333 | 375.9 | 0.13724 | 552 | 2.1749
734.6 | 0.06 | 376 | 0.13774 | 552.2 | 2.1818
734.848 | 0.056667 | 376.1 | 0.13824 | 552.4 | 2.1858
735.096 | 0.06 | 376.2 | 0.13736 | 552.6 | 2.1833
735.344 | 0.053333 | 376.3 | 0.13686 | 552.8 | 2.1819
735.592 | 0.056667 | 376.4 | 0.13636 | 553 | 2.1796
735.84 | 0.07 | 376.5 | 0.14214 | 553.2 | 2.18
736.088 | 0.066667 | 376.6 | 0.14415 | 553.4 | 2.1839
736.336 | 0.06 | 376.7 | 0.14617 | 553.6 | 2.1892
736.583 | 0.06 | 376.8 | 0.14232 | 553.8 | 2.1934
736.831 | 0.07 | 376.9 | 0.14283 | 554 | 2.194
737.079 | 0.066667 | 377 | 0.14333 | 554.2 | 2.1932
737.326 | 0.06 | 377.1 | 0.14383 | 554.4 | 2.1912
737.574 | 0.06 | 377.2 | 0.14667 | 554.6 | 2.1903
737.821 | 0.066667 | 377.3 | 0.14666 | 554.8 | 2.1933
738.068 | 0.06 | 377.4 | 0.14666 | 555 | 2.1971
738.316 | 0.11 | 377.5 | 0.14755 | 555.2 | 2.1994
738.563 | 0.11 | 377.6 | 0.14907 | 555.4 | 2.1944
738.81 | 0.076667 | 377.7 | 0.15058 | 555.6 | 2.1866
739.057 | 0.05 | 377.8 | 0.15251 | 555.8 | 2.1794
739.304 | 0.063333 | 377.9 | 0.15352 | 556 | 2.1779
739.552 | 0.056667 | 378 | 0.15452 | 556.2 | 2.1791
739.799 | 0.053333 | 378.1 | 0.15553 | 556.4 | 2.1798
740.045 | 0.04 | 378.2 | 0.16024 | 556.6 | 2.18
740.292 | 0.053333 | 378.3 | 0.16276 | 556.8 | 2.181
740.539 | 0.05 | 378.4 | 0.16528 | 557 | 2.1831
740.786 | 0.056667 | 378.5 | 0.16145 | 557.2 | 2.1868
741.033 | 0.07 | 378.6 | 0.16196 | 557.4 | 2.1954
741.279 | 0.05 | 378.7 | 0.16246 | 557.6 | 2.2027
741.526 | 0.06 | 378.8 | 0.16574 | 557.8 | 2.2027
741.773 | 0.056667 | 378.9 | 0.16625 | 558 | 2.2095
742.019 | 0.046667 | 379 | 0.16675 | 558.2 | 2.2146
742.265 | 0.06 | 379.1 | 0.16677 | 558.4 | 2.2196
742.512 | 0.06 | 379.2 | 0.16828 | 558.6 | 2.2249
742.758 | 0.053333 | 379.3 | 0.16979 | 558.8 | 2.2216
743.004 | 0.063333 | 379.4 | 0.1713 | 559 | 2.2181
743.251 | 0.05 | 379.5 | 0.17 | 559.2 | 2.218
743.497 | 0.063333 | 379.6 | 0.17 | 559.4 | 2.2206
743.743 | 0.053333 | 379.7 | 0.17 | 559.6 | 2.2275
743.989 | 0.033333 | 379.8 | 0.17134 | 559.8 | 2.2329
744.235 | 0.043333 | 379.9 | 0.17185 | 560 | 2.2402
744.481 | 0.043333 | 380 | 0.17235 | 560.2 | 2.2448
744.727 | 0.05 | 380.1 | 0.17469 | 560.4 | 2.2432
744.973 | 0.056667 | 380.2 | 0.1762 | 560.6 | 2.2457
745.219 | 0.026667 | 380.3 | 0.17772 | 560.8 | 2.2488
745.464 | 0.05 | 380.4 | 0.17923 | 561 | 2.2565
745.71 | 0.05 | 380.5 | 0.1771 | 561.2 | 2.2635
745.956 | 0.056667 | 380.6 | 0.17761 | 561.4 | 2.2685
746.201 | 0.036667 | 380.7 | 0.17811 | 561.6 | 2.2714
746.447 | 0.053333 | 380.8 | 0.18308 | 561.8 | 2.2711
746.692 | 0.04 | 380.9 | 0.18459 | 562 | 2.2772
746.938 | 0.053333 | 381 | 0.18611 | 562.2 | 2.2824
747.183 | 0.046667 | 381.1 | 0.18609 | 562.4 | 2.2826
747.428 | 0.036667 | 381.2 | 0.18812 | 562.6 | 2.2847
747.673 | 0.056667 | 381.3 | 0.19014 | 562.8 | 2.2836
747.919 | 0.053333 | 381.4 | 0.19216 | 563 | 2.2822
748.164 | 0.04 | 381.5 | 0.1948 | 563.2 | 2.285
748.409 | 0.04 | 381.6 | 0.19632 | 563.4 | 2.2902
748.654 | 0.033333 | 381.7 | 0.19783 | 563.6 | 2.2992
748.899 | 0.04 | 381.8 | 0.19914 | 563.8 | 2.3009
749.144 | 0.043333 | 381.9 | 0.20116 | 564 | 2.3049
749.389 | 0.04 | 382 | 0.20319 | 564.2 | 2.3068
749.633 | 0.05 | 382.1 | 0.20241 | 564.4 | 2.3091
749.878 | 0.046667 | 382.2 | 0.20292 | 564.6 | 2.3103
750.123 | 0.04 | 382.3 | 0.20342 | 564.8 | 2.3115
750.368 | 0.1 | 382.4 | 0.20218 | 565 | 2.3162
750.612 | 0.093333 | 382.5 | 0.20167 | 565.2 | 2.3223
750.857 | 0.06 | 382.6 | 0.20117 | 565.4 | 2.3264
751.101 | 0.053333 | 382.7 | 0.20066 | 565.6 | 2.3241
751.346 | 0.066667 | 382.8 | 0.20222 | 565.8 | 2.325
751.59 | 0.096667 | 382.9 | 0.20222 | 566 | 2.3229
751.834 | 0.056667 | 383 | 0.20222 | 566.2 | 2.3251
752.079 | 0.05 | 383.1 | 0.20519 | 566.4 | 2.3338
752.323 | 0.04 | 383.2 | 0.20671 | 566.6 | 2.3374
752.567 | 0.036667 | 383.3 | 0.20823 | 566.8 | 2.3425
752.811 | 0.036667 | 383.4 | 0.21162 | 567 | 2.3463
753.055 | 0.033333 | 383.5 | 0.21415 | 567.2 | 2.3447
753.299 | 0.03 | 383.6 | 0.21668 | 567.4 | 2.3444
753.543 | 0.026667 | 383.7 | 0.21922 | 567.6 | 2.3467
753.787 | 0.026667 | 383.8 | 0.21805 | 567.8 | 2.3524
754.031 | 0.03 | 383.9 | 0.21957 | 568 | 2.3575
754.275 | 0.02 | 384 | 0.22109 | 568.2 | 2.3584
754.518 | 0.033333 | 384.1 | 0.22 | 568.4 | 2.3545
754.762 | 0.026667 | 384.2 | 0.22 | 568.6 | 2.3586
755.006 | 0.023333 | 384.3 | 0.22 | 568.8 | 2.3611
755.249 | 0.036667 | 384.4 | 0.22383 | 569 | 2.3567
755.493 | 0.023333 | 384.5 | 0.22535 | 569.2 | 2.3587
755.736 | 0.03 | 384.6 | 0.22687 | 569.4 | 2.3588
755.98 | 0.03 | 384.7 | 0.22673 | 569.6 | 2.361
756.223 | 0.02 | 384.8 | 0.22825 | 569.8 | 2.365
756.466 | 0.036667 | 384.9 | 0.22977 | 570 | 2.3674
756.709 | 0.016667 | 385 | 0.23129 | 570.2 | 2.3709
756.953 | 0.03 | 385.1 | 0.23 | 570.4 | 2.3755
757.196 | 0.036667 | 385.2 | 0.23 | 570.6 | 2.3812
757.439 | 0.026667 | 385.3 | 0.23 | 570.8 | 2.3841
757.682 | 0.02 | 385.4 | 0.23673 | 571 | 2.3773
757.925 | 0.02 | 385.5 | 0.23927 | 571.2 | 2.3735
758.168 | 0.036667 | 385.6 | 0.2418 | 571.4 | 2.3709
758.411 | 0.0066667 | 385.7 | 0.23916 | 571.6 | 2.3765
758.653 | 0.026667 | 385.8 | 0.24068 | 571.8 | 2.3789
758.896 | 0.01 | 385.9 | 0.2422 | 572 | 2.383
759.139 | 0.023333 | 386 | 0.24372 | 572.2 | 2.3865
759.382 | 0.023333 | 386.1 | 0.24601 | 572.4 | 2.3889
759.624 | 0.013333 | 386.2 | 0.24651 | 572.6 | 2.3935
759.867 | 0.026667 | 386.3 | 0.24702 | 572.8 | 2.3956
760.109 | 0.026667 | 386.4 | 0.24759 | 573 | 2.3933
760.352 | 0.03 | 386.5 | 0.24911 | 573.2 | 2.3907
760.594 | 0.023333 | 386.6 | 0.25063 | 573.4 | 2.3882
760.836 | 0.016667 | 386.7 | 0.25 | 573.6 | 2.3894
761.079 | 0.02 | 386.8 | 0.25 | 573.8 | 2.3855
761.321 | 0.016667 | 386.9 | 0.25 | 574 | 2.3884
761.563 | 0.02 | 387 | 0.25338 | 574.2 | 2.3931
761.805 | 0.016667 | 387.1 | 0.25491 | 574.4 | 2.4026
762.047 | 0.02 | 387.2 | 0.25643 | 574.6 | 2.3989
762.289 | 0.013333 | 387.3 | 0.25795 | 574.8 | 2.3921
762.531 | 0.016667 | 387.4 | 0.25781 | 575 | 2.3887
762.773 | 0.02 | 387.5 | 0.25933 | 575.2 | 2.3917
763.015 | 0.023333 | 387.6 | 0.26086 | 575.4 | 2.3991
763.257 | 0.03 | 387.7 | 0.2654 | 575.6 | 2.4087
763.498 | 0.17 | 387.8 | 0.26743 | 575.8 | 2.4286
763.74 | 0.19 | 387.9 | 0.26946 | 576 | 2.525
763.982 | 0.08 | 388 | 0.27169 | 576.2 | 2.7293
764.223 | 0.043333 | 388.1 | 0.27474 | 576.4 | 3.1111
764.465 | 0.026667 | 388.2 | 0.27779 | 576.6 | 3.4925
764.706 | 0.03 | 388.3 | 0.28084 | 576.8 | 3.9574
764.948 | 0.026667 | 388.4 | 0.2787 | 577 | 4.4932
765.189 | 0.023333 | 388.5 | 0.27972 | 577.2 | 4.5117
765.43 | 0.0066667 | 388.6 | 0.28074 | 577.4 | 3.95
765.672 | 0.016667 | 388.7 | 0.28286 | 577.6 | 3.2545
765.913 | 0.016667 | 388.8 | 0.28388 | 577.8 | 2.7787
766.154 | 0.016667 | 388.9 | 0.28489 | 578 | 2.5408
766.395 | 0.03 | 389 | 0.28609 | 578.2 | 2.6215
766.636 | 0.0066667 | 389.1 | 0.28761 | 578.4 | 2.8676
766.877 | 0.01 | 389.2 | 0.28914 | 578.6 | 3.2213
767.118 | 0.023333 | 389.3 | 0.29122 | 578.8 | 3.655
767.359 | 0.01 | 389.4 | 0.29274 | 579 | 4.0691
767.6 | 0.02 | 389.5 | 0.29427 | 579.2 | 4.4305
767.841 | 0.01 | 389.6 | 0.29579 | 579.4 | 4.1287
768.081 | 0.02 | 389.7 | 0.3013 | 579.6 | 3.5497
768.322 | 0.02 | 389.8 | 0.30435 | 579.8 | 2.9225
768.563 | 0.016667 | 389.9 | 0.30739 | 580 | 2.5634
768.803 | 0.013333 | 390 | 0.31483 | 580.2 | 2.4716
769.044 | 0.01 | 390.1 | 0.31992 | 580.4 | 2.4617
769.284 | 0.0066667 | 390.2 | 0.32501 | 580.6 | 2.4457
769.525 | 0.02 | 390.3 | 0.32036 | 580.8 | 2.4378
769.765 | 0.016667 | 390.4 | 0.32189 | 581 | 2.4269
770.005 | 0.01 | 390.5 | 0.32341 | 581.2 | 2.4191
770.246 | 0.02 | 390.6 | 0.32493 | 581.4 | 2.4153
770.486 | 0.01 | 390.7 | 0.32667 | 581.6 | 2.4105
770.726 | -0.0033333 | 390.8 | 0.32667 | 581.8 | 2.4013
770.966 | 0.0033333 | 390.9 | 0.32667 | 582 | 2.4025
771.206 | 0.013333 | 391 | 0.32625 | 582.2 | 2.4052
771.446 | 0.0033333 | 391.1 | 0.32727 | 582.4 | 2.4097
771.686 | -0.0033333 | 391.2 | 0.32829 | 582.6 | 2.4134
771.926 | 0.01 | 391.3 | 0.32757 | 582.8 | 2.411
772.166 | 0.023333 | 391.4 | 0.32706 | 583 | 2.4094
772.406 | 0.06 | 391.5 | 0.32655 | 583.2 | 2.4122
772.645 | 0.076667 | 391.6 | 0.33021 | 583.4 | 2.4104
772.885 | 0.04 | 391.7 | 0.33174 | 583.6 | 2.4091
773.125 | 0.036667 | 391.8 | 0.33326 | 583.8 | 2.4048
773.364 | 0.013333 | 391.9 | 0.33479 | 584 | 2.3993
773.604 | 0.016667 | 392 | 0.33775 | 584.2 | 2.3979
773.843 | 0.026667 | 392.1 | 0.3403 | 584.4 | 2.3955
774.083 | 0.0033333 | 392.2 | 0.34284 | 584.6 | 2.3881
774.322 | -0.0033333 | 392.3 | 0.3403 | 584.8 | 2.3797
774.561 | 0.013333 | 392.4 | 0.34081 | 585 | 2.3724
774.801 | 0.0066667 | 392.5 | 0.34132 | 585.2 | 2.3756
775.04 | 0 | 392.6 | 0.34477 | 585.4 | 2.3701
775.279 | -0.0033333 | 392.7 | 0.34579 | 585.6 | 2.3641
775.518 | 0.01 | 392.8 | 0.34681 | 585.8 | 2.3572
775.757 | 0.01 | 392.9 | 0.34788 | 586 | 2.356
775.996 | 0 | 393 | 0.34992 | 586.2 | 2.3548
776.235 | -0.0033333 | 393.1 | 0.35196 | 586.4 | 2.3567
776.474 | 0.0066667 | 393.2 | 0.354 | 586.6 | 2.3636
776.713 | 0 | 393.3 | 0.35564 | 586.8 | 2.3745
776.951 | 0.016667 | 393.4 | 0.35717 | 587 | 2.3911
777.19 | -0.0033333 | 393.5 | 0.3587 | 587.2 | 2.4095
777.429 | -0.0033333 | 393.6 | 0.35967 | 587.4 | 2.4202
777.667 | 0.013333 | 393.7 | 0.3612 | 587.6 | 2.4159
777.906 | 0.013333 | 393.8 | 0.36273 | 587.8 | 2.3968
778.145 | 0.0033333 | 393.9 | 0.36605 | 588 | 2.3648
778.383 | 0.0033333 | 394 | 0.36809 | 588.2 | 2.3374
778.621 | 0.0066667 | 394.1 | 0.37013 | 588.4 | 2.3132
778.86 | 0 | 394.2 | 0.37218 | 588.6 | 2.2974
779.098 | 0 | 394.3 | 0.37632 | 588.8 | 2.2825
779.336 | 0.0066667 | 394.4 | 0.37939 | 589 | 2.2749
779.575 | 0.0033333 | 394.5 | 0.38245 | 589.2 | 2.2705
779.813 | -0.0033333 | 394.6 | 0.37962 | 589.4 | 2.27
780.051 | 0.0033333 | 394.7 | 0.38064 | 589.6 | 2.2631
780.289 | 0 | 394.8 | 0.38166 | 589.8 | 2.2564
780.527 | -0.016667 | 394.9 | 0.38624 | 590 | 2.2548
780.765 | -0.0033333 | 395 | 0.38777 | 590.2 | 2.251
781.003 | 0.016667 | 395.1 | 0.38931 | 590.4 | 2.2444
781.241 | -0.013333 | 395.2 | 0.39047 | 590.6 | 2.2374
781.478 | 0.0066667 | 395.3 | 0.39302 | 590.8 | 2.2373
781.716 | 0.01 | 395.4 | 0.39557 | 591 | 2.2331
781.954 | 0.016667 | 395.5 | 0.39813 | 591.2 | 2.2275
782.191 | 0.0033333 | 395.6 | 0.3976 | 591.4 | 2.2145
782.429 | 0.0066667 | 395.7 | 0.39862 | 591.6 | 2.2162
782.667 | -0.0066667 | 395.8 | 0.39964 | 591.8 | 2.2194
782.904 | -0.01 | 395.9 | 0.405 | 592 | 2.2212
783.142 | 0 | 396 | 0.40755 | 592.2 | 2.2189
783.379 | -0.01 | 396.1 | 0.4101 | 592.4 | 2.2216
783.616 | 0.0066667 | 396.2 | 0.41231 | 592.6 | 2.2277
783.853 | -0.0066667 | 396.3 | 0.41537 | 592.8 | 2.2357
784.091 | -0.023333 | 396.4 | 0.41844 | 593 | 2.2432
784.328 | 0.013333 | 396.5 | 0.42144 | 593.2 | 2.2415
784.565 | -0.0066667 | 396.6 | 0.424 | 593.4 | 2.2343
784.802 | -0.013333 | 396.7 | 0.42656 | 593.6 | 2.2212
785.039 | -0.0033333 | 396.8 | 0.42911 | 593.8 | 2.198
785.276 | -0.0066667 | 396.9 | 0.42956 | 594 | 2.1732
785.513 | -0.013333 | 397 | 0.4316 | 594.2 | 2.1507
785.75 | -0.0033333 | 397.1 | 0.43365 | 594.4 | 2.1371
785.987 | -0.016667 | 397.2 | 0.43538 | 594.6 | 2.129
786.223 | -0.016667 | 397.3 | 0.43691 | 594.8 | 2.1202
786.46 | -0.02 | 397.4 | 0.43845 | 595 | 2.1108
786.697 | -0.013333 | 397.5 | 0.43814 | 595.2 | 2.1024
786.933 | -0.013333 | 397.6 | 0.43916 | 595.4 | 2.0961
787.17 | -0.0033333 | 397.7 | 0.44018 | 595.6 | 2.0937
787.406 | -0.0033333 | 397.8 | 0.44694 | 595.8 | 2.0852
787.643 | -0.026667 | 397.9 | 0.45001 | 596 | 2.0787
787.879 | -0.01 | 398 | 0.45308 | 596.2 | 2.0724
788.116 | -0.0033333 | 398.1 | 0.45615 | 596.4 | 2.0659
788.352 | -0.023333 | 398.2 | 0.45895 | 596.6 | 2.0593
788.588 | 0 | 398.3 | 0.46304 | 596.8 | 2.0512
788.824 | -0.016667 | 398.4 | 0.46713 | 597 | 2.0441
789.06 | -0.023333 | 398.5 | 0.46644 | 597.2 | 2.0369
789.297 | -0.01 | 398.6 | 0.46797 | 597.4 | 2.0288
789.533 | -0.016667 | 398.7 | 0.46951 | 597.6 | 2.0234
789.769 | -0.023333 | 398.8 | 0.47287 | 597.8 | 2.0256
790.004 | -0.01 | 398.9 | 0.47492 | 598 | 2.0215
790.24 | 0.016667 | 399 | 0.47697 | 598.2 | 2.0202
790.476 | -0.023333 | 399.1 | 0.4779 | 598.4 | 2.0223
790.712 | -0.013333 | 399.2 | 0.47995 | 598.6 | 2.0152
790.948 | -0.02 | 399.3 | 0.48199 | 598.8 | 2.0124
791.183 | -0.023333 | 399.4 | 0.48404 | 599 | 2.0123
791.419 | -0.0066667 | 399.5 | 0.48568 | 599.2 | 2.0199
791.655 | -0.013333 | 399.6 | 0.48721 | 599.4 | 2.0199
791.89 | -0.016667 | 399.7 | 0.48875 | 599.6 | 2.012
792.126 | -0.0033333 | 399.8 | 0.49391 | 599.8 | 1.9921
792.361 | -0.0066667 | 399.9 | 0.49699 | 600 | 1.9842
792.596 | -0.013333 | 400 | 0.50007 | 600.2 | 1.9753
792.832 | -0.013333 | 400.1 | 0.50202 | 600.4 | 1.9673
793.067 | -0.013333 | 400.2 | 0.50509 | 600.6 | 1.9601
793.302 | -0.016667 | 400.3 | 0.50817 | 600.8 | 1.9495
793.537 | -0.013333 | 400.4 | 0.51122 | 601 | 1.9375
793.772 | -0.0066667 | 400.5 | 0.51378 | 601.2 | 1.9231
794.007 | -0.016667 | 400.6 | 0.51634 | 601.4 | 1.9192
794.242 | -0.013333 | 400.7 | 0.5189 | 601.6 | 1.9081
794.477 | -0.023333 | 400.8 | 0.52243 | 601.8 | 1.8967
794.712 | -0.023333 | 400.9 | 0.52551 | 602 | 1.8803
794.947 | 0.033333 | 401 | 0.52858 | 602.2 | 1.8769
795.182 | 0.04 | 401.1 | 0.53583 | 602.4 | 1.8744
795.417 | 0.023333 | 401.2 | 0.54044 | 602.6 | 1.8641
795.651 | 0 | 401.3 | 0.54506 | 602.8 | 1.8574
795.886 | -0.013333 | 401.4 | 0.54423 | 603 | 1.8518
796.121 | -0.0066667 | 401.5 | 0.5473 | 603.2 | 1.8496
796.355 | -0.01 | 401.6 | 0.55038 | 603.4 | 1.8441
796.59 | -0.036667 | 401.7 | 0.55339 | 603.6 | 1.8375
796.824 | -0.02 | 401.8 | 0.55493 | 603.8 | 1.8289
797.059 | -0.02 | 401.9 | 0.55647 | 604 | 1.8207
797.293 | -0.02 | 402 | 0.55801 | 604.2 | 1.819
797.527 | -0.016667 | 402.1 | 0.5685 | 604.4 | 1.8117
797.761 | -0.03 | 402.2 | 0.57362 | 604.6 | 1.8058
797.996 | -0.013333 | 402.3 | 0.57875 | 604.8 | 1.8007
798.23 | -0.0066667 | 402.4 | 0.58446 | 605 | 1.7978
798.464 | -0.013333 | 402.5 | 0.59062 | 605.2 | 1.7936
798.698 | -0.023333 | 402.6 | 0.59679 | 605.4 | 1.7863
798.932 | -0.03 | 402.7 | 0.60658 | 605.6 | 1.777
799.166 | -0.013333 | 402.8 | 0.61377 | 605.8 | 1.7755
799.4 | -0.02 | 402.9 | 0.62096 | 606 | 1.7755
799.634 | -0.026667 | 403 | 0.6657 | 606.2 | 1.7733
799.867 | -0.016667 | 403.1 | 0.69083 | 606.4 | 1.7675
800.101 | -0.04 | 403.2 | 0.71596 | 606.6 | 1.7571
800.335 | -0.023333 | 403.3 | 0.74109 | 606.8 | 1.7484
800.568 | -0.0066667 | 403.4 | 1.0858 | 607 | 1.7425
800.802 | -0.0033333 | 403.5 | 1.2244 | 607.2 | 1.7399
801.036 | 0.013333 | 403.6 | 1.3631 | 607.4 | 1.7365
801.269 | -0.01 | 403.7 | 2.3182 | 607.6 | 1.7287
801.503 | 0.0033333 | 403.8 | 2.796 | 607.8 | 1.7279
801.736 | 0.053333 | 403.9 | 3.2738 | 608 | 1.7264
801.969 | 0.06 | 404 | 4.7157 | 608.2 | 1.7267
802.203 | 0.033333 | 404.1 | 5.6878 | 608.4 | 1.7227
802.436 | -0.0033333 | 404.2 | 6.66 | 608.6 | 1.7194
802.669 | -0.01 | 404.3 | 7.1963 | 608.8 | 1.7214
802.902 | 0.0066667 | 404.4 | 8.0966 | 609 | 1.7267
803.135 | -0.013333 | 404.5 | 8.9968 | 609.2 | 1.7343
803.368 | -0.016667 | 404.6 | 9.897 | 609.4 | 1.7501
803.601 | 0 | 404.7 | 7.296 | 609.6 | 1.7689
803.834 | -0.016667 | 404.8 | 6.9067 | 609.8 | 1.7902
804.067 | 0.0066667 | 404.9 | 6.5174 | 610 | 1.8355
804.3 | -0.0033333 | 405 | 4.6788 | 610.2 | 1.9063
804.533 | 0 | 405.1 | 3.3835 | 610.4 | 2.0041
804.765 | 0 | 405.2 | 2.0881 | 610.6 | 2.1459
804.998 | -0.01 | 405.3 | 2.2088 | 610.8 | 2.3326
805.231 | 0.0033333 | 405.4 | 1.5025 | 611 | 2.5493
805.463 | 0 | 405.5 | 0.7962 | 611.2 | 2.7372
805.696 | -0.01 | 405.6 | 0.89741 | 611.4 | 2.8325
805.928 | -0.01 | 405.7 | 0.80481 | 611.6 | 2.7764
806.161 | 0.0033333 | 405.8 | 0.71222 | 611.8 | 2.6044
806.393 | 0.0066667 | 405.9 | 0.61963 | 612 | 2.3622
806.625 | -0.046667 | 406 | 0.69537 | 612.2 | 2.1741
806.858 | -0.056667 | 406.1 | 0.68611 | 612.4 | 2.0504
807.09 | -0.033333 | 406.2 | 0.67685 | 612.6 | 1.9646
807.322 | -0.026667 | 406.3 | 0.69525 | 612.8 | 1.9094
807.554 | -0.036667 | 406.4 | 0.69833 | 613 | 1.868
807.786 | -0.04 | 406.5 | 0.70142 | 613.2 | 1.837
|  | 406.6 | 0.7358 | 613.4 | 1.815
|  | 406.7 | 0.75383 | 613.6 | 1.7966
|  | 406.8 | 0.77185 | 613.8 | 1.7741
|  | 406.9 | 0.83976 | 614 | 1.7556
|  | 407 | 0.88356 | 614.2 | 1.7472
|  | 407.1 | 0.92736 | 614.4 | 1.7393
|  | 407.2 | 0.97116 | 614.6 | 1.7271
|  | 407.3 | 1.0912 | 614.8 | 1.7083
|  | 407.4 | 1.1658 | 615 | 1.6877
|  | 407.5 | 1.2404 | 615.2 | 1.6581
|  | 407.6 | 1.2128 | 615.4 | 1.6268
|  | 407.7 | 1.2567 | 615.6 | 1.5958
|  | 407.8 | 1.3005 | 615.8 | 1.5654
|  | 407.9 | 1.1359 | 616 | 1.5424
|  | 408 | 1.0808 | 616.2 | 1.5231
|  | 408.1 | 1.0256 | 616.4 | 1.5085
|  | 408.2 | 0.97043 | 616.6 | 1.4952
|  | 408.3 | 0.89056 | 616.8 | 1.4838
|  | 408.4 | 0.81068 | 617 | 1.4725
|  | 408.5 | 0.73081 | 617.2 | 1.4597
|  | 408.6 | 0.78508 | 617.4 | 1.4425
|  | 408.7 | 0.75825 | 617.6 | 1.4286
|  | 408.8 | 0.73142 | 617.8 | 1.4268
|  | 408.9 | 0.75667 | 618 | 1.4233
|  | 409 | 0.75667 | 618.2 | 1.418
|  | 409.1 | 0.75667 | 618.4 | 1.4096
|  | 409.2 | 0.75745 | 618.6 | 1.4031
|  | 409.3 | 0.75899 | 618.8 | 1.3963
|  | 409.4 | 0.76054 | 619 | 1.3877
|  | 409.5 | 0.76375 | 619.2 | 1.381
|  | 409.6 | 0.7653 | 619.4 | 1.3825
|  | 409.7 | 0.76685 | 619.6 | 1.3744
|  | 409.8 | 0.77012 | 619.8 | 1.367
|  | 409.9 | 0.77322 | 620 | 1.3603
|  | 410 | 0.77632 | 620.2 | 1.3546
|  | 410.1 | 0.77941 | 620.4 | 1.351
|  | 410.2 | 0.78251 | 620.6 | 1.346
|  | 410.3 | 0.7856 | 620.8 | 1.3404
|  | 410.4 | 0.7887 | 621 | 1.3311
|  | 410.5 | 0.79039 | 621.2 | 1.3241
|  | 410.6 | 0.79297 | 621.4 | 1.3175
|  | 410.7 | 0.79555 | 621.6 | 1.3081
|  | 410.8 | 0.79869 | 621.8 | 1.2993
|  | 410.9 | 0.80127 | 622 | 1.2897
|  | 411 | 0.80386 | 622.2 | 1.2833
|  | 411.1 | 0.80471 | 622.4 | 1.2774
|  | 411.2 | 0.80677 | 622.6 | 1.2698
|  | 411.3 | 0.80884 | 622.8 | 1.2634
|  | 411.4 | 0.81091 | 623 | 1.2613
|  | 411.5 | 0.80935 | 623.2 | 1.2617
|  | 411.6 | 0.80987 | 623.4 | 1.2647
|  | 411.7 | 0.81039 | 623.6 | 1.2622
|  | 411.8 | 0.81729 | 623.8 | 1.245
|  | 411.9 | 0.81988 | 624 | 1.2316
|  | 412 | 0.82246 | 624.2 | 1.2256
|  | 412.1 | 0.82361 | 624.4 | 1.2256
|  | 412.2 | 0.82671 | 624.6 | 1.2233
|  | 412.3 | 0.82981 | 624.8 | 1.2184
|  | 412.4 | 0.83268 | 625 | 1.2104
|  | 412.5 | 0.83475 | 625.2 | 1.2031
|  | 412.6 | 0.83682 | 625.4 | 1.1958
|  | 412.7 | 0.83888 | 625.6 | 1.1862
|  | 412.8 | 0.83822 | 625.8 | 1.1752
|  | 412.9 | 0.83977 | 626 | 1.1629
|  | 413 | 0.84133 | 626.2 | 1.1587
|  | 413.1 | 0.84908 | 626.4 | 1.1622
|  | 413.2 | 0.85218 | 626.6 | 1.1626
|  | 413.3 | 0.85528 | 626.8 | 1.1511
|  | 413.4 | 0.85505 | 627 | 1.1457
|  | 413.5 | 0.85815 | 627.2 | 1.1405
|  | 413.6 | 0.86125 | 627.4 | 1.1353
|  | 413.7 | 0.86256 | 627.6 | 1.1218
|  | 413.8 | 0.8636 | 627.8 | 1.1177
|  | 413.9 | 0.86463 | 628 | 1.1116
|  | 414 | 0.86806 | 628.2 | 1.1019
|  | 414.1 | 0.87065 | 628.4 | 1.099
|  | 414.2 | 0.87324 | 628.6 | 1.0951
|  | 414.3 | 0.87583 | 628.8 | 1.0964
|  | 414.4 | 0.87157 | 629 | 1.0955
|  | 414.5 | 0.87209 | 629.2 | 1.0886
|  | 414.6 | 0.87261 | 629.4 | 1.0835
|  | 414.7 | 0.8759 | 629.6 | 1.0768
|  | 414.8 | 0.87642 | 629.8 | 1.0727
|  | 414.9 | 0.87694 | 630 | 1.0745
|  | 415 | 0.88274 | 630.2 | 1.0835
|  | 415.1 | 0.88636 | 630.4 | 1.0959
|  | 415.2 | 0.88999 | 630.6 | 1.1104
|  | 415.3 | 0.89184 | 630.8 | 1.1239
|  | 415.4 | 0.89495 | 631 | 1.1344
|  | 415.5 | 0.89806 | 631.2 | 1.133
|  | 415.6 | 0.89891 | 631.4 | 1.117
|  | 415.7 | 0.90047 | 631.6 | 1.0907
|  | 415.8 | 0.90202 | 631.8 | 1.0569
|  | 415.9 | 0.90358 | 632 | 1.0262
|  | 416 | 0.90569 | 632.2 | 1.0011
|  | 416.1 | 0.90724 | 632.4 | 0.97787
|  | 416.2 | 0.90879 | 632.6 | 0.96236
|  | 416.3 | 0.91262 | 632.8 | 0.95827
|  | 416.4 | 0.91521 | 633 | 0.94695
|  | 416.5 | 0.91781 | 633.2 | 0.93843
|  | 416.6 | 0.92114 | 633.4 | 0.93016
|  | 416.7 | 0.92425 | 633.6 | 0.92601
|  | 416.8 | 0.92736 | 633.8 | 0.92528
|  | 416.9 | 0.92691 | 634 | 0.91879
|  | 417 | 0.92846 | 634.2 | 0.91136
|  | 417.1 | 0.93002 | 634.4 | 0.90695
|  | 417.2 | 0.93157 | 634.6 | 0.90281
|  | 417.3 | 0.9332 | 634.8 | 0.89896
|  | 417.4 | 0.93424 | 635 | 0.89503
|  | 417.5 | 0.93528 | 635.2 | 0.89667
|  | 417.6 | 0.94116 | 635.4 | 0.89445
|  | 417.7 | 0.94428 | 635.6 | 0.88779
|  | 417.8 | 0.94739 | 635.8 | 0.87164
|  | 417.9 | 0.95076 | 636 | 0.86338
|  | 418 | 0.95439 | 636.2 | 0.85928
|  | 418.1 | 0.95802 | 636.4 | 0.85947
|  | 418.2 | 0.96094 | 636.6 | 0.8572
|  | 418.3 | 0.96405 | 636.8 | 0.85356
|  | 418.4 | 0.96716 | 637 | 0.84693
|  | 418.5 | 0.96449 | 637.2 | 0.83936
|  | 418.6 | 0.96501 | 637.4 | 0.8327
|  | 418.7 | 0.96553 | 637.6 | 0.8259
|  | 418.8 | 0.96605 | 637.8 | 0.81903
|  | 418.9 | 0.9772 | 638 | 0.81208
|  | 419 | 0.98032 | 638.2 | 0.80627
|  | 419.1 | 0.98343 | 638.4 | 0.81016
|  | 419.2 | 0.98392 | 638.6 | 0.80977
|  | 419.3 | 0.98808 | 638.8 | 0.80473
|  | 419.4 | 0.99224 | 639 | 0.79513
|  | 419.5 | 0.99407 | 639.2 | 0.78796
|  | 419.6 | 0.99562 | 639.4 | 0.78706
|  | 419.7 | 0.99718 | 639.6 | 0.78479
|  | 419.8 | 1.0008 | 639.8 | 0.77573
|  | 419.9 | 1.0039 | 640 | 0.77411
|  | 420 | 1.007 | 640.2 | 0.77044
|  | 420.1 | 1.0067 | 640.4 | 0.76593
|  | 420.2 | 1.0083 | 640.6 | 0.76346
|  | 420.3 | 1.0099 | 640.8 | 0.75749
|  | 420.4 | 1.0114 | 641 | 0.75332
|  | 420.5 | 1.01 | 641.2 | 0.7494
|  | 420.6 | 1.01 | 641.4 | 0.74442
|  | 420.7 | 1.01 | 641.6 | 0.74181
|  | 420.8 | 1.0129 | 641.8 | 0.73864
|  | 420.9 | 1.0139 | 642 | 0.73588
|  | 421 | 1.015 | 642.2 | 0.73414
|  | 421.1 | 1.0203 | 642.4 | 0.72752
|  | 421.2 | 1.0234 | 642.6 | 0.72164
|  | 421.3 | 1.0265 | 642.8 | 0.71978
|  | 421.4 | 1.026 | 643 | 0.71955
|  | 421.5 | 1.0281 | 643.2 | 0.71674
|  | 421.6 | 1.0302 | 643.4 | 0.71125
|  | 421.7 | 1.03 | 643.6 | 0.70481
|  | 421.8 | 1.03 | 643.8 | 0.70146
|  | 421.9 | 1.03 | 644 | 0.69015
|  | 422 | 1.03 | 644.2 | 0.68273
|  | 422.1 | 1.0393 | 644.4 | 0.68786
|  | 422.2 | 1.0424 | 644.6 | 0.68836
|  | 422.3 | 1.0455 | 644.8 | 0.6871
|  | 422.4 | 1.0467 | 645 | 0.68057
|  | 422.5 | 1.0504 | 645.2 | 0.67399
|  | 422.6 | 1.054 | 645.4 | 0.67079
|  | 422.7 | 1.0591 | 645.6 | 0.66683
|  | 422.8 | 1.0622 | 645.8 | 0.66166
|  | 422.9 | 1.0654 | 646 | 0.65429
|  | 423 | 1.0638 | 646.2 | 0.65074
|  | 423.1 | 1.0659 | 646.4 | 0.65301
|  | 423.2 | 1.068 | 646.6 | 0.65446
|  | 423.3 | 1.079 | 646.8 | 0.64927
|  | 423.4 | 1.0837 | 647 | 0.64335
|  | 423.5 | 1.0884 | 647.2 | 0.63635
|  | 423.6 | 1.0931 | 647.4 | 0.6323
|  | 423.7 | 1.0901 | 647.6 | 0.63042
|  | 423.8 | 1.0943 | 647.8 | 0.62846
|  | 423.9 | 1.0985 | 648 | 0.62754
|  | 424 | 1.1064 | 648.2 | 0.62447
|  | 424.1 | 1.1095 | 648.4 | 0.61671
|  | 424.2 | 1.1127 | 648.6 | 0.6161
|  | 424.3 | 1.1169 | 648.8 | 0.60959
|  | 424.4 | 1.1227 | 649 | 0.6043
|  | 424.5 | 1.1284 | 649.2 | 0.60519
|  | 424.6 | 1.1233 | 649.4 | 0.60162
|  | 424.7 | 1.1233 | 649.6 | 0.60406
|  | 424.8 | 1.1233 | 649.8 | 0.6072
|  | 424.9 | 1.1267 | 650 | 0.61484
|  | 425 | 1.1267 | 650.2 | 0.619
|  | 425.1 | 1.1267 | 650.4 | 0.62186
|  | 425.2 | 1.1267 | 650.6 | 0.62304
|  | 425.3 | 1.1233 | 650.8 | 0.62027
|  | 425.4 | 1.1233 | 651 | 0.60486
|  | 425.5 | 1.1233 | 651.2 | 0.59494
|  | 425.6 | 1.1296 | 651.4 | 0.58817
|  | 425.7 | 1.1306 | 651.6 | 0.57777
|  | 425.8 | 1.1317 | 651.8 | 0.57059
|  | 425.9 | 1.1337 | 652 | 0.56319
|  | 426 | 1.1368 | 652.2 | 0.5665
|  | 426.1 | 1.1399 | 652.4 | 0.57027
|  | 426.2 | 1.134 | 652.6 | 0.56752
|  | 426.3 | 1.1325 | 652.8 | 0.56387
|  | 426.4 | 1.1309 | 653 | 0.56088
|  | 426.5 | 1.1287 | 653.2 | 0.55528
|  | 426.6 | 1.1255 | 653.4 | 0.55093
|  | 426.7 | 1.1224 | 653.6 | 0.5456
|  | 426.8 | 1.1193 | 653.8 | 0.54154
|  | 426.9 | 1.126 | 654 | 0.53666
|  | 427 | 1.1265 | 654.2 | 0.53042
|  | 427.1 | 1.1271 | 654.4 | 0.52709
|  | 427.2 | 1.141 | 654.6 | 0.52452
|  | 427.3 | 1.1473 | 654.8 | 0.52053
|  | 427.4 | 1.1535 | 655 | 0.51703
|  | 427.5 | 1.1565 | 655.2 | 0.51333
|  | 427.6 | 1.1627 | 655.4 | 0.50989
|  | 427.7 | 1.169 | 655.6 | 0.50434
|  | 427.8 | 1.171 | 655.8 | 0.50256
|  | 427.9 | 1.1741 | 656 | 0.50345
|  | 428 | 1.1773 | 656.2 | 0.50409
|  | 428.1 | 1.1804 | 656.4 | 0.50048
|  | 428.2 | 1.1835 | 656.6 | 0.49193
|  | 428.3 | 1.1867 | 656.8 | 0.48542
|  | 428.4 | 1.1898 | 657 | 0.48219
|  | 428.5 | 1.1881 | 657.2 | 0.48111
|  | 428.6 | 1.1897 | 657.4 | 0.48185
|  | 428.7 | 1.1913 | 657.6 | 0.48446
|  | 428.8 | 1.199 | 657.8 | 0.48729
|  | 428.9 | 1.2022 | 658 | 0.48757
|  | 429 | 1.2053 | 658.2 | 0.48097
|  | 429.1 | 1.2051 | 658.4 | 0.47721
|  | 429.2 | 1.2082 | 658.6 | 0.47566
|  | 429.3 | 1.2114 | 658.8 | 0.47333
|  | 429.4 | 1.2179 | 659 | 0.46516
|  | 429.5 | 1.221 | 659.2 | 0.45736
|  | 429.6 | 1.2241 | 659.4 | 0.45496
|  | 429.7 | 1.2312 | 659.6 | 0.45955
|  | 429.8 | 1.2375 | 659.8 | 0.4543
|  | 429.9 | 1.2438 | 660 | 0.44914
|  | 430 | 1.2445 | 660.2 | 0.44575
|  | 430.1 | 1.2482 | 660.4 | 0.44655
|  | 430.2 | 1.2518 | 660.6 | 0.44416
|  | 430.3 | 1.2555 | 660.8 | 0.44257
|  | 430.4 | 1.2697 | 661 | 0.44084
|  | 430.5 | 1.276 | 661.2 | 0.43683
|  | 430.6 | 1.2822 | 661.4 | 0.43716
|  | 430.7 | 1.2844 | 661.6 | 0.44062
|  | 430.8 | 1.2912 | 661.8 | 0.44486
|  | 430.9 | 1.2981 | 662 | 0.44657
|  | 431 | 1.3004 | 662.2 | 0.44844
|  | 431.1 | 1.3035 | 662.4 | 0.44597
|  | 431.2 | 1.3067 | 662.6 | 0.44115
|  | 431.3 | 1.3062 | 662.8 | 0.43452
|  | 431.4 | 1.3083 | 663 | 0.43122
|  | 431.5 | 1.3104 | 663.2 | 0.42518
|  | 431.6 | 1.303 | 663.4 | 0.41952
|  | 431.7 | 1.2998 | 663.6 | 0.4149
|  | 431.8 | 1.2967 | 663.8 | 0.41183
|  | 431.9 | 1.2936 | 664 | 0.40576
|  | 432 | 1.2889 | 664.2 | 0.40086
|  | 432.1 | 1.2842 | 664.4 | 0.39981
|  | 432.2 | 1.2795 | 664.6 | 0.39506
|  | 432.3 | 1.2912 | 664.8 | 0.39233
|  | 432.4 | 1.2928 | 665 | 0.38938
|  | 432.5 | 1.2943 | 665.2 | 0.38956
|  | 432.6 | 1.3036 | 665.4 | 0.38804
|  | 432.7 | 1.3099 | 665.6 | 0.38213
|  | 432.8 | 1.3162 | 665.8 | 0.38079
|  | 432.9 | 1.3317 | 666 | 0.38074
|  | 433 | 1.3417 | 666.2 | 0.37778
|  | 433.1 | 1.3516 | 666.4 | 0.38075
|  | 433.2 | 1.3675 | 666.6 | 0.38131
|  | 433.3 | 1.3817 | 666.8 | 0.38193
|  | 433.4 | 1.3959 | 667 | 0.38028
|  | 433.5 | 1.4208 | 667.2 | 0.37568
|  | 433.6 | 1.4403 | 667.4 | 0.3712
|  | 433.7 | 1.4597 | 667.6 | 0.36711
|  | 433.8 | 1.4791 | 667.8 | 0.3646
|  | 433.9 | 1.5248 | 668 | 0.36344
|  | 434 | 1.5547 | 668.2 | 0.36217
|  | 434.1 | 1.5847 | 668.4 | 0.36038
|  | 434.2 | 1.6749 | 668.6 | 0.35819
|  | 434.3 | 1.7284 | 668.8 | 0.35707
|  | 434.4 | 1.782 | 669 | 0.35428
|  | 434.5 | 2.3152 | 669.2 | 0.35167
|  | 434.6 | 2.5581 | 669.4 | 0.35076
|  | 434.7 | 2.801 | 669.6 | 0.35189
|  | 434.8 | 4.562 | 669.8 | 0.34797
|  | 434.9 | 5.4947 | 670 | 0.34376
|  | 435 | 6.4274 | 670.2 | 0.34221
|  | 435.1 | 8.6053 | 670.4 | 0.34111
|  | 435.2 | 10.3198 | 670.6 | 0.3406
|  | 435.3 | 12.0343 | 670.8 | 0.33902
|  | 435.4 | 12.8004 | 671 | 0.33681
|  | 435.5 | 14.1991 | 671.2 | 0.33277
|  | 435.6 | 15.5977 | 671.4 | 0.33063
|  | 435.7 | 16.9964 | 671.6 | 0.33309
|  | 435.8 | 16.0928 | 671.8 | 0.33552
|  | 435.9 | 16.5039 | 672 | 0.33334
|  | 436 | 16.9149 | 672.2 | 0.32737
|  | 436.1 | 12.6395 | 672.4 | 0.32197
|  | 436.2 | 11.2053 | 672.6 | 0.31987
|  | 436.3 | 9.7712 | 672.8 | 0.32
|  | 436.4 | 7.5824 | 673 | 0.32
|  | 436.5 | 5.2827 | 673.2 | 0.31652
|  | 436.6 | 2.983 | 673.4 | 0.31358
|  | 436.7 | 3.3924 | 673.6 | 0.31386
|  | 436.8 | 2.4434 | 673.8 | 0.31222
|  | 436.9 | 1.4944 | 674 | 0.30638
|  | 437 | 1.6647 | 674.2 | 0.30231
|  | 437.1 | 1.5758 | 674.4 | 0.30193
|  | 437.2 | 1.487 | 674.6 | 0.30489
|  | 437.3 | 1.3981 | 674.8 | 0.30266
|  | 437.4 | 1.4828 | 675 | 0.30016
|  | 437.5 | 1.4733 | 675.2 | 0.2981
|  | 437.6 | 1.4638 | 675.4 | 0.29568
|  | 437.7 | 1.4711 | 675.6 | 0.29147
|  | 437.8 | 1.4711 | 675.8 | 0.28912
|  | 437.9 | 1.4711 | 676 | 0.28745
|  | 438 | 1.4751 | 676.2 | 0.28374
|  | 438.1 | 1.4762 | 676.4 | 0.28125
|  | 438.2 | 1.4772 | 676.6 | 0.27862
|  | 438.3 | 1.4857 | 676.8 | 0.27916
|  | 438.4 | 1.4905 | 677 | 0.28233
|  | 438.5 | 1.4952 | 677.2 | 0.28191
|  | 438.6 | 1.4983 | 677.4 | 0.27891
|  | 438.7 | 1.503 | 677.6 | 0.27712
|  | 438.8 | 1.5077 | 677.8 | 0.27704
|  | 438.9 | 1.5105 | 678 | 0.27555
|  | 439 | 1.5137 | 678.2 | 0.27182
|  | 439.1 | 1.5169 | 678.4 | 0.26717
|  | 439.2 | 1.5167 | 678.6 | 0.26386
|  | 439.3 | 1.5183 | 678.8 | 0.26129
|  | 439.4 | 1.5198 | 679 | 0.26424
|  | 439.5 | 1.5214 | 679.2 | 0.26903
|  | 439.6 | 1.5294 | 679.4 | 0.27049
|  | 439.7 | 1.5325 | 679.6 | 0.26488
|  | 439.8 | 1.5357 | 679.8 | 0.26144
|  | 439.9 | 1.5355 | 680 | 0.25963
|  | 440 | 1.5387 | 680.2 | 0.26058
|  | 440.1 | 1.5418 | 680.4 | 0.26174
|  | 440.2 | 1.547 | 680.6 | 0.26111
|  | 440.3 | 1.5496 | 680.8 | 0.25929
|  | 440.4 | 1.5522 | 681 | 0.25851
|  | 440.5 | 1.5495 | 681.2 | 0.25656
|  | 440.6 | 1.5511 | 681.4 | 0.24935
|  | 440.7 | 1.5526 | 681.6 | 0.24781
|  | 440.8 | 1.5557 | 681.8 | 0.25168
|  | 440.9 | 1.5562 | 682 | 0.25527
|  | 441 | 1.5567 | 682.2 | 0.25078
|  | 441.1 | 1.5533 | 682.4 | 0.24629
|  | 441.2 | 1.5533 | 682.6 | 0.24234
|  | 441.3 | 1.5533 | 682.8 | 0.2412
|  | 441.4 | 1.5533 | 683 | 0.24364
|  | 441.5 | 1.5535 | 683.2 | 0.24904
|  | 441.6 | 1.5524 | 683.4 | 0.24832
|  | 441.7 | 1.5514 | 683.6 | 0.24479
|  | 441.8 | 1.5556 | 683.8 | 0.24011
|  | 441.9 | 1.5572 | 684 | 0.23586
|  | 442 | 1.5588 | 684.2 | 0.24062
|  | 442.1 | 1.5627 | 684.4 | 0.23978
|  | 442.2 | 1.5653 | 684.6 | 0.23667
|  | 442.3 | 1.5679 | 684.8 | 0.23447
|  | 442.4 | 1.57 | 685 | 0.23222
|  | 442.5 | 1.5727 | 685.2 | 0.23128
|  | 442.6 | 1.5753 | 685.4 | 0.23197
|  | 442.7 | 1.5785 | 685.6 | 0.23214
|  | 442.8 | 1.5812 | 685.8 | 0.22615
|  | 442.9 | 1.5838 | 686 | 0.22489
|  | 443 | 1.5882 | 686.2 | 0.22501
|  | 443.1 | 1.5919 | 686.4 | 0.22769
|  | 443.2 | 1.5956 | 686.6 | 0.2314
|  | 443.3 | 1.5993 | 686.8 | 0.23606
|  | 443.4 | 1.5976 | 687 | 0.2389
|  | 443.5 | 1.5997 | 687.2 | 0.23853
|  | 443.6 | 1.6018 | 687.4 | 0.23619
|  | 443.7 | 1.6093 | 687.6 | 0.24249
|  | 443.8 | 1.6124 | 687.8 | 0.24303
|  | 443.9 | 1.6156 | 688 | 0.23951
|  | 444 | 1.6096 | 688.2 | 0.23132
|  | 444.1 | 1.6107 | 688.4 | 0.22197
|  | 444.2 | 1.6117 | 688.6 | 0.21604
|  | 444.3 | 1.6114 | 688.8 | 0.21495
|  | 444.4 | 1.6098 | 689 | 0.21626
|  | 444.5 | 1.6082 | 689.2 | 0.21649
|  | 444.6 | 1.6104 | 689.4 | 0.21026
|  | 444.7 | 1.6114 | 689.6 | 0.20488
|  | 444.8 | 1.6125 | 689.8 | 0.20393
|  | 444.9 | 1.6111 | 690 | 0.20772
|  | 445 | 1.6111 | 690.2 | 0.21355
|  | 445.1 | 1.6111 | 690.4 | 0.24823
|  | 445.2 | 1.6122 | 690.6 | 0.29068
|  | 445.3 | 1.6122 | 690.8 | 0.307
|  | 445.4 | 1.6122 | 691 | 0.26329
|  | 445.5 | 1.6122 | 691.2 | 0.22608
|  | 445.6 | 1.6127 | 691.4 | 0.20094
|  | 445.7 | 1.6132 | 691.6 | 0.19578
|  | 445.8 | 1.6138 | 691.8 | 0.1976
|  | 445.9 | 1.6164 | 692 | 0.19756
|  | 446 | 1.6174 | 692.2 | 0.19082
|  | 446.1 | 1.6185 | 692.4 | 0.18808
|  | 446.2 | 1.6141 | 692.6 | 0.18894
|  | 446.3 | 1.6136 | 692.8 | 0.19116
|  | 446.4 | 1.6131 | 693 | 0.19404
|  | 446.5 | 1.6126 | 693.2 | 0.19658
|  | 446.6 | 1.611 | 693.4 | 0.20062
|  | 446.7 | 1.6094 | 693.6 | 0.20469
|  | 446.8 | 1.61 | 693.8 | 0.2082
|  | 446.9 | 1.61 | 694 | 0.20207
|  | 447 | 1.61 | 694.2 | 0.19472
|  | 447.1 | 1.61 | 694.4 | 0.18845
|  | 447.2 | 1.61 | 694.6 | 0.18364
|  | 447.3 | 1.61 | 694.8 | 0.17944
|  | 447.4 | 1.6134 | 695 | 0.17938
|  | 447.5 | 1.615 | 695.2 | 0.18049
|  | 447.6 | 1.6166 | 695.4 | 0.17829
|  | 447.7 | 1.6182 | 695.6 | 0.17513
|  | 447.8 | 1.6229 | 695.8 | 0.16956
|  | 447.9 | 1.6261 | 696 | 0.17304
|  | 448 | 1.6293 | 696.2 | 0.183
|  | 448.1 | 1.6294 | 696.4 | 0.19467
|  | 448.2 | 1.6315 | 696.6 | 0.19333
|  | 448.3 | 1.6337 | 696.8 | 0.18475
|  | 448.4 | 1.6311 | 697 | 0.17245
|  | 448.5 | 1.6311 | 697.2 | 0.16699
|  | 448.6 | 1.6311 | 697.4 | 0.16082
|  | 448.7 | 1.6391 | 697.6 | 0.16032
|  | 448.8 | 1.6417 | 697.8 | 0.16603
|  | 448.9 | 1.6444 | 698 | 0.17036
|  | 449 | 1.6406 | 698.2 | 0.16443
|  | 449.1 | 1.6421 | 698.4 | 0.15631
|  | 449.2 | 1.6437 | 698.6 | 0.1524
|  | 449.3 | 1.6467 | 698.8 | 0.15346
|  | 449.4 | 1.6467 | 699 | 0.15455
|  | 449.5 | 1.6467 | 699.2 | 0.15239
|  | 449.6 | 1.6468 | 699.4 | 0.15892
|  | 449.7 | 1.6484 | 699.6 | 0.15905
|  | 449.8 | 1.65 | 699.8 | 0.15728
|  | 449.9 | 1.6516 | 700 | 0.15737
|  | 450 | 1.65 | 700.2 | 0.15903
|  | 450.1 | 1.65 | 700.4 | 0.16033
|  | 450.2 | 1.65 | 700.6 | 0.158
|  | 450.3 | 1.6592 | 700.8 | 0.15493
|  | 450.4 | 1.6624 | 701 | 0.15144
|  | 450.5 | 1.6656 | 701.2 | 0.15407
|  | 450.6 | 1.6655 | 701.4 | 0.15342
|  | 450.7 | 1.6686 | 701.6 | 0.14915
|  | 450.8 | 1.6718 | 701.8 | 0.14516
|  | 450.9 | 1.6714 | 702 | 0.15034
|  | 451 | 1.6719 | 702.2 | 0.15037
|  | 451.1 | 1.6725 | 702.4 | 0.15224
|  | 451.2 | 1.6777 | 702.6 | 0.15376
|  | 451.3 | 1.6804 | 702.8 | 0.15556
|  | 451.4 | 1.683 | 703 | 0.15484
|  | 451.5 | 1.684 | 703.2 | 0.1539
|  | 451.6 | 1.6867 | 703.4 | 0.15143
|  | 451.7 | 1.6894 | 703.6 | 0.14759
|  | 451.8 | 1.6842 | 703.8 | 0.14795
|  | 451.9 | 1.6826 | 704 | 0.14479
|  | 452 | 1.681 | 704.2 | 0.14662
|  | 452.1 | 1.6794 | 704.4 | 0.14801
|  | 452.2 | 1.6856 | 704.6 | 0.14803
|  | 452.3 | 1.6856 | 704.8 | 0.15193
|  | 452.4 | 1.6856 | 705 | 0.14812
|  | 452.5 | 1.6858 | 705.2 | 0.14721
|  | 452.6 | 1.6874 | 705.4 | 0.14811
|  | 452.7 | 1.689 | 705.6 | 0.15432
|  | 452.8 | 1.6919 | 705.8 | 0.16124
|  | 452.9 | 1.693 | 706 | 0.16979
|  | 453 | 1.694 | 706.2 | 0.18027
|  | 453.1 | 1.6943 | 706.4 | 0.19391
|  | 453.2 | 1.6959 | 706.6 | 0.21007
|  | 453.3 | 1.6975 | 706.8 | 0.21699
|  | 453.4 | 1.7021 | 707 | 0.21213
|  | 453.5 | 1.7043 | 707.2 | 0.20271
|  | 453.6 | 1.7064 | 707.4 | 0.19331
|  | 453.7 | 1.7063 | 707.6 | 0.18508
|  | 453.8 | 1.7084 | 707.8 | 0.18229
|  | 453.9 | 1.7106 | 708 | 0.18148
|  | 454 | 1.71 | 708.2 | 0.17794
|  | 454.1 | 1.71 | 708.4 | 0.17297
|  | 454.2 | 1.71 | 708.6 | 0.18425
|  | 454.3 | 1.7135 | 708.8 | 0.20579
|  | 454.4 | 1.7151 | 709 | 0.22917
|  | 454.5 | 1.7167 | 709.2 | 0.22514
|  | 454.6 | 1.7183 | 709.4 | 0.20331
|  | 454.7 | 1.7231 | 709.6 | 0.17949
|  | 454.8 | 1.7263 | 709.8 | 0.15791
|  | 454.9 | 1.7295 | 710 | 0.14322
|  | 455 | 1.728 | 710.2 | 0.13572
|  | 455.1 | 1.7296 | 710.4 | 0.12987
|  | 455.2 | 1.7312 | 710.6 | 0.12852
|  | 455.3 | 1.73 | 710.8 | 0.13347
|  | 455.4 | 1.73 | 711 | 0.13753
|  | 455.5 | 1.73 | 711.2 | 0.13552
|  | 455.6 | 1.73 | 711.4 | 0.137
|  | 455.7 | 1.73 | 711.6 | 0.14926
|  | 455.8 | 1.73 | 711.8 | 0.16736
|  | 455.9 | 1.7341 | 712 | 0.18327
|  | 456 | 1.7357 | 712.2 | 0.19206
|  | 456.1 | 1.7373 | 712.4 | 0.19291
|  | 456.2 | 1.7411 | 712.6 | 0.1885
|  | 456.3 | 1.7443 | 712.8 | 0.17534
|  | 456.4 | 1.7475 | 713 | 0.15377
|  | 456.5 | 1.747 | 713.2 | 0.13668
|  | 456.6 | 1.7486 | 713.4 | 0.12528
|  | 456.7 | 1.7502 | 713.6 | 0.12064
|  | 456.8 | 1.7535 | 713.8 | 0.11833
|  | 456.9 | 1.7551 | 714 | 0.11934
|  | 457 | 1.7567 | 714.2 | 0.1149
|  | 457.1 | 1.7583 | 714.4 | 0.1158
|  | 457.2 | 1.7582 | 714.6 | 0.11442
|  | 457.3 | 1.7598 | 714.8 | 0.11283
|  | 457.4 | 1.7614 | 715 | 0.10886
|  | 457.5 | 1.76 | 715.2 | 0.10227
|  | 457.6 | 1.76 | 715.4 | 0.10231
|  | 457.7 | 1.76 | 715.6 | 0.1032
|  | 457.8 | 1.7691 | 715.8 | 0.10327
|  | 457.9 | 1.7723 | 716 | 0.10071
|  | 458 | 1.7755 | 716.2 | 0.10297
|  | 458.1 | 1.7725 | 716.4 | 0.10245
|  | 458.2 | 1.7746 | 716.6 | 0.1047
|  | 458.3 | 1.7767 | 716.8 | 0.10701
|  | 458.4 | 1.7709 | 717 | 0.10882
|  | 458.5 | 1.7682 | 717.2 | 0.10819
|  | 458.6 | 1.7655 | 717.4 | 0.10364
|  | 458.7 | 1.7726 | 717.6 | 0.10138
|  | 458.8 | 1.7737 | 717.8 | 0.1014
|  | 458.9 | 1.7748 | 718 | 0.1031
|  | 459 | 1.7727 | 718.2 | 0.1023
|  | 459.1 | 1.7743 | 718.4 | 0.10295
|  | 459.2 | 1.7759 | 718.6 | 0.10557
|  | 459.3 | 1.7826 | 718.8 | 0.10682
|  | 459.4 | 1.7847 | 719 | 0.10161
|  | 459.5 | 1.7868 | 719.2 | 0.096751
|  | 459.6 | 1.7868 | 719.4 | 0.096061
|  | 459.7 | 1.79 | 719.6 | 0.096006
|  | 459.8 | 1.7932 | 719.8 | 0.092855
|  | 459.9 | 1.7964 | 720 | 0.09556
|  | 460 | 1.7966 | 720.2 | 0.096613
|  | 460.1 | 1.7976 | 720.4 | 0.094187
|  | 460.2 | 1.7987 | 720.6 | 0.089157
|  | 460.3 | 1.794 | 720.8 | 0.08889
|  | 460.4 | 1.7935 | 721 | 0.09459
|  | 460.5 | 1.7929 | 721.2 | 0.097102
|  | 460.6 | 1.7952 | 721.4 | 0.095331
|  | 460.7 | 1.7946 | 721.6 | 0.091722
|  | 460.8 | 1.7941 | 721.8 | 0.089058
|  | 460.9 | 1.7994 | 722 | 0.084137
|  | 461 | 1.8021 | 722.2 | 0.081264
|  | 461.1 | 1.8047 | 722.4 | 0.082088
|  | 461.2 | 1.8014 | 722.6 | 0.083772
|  | 461.3 | 1.8019 | 722.8 | 0.082119
|  | 461.4 | 1.8024 | 723 | 0.080578
|  | 461.5 | 1.8071 | 723.2 | 0.083605
|  | 461.6 | 1.8081 | 723.4 | 0.086606
|  | 461.7 | 1.8092 | 723.6 | 0.087139
|  | 461.8 | 1.8057 | 723.8 | 0.084799
|  | 461.9 | 1.8062 | 724 | 0.083637
|  | 462 | 1.8068 | 724.2 | 0.083213
|  | 462.1 | 1.8124 | 724.4 | 0.084386
|  | 462.2 | 1.8135 | 724.6 | 0.087824
|  | 462.3 | 1.8145 | 724.8 | 0.092612
|  | 462.4 | 1.8134 | 725 | 0.089211
|  | 462.5 | 1.8156 | 725.2 | 0.082322
|  | 462.6 | 1.8177 | 725.4 | 0.0778
|  | 462.7 | 1.8199 | 725.6 | 0.07755
|  | 462.8 | 1.8237 | 725.8 | 0.077936
|  | 462.9 | 1.8253 | 726 | 0.078889
|  | 463 | 1.827 | 726.2 | 0.080687
|  | 463.1 | 1.8274 | 726.4 | 0.079397
|  | 463.2 | 1.8295 | 726.6 | 0.07671
|  | 463.3 | 1.8316 | 726.8 | 0.075609
|  | 463.4 | 1.83 | 727 | 0.074444
|  | 463.5 | 1.83 | 727.2 | 0.078994
|  | 463.6 | 1.83 | 727.4 | 0.080322
|  | 463.7 | 1.8329 | 727.6 | 0.078498
|  | 463.8 | 1.834 | 727.8 | 0.074009
|  | 463.9 | 1.835 | 728 | 0.074447
|  | 464 | 1.8364 | 728.2 | 0.075518
|  | 464.1 | 1.838 | 728.4 | 0.075629
|  | 464.2 | 1.8396 | 728.6 | 0.074795
|  | 464.3 | 1.8378 | 728.8 | 0.07644
|  | 464.4 | 1.8378 | 729 | 0.080217
|  | 464.5 | 1.8378 | 729.2 | 0.07486
|  | 464.6 | 1.8427 | 729.4 | 0.067994
|  | 464.7 | 1.8443 | 729.6 | 0.066359
|  | 464.8 | 1.8459 | 729.8 | 0.075021
|  | 464.9 | 1.8448 | 730 | 0.077771
|  | 465 | 1.8464 | 730.2 | 0.075692
|  | 465.1 | 1.848 | 730.4 | 0.072544
|  | 465.2 | 1.8501 | 730.6 | 0.071807
|  | 465.3 | 1.8506 | 730.8 | 0.071049
|  | 465.4 | 1.8511 | 731 | 0.070254
|  | 465.5 | 1.8534 | 731.2 | 0.067229
|  | 465.6 | 1.8555 | 731.4 | 0.066978
|  | 465.7 | 1.8577 | 731.6 | 0.068141
|  | 465.8 | 1.8598 | 731.8 | 0.070139
|  | 465.9 | 1.854 | 732 | 0.072001
|  | 466 | 1.8534 | 732.2 | 0.07297
|  | 466.1 | 1.8529 | 732.4 | 0.068885
|  | 466.2 | 1.8613 | 732.6 | 0.062192
|  | 466.3 | 1.8629 | 732.8 | 0.058899
|  | 466.4 | 1.8645 | 733 | 0.065928
|  | 466.5 | 1.8582 | 733.2 | 0.061499
|  | 466.6 | 1.8587 | 733.4 | 0.064051
|  | 466.7 | 1.8592 | 733.6 | 0.066657
|  | 466.8 | 1.8644 | 733.8 | 0.065227
|  | 466.9 | 1.8644 | 734 | 0.06279
|  | 467 | 1.8644 | 734.2 | 0.062173
|  | 467.1 | 1.8686 | 734.4 | 0.062222
|  | 467.2 | 1.8724 | 734.6 | 0.060146
|  | 467.3 | 1.8762 | 734.8 | 0.058286
|  | 467.4 | 1.8749 | 735 | 0.058139
|  | 467.5 | 1.876 | 735.2 | 0.055968
|  | 467.6 | 1.877 | 735.4 | 0.055735
|  | 467.7 | 1.878 | 735.6 | 0.059839
|  | 467.8 | 1.8785 | 735.8 | 0.065424
|  | 467.9 | 1.879 | 736 | 0.066213
|  | 468 | 1.879 | 736.2 | 0.063297
|  | 468.1 | 1.8796 | 736.4 | 0.060881
|  | 468.2 | 1.8801 | 736.6 | 0.063123
|  | 468.3 | 1.88 | 736.8 | 0.066629
|  | 468.4 | 1.88 | 737 | 0.066478
|  | 468.5 | 1.88 | 737.2 | 0.063105
|  | 468.6 | 1.8834 | 737.4 | 0.060734
|  | 468.7 | 1.885 | 737.6 | 0.062128
|  | 468.8 | 1.8866 | 737.8 | 0.063804
|  | 468.9 | 1.8882 | 738 | 0.068703
|  | 469 | 1.8833 | 738.2 | 0.090425
|  | 469.1 | 1.8833 | 738.4 | 0.10434
|  | 469.2 | 1.8833 | 738.6 | 0.099334
|  | 469.3 | 1.8835 | 738.8 | 0.080166
|  | 469.4 | 1.8824 | 739 | 0.061981
|  | 469.5 | 1.8814 | 739.2 | 0.059474
|  | 469.6 | 1.8857 | 739.4 | 0.058008
|  | 469.7 | 1.8873 | 739.6 | 0.056959
|  | 469.8 | 1.8889 | 739.8 | 0.050437
|  | 469.9 | 1.8885 | 740 | 0.046221
|  | 470 | 1.8896 | 740.2 | 0.048575
|  | 470.1 | 1.8907 | 740.4 | 0.049964
|  | 470.2 | 1.8914 | 740.6 | 0.053384
|  | 470.3 | 1.892 | 740.8 | 0.059127
|  | 470.4 | 1.8925 | 741 | 0.063013
|  | 470.5 | 1.887 | 741.2 | 0.058066
|  | 470.6 | 1.8854 | 741.4 | 0.057557
|  | 470.7 | 1.8837 | 741.6 | 0.056664
|  | 470.8 | 1.8916 | 741.8 | 0.054204
|  | 470.9 | 1.8926 | 742 | 0.051431
|  | 471 | 1.8937 | 742.2 | 0.055719
|  | 471.1 | 1.8918 | 742.4 | 0.059195
|  | 471.2 | 1.894 | 742.6 | 0.057905
|  | 471.3 | 1.8961 | 742.8 | 0.058086
|  | 471.4 | 1.8978 | 743 | 0.058428
|  | 471.5 | 1.8978 | 743.2 | 0.056006
|  | 471.6 | 1.8978 | 743.4 | 0.057782
|  | 471.7 | 1.8979 | 743.6 | 0.056253
|  | 471.8 | 1.899 | 743.8 | 0.047961
|  | 471.9 | 1.9001 | 744 | 0.041538
|  | 472 | 1.9034 | 744.2 | 0.040481
|  | 472.1 | 1.9051 | 744.4 | 0.043496
|  | 472.2 | 1.9067 | 744.6 | 0.047168
|  | 472.3 | 1.9083 | 744.8 | 0.05173
|  | 472.4 | 1.9082 | 745 | 0.045425
|  | 472.5 | 1.9099 | 745.2 | 0.038478
|  | 472.6 | 1.9115 | 745.4 | 0.043073
|  | 472.7 | 1.91 | 745.6 | 0.048665
|  | 472.8 | 1.91 | 745.8 | 0.053884
|  | 472.9 | 1.91 | 746 | 0.048412
|  | 473 | 1.9054 | 746.2 | 0.044413
|  | 473.1 | 1.9037 | 746.4 | 0.046491
|  | 473.2 | 1.9021 | 746.6 | 0.045656
|  | 473.3 | 1.9037 | 746.8 | 0.048893
|  | 473.4 | 1.9026 | 747 | 0.048429
|  | 473.5 | 1.9015 | 747.2 | 0.045143
|  | 473.6 | 1.9055 | 747.4 | 0.042629
|  | 473.7 | 1.9071 | 747.6 | 0.049369
|  | 473.8 | 1.9087 | 747.8 | 0.053184
|  | 473.9 | 1.9084 | 748 | 0.049221
|  | 474 | 1.9095 | 748.2 | 0.042685
|  | 474.1 | 1.9105 | 748.4 | 0.038435
|  | 474.2 | 1.9127 | 748.6 | 0.036499
|  | 474.3 | 1.9138 | 748.8 | 0.038155
|  | 474.4 | 1.9149 | 749 | 0.04095
|  | 474.5 | 1.9135 | 749.2 | 0.041738
|  | 474.6 | 1.9141 | 749.4 | 0.04389
|  | 474.7 | 1.9146 | 749.6 | 0.046582
|  | 474.8 | 1.9156 | 749.8 | 0.046465
|  | 474.9 | 1.9156 | 750 | 0.043066
|  | 475 | 1.9156 | 750.2 | 0.064372
|  | 475.1 | 1.9169 | 750.4 | 0.085006
|  | 475.2 | 1.918 | 750.6 | 0.090877
|  | 475.3 | 1.9191 | 750.8 | 0.070566
|  | 475.4 | 1.9226 | 751 | 0.057159
|  | 475.5 | 1.9248 | 751.2 | 0.057679
|  | 475.6 | 1.9269 | 751.4 | 0.075215
|  | 475.7 | 1.9233 | 751.6 | 0.076701
|  | 475.8 | 1.9233 | 751.8 | 0.067208
|  | 475.9 | 1.9233 | 752 | 0.053884
|  | 476 | 1.9222 | 752.2 | 0.044754
|  | 476.1 | 1.92 | 752.4 | 0.039617
|  | 476.2 | 1.9179 | 752.6 | 0.037269
|  | 476.3 | 1.9157 | 752.8 | 0.036001
|  | 476.4 | 1.9189 | 753 | 0.033977
|  | 476.5 | 1.9189 | 753.2 | 0.031352
|  | 476.6 | 1.9189 | 753.4 | 0.02862
|  | 476.7 | 1.9225 | 753.6 | 0.027116
|  | 476.8 | 1.9241 | 753.8 | 0.027707
|  | 476.9 | 1.9258 | 754 | 0.027465
|  | 477 | 1.9222 | 754.2 | 0.02527
|  | 477.1 | 1.9222 | 754.4 | 0.028632
|  | 477.2 | 1.9222 | 754.6 | 0.028748
|  | 477.3 | 1.9259 | 754.8 | 0.027032
|  | 477.4 | 1.9265 | 755 | 0.026225
|  | 477.5 | 1.927 | 755.2 | 0.030491
|  | 477.6 | 1.9233 | 755.4 | 0.028482
|  | 477.7 | 1.9233 | 755.6 | 0.028528
|  | 477.8 | 1.9233 | 755.8 | 0.028607
|  | 477.9 | 1.9295 | 756 | 0.026821
|  | 478 | 1.9306 | 756.2 | 0.025429
|  | 478.1 | 1.9317 | 756.4 | 0.028491
|  | 478.2 | 1.9297 | 756.6 | 0.023525
|  | 478.3 | 1.9313 | 756.8 | 0.024852
|  | 478.4 | 1.933 | 757 | 0.029576
|  | 478.5 | 1.9318 | 757.2 | 0.03177
|  | 478.6 | 1.9307 | 757.4 | 0.028599
|  | 478.7 | 1.9296 | 757.6 | 0.023306
|  | 478.8 | 1.9326 | 757.8 | 0.020604
|  | 478.9 | 1.9337 | 758 | 0.026305
|  | 479 | 1.9348 | 758.2 | 0.023093
|  | 479.1 | 1.9322 | 758.4 | 0.017907
|  | 479.2 | 1.9322 | 758.6 | 0.018555
|  | 479.3 | 1.9322 | 758.8 | 0.016727
|  | 479.4 | 1.932 | 759 | 0.019287
|  | 479.5 | 1.9309 | 759.2 | 0.020823
|  | 479.6 | 1.9298 | 759.4 | 0.020163
|  | 479.7 | 1.9276 | 759.6 | 0.018232
|  | 479.8 | 1.9265 | 759.8 | 0.022366
|  | 479.9 | 1.9254 | 760 | 0.02589
|  | 480 | 1.9266 | 760.2 | 0.028711
|  | 480.1 | 1.9261 | 760.4 | 0.027025
|  | 480.2 | 1.9255 | 760.6 | 0.023077
|  | 480.3 | 1.9301 | 760.8 | 0.019212
|  | 480.4 | 1.9328 | 761 | 0.01874
|  | 480.5 | 1.9355 | 761.2 | 0.017779
|  | 480.6 | 1.9382 | 761.4 | 0.018519
|  | 480.7 | 1.9344 | 761.6 | 0.018148
|  | 480.8 | 1.9355 | 761.8 | 0.018148
|  | 480.9 | 1.9365 | 762 | 0.018026
|  | 481 | 1.9378 | 762.2 | 0.016169
|  | 481.1 | 1.9378 | 762.4 | 0.015902
|  | 481.2 | 1.9378 | 762.6 | 0.017275
|  | 481.3 | 1.934 | 762.8 | 0.020372
|  | 481.4 | 1.9334 | 763 | 0.023692
|  | 481.5 | 1.9329 | 763.2 | 0.047354
|  | 481.6 | 1.9367 | 763.4 | 0.11203
|  | 481.7 | 1.9367 | 763.6 | 0.16366
|  | 481.8 | 1.9367 | 763.8 | 0.14973
|  | 481.9 | 1.9393 | 764 | 0.094763
|  | 482 | 1.9415 | 764.2 | 0.049386
|  | 482.1 | 1.9437 | 764.4 | 0.034797
|  | 482.2 | 1.9411 | 764.6 | 0.029594
|  | 482.3 | 1.9411 | 764.8 | 0.027461
|  | 482.4 | 1.9411 | 765 | 0.026095
|  | 482.5 | 1.9436 | 765.2 | 0.019093
|  | 482.6 | 1.9442 | 765.4 | 0.013004
|  | 482.7 | 1.9447 | 765.6 | 0.013912
|  | 482.8 | 1.9436 | 765.8 | 0.015992
|  | 482.9 | 1.9442 | 766 | 0.017132
|  | 483 | 1.9447 | 766.2 | 0.0214
|  | 483.1 | 1.9483 | 766.4 | 0.019474
|  | 483.2 | 1.9494 | 766.6 | 0.014008
|  | 483.3 | 1.9505 | 766.8 | 0.011926
|  | 483.4 | 1.952 | 767 | 0.017586
|  | 483.5 | 1.9542 | 767.2 | 0.016644
|  | 483.6 | 1.9563 | 767.4 | 0.016247
|  | 483.7 | 1.9557 | 767.6 | 0.015416
|  | 483.8 | 1.9563 | 767.8 | 0.014444
|  | 483.9 | 1.9568 | 768 | 0.016794
|  | 484 | 1.9567 | 768.2 | 0.019132
|  | 484.1 | 1.9567 | 768.4 | 0.019072
|  | 484.2 | 1.9567 | 768.6 | 0.01619
|  | 484.3 | 1.9579 | 768.8 | 0.013378
|  | 484.4 | 1.9584 | 769 | 0.010607
|  | 484.5 | 1.959 | 769.2 | 0.0098669
|  | 484.6 | 1.9602 | 769.4 | 0.01463
|  | 484.7 | 1.9613 | 769.6 | 0.016939
|  | 484.8 | 1.9624 | 769.8 | 0.015221
|  | 484.9 | 1.9646 | 770 | 0.013206
|  | 485 | 1.9663 | 770.2 | 0.015486
|  | 485.1 | 1.9679 | 770.4 | 0.011627
|  | 485.2 | 1.9713 | 770.6 | 0.0033472
|  | 485.3 | 1.974 | 770.8 | 0.00056944
|  | 485.4 | 1.9767 | 771 | 0.005213
|  | 485.5 | 1.9778 | 771.2 | 0.0094444
|  | 485.6 | 1.98 | 771.4 | 0.0053009
|  | 485.7 | 1.9822 | 771.6 | 0.0017315
|  | 485.8 | 1.9843 | 771.8 | 0.0049167
|  | 485.9 | 1.9811 | 772 | 0.013361
|  | 486 | 1.9811 | 772.2 | 0.03313
|  | 486.1 | 1.9811 | 772.4 | 0.054798
|  | 486.2 | 1.9806 | 772.6 | 0.065139
|  | 486.3 | 1.9801 | 772.8 | 0.052961
|  | 486.4 | 1.9795 | 773 | 0.041528
|  | 486.5 | 1.9816 | 773.2 | 0.026416
|  | 486.6 | 1.9821 | 773.4 | 0.01918
|  | 486.7 | 1.9827 | 773.6 | 0.016841
|  | 486.8 | 1.9858 | 773.8 | 0.019966
|  | 486.9 | 1.9874 | 774 | 0.010988
|  | 487 | 1.989 | 774.2 | 0.0015374
|  | 487.1 | 1.984 | 774.4 | 0.0031722
|  | 487.2 | 1.9835 | 774.6 | 0.0076585
|  | 487.3 | 1.9829 | 774.8 | 0.0076258
|  | 487.4 | 1.9822 | 775 | 0.0016252
|  | 487.5 | 1.9805 | 775.2 | 9.3755e-05
|  | 487.6 | 1.9789 | 775.4 | 0.0047536
|  | 487.7 | 1.98 | 775.6 | 0.0084403
|  | 487.8 | 1.98 | 775.8 | 0.0066915
|  | 487.9 | 1.98 | 776 | 0.0017023
|  | 488 | 1.98 | 776.2 | -0.00070975
|  | 488.1 | 1.98 | 776.4 | 0.0022408
|  | 488.2 | 1.98 | 776.6 | 0.0019898
|  | 488.3 | 1.9715 | 776.8 | 0.0079972
|  | 488.4 | 1.9682 | 777 | 0.0063993
|  | 488.5 | 1.965 | 777.2 | 0.0016811
|  | 488.6 | 1.9692 | 777.4 | -0.00066985
|  | 488.7 | 1.9675 | 777.6 | 0.0072786
|  | 488.8 | 1.9659 | 777.8 | 0.01241
|  | 488.9 | 1.9647 | 778 | 0.0099588
|  | 489 | 1.9652 | 778.2 | 0.0049517
|  | 489.1 | 1.9658 | 778.4 | 0.0044117
|  | 489.2 | 1.9691 | 778.6 | 0.0046701
|  | 489.3 | 1.9697 | 778.8 | 0.0021742
|  | 489.4 | 1.9702 | 779 | 0.0013893
|  | 489.5 | 1.9693 | 779.2 | 0.0036508
|  | 489.6 | 1.9704 | 779.4 | 0.0044344
|  | 489.7 | 1.9715 | 779.6 | 0.0019204
|  | 489.8 | 1.9696 | 779.8 | -0.00059807
|  | 489.9 | 1.9685 | 780 | 0.00085512
|  | 490 | 1.9675 | 780.2 | -0.0011734
|  | 490.1 | 1.9613 | 780.4 | -0.0091083
|  | 490.2 | 1.958 | 780.6 | -0.0094553
|  | 490.3 | 1.9547 | 780.8 | 0.00060146
|  | 490.4 | 1.9611 | 781 | 0.0072168
|  | 490.5 | 1.9611 | 781.2 | -0.0013655
|  | 490.6 | 1.9611 | 781.4 | 0.0014073
|  | 490.7 | 1.9702 | 781.6 | 0.0070847
|  | 490.8 | 1.9768 | 781.8 | 0.013282
|  | 490.9 | 1.9834 | 782 | 0.010759
|  | 491 | 1.9828 | 782.2 | 0.0077537
|  | 491.1 | 1.9867 | 782.4 | 0.0035748
|  | 491.2 | 1.9905 | 782.6 | -0.0024551
|  | 491.3 | 1.9971 | 782.8 | -0.0080137
|  | 491.4 | 2.0004 | 783 | -0.0061801
|  | 491.5 | 2.0036 | 783.2 | -0.0053775
|  | 491.6 | 2.0117 | 783.4 | -0.0022957
|  | 491.7 | 2.0188 | 783.6 | -5.1571e-05
|  | 491.8 | 2.0259 | 783.8 | -0.0051691
|  | 491.9 | 2.0053 | 784 | -0.010977
|  | 492 | 1.9993 | 784.2 | -0.00096067
|  | 492.1 | 1.9933 | 784.4 | 0.00023678
|  | 492.2 | 1.9872 | 784.6 | -0.0045054
|  | 492.3 | 1.9685 | 784.8 | -0.010783
|  | 492.4 | 1.9537 | 785 | -0.0067979
|  | 492.5 | 1.9389 | 785.2 | -0.0069058
|  | 492.6 | 1.947 | 785.4 | -0.010394
|  | 492.7 | 1.9415 | 785.6 | -0.0087107
|  | 492.8 | 1.936 | 785.8 | -0.01003
|  | 492.9 | 1.9309 | 786 | -0.013395
|  | 493 | 1.9271 | 786.2 | -0.017326
|  | 493.1 | 1.9232 | 786.4 | -0.017909
|  | 493.2 | 1.9285 | 786.6 | -0.015859
|  | 493.3 | 1.9269 | 786.8 | -0.014098
|  | 493.4 | 1.9253 | 787 | -0.0092504
|  | 493.5 | 1.9297 | 787.2 | -0.0054694
|  | 493.6 | 1.9318 | 787.4 | -0.007146
|  | 493.7 | 1.934 | 787.6 | -0.01685
|  | 493.8 | 1.9296 | 787.8 | -0.014006
|  | 493.9 | 1.929 | 788 | -0.0073826
|  | 494 | 1.9285 | 788.2 | -0.011486
|  | 494.1 | 1.9234 | 788.4 | -0.011268
|  | 494.2 | 1.9207 | 788.6 | -0.01092
|  | 494.3 | 1.9179 | 788.8 | -0.01306
|  | 494.4 | 1.9203 | 789 | -0.01841
|  | 494.5 | 1.9192 | 789.2 | -0.015807
|  | 494.6 | 1.9181 | 789.4 | -0.015207
|  | 494.7 | 1.9102 | 789.6 | -0.017913
|  | 494.8 | 1.9075 | 789.8 | -0.017473
|  | 494.9 | 1.9047 | 790 | -0.0085467
|  | 495 | 1.9118 | 790.2 | 0.0021286
|  | 495.1 | 1.9113 | 790.4 | -0.008506
|  | 495.2 | 1.9108 | 790.6 | -0.014548
|  | 495.3 | 1.9129 | 790.8 | -0.018911
|  | 495.4 | 1.9157 | 791 | -0.019895
|  | 495.5 | 1.9184 | 791.2 | -0.017226
|  | 495.6 | 1.9094 | 791.4 | -0.012205
|  | 495.7 | 1.9067 | 791.6 | -0.012174
|  | 495.8 | 1.9039 | 791.8 | -0.013133
|  | 495.9 | 1.9011 | 792 | -0.00878
|  | 496 | 1.8972 | 792.2 | -0.0063821
|  | 496.1 | 1.8934 | 792.4 | -0.0081753
|  | 496.2 | 1.8992 | 792.6 | -0.011492
|  | 496.3 | 1.8997 | 792.8 | -0.013301
|  | 496.4 | 1.9003 | 793 | -0.013782
|  | 496.5 | 1.8953 | 793.2 | -0.015388
|  | 496.6 | 1.8947 | 793.4 | -0.014995
|  | 496.7 | 1.8942 | 793.6 | -0.01136
|  | 496.8 | 1.8947 | 793.8 | -0.011435
|  | 496.9 | 1.8931 | 794 | -0.013504
|  | 497 | 1.8914 | 794.2 | -0.015701
|  | 497.1 | 1.8952 | 794.4 | -0.019703
|  | 497.2 | 1.8969 | 794.6 | -0.022617
|  | 497.3 | 1.8985 | 794.8 | -0.00011899
|  | 497.4 | 1.8901 | 795 | 0.026438
|  | 497.5 | 1.8879 | 795.2 | 0.032945
|  | 497.6 | 1.8857 | 795.4 | 0.023767
|  | 497.7 | 1.8866 | 795.6 | 0.0063369
|  | 497.8 | 1.8839 | 795.8 | -0.0052077
|  | 497.9 | 1.8811 | 796 | -0.0082837
|  | 498 | 1.8851 | 796.2 | -0.0092866
|  | 498.1 | 1.8868 | 796.4 | -0.018799
|  | 498.2 | 1.8884 | 796.6 | -0.024687
|  | 498.3 | 1.879 | 796.8 | -0.024724
|  | 498.4 | 1.8762 | 797 | -0.020729
|  | 498.5 | 1.8734 | 797.2 | -0.019613
|  | 498.6 | 1.8744 | 797.4 | -0.018127
|  | 498.7 | 1.8711 | 797.6 | -0.022372
|  | 498.8 | 1.8678 | 797.8 | -0.02155
|  | 498.9 | 1.8691 | 798 | -0.015645
|  | 499 | 1.8696 | 798.2 | -0.0096944
|  | 499.1 | 1.8702 | 798.4 | -0.01265
|  | 499.2 | 1.8636 | 798.6 | -0.019103
|  | 499.3 | 1.8608 | 798.8 | -0.025855
|  | 499.4 | 1.8581 | 799 | -0.022939
|  | 499.5 | 1.8632 | 799.2 | -0.019088
|  | 499.6 | 1.8626 | 799.4 | -0.019069
|  | 499.7 | 1.8621 | 799.6 | -0.023062
|  | 499.8 | 1.8603 | 799.8 | -0.023093
|  | 499.9 | 1.8614 | 800 | -0.031572
|  | 500 | 1.8625 | 800.2 | -0.03067
|  | 500.1 | 1.8607 | 800.4 | -0.019901
|  | 500.2 | 1.859 | 800.6 | -0.0089209
|  | 500.3 | 1.8574 | 800.8 | -0.00045673
|  | 500.4 | 1.8552 | 801 | 0.0046524
|  | 500.5 | 1.8535 | 801.2 | -0.00056103
|  | 500.6 | 1.8519 | 801.4 | -0.00056845
|  | 500.7 | 1.8557 | 801.6 | 0.022784
|  | 500.8 | 1.8562 | 801.8 | 0.04927
|  | 500.9 | 1.8568 | 802 | 0.049897
|  | 501 | 1.8497 | 802.2 | 0.033029
|  | 501.1 | 1.848 | 802.4 | 0.0069193
|  | 501.2 | 1.8464 | 802.6 | -0.0031418
|  | 501.3 | 1.8425 | 802.8 | 0.00058894
|  | 501.4 | 1.8375 | 803 | -0.0034979
|  | 501.5 | 1.8325 | 803.2 | -0.011348
|  | 501.6 | 1.8343 | 803.4 | -0.010449
|  | 501.7 | 1.8332 | 803.6 | -0.0068987
|  | 501.8 | 1.8321 | 803.8 | -0.0082038
|  | 501.9 | 1.8323 | 804 | -0.0022826
|  | 502 | 1.8329 | 804.2 | -0.00063901
|  | 502.1 | 1.8334 | 804.4 | 5.0866e-05
|  | 502.2 | 1.8333 | 804.6 | -0.00076298
|  | 502.3 | 1.8333 | 804.8 | -0.0034072
|  | 502.4 | 1.8333 | 805 | -0.0033107
|  | 502.5 | 1.8298 | 805.2 | -0.00099977
|  | 502.6 | 1.8282 | 805.4 | -0.00085989
|  | 502.7 | 1.8265 | 805.6 | -0.0053921
|  | 502.8 | 1.8288 | 805.8 | -0.008912
|  | 502.9 | 1.8283 | 806 | -0.0050643
|  | 503 | 1.8277 | 806.2 | 0.0019256
|  | 503.1 | 1.8244 | 806.4 | -0.010578
|  | 503.2 | 1.8238 | 806.6 | -0.033556
|  | 503.3 | 1.8233 | 806.8 | -0.047577
|  | 503.4 | 1.8255 | 807 | -0.041356
|  | 503.5 | 1.825 | 807.2 | -0.031777
|  | 503.6 | 1.8244 | 807.4 | -0.030532
|  | 503.7 | 1.8188 | 807.6 | -0.035969
|  | 503.8 | 1.8171
|  | 503.9 | 1.8155
|  | 504 | 1.8188
|  | 504.1 | 1.8177
|  | 504.2 | 1.8166
|  | 504.3 | 1.8099
|  | 504.4 | 1.8083
|  | 504.5 | 1.8066
|  | 504.6 | 1.8111
|  | 504.7 | 1.81
|  | 504.8 | 1.8089
|  | 504.9 | 1.8033
|  | 505 | 1.8028
|  | 505.1 | 1.8022
|  | 505.2 | 1.8033
|  | 505.3 | 1.8011
|  | 505.4 | 1.7989
|  | 505.5 | 1.8
|  | 505.6 | 1.8011
|  | 505.7 | 1.8022
|  | 505.8 | 1.8033
|  | 505.9 | 1.7961
|  | 506 | 1.7944
|  | 506.1 | 1.7928
|  | 506.2 | 1.7956
|  | 506.3 | 1.7945
|  | 506.4 | 1.7933
|  | 506.5 | 1.7872
|  | 506.6 | 1.7856
|  | 506.7 | 1.7839
|  | 506.8 | 1.7878
|  | 506.9 | 1.7867
|  | 507 | 1.7856
|  | 507.1 | 1.785
|  | 507.2 | 1.7855
|  | 507.3 | 1.7861
|  | 507.4 | 1.7828
|  | 507.5 | 1.7811
|  | 507.6 | 1.7795
|  | 507.7 | 1.7806
|  | 507.8 | 1.78
|  | 507.9 | 1.7795
|  | 508 | 1.78
|  | 508.1 | 1.78
|  | 508.2 | 1.78
|  | 508.3 | 1.7849
|  | 508.4 | 1.7866
|  | 508.5 | 1.7883
|  | 508.6 | 1.7784
|  | 508.7 | 1.7767
|  | 508.8 | 1.7751
|  | 508.9 | 1.7751
|  | 509 | 1.7723
|  | 509.1 | 1.7695
|  | 509.2 | 1.7711
|  | 509.3 | 1.7711
|  | 509.4 | 1.7711
|  | 509.5 | 1.7722
|  | 509.6 | 1.7722
|  | 509.7 | 1.7722
|  | 509.8 | 1.7695
|  | 509.9 | 1.7689
|  | 510 | 1.7683
|  | 510.1 | 1.7645
|  | 510.2 | 1.7623
|  | 510.3 | 1.76
|  | 510.4 | 1.7611
|  | 510.5 | 1.76
|  | 510.6 | 1.7589
|  | 510.7 | 1.7617
|  | 510.8 | 1.7622
|  | 510.9 | 1.7628
|  | 511 | 1.7578
|  | 511.1 | 1.7567
|  | 511.2 | 1.7555
|  | 511.3 | 1.7483
|  | 511.4 | 1.7444
|  | 511.5 | 1.7467
|  | 511.6 | 1.7455
|  | 511.7 | 1.7444
|  | 511.8 | 1.7456
|  | 511.9 | 1.7473
|  | 512 | 1.7489
|  | 512.1 | 1.7455
|  | 512.2 | 1.7444
|  | 512.3 | 1.7433
|  | 512.4 | 1.7467
|  | 512.5 | 1.7467
|  | 512.6 | 1.7467
|  | 512.7 | 1.7467
|  | 512.8 | 1.7484
|  | 512.9 | 1.7501
|  | 513 | 1.7432
|  | 513.1 | 1.7398
|  | 513.2 | 1.7365
|  | 513.3 | 1.7353
|  | 513.4 | 1.7314
|  | 513.5 | 1.7275
|  | 513.6 | 1.7323
|  | 513.7 | 1.734
|  | 513.8 | 1.7357
|  | 513.9 | 1.7323
|  | 514 | 1.7328
|  | 514.1 | 1.7334
|  | 514.2 | 1.7297
|  | 514.3 | 1.7263
|  | 514.4 | 1.723
|  | 514.5 | 1.7255
|  | 514.6 | 1.7249
|  | 514.7 | 1.7244
|  | 514.8 | 1.7234
|  | 514.9 | 1.724
|  | 515 | 1.7245
|  | 515.1 | 1.7267
|  | 515.2 | 1.7273
|  | 515.3 | 1.7279
|  | 515.4 | 1.7244
|  | 515.5 | 1.7238
|  | 515.6 | 1.7232
|  | 515.7 | 1.723
|  | 515.8 | 1.7214
|  | 515.9 | 1.7197
|  | 516 | 1.7212
|  | 516.1 | 1.7218
|  | 516.2 | 1.7223
|  | 516.3 | 1.7211
|  | 516.4 | 1.7211
|  | 516.5 | 1.7211
|  | 516.6 | 1.721
|  | 516.7 | 1.7204
|  | 516.8 | 1.7199
|  | 516.9 | 1.7162
|  | 517 | 1.7146
|  | 517.1 | 1.7129
|  | 517.2 | 1.7129
|  | 517.3 | 1.7112
|  | 517.4 | 1.7095
|  | 517.5 | 1.71
|  | 517.6 | 1.71
|  | 517.7 | 1.71
|  | 517.8 | 1.71
|  | 517.9 | 1.71
|  | 518 | 1.71
|  | 518.1 | 1.71
|  | 518.2 | 1.71
|  | 518.3 | 1.71
|  | 518.4 | 1.71
|  | 518.5 | 1.71
|  | 518.6 | 1.71
|  | 518.7 | 1.706
|  | 518.8 | 1.7044
|  | 518.9 | 1.7027
|  | 519 | 1.7027
|  | 519.1 | 1.701
|  | 519.2 | 1.6993
|  | 519.3 | 1.704
|  | 519.4 | 1.7057
|  | 519.5 | 1.7074
|  | 519.6 | 1.7074
|  | 519.7 | 1.7091
|  | 519.8 | 1.7108
|  | 519.9 | 1.7059
|  | 520 | 1.7042
|  | 520.1 | 1.7025
|  | 520.2 | 1.7053
|  | 520.3 | 1.7047
|  | 520.4 | 1.7041
|  | 520.5 | 1.7022
|  | 520.6 | 1.7022
|  | 520.7 | 1.7022
|  | 520.8 | 1.6973
|  | 520.9 | 1.6945
|  | 521 | 1.6917
|  | 521.1 | 1.6923
|  | 521.2 | 1.6907
|  | 521.3 | 1.689
|  | 521.4 | 1.6915
|  | 521.5 | 1.692
|  | 521.6 | 1.6926
|  | 521.7 | 1.6955
|  | 521.8 | 1.6972
|  | 521.9 | 1.6989
|  | 522 | 1.7015
|  | 522.1 | 1.7038
|  | 522.2 | 1.706
|  | 522.3 | 1.6977
|  | 522.4 | 1.696
|  | 522.5 | 1.6943
|  | 522.6 | 1.6935
|  | 522.7 | 1.6907
|  | 522.8 | 1.6879
|  | 522.9 | 1.6854
|  | 523 | 1.6837
|  | 523.1 | 1.682
|  | 523.2 | 1.6898
|  | 523.3 | 1.6909
|  | 523.4 | 1.6921
|  | 523.5 | 1.6856
|  | 523.6 | 1.6856
|  | 523.7 | 1.6856
|  | 523.8 | 1.6927
|  | 523.9 | 1.6933
|  | 524 | 1.6938
|  | 524.1 | 1.6915
|  | 524.2 | 1.6932
|  | 524.3 | 1.6949
|  | 524.4 | 1.695
|  | 524.5 | 1.6945
|  | 524.6 | 1.6939
|  | 524.7 | 1.6939
|  | 524.8 | 1.6944
|  | 524.9 | 1.6933
|  | 525 | 1.6922
|  | 525.1 | 1.6911
|  | 525.2 | 1.6945
|  | 525.3 | 1.6956
|  | 525.4 | 1.6968
|  | 525.5 | 1.6968
|  | 525.6 | 1.6985
|  | 525.7 | 1.7002
|  | 525.8 | 1.7
|  | 525.9 | 1.7
|  | 526 | 1.7
|  | 526.1 | 1.6964
|  | 526.2 | 1.6947
|  | 526.3 | 1.693
|  | 526.4 | 1.6967
|  | 526.5 | 1.6967
|  | 526.6 | 1.6967
|  | 526.7 | 1.6933
|  | 526.8 | 1.6933
|  | 526.9 | 1.6933
|  | 527 | 1.6954
|  | 527.1 | 1.6948
|  | 527.2 | 1.6943
|  | 527.3 | 1.6988
|  | 527.4 | 1.7016
|  | 527.5 | 1.7044
|  | 527.6 | 1.6973
|  | 527.7 | 1.6962
|  | 527.8 | 1.6951
|  | 527.9 | 1.6984
|  | 528 | 1.6973
|  | 528.1 | 1.6961
|  | 528.2 | 1.6975
|  | 528.3 | 1.6992
|  | 528.4 | 1.7009
|  | 528.5 | 1.7014
|  | 528.6 | 1.702
|  | 528.7 | 1.7025
|  | 528.8 | 1.7083
|  | 528.9 | 1.7111
|  | 529 | 1.714
|  | 529.1 | 1.7063
|  | 529.2 | 1.7057
|  | 529.3 | 1.7052
|  | 529.4 | 1.7037
|  | 529.5 | 1.7008
|  | 529.6 | 1.698
|  | 529.7 | 1.7091
|  | 529.8 | 1.7125
|  | 529.9 | 1.7159
|  | 530 | 1.7082
|  | 530.1 | 1.7088
|  | 530.2 | 1.7093
|  | 530.3 | 1.7144
|  | 530.4 | 1.7144
|  | 530.5 | 1.7144
|  | 530.6 | 1.7121
|  | 530.7 | 1.7132
|  | 530.8 | 1.7144
|  | 530.9 | 1.715
|  | 531 | 1.7145
|  | 531.1 | 1.7139
|  | 531.2 | 1.7122
|  | 531.3 | 1.7122
|  | 531.4 | 1.7133
|  | 531.5 | 1.7128
|  | 531.6 | 1.7122
|  | 531.7 | 1.7169
|  | 531.8 | 1.7197
|  | 531.9 | 1.7226
|  | 532 | 1.7237
|  | 532.1 | 1.7265
|  | 532.2 | 1.7294
|  | 532.3 | 1.7253
|  | 532.4 | 1.7242
|  | 532.5 | 1.7231
|  | 532.6 | 1.7242
|  | 532.7 | 1.723
|  | 532.8 | 1.7219
|  | 532.9 | 1.7288
|  | 533 | 1.7322
|  | 533.1 | 1.7356
|  | 533.2 | 1.7315
|  | 533.3 | 1.7327
|  | 533.4 | 1.7338
|  | 533.5 | 1.7407
|  | 533.6 | 1.7424
|  | 533.7 | 1.7441
|  | 533.8 | 1.7394
|  | 533.9 | 1.7406
|  | 534 | 1.7417
|  | 534.1 | 1.7413
|  | 534.2 | 1.7396
|  | 534.3 | 1.7379
|  | 534.4 | 1.7436
|  | 534.5 | 1.7459
|  | 534.6 | 1.7481
|  | 534.7 | 1.7433
|  | 534.8 | 1.7433
|  | 534.9 | 1.7433
|  | 535 | 1.7467
|  | 535.1 | 1.7467
|  | 535.2 | 1.7467
|  | 535.3 | 1.748
|  | 535.4 | 1.7497
|  | 535.5 | 1.7514
|  | 535.6 | 1.7548
|  | 535.7 | 1.7565
|  | 535.8 | 1.7582
|  | 535.9 | 1.7615
|  | 536 | 1.7643
|  | 536.1 | 1.7671
|  | 536.2 | 1.7672
|  | 536.3 | 1.7689
|  | 536.4 | 1.7723
|  | 536.5 | 1.7746
|  | 536.6 | 1.7768
|  | 536.7 | 1.7768
|  | 536.8 | 1.7785
|  | 536.9 | 1.7803
|  | 537 | 1.7872
|  | 537.1 | 1.7906
|  | 537.2 | 1.7941
|  | 537.3 | 1.7966
|  | 537.4 | 1.8012
|  | 537.5 | 1.8057
|  | 537.6 | 1.8099
|  | 537.7 | 1.8133
|  | 537.8 | 1.8167
|  | 537.9 | 1.8164
|  | 538 | 1.8187
|  | 538.1 | 1.8209
|  | 538.2 | 1.8282
|  | 538.3 | 1.8316
|  | 538.4 | 1.835
|  | 538.5 | 1.8407
|  | 538.6 | 1.8464
|  | 538.7 | 1.8521
|  | 538.8 | 1.8488
|  | 538.9 | 1.8505
|  | 539 | 1.8522
|  | 539.1 | 1.8596
|  | 539.2 | 1.8625
|  | 539.3 | 1.8653
|  | 539.4 | 1.8704
|  | 539.5 | 1.8755
|  | 539.6 | 1.8806
|  | 539.7 | 1.8889
|  | 539.8 | 1.8946
|  | 539.9 | 1.9003
|  | 540 | 1.907
|  | 540.1 | 1.9139
|  | 540.2 | 1.9207
|  | 540.3 | 1.9286
|  | 540.4 | 1.9355
|  | 540.5 | 1.9445
|  | 540.6 | 1.9531
|  | 540.7 | 1.9616
|  | 540.8 | 1.9789
|  | 540.9 | 1.992
|  | 541 | 2.0051
|  | 541.1 | 2.0191
|  | 541.2 | 2.0345
|  | 541.3 | 2.05
|  | 541.4 | 2.0745
|  | 541.5 | 2.0933
|  | 541.6 | 2.1121
|  | 541.7 | 2.1298
|  | 541.8 | 2.1503
|  | 541.9 | 2.1708
|  | 542 | 2.1667
|  | 542.1 | 2.1752
|  | 542.2 | 2.1838
|  | 542.3 | 2.18
|  | 542.4 | 2.18
|  | 542.5 | 2.18
|  | 542.6 | 2.1843
|  | 542.7 | 2.186
|  | 542.8 | 2.1877
|  | 542.9 | 2.1789
|  | 543 | 2.1772
|  | 543.1 | 2.1755
|  | 543.2 | 2.1742
|  | 543.3 | 2.1708
|  | 543.4 | 2.1674
|  | 543.5 | 2.1887
|  | 543.6 | 2.1956
|  | 543.7 | 2.2024
|  | 543.8 | 2.2026
|  | 543.9 | 2.2095
|  | 544 | 2.2163
|  | 544.1 | 2.2199
|  | 544.2 | 2.2234
|  | 544.3 | 2.2314
|  | 544.4 | 2.2388
|  | 544.5 | 2.2462
|  | 544.6 | 2.2873
|  | 544.7 | 2.3113
|  | 544.8 | 2.3353
|  | 544.9 | 2.5782
|  | 545 | 2.7114
|  | 545.1 | 2.8445
|  | 545.2 | 3.925
|  | 545.3 | 4.5238
|  | 545.4 | 5.1227
|  | 545.5 | 6.6317
|  | 545.6 | 7.7791
|  | 545.7 | 8.9265
|  | 545.8 | 9.9618
|  | 545.9 | 11.1348
|  | 546 | 12.3077
|  | 546.1 | 14.1201
|  | 546.2 | 15.4902
|  | 546.3 | 16.8604
|  | 546.4 | 12.674
|  | 546.5 | 12.0207
|  | 546.6 | 11.3674
|  | 546.7 | 8.0806
|  | 546.8 | 5.5692
|  | 546.9 | 3.0577
|  | 547 | 3.8689
|  | 547.1 | 2.7509
|  | 547.2 | 1.6329
|  | 547.3 | 2.4184
|  | 547.4 | 2.2672
|  | 547.5 | 2.116
|  | 547.6 | 2.2099
|  | 547.7 | 2.1749
|  | 547.8 | 2.1656
|  | 547.9 | 2.1518
|  | 548 | 2.1381
|  | 548.1 | 2.1427
|  | 548.2 | 2.1393
|  | 548.3 | 2.1358
|  | 548.4 | 2.1359
|  | 548.5 | 2.137
|  | 548.6 | 2.1382
|  | 548.7 | 2.1349
|  | 548.8 | 2.1332
|  | 548.9 | 2.1315
|  | 549 | 2.1358
|  | 549.1 | 2.1364
|  | 549.2 | 2.137
|  | 549.3 | 2.1333
|  | 549.4 | 2.1333
|  | 549.5 | 2.1333
|  | 549.6 | 2.1381
|  | 549.7 | 2.1387
|  | 549.8 | 2.1393
|  | 549.9 | 2.1483
|  | 550 | 2.1534
|  | 550.1 | 2.1586
|  | 550.2 | 2.1521
|  | 550.3 | 2.1532
|  | 550.4 | 2.1544
|  | 550.5 | 2.1517
|  | 550.6 | 2.15
|  | 550.7 | 2.15
|  | 550.8 | 2.15
|  | 550.9 | 2.15
|  | 551 | 2.1547
|  | 551.1 | 2.157
|  | 551.2 | 2.1593
|  | 551.3 | 2.1582
|  | 551.4 | 2.1599
|  | 551.5 | 2.1616
|  | 551.6 | 2.1648
|  | 551.7 | 2.166
|  | 551.8 | 2.1671
|  | 551.9 | 2.1715
|  | 552 | 2.1749
|  | 552.1 | 2.1784
|  | 552.2 | 2.1818
|  | 552.3 | 2.1853
|  | 552.4 | 2.1887
|  | 552.5 | 2.1833
|  | 552.6 | 2.1833
|  | 552.7 | 2.1833
|  | 552.8 | 2.1821
|  | 552.9 | 2.1803
|  | 553 | 2.1786
|  | 553.1 | 2.18
|  | 553.2 | 2.18
|  | 553.3 | 2.18
|  | 553.4 | 2.185
|  | 553.5 | 2.1867
|  | 553.6 | 2.189
|  | 553.7 | 2.1919
|  | 553.8 | 2.1948
|  | 553.9 | 2.1934
|  | 554 | 2.194
|  | 554.1 | 2.1946
|  | 554.2 | 2.193
|  | 554.3 | 2.1919
|  | 554.4 | 2.1907
|  | 554.5 | 2.1909
|  | 554.6 | 2.1903
|  | 554.7 | 2.1897
|  | 554.8 | 2.1941
|  | 554.9 | 2.1959
|  | 555 | 2.1976
|  | 555.1 | 2.1977
|  | 555.2 | 2.1994
|  | 555.3 | 2.2011
|  | 555.4 | 2.1925
|  | 555.5 | 2.1896
|  | 555.6 | 2.1867
|  | 555.7 | 2.1834
|  | 555.8 | 2.1794
|  | 555.9 | 2.1754
|  | 556 | 2.1795
|  | 556.1 | 2.1789
|  | 556.2 | 2.1789
|  | 556.3 | 2.1795
|  | 556.4 | 2.1801
|  | 556.5 | 2.18
|  | 556.6 | 2.18
|  | 556.7 | 2.18
|  | 556.8 | 2.1812
|  | 556.9 | 2.1818
|  | 557 | 2.1824
|  | 557.1 | 2.1851
|  | 557.2 | 2.1868
|  | 557.3 | 2.1885
|  | 557.4 | 2.1966
|  | 557.5 | 2.2012
|  | 557.6 | 2.2058
|  | 557.7 | 2.201
|  | 557.8 | 2.2027
|  | 557.9 | 2.2045
|  | 558 | 2.2112
|  | 558.1 | 2.2129
|  | 558.2 | 2.2147
|  | 558.3 | 2.2161
|  | 558.4 | 2.2196
|  | 558.5 | 2.2231
|  | 558.6 | 2.2249
|  | 558.7 | 2.2267
|  | 558.8 | 2.2199
|  | 558.9 | 2.2182
|  | 559 | 2.2165
|  | 559.1 | 2.2197
|  | 559.2 | 2.218
|  | 559.3 | 2.2163
|  | 559.4 | 2.221
|  | 559.5 | 2.2244
|  | 559.6 | 2.2279
|  | 559.7 | 2.23
|  | 559.8 | 2.2329
|  | 559.9 | 2.2358
|  | 560 | 2.2407
|  | 560.1 | 2.2442
|  | 560.2 | 2.2476
|  | 560.3 | 2.2426
|  | 560.4 | 2.2432
|  | 560.5 | 2.2437
|  | 560.6 | 2.2467
|  | 560.7 | 2.2467
|  | 560.8 | 2.2467
|  | 560.9 | 2.2531
|  | 561 | 2.2565
|  | 561.1 | 2.26
|  | 561.2 | 2.2635
|  | 561.3 | 2.2669
|  | 561.4 | 2.2681
|  | 561.5 | 2.2704
|  | 561.6 | 2.2727
|  | 561.7 | 2.2711
|  | 561.8 | 2.2711
|  | 561.9 | 2.2711
|  | 562 | 2.2788
|  | 562.1 | 2.2817
|  | 562.2 | 2.2846
|  | 562.3 | 2.2809
|  | 562.4 | 2.2826
|  | 562.5 | 2.2843
|  | 562.6 | 2.2852
|  | 562.7 | 2.2846
|  | 562.8 | 2.284
|  | 562.9 | 2.2822
|  | 563 | 2.2822
|  | 563.1 | 2.2822
|  | 563.2 | 2.2861
|  | 563.3 | 2.2866
|  | 563.4 | 2.2872
|  | 563.5 | 2.2969
|  | 563.6 | 2.3015
|  | 563.7 | 2.2991
|  | 563.8 | 2.3009
|  | 563.9 | 2.3026
|  | 564 | 2.3057
|  | 564.1 | 2.3063
|  | 564.2 | 2.3069
|  | 564.3 | 2.3073
|  | 564.4 | 2.3091
|  | 564.5 | 2.3108
|  | 564.6 | 2.31
|  | 564.7 | 2.31
|  | 564.8 | 2.31
|  | 564.9 | 2.3145
|  | 565 | 2.3162
|  | 565.1 | 2.318
|  | 565.2 | 2.3227
|  | 565.3 | 2.3262
|  | 565.4 | 2.3297
|  | 565.5 | 2.3233
|  | 565.6 | 2.3233
|  | 565.7 | 2.3255
|  | 565.8 | 2.325
|  | 565.9 | 2.3244
|  | 566 | 2.3222
|  | 566.1 | 2.3222
|  | 566.2 | 2.3222
|  | 566.3 | 2.3309
|  | 566.4 | 2.3338
|  | 566.5 | 2.3367
|  | 566.6 | 2.336
|  | 566.7 | 2.3395
|  | 566.8 | 2.343
|  | 566.9 | 2.3451
|  | 567 | 2.3463
|  | 567.1 | 2.3474
|  | 567.2 | 2.3433
|  | 567.3 | 2.3433
|  | 567.4 | 2.3433
|  | 567.5 | 2.3467
|  | 567.6 | 2.3467
|  | 567.7 | 2.3467
|  | 567.8 | 2.3535
|  | 567.9 | 2.357
|  | 568 | 2.3569
|  | 568.1 | 2.3586
|  | 568.2 | 2.3604
|  | 568.3 | 2.3562
|  | 568.4 | 2.3545
|  | 568.5 | 2.3527
|  | 568.6 | 2.3607
|  | 568.7 | 2.3624
|  | 568.8 | 2.3642
|  | 568.9 | 2.3567
|  | 569 | 2.3567
|  | 569.1 | 2.3567
|  | 569.2 | 2.3603
|  | 569.3 | 2.3591
|  | 569.4 | 2.358
|  | 569.5 | 2.3592
|  | 569.6 | 2.361
|  | 569.7 | 2.3627
|  | 569.8 | 2.3656
|  | 569.9 | 2.3667
|  | 570 | 2.3669
|  | 570.1 | 2.3686
|  | 570.2 | 2.3704
|  | 570.3 | 2.3738
|  | 570.4 | 2.3755
|  | 570.5 | 2.3773
|  | 570.6 | 2.3814
|  | 570.7 | 2.3849
|  | 570.8 | 2.3884
|  | 570.9 | 2.3791
|  | 571 | 2.3773
|  | 571.1 | 2.3756
|  | 571.2 | 2.3743
|  | 571.3 | 2.3708
|  | 571.4 | 2.3673
|  | 571.5 | 2.3748
|  | 571.6 | 2.3765
|  | 571.7 | 2.3783
|  | 571.8 | 2.3784
|  | 571.9 | 2.3801
|  | 572 | 2.3836
|  | 572.1 | 2.3853
|  | 572.2 | 2.3871
|  | 572.3 | 2.3872
|  | 572.4 | 2.3889
|  | 572.5 | 2.3907
|  | 572.6 | 2.3941
|  | 572.7 | 2.3958
|  | 572.8 | 2.3976
|  | 572.9 | 2.3933
|  | 573 | 2.3933
|  | 573.1 | 2.3933
|  | 573.2 | 2.3905
|  | 573.3 | 2.3882
|  | 573.4 | 2.3859
|  | 573.5 | 2.3905
|  | 573.6 | 2.3911
|  | 573.7 | 2.3866
|  | 573.8 | 2.3855
|  | 573.9 | 2.3843
|  | 574 | 2.3901
|  | 574.1 | 2.3907
|  | 574.2 | 2.3913
|  | 574.3 | 2.3973
|  | 574.4 | 2.4026
|  | 574.5 | 2.4078
|  | 574.6 | 2.3955
|  | 574.7 | 2.3932
|  | 574.8 | 2.3909
|  | 574.9 | 2.3922
|  | 575 | 2.3887
|  | 575.1 | 2.3852
|  | 575.2 | 2.3934
|  | 575.3 | 2.3964
|  | 575.4 | 2.3993
|  | 575.5 | 2.4017
|  | 575.6 | 2.4046
|  | 575.7 | 2.4198
|  | 575.8 | 2.4286
|  | 575.9 | 2.4374
|  | 576 | 2.541
|  | 576.1 | 2.5966
|  | 576.2 | 2.6523
|  | 576.3 | 2.9389
|  | 576.4 | 3.1111
|  | 576.5 | 3.2833
|  | 576.6 | 3.4894
|  | 576.7 | 3.7048
|  | 576.8 | 3.9202
|  | 576.9 | 4.2473
|  | 577 | 4.4932
|  | 577.1 | 4.7391
|  | 577.2 | 4.3739
|  | 577.3 | 4.4221
|  | 577.4 | 3.9003
|  | 577.5 | 3.5276
|  | 577.6 | 3.155
|  | 577.7 | 3.081
|  | 577.8 | 2.7787
|  | 577.9 | 2.4765
|  | 578 | 2.5853
|  | 578.1 | 2.5607
|  | 578.2 | 2.536
|  | 578.3 | 2.7677
|  | 578.4 | 2.8676
|  | 578.5 | 2.9675
|  | 578.6 | 3.2491
|  | 578.7 | 3.4473
|  | 578.8 | 3.6454
|  | 578.9 | 3.8724
|  | 579 | 4.0884
|  | 579.1 | 4.2465
|  | 579.2 | 4.4305
|  | 579.3 | 4.6146
|  | 579.4 | 3.994
|  | 579.5 | 3.7775
|  | 579.6 | 3.5609
|  | 579.7 | 3.3106
|  | 579.8 | 2.9225
|  | 579.9 | 2.5344
|  | 580 | 2.6319
|  | 580.1 | 2.5238
|  | 580.2 | 2.4157
|  | 580.3 | 2.4752
|  | 580.4 | 2.4617
|  | 580.5 | 2.4481
|  | 580.6 | 2.4465
|  | 580.7 | 2.4424
|  | 580.8 | 2.4388
|  | 580.9 | 2.4323
|  | 581 | 2.4258
|  | 581.1 | 2.4227
|  | 581.2 | 2.4191
|  | 581.3 | 2.4156
|  | 581.4 | 2.416
|  | 581.5 | 2.4143
|  | 581.6 | 2.4125
|  | 581.7 | 2.4048
|  | 581.8 | 2.4013
|  | 581.9 | 2.3978
|  | 582 | 2.4051
|  | 582.1 | 2.4046
|  | 582.2 | 2.404
|  | 582.3 | 2.407
|  | 582.4 | 2.4099
|  | 582.5 | 2.4123
|  | 582.6 | 2.4134
|  | 582.7 | 2.4146
|  | 582.8 | 2.4098
|  | 582.9 | 2.4086
|  | 583 | 2.4074
|  | 583.1 | 2.4122
|  | 583.2 | 2.4122
|  | 583.3 | 2.4122
|  | 583.4 | 2.4092
|  | 583.5 | 2.4098
|  | 583.6 | 2.4104
|  | 583.7 | 2.4072
|  | 583.8 | 2.4048
|  | 583.9 | 2.4024
|  | 584 | 2.399
|  | 584.1 | 2.3966
|  | 584.2 | 2.3988
|  | 584.3 | 2.3982
|  | 584.4 | 2.3977
|  | 584.5 | 2.3905
|  | 584.6 | 2.3881
|  | 584.7 | 2.3858
|  | 584.8 | 2.3797
|  | 584.9 | 2.3738
|  | 585 | 2.3679
|  | 585.1 | 2.3756
|  | 585.2 | 2.3756
|  | 585.3 | 2.3756
|  | 585.4 | 2.368
|  | 585.5 | 2.3668
|  | 585.6 | 2.3656
|  | 585.7 | 2.3599
|  | 585.8 | 2.3551
|  | 585.9 | 2.3566
|  | 586 | 2.356
|  | 586.1 | 2.3554
|  | 586.2 | 2.3544
|  | 586.3 | 2.3544
|  | 586.4 | 2.3544
|  | 586.5 | 2.3612
|  | 586.6 | 2.3636
|  | 586.7 | 2.3659
|  | 586.8 | 2.3756
|  | 586.9 | 2.3821
|  | 587 | 2.3886
|  | 587.1 | 2.4027
|  | 587.2 | 2.4121
|  | 587.3 | 2.4137
|  | 587.4 | 2.4202
|  | 587.5 | 2.4267
|  | 587.6 | 2.4132
|  | 587.7 | 2.4078
|  | 587.8 | 2.4025
|  | 587.9 | 2.3802
|  | 588 | 2.3648
|  | 588.1 | 2.3494
|  | 588.2 | 2.3388
|  | 588.3 | 2.324
|  | 588.4 | 2.3092
|  | 588.5 | 2.3063
|  | 588.6 | 2.2974
|  | 588.7 | 2.2885
|  | 588.8 | 2.2831
|  | 588.9 | 2.276
|  | 589 | 2.2761
|  | 589.1 | 2.2725
|  | 589.2 | 2.269
|  | 589.3 | 2.27
|  | 589.4 | 2.27
|  | 589.5 | 2.27
|  | 589.6 | 2.2614
|  | 589.7 | 2.2579
|  | 589.8 | 2.2543
|  | 589.9 | 2.2572
|  | 590 | 2.2548
|  | 590.1 | 2.2524
|  | 590.2 | 2.2506
|  | 590.3 | 2.25
|  | 590.4 | 2.2439
|  | 590.5 | 2.2392
|  | 590.6 | 2.2344
|  | 590.7 | 2.2385
|  | 590.8 | 2.2373
|  | 590.9 | 2.2362
|  | 591 | 2.2319
|  | 591.1 | 2.2313
|  | 591.2 | 2.2307
|  | 591.3 | 2.2204
|  | 591.4 | 2.2145
|  | 591.5 | 2.2085
|  | 591.6 | 2.22
|  | 591.7 | 2.22
|  | 591.8 | 2.2179
|  | 591.9 | 2.2203
|  | 592 | 2.2227
|  | 592.1 | 2.2207
|  | 592.2 | 2.2189
|  | 592.3 | 2.2171
|  | 592.4 | 2.223
|  | 592.5 | 2.2248
|  | 592.6 | 2.2266
|  | 592.7 | 2.2316
|  | 592.8 | 2.2357
|  | 592.9 | 2.2399
|  | 593 | 2.243
|  | 593.1 | 2.2466
|  | 593.2 | 2.2399
|  | 593.3 | 2.2381
|  | 593.4 | 2.2364
|  | 593.5 | 2.2283
|  | 593.6 | 2.2212
|  | 593.7 | 2.2141
|  | 593.8 | 2.196
|  | 593.9 | 2.1841
|  | 594 | 2.1722
|  | 594.1 | 2.1632
|  | 594.2 | 2.1507
|  | 594.3 | 2.1382
|  | 594.4 | 2.1395
|  | 594.5 | 2.1336
|  | 594.6 | 2.1288
|  | 594.7 | 2.1246
|  | 594.8 | 2.1204
|  | 594.9 | 2.1156
|  | 595 | 2.1108
|  | 595.1 | 2.106
|  | 595.2 | 2.1027
|  | 595.3 | 2.0985
|  | 595.4 | 2.0943
|  | 595.5 | 2.0955
|  | 595.6 | 2.0937
|  | 595.7 | 2.092
|  | 595.8 | 2.0837
|  | 595.9 | 2.0801
|  | 596 | 2.0799
|  | 596.1 | 2.0763
|  | 596.2 | 2.0727
|  | 596.3 | 2.0683
|  | 596.4 | 2.0659
|  | 596.5 | 2.0635
|  | 596.6 | 2.0592
|  | 596.7 | 2.0551
|  | 596.8 | 2.0509
|  | 596.9 | 2.0477
|  | 597 | 2.0441
|  | 597.1 | 2.0405
|  | 597.2 | 2.0369
|  | 597.3 | 2.0334
|  | 597.4 | 2.0286
|  | 597.5 | 2.0245
|  | 597.6 | 2.0203
|  | 597.7 | 2.0256
|  | 597.8 | 2.0256
|  | 597.9 | 2.0256
|  | 598 | 2.0197
|  | 598.1 | 2.0191
|  | 598.2 | 2.0185
|  | 598.3 | 2.0229
|  | 598.4 | 2.0223
|  | 598.5 | 2.0217
|  | 598.6 | 2.0128
|  | 598.7 | 2.0111
|  | 598.8 | 2.0142
|  | 598.9 | 2.0118
|  | 599 | 2.0094
|  | 599.1 | 2.0157
|  | 599.2 | 2.0199
|  | 599.3 | 2.0241
|  | 599.4 | 2.0178
|  | 599.5 | 2.0178
|  | 599.6 | 2.0178
|  | 599.7 | 2.0005
|  | 599.8 | 1.9921
|  | 599.9 | 1.9838
|  | 600 | 1.9871
|  | 600.1 | 1.9817
|  | 600.2 | 1.9739
|  | 600.3 | 1.9703
|  | 600.4 | 1.9667
|  | 600.5 | 1.9649
|  | 600.6 | 1.9601
|  | 600.7 | 1.9553
|  | 600.8 | 1.9487
|  | 600.9 | 1.9445
|  | 601 | 1.9403
|  | 601.1 | 1.9278
|  | 601.2 | 1.9194
|  | 601.3 | 1.9222
|  | 601.4 | 1.9192
|  | 601.5 | 1.9162
|  | 601.6 | 1.9058
|  | 601.7 | 1.9022
|  | 601.8 | 1.8986
|  | 601.9 | 1.8893
|  | 602 | 1.8803
|  | 602.1 | 1.8714
|  | 602.2 | 1.8803
|  | 602.3 | 1.8791
|  | 602.4 | 1.8779
|  | 602.5 | 1.8663
|  | 602.6 | 1.8633
|  | 602.7 | 1.8628
|  | 602.8 | 1.8574
|  | 602.9 | 1.852
|  | 603 | 1.852
|  | 603.1 | 1.8514
|  | 603.2 | 1.8508
|  | 603.3 | 1.8465
|  | 603.4 | 1.8441
|  | 603.5 | 1.8417
|  | 603.6 | 1.8373
|  | 603.7 | 1.8337
|  | 603.8 | 1.8301
|  | 603.9 | 1.8231
|  | 604 | 1.8183
|  | 604.1 | 1.8208
|  | 604.2 | 1.819
|  | 604.3 | 1.8172
|  | 604.4 | 1.8101
|  | 604.5 | 1.8077
|  | 604.6 | 1.8053
|  | 604.7 | 1.8043
|  | 604.8 | 1.8007
|  | 604.9 | 1.7971
|  | 605 | 1.7984
|  | 605.1 | 1.7978
|  | 605.2 | 1.7931
|  | 605.3 | 1.7901
|  | 605.4 | 1.787
|  | 605.5 | 1.7818
|  | 605.6 | 1.777
|  | 605.7 | 1.7722
|  | 605.8 | 1.7775
|  | 605.9 | 1.7768
|  | 606 | 1.7762
|  | 606.1 | 1.7733
|  | 606.2 | 1.7733
|  | 606.3 | 1.7733
|  | 606.4 | 1.7664
|  | 606.5 | 1.7628
|  | 606.6 | 1.7567
|  | 606.7 | 1.7519
|  | 606.8 | 1.7471
|  | 606.9 | 1.7462
|  | 607 | 1.7425
|  | 607.1 | 1.7389
|  | 607.2 | 1.7407
|  | 607.3 | 1.7401
|  | 607.4 | 1.7395
|  | 607.5 | 1.73
|  | 607.6 | 1.7264
|  | 607.7 | 1.7297
|  | 607.8 | 1.7279
|  | 607.9 | 1.7261
|  | 608 | 1.726
|  | 608.1 | 1.7272
|  | 608.2 | 1.7284
|  | 608.3 | 1.7245
|  | 608.4 | 1.7227
|  | 608.5 | 1.7209
|  | 608.6 | 1.7196
|  | 608.7 | 1.7178
|  | 608.8 | 1.7223
|  | 608.9 | 1.7241
|  | 609 | 1.7259
|  | 609.1 | 1.7301
|  | 609.2 | 1.7343
|  | 609.3 | 1.7386
|  | 609.4 | 1.7519
|  | 609.5 | 1.7597
|  | 609.6 | 1.7676
|  | 609.7 | 1.7794
|  | 609.8 | 1.7902
|  | 609.9 | 1.8011
|  | 610 | 1.8421
|  | 610.1 | 1.8632
|  | 610.2 | 1.9089
|  | 610.3 | 1.9468
|  | 610.4 | 1.9848
|  | 610.5 | 2.0807
|  | 610.6 | 2.1459
|  | 610.7 | 2.2111
|  | 610.8 | 2.3436
|  | 610.9 | 2.4431
|  | 611 | 2.5426
|  | 611.1 | 2.6624
|  | 611.2 | 2.7733
|  | 611.3 | 2.7758
|  | 611.4 | 2.8325
|  | 611.5 | 2.8893
|  | 611.6 | 2.7466
|  | 611.7 | 2.6934
|  | 611.8 | 2.6403
|  | 611.9 | 2.4793
|  | 612 | 2.3622
|  | 612.1 | 2.245
|  | 612.2 | 2.1875
|  | 612.3 | 2.0897
|  | 612.4 | 2.0571
|  | 612.5 | 2.0046
|  | 612.6 | 1.952
|  | 612.7 | 1.9372
|  | 612.8 | 1.9094
|  | 612.9 | 1.8817
|  | 613 | 1.8705
|  | 613.1 | 1.8518
|  | 613.2 | 1.833
|  | 613.3 | 1.8262
|  | 613.4 | 1.8129
|  | 613.5 | 1.8057
|  | 613.6 | 1.7966
|  | 613.7 | 1.7876
|  | 613.8 | 1.7728
|  | 613.9 | 1.7619
|  | 614 | 1.751
|  | 614.1 | 1.7538
|  | 614.2 | 1.7472
|  | 614.3 | 1.7405
|  | 614.4 | 1.7396
|  | 614.5 | 1.7378
|  | 614.6 | 1.7259
|  | 614.7 | 1.7175
|  | 614.8 | 1.709
|  | 614.9 | 1.6986
|  | 615 | 1.6877
|  | 615.1 | 1.6768
|  | 615.2 | 1.656
|  | 615.3 | 1.6415
|  | 615.4 | 1.6269
|  | 615.5 | 1.612
|  | 615.6 | 1.5956
|  | 615.7 | 1.5799
|  | 615.8 | 1.5654
|  | 615.9 | 1.5509
|  | 616 | 1.544
|  | 616.1 | 1.5325
|  | 616.2 | 1.5209
|  | 616.3 | 1.5158
|  | 616.4 | 1.5085
|  | 616.5 | 1.5012
|  | 616.6 | 1.4956
|  | 616.7 | 1.4889
|  | 616.8 | 1.484
|  | 616.9 | 1.4786
|  | 617 | 1.4731
|  | 617.1 | 1.4658
|  | 617.2 | 1.4597
|  | 617.3 | 1.4537
|  | 617.4 | 1.4414
|  | 617.5 | 1.4323
|  | 617.6 | 1.4232
|  | 617.7 | 1.4303
|  | 617.8 | 1.4266
|  | 617.9 | 1.4233
|  | 618 | 1.4233
|  | 618.1 | 1.4233
|  | 618.2 | 1.4174
|  | 618.3 | 1.4132
|  | 618.4 | 1.409
|  | 618.5 | 1.4067
|  | 618.6 | 1.4031
|  | 618.7 | 1.3994
|  | 618.8 | 1.3963
|  | 618.9 | 1.3933
|  | 619 | 1.3873
|  | 619.1 | 1.3824
|  | 619.2 | 1.3776
|  | 619.3 | 1.3831
|  | 619.4 | 1.3825
|  | 619.5 | 1.3819
|  | 619.6 | 1.3718
|  | 619.7 | 1.3694
|  | 619.8 | 1.367
|  | 619.9 | 1.3646
|  | 620 | 1.3598
|  | 620.1 | 1.3564
|  | 620.2 | 1.3546
|  | 620.3 | 1.3528
|  | 620.4 | 1.3513
|  | 620.5 | 1.3488
|  | 620.6 | 1.3464
|  | 620.7 | 1.3429
|  | 620.8 | 1.3404
|  | 620.9 | 1.338
|  | 621 | 1.33
|  | 621.1 | 1.3252
|  | 621.2 | 1.325
|  | 621.3 | 1.322
|  | 621.4 | 1.3189
|  | 621.5 | 1.3117
|  | 621.6 | 1.3081
|  | 621.7 | 1.3044
|  | 621.8 | 1.2995
|  | 621.9 | 1.294
|  | 622 | 1.2886
|  | 622.1 | 1.2865
|  | 622.2 | 1.2828
|  | 622.3 | 1.2804
|  | 622.4 | 1.2774
|  | 622.5 | 1.2743
|  | 622.6 | 1.2693
|  | 622.7 | 1.2657
|  | 622.8 | 1.262
|  | 622.9 | 1.2626
|  | 623 | 1.2602
|  | 623.1 | 1.2611
|  | 623.2 | 1.2617
|  | 623.3 | 1.2623
|  | 623.4 | 1.265
|  | 623.5 | 1.2668
|  | 623.6 | 1.2686
|  | 623.7 | 1.2511
|  | 623.8 | 1.245
|  | 623.9 | 1.2389
|  | 624 | 1.2325
|  | 624.1 | 1.2234
|  | 624.2 | 1.2267
|  | 624.3 | 1.2267
|  | 624.4 | 1.2267
|  | 624.5 | 1.2233
|  | 624.6 | 1.2233
|  | 624.7 | 1.2233
|  | 624.8 | 1.2177
|  | 624.9 | 1.2141
|  | 625 | 1.2104
|  | 625.1 | 1.2067
|  | 625.2 | 1.2031
|  | 625.3 | 1.1994
|  | 625.4 | 1.1958
|  | 625.5 | 1.1921
|  | 625.6 | 1.1857
|  | 625.7 | 1.1808
|  | 625.8 | 1.1759
|  | 625.9 | 1.169
|  | 626 | 1.1629
|  | 626.1 | 1.1568
|  | 626.2 | 1.1609
|  | 626.3 | 1.1585
|  | 626.4 | 1.1628
|  | 626.5 | 1.1653
|  | 626.6 | 1.1677
|  | 626.7 | 1.1547
|  | 626.8 | 1.1511
|  | 626.9 | 1.1474
|  | 627 | 1.1472
|  | 627.1 | 1.1423
|  | 627.2 | 1.14
|  | 627.3 | 1.1393
|  | 627.4 | 1.1387
|  | 627.5 | 1.1279
|  | 627.6 | 1.1218
|  | 627.7 | 1.1156
|  | 627.8 | 1.1203
|  | 627.9 | 1.1173
|  | 628 | 1.1142
|  | 628.1 | 1.1034
|  | 628.2 | 1.0997
|  | 628.3 | 1.1027
|  | 628.4 | 1.099
|  | 628.5 | 1.0954
|  | 628.6 | 1.0947
|  | 628.7 | 1.0953
|  | 628.8 | 1.0959
|  | 628.9 | 1.0978
|  | 629 | 1.0978
|  | 629.1 | 1.0911
|  | 629.2 | 1.0886
|  | 629.3 | 1.0862
|  | 629.4 | 1.0843
|  | 629.5 | 1.08
|  | 629.6 | 1.0757
|  | 629.7 | 1.0745
|  | 629.8 | 1.0727
|  | 629.9 | 1.0708
|  | 630 | 1.0761
|  | 630.1 | 1.0767
|  | 630.2 | 1.0841
|  | 630.3 | 1.0896
|  | 630.4 | 1.0951
|  | 630.5 | 1.1031
|  | 630.6 | 1.1104
|  | 630.7 | 1.1178
|  | 630.8 | 1.1236
|  | 630.9 | 1.1303
|  | 631 | 1.1371
|  | 631.1 | 1.1358
|  | 631.2 | 1.1395
|  | 631.3 | 1.1238
|  | 631.4 | 1.117
|  | 631.5 | 1.1102
|  | 631.6 | 1.0889
|  | 631.7 | 1.073
|  | 631.8 | 1.057
|  | 631.9 | 1.0406
|  | 632 | 1.024
|  | 632.1 | 1.014
|  | 632.2 | 1.0011
|  | 632.3 | 0.98818
|  | 632.4 | 0.97794
|  | 632.5 | 0.96748
|  | 632.6 | 0.95703
|  | 632.7 | 0.96256
|  | 632.8 | 0.95827
|  | 632.9 | 0.95398
|  | 633 | 0.94527
|  | 633.1 | 0.94159
|  | 633.2 | 0.93962
|  | 633.3 | 0.93408
|  | 633.4 | 0.92855
|  | 633.5 | 0.92786
|  | 633.6 | 0.92601
|  | 633.7 | 0.92417
|  | 633.8 | 0.92614
|  | 633.9 | 0.92553
|  | 634 | 0.91726
|  | 634.1 | 0.91357
|  | 634.2 | 0.90988
|  | 634.3 | 0.91064
|  | 634.4 | 0.90695
|  | 634.5 | 0.90326
|  | 634.6 | 0.90289
|  | 634.7 | 0.90227
|  | 634.8 | 0.89884
|  | 634.9 | 0.89576
|  | 635 | 0.89268
|  | 635.1 | 0.89667
|  | 635.2 | 0.89667
|  | 635.3 | 0.89667
|  | 635.4 | 0.89333
|  | 635.5 | 0.89334
|  | 635.6 | 0.89334
|  | 635.7 | 0.8767
|  | 635.8 | 0.86931
|  | 635.9 | 0.86893
|  | 636 | 0.86338
|  | 636.1 | 0.85783
|  | 636.2 | 0.86
|  | 636.3 | 0.86
|  | 636.4 | 0.86
|  | 636.5 | 0.85841
|  | 636.6 | 0.85779
|  | 636.7 | 0.85541
|  | 636.8 | 0.85356
|  | 636.9 | 0.8517
|  | 637 | 0.84639
|  | 637.1 | 0.84269
|  | 637.2 | 0.83899
|  | 637.3 | 0.8364
|  | 637.4 | 0.8327
|  | 637.5 | 0.829
|  | 637.6 | 0.8259
|  | 637.7 | 0.82281
|  | 637.8 | 0.819
|  | 637.9 | 0.8153
|  | 638 | 0.81159
|  | 638.1 | 0.80935
|  | 638.2 | 0.80627
|  | 638.3 | 0.80318
|  | 638.4 | 0.81272
|  | 638.5 | 0.81457
|  | 638.6 | 0.80767
|  | 638.7 | 0.80706
|  | 638.8 | 0.80644
|  | 638.9 | 0.8007
|  | 639 | 0.79513
|  | 639.1 | 0.78956
|  | 639.2 | 0.78871
|  | 639.3 | 0.78562
|  | 639.4 | 0.78778
|  | 639.5 | 0.78778
|  | 639.6 | 0.78778
|  | 639.7 | 0.77882
|  | 639.8 | 0.77573
|  | 639.9 | 0.77264
|  | 640 | 0.77608
|  | 640.1 | 0.77361
|  | 640.2 | 0.77114
|  | 640.3 | 0.76655
|  | 640.4 | 0.76532
|  | 640.5 | 0.76593
|  | 640.6 | 0.76346
|  | 640.7 | 0.76099
|  | 640.8 | 0.75666
|  | 640.9 | 0.75481
|  | 641 | 0.75295
|  | 641.1 | 0.7522
|  | 641.2 | 0.74973
|  | 641.3 | 0.74627
|  | 641.4 | 0.74442
|  | 641.5 | 0.74256
|  | 641.6 | 0.74237
|  | 641.7 | 0.74051
|  | 641.8 | 0.73865
|  | 641.9 | 0.73675
|  | 642 | 0.73551
|  | 642.1 | 0.73538
|  | 642.2 | 0.73414
|  | 642.3 | 0.7329
|  | 642.4 | 0.72638
|  | 642.5 | 0.72329
|  | 642.6 | 0.72019
|  | 642.7 | 0.72145
|  | 642.8 | 0.71897
|  | 642.9 | 0.71893
|  | 643 | 0.71955
|  | 643.1 | 0.72017
|  | 643.2 | 0.71596
|  | 643.3 | 0.7141
|  | 643.4 | 0.71224
|  | 643.5 | 0.70742
|  | 643.6 | 0.70369
|  | 643.7 | 0.70332
|  | 643.8 | 0.70146
|  | 643.9 | 0.6996
|  | 644 | 0.68822
|  | 644.1 | 0.68264
|  | 644.2 | 0.67706
|  | 644.3 | 0.68849
|  | 644.4 | 0.68786
|  | 644.5 | 0.68724
|  | 644.6 | 0.68738
|  | 644.7 | 0.69048
|  | 644.8 | 0.68696
|  | 644.9 | 0.68386
|  | 645 | 0.68075
|  | 645.1 | 0.6771
|  | 645.2 | 0.67399
|  | 645.3 | 0.67089
|  | 645.4 | 0.67167
|  | 645.5 | 0.66981
|  | 645.6 | 0.66627
|  | 645.7 | 0.66441
|  | 645.8 | 0.66254
|  | 645.9 | 0.65802
|  | 646 | 0.65429
|  | 646.1 | 0.65056
|  | 646.2 | 0.65176
|  | 646.3 | 0.6499
|  | 646.4 | 0.65364
|  | 646.5 | 0.6555
|  | 646.6 | 0.65736
|  | 646.7 | 0.65051
|  | 646.8 | 0.64927
|  | 646.9 | 0.64803
|  | 647 | 0.64318
|  | 647.1 | 0.63883
|  | 647.2 | 0.63636
|  | 647.3 | 0.63387
|  | 647.4 | 0.63137
|  | 647.5 | 0.63167
|  | 647.6 | 0.63042
|  | 647.7 | 0.62918
|  | 647.8 | 0.62841
|  | 647.9 | 0.62779
|  | 648 | 0.62772
|  | 648.1 | 0.6271
|  | 648.2 | 0.62648
|  | 648.3 | 0.61982
|  | 648.4 | 0.61671
|  | 648.5 | 0.6136
|  | 648.6 | 0.61797
|  | 648.7 | 0.61673
|  | 648.8 | 0.60758
|  | 648.9 | 0.60447
|  | 649 | 0.60136
|  | 649.1 | 0.60706
|  | 649.2 | 0.60519
|  | 649.3 | 0.60333
|  | 649.4 | 0.60045
|  | 649.5 | 0.60107
|  | 649.6 | 0.60556
|  | 649.7 | 0.60556
|  | 649.8 | 0.60556
|  | 649.9 | 0.61047
|  | 650 | 0.61484
|  | 650.1 | 0.6192
|  | 650.2 | 0.61796
|  | 650.3 | 0.61983
|  | 650.4 | 0.62225
|  | 650.5 | 0.6235
|  | 650.6 | 0.62474
|  | 650.7 | 0.62089
|  | 650.8 | 0.62027
|  | 650.9 | 0.61964
|  | 651 | 0.60184
|  | 651.1 | 0.5931
|  | 651.2 | 0.59774
|  | 651.3 | 0.594
|  | 651.4 | 0.59025
|  | 651.5 | 0.58026
|  | 651.6 | 0.57777
|  | 651.7 | 0.57527
|  | 651.8 | 0.57137
|  | 651.9 | 0.56513
|  | 652 | 0.55889
|  | 652.1 | 0.56556
|  | 652.2 | 0.56556
|  | 652.3 | 0.5684
|  | 652.4 | 0.57027
|  | 652.5 | 0.57214
|  | 652.6 | 0.56584
|  | 652.7 | 0.56459
|  | 652.8 | 0.56334
|  | 652.9 | 0.56369
|  | 653 | 0.56181
|  | 653.1 | 0.55715
|  | 653.2 | 0.55528
|  | 653.3 | 0.5534
|  | 653.4 | 0.55125
|  | 653.5 | 0.54812
|  | 653.6 | 0.545
|  | 653.7 | 0.54367
|  | 653.8 | 0.5418
|  | 653.9 | 0.53916
|  | 654 | 0.53666
|  | 654.1 | 0.53415
|  | 654.2 | 0.53012
|  | 654.3 | 0.52699
|  | 654.4 | 0.52777
|  | 654.5 | 0.52651
|  | 654.6 | 0.52526
|  | 654.7 | 0.52178
|  | 654.8 | 0.52053
|  | 654.9 | 0.51928
|  | 655 | 0.51716
|  | 655.1 | 0.51465
|  | 655.2 | 0.5133
|  | 655.3 | 0.51204
|  | 655.4 | 0.51079
|  | 655.5 | 0.50685
|  | 655.6 | 0.50434
|  | 655.7 | 0.50184
|  | 655.8 | 0.50355
|  | 655.9 | 0.5023
|  | 656 | 0.5034
|  | 656.1 | 0.50465
|  | 656.2 | 0.5059
|  | 656.3 | 0.50173
|  | 656.4 | 0.50048
|  | 656.5 | 0.49923
|  | 656.6 | 0.49079
|  | 656.7 | 0.48577
|  | 656.8 | 0.48649
|  | 656.9 | 0.48398
|  | 657 | 0.48148
|  | 657.1 | 0.48111
|  | 657.2 | 0.48111
|  | 657.3 | 0.48111
|  | 657.4 | 0.48222
|  | 657.5 | 0.48222
|  | 657.6 | 0.48464
|  | 657.7 | 0.48652
|  | 657.8 | 0.4884
|  | 657.9 | 0.48695
|  | 658 | 0.48757
|  | 658.1 | 0.4882
|  | 658.2 | 0.47924
|  | 658.3 | 0.47547
|  | 658.4 | 0.47871
|  | 658.5 | 0.47745
|  | 658.6 | 0.4762
|  | 658.7 | 0.47333
|  | 658.8 | 0.47333
|  | 658.9 | 0.47333
|  | 659 | 0.46359
|  | 659.1 | 0.45857
|  | 659.2 | 0.45832
|  | 659.3 | 0.45518
|  | 659.4 | 0.45203
|  | 659.5 | 0.45767
|  | 659.6 | 0.45955
|  | 659.7 | 0.46144
|  | 659.8 | 0.45167
|  | 659.9 | 0.44979
|  | 660 | 0.45039
|  | 660.1 | 0.44724
|  | 660.2 | 0.4441
|  | 660.3 | 0.44592
|  | 660.4 | 0.44655
|  | 660.5 | 0.44718
|  | 660.6 | 0.44327
|  | 660.7 | 0.44202
|  | 660.8 | 0.44316
|  | 660.9 | 0.44253
|  | 661 | 0.4419
|  | 661.1 | 0.43809
|  | 661.2 | 0.43683
|  | 661.3 | 0.43557
|  | 661.4 | 0.43827
|  | 661.5 | 0.43764
|  | 661.6 | 0.44086
|  | 661.7 | 0.44337
|  | 661.8 | 0.44589
|  | 661.9 | 0.44532
|  | 662 | 0.44658
|  | 662.1 | 0.44781
|  | 662.2 | 0.44844
|  | 662.3 | 0.44907
|  | 662.4 | 0.44505
|  | 662.5 | 0.44379
|  | 662.6 | 0.44253
|  | 662.7 | 0.43712
|  | 662.8 | 0.43333
|  | 662.9 | 0.43311
|  | 663 | 0.43122
|  | 663.1 | 0.42933
|  | 663.2 | 0.42436
|  | 663.3 | 0.42184
|  | 663.4 | 0.41933
|  | 663.5 | 0.41738
|  | 663.6 | 0.41422
|  | 663.7 | 0.41309
|  | 663.8 | 0.41183
|  | 663.9 | 0.41057
|  | 664 | 0.40493
|  | 664.1 | 0.40177
|  | 664.2 | 0.39862
|  | 664.3 | 0.40218
|  | 664.4 | 0.40092
|  | 664.5 | 0.39633
|  | 664.6 | 0.39506
|  | 664.7 | 0.3938
|  | 664.8 | 0.39285
|  | 664.9 | 0.39033
|  | 665 | 0.3878
|  | 665.1 | 0.39
|  | 665.2 | 0.39
|  | 665.3 | 0.38867
|  | 665.4 | 0.38804
|  | 665.5 | 0.38741
|  | 665.6 | 0.38108
|  | 665.7 | 0.37792
|  | 665.8 | 0.38222
|  | 665.9 | 0.38223
|  | 666 | 0.38223
|  | 666.1 | 0.37778
|  | 666.2 | 0.37778
|  | 666.3 | 0.37779
|  | 666.4 | 0.38222
|  | 666.5 | 0.38223
|  | 666.6 | 0.38022
|  | 666.7 | 0.38149
|  | 666.8 | 0.38275
|  | 666.9 | 0.38154
|  | 667 | 0.38028
|  | 667.1 | 0.37902
|  | 667.2 | 0.37495
|  | 667.3 | 0.37306
|  | 667.4 | 0.37154
|  | 667.5 | 0.369
|  | 667.6 | 0.36647
|  | 667.7 | 0.36586
|  | 667.8 | 0.3646
|  | 667.9 | 0.36333
|  | 668 | 0.36381
|  | 668.1 | 0.36318
|  | 668.2 | 0.36199
|  | 668.3 | 0.36135
|  | 668.4 | 0.36072
|  | 668.5 | 0.35906
|  | 668.6 | 0.35779
|  | 668.7 | 0.35771
|  | 668.8 | 0.35707
|  | 668.9 | 0.35644
|  | 669 | 0.35384
|  | 669.1 | 0.35257
|  | 669.2 | 0.3513
|  | 669.3 | 0.35114
|  | 669.4 | 0.34988
|  | 669.5 | 0.35125
|  | 669.6 | 0.35189
|  | 669.7 | 0.35252
|  | 669.8 | 0.34664
|  | 669.9 | 0.34473
|  | 670 | 0.34283
|  | 670.1 | 0.34371
|  | 670.2 | 0.34181
|  | 670.3 | 0.34111
|  | 670.4 | 0.34111
|  | 670.5 | 0.34111
|  | 670.6 | 0.34066
|  | 670.7 | 0.34002
|  | 670.8 | 0.33883
|  | 670.9 | 0.3382
|  | 671 | 0.33756
|  | 671.1 | 0.33467
|  | 671.2 | 0.33277
|  | 671.3 | 0.33086
|  | 671.4 | 0.33115
|  | 671.5 | 0.32988
|  | 671.6 | 0.33375
|  | 671.7 | 0.33565
|  | 671.8 | 0.33756
|  | 671.9 | 0.33333
|  | 672 | 0.33334
|  | 672.1 | 0.33334
|  | 672.2 | 0.3263
|  | 672.3 | 0.32248
|  | 672.4 | 0.32267
|  | 672.5 | 0.32076
|  | 672.6 | 0.31885
|  | 672.7 | 0.32
|  | 672.8 | 0.32
|  | 672.9 | 0.32
|  | 673 | 0.32
|  | 673.1 | 0.32
|  | 673.2 | 0.31573
|  | 673.3 | 0.31382
|  | 673.4 | 0.31191
|  | 673.5 | 0.315
|  | 673.6 | 0.31437
|  | 673.7 | 0.31222
|  | 673.8 | 0.31222
|  | 673.9 | 0.31222
|  | 674 | 0.30537
|  | 674.1 | 0.30156
|  | 674.2 | 0.30332
|  | 674.3 | 0.30205
|  | 674.4 | 0.30078
|  | 674.5 | 0.30297
|  | 674.6 | 0.30489
|  | 674.7 | 0.3068
|  | 674.8 | 0.30123
|  | 674.9 | 0.29995
|  | 675 | 0.3009
|  | 675.1 | 0.29963
|  | 675.2 | 0.29835
|  | 675.3 | 0.29632
|  | 675.4 | 0.29568
|  | 675.5 | 0.29504
|  | 675.6 | 0.29095
|  | 675.7 | 0.2884
|  | 675.8 | 0.2898
|  | 675.9 | 0.28916
|  | 676 | 0.28853
|  | 676.1 | 0.28466
|  | 676.2 | 0.28339
|  | 676.3 | 0.28316
|  | 676.4 | 0.28125
|  | 676.5 | 0.27933
|  | 676.6 | 0.27858
|  | 676.7 | 0.27794
|  | 676.8 | 0.27731
|  | 676.9 | 0.28222
|  | 677 | 0.2835
|  | 677.1 | 0.28127
|  | 677.2 | 0.28191
|  | 677.3 | 0.28254
|  | 677.4 | 0.27836
|  | 677.5 | 0.27581
|  | 677.6 | 0.27778
|  | 677.7 | 0.27778
|  | 677.8 | 0.27777
|  | 677.9 | 0.27555
|  | 678 | 0.27555
|  | 678.1 | 0.27555
|  | 678.2 | 0.27124
|  | 678.3 | 0.26868
|  | 678.4 | 0.26738
|  | 678.5 | 0.26546
|  | 678.6 | 0.26354
|  | 678.7 | 0.26257
|  | 678.8 | 0.26129
|  | 678.9 | 0.26001
|  | 679 | 0.26571
|  | 679.1 | 0.26699
|  | 679.2 | 0.26877
|  | 679.3 | 0.27133
|  | 679.4 | 0.27388
|  | 679.5 | 0.26628
|  | 679.6 | 0.26436
|  | 679.7 | 0.26399
|  | 679.8 | 0.26144
|  | 679.9 | 0.25888
|  | 680 | 0.26
|  | 680.1 | 0.26
|  | 680.2 | 0.26
|  | 680.3 | 0.26173
|  | 680.4 | 0.26237
|  | 680.5 | 0.26111
|  | 680.6 | 0.26111
|  | 680.7 | 0.26111
|  | 680.8 | 0.25903
|  | 680.9 | 0.25774
|  | 681 | 0.25889
|  | 681.1 | 0.25889
|  | 681.2 | 0.25889
|  | 681.3 | 0.25191
|  | 681.4 | 0.24935
|  | 681.5 | 0.24678
|  | 681.6 | 0.24928
|  | 681.7 | 0.24736
|  | 681.8 | 0.25223
|  | 681.9 | 0.25544
|  | 682 | 0.25864
|  | 682.1 | 0.25174
|  | 682.2 | 0.2511
|  | 682.3 | 0.24951
|  | 682.4 | 0.24629
|  | 682.5 | 0.24308
|  | 682.6 | 0.24262
|  | 682.7 | 0.24133
|  | 682.8 | 0.24005
|  | 682.9 | 0.24222
|  | 683 | 0.24222
|  | 683.1 | 0.24648
|  | 683.2 | 0.24904
|  | 683.3 | 0.25161
|  | 683.4 | 0.24667
|  | 683.5 | 0.24667
|  | 683.6 | 0.24513
|  | 683.7 | 0.24256
|  | 683.8 | 0.23999
|  | 683.9 | 0.23778
|  | 684 | 0.23586
|  | 684.1 | 0.23393
|  | 684.2 | 0.243
|  | 684.3 | 0.24493
|  | 684.4 | 0.23753
|  | 684.5 | 0.2369
|  | 684.6 | 0.23626
|  | 684.7 | 0.23687
|  | 684.8 | 0.2343
|  | 684.9 | 0.23222
|  | 685 | 0.23222
|  | 685.1 | 0.23223
|  | 685.2 | 0.23145
|  | 685.3 | 0.23016
|  | 685.4 | 0.23224
|  | 685.5 | 0.23352
|  | 685.6 | 0.23481
|  | 685.7 | 0.22808
|  | 685.8 | 0.22615
|  | 685.9 | 0.22422
|  | 686 | 0.22619
|  | 686.1 | 0.22426
|  | 686.2 | 0.22474
|  | 686.3 | 0.22603
|  | 686.4 | 0.22732
|  | 686.5 | 0.22972
|  | 686.6 | 0.23101
|  | 686.7 | 0.23348
|  | 686.8 | 0.23606
|  | 686.9 | 0.23864
|  | 687 | 0.23838
|  | 687.1 | 0.23967
|  | 687.2 | 0.24096
|  | 687.3 | 0.23497
|  | 687.4 | 0.23304
|  | 687.5 | 0.24056
|  | 687.6 | 0.24249
|  | 687.7 | 0.24442
|  | 687.8 | 0.24136
|  | 687.9 | 0.2433
|  | 688 | 0.23955
|  | 688.1 | 0.23568
|  | 688.2 | 0.23181
|  | 688.3 | 0.22648
|  | 688.4 | 0.22197
|  | 688.5 | 0.21746
|  | 688.6 | 0.21695
|  | 688.7 | 0.21373
|  | 688.8 | 0.21556
|  | 688.9 | 0.21555
|  | 689 | 0.21555
|  | 689.1 | 0.21767
|  | 689.2 | 0.21896
|  | 689.3 | 0.21284
|  | 689.4 | 0.21026
|  | 689.5 | 0.20768
|  | 689.6 | 0.20541
|  | 689.7 | 0.20154
|  | 689.8 | 0.20448
|  | 689.9 | 0.20578
|  | 690 | 0.20707
|  | 690.1 | 0.21033
|  | 690.2 | 0.21355
|  | 690.3 | 0.21678
|  | 690.4 | 0.25558
|  | 690.5 | 0.27233
|  | 690.6 | 0.28855
|  | 690.7 | 0.31116
|  | 690.8 | 0.33377
|  | 690.9 | 0.27606
|  | 691 | 0.26249
|  | 691.1 | 0.25133
|  | 691.2 | 0.22608
|  | 691.3 | 0.20082
|  | 691.4 | 0.20424
|  | 691.5 | 0.19777
|  | 691.6 | 0.1913
|  | 691.7 | 0.19826
|  | 691.8 | 0.19762
|  | 691.9 | 0.19692
|  | 692 | 0.19756
|  | 692.1 | 0.19821
|  | 692.2 | 0.18907
|  | 692.3 | 0.18518
|  | 692.4 | 0.18985
|  | 692.5 | 0.1892
|  | 692.6 | 0.18856
|  | 692.7 | 0.18907
|  | 692.8 | 0.19101
|  | 692.9 | 0.19339
|  | 693 | 0.19404
|  | 693.1 | 0.19469
|  | 693.2 | 0.19655
|  | 693.3 | 0.19849
|  | 693.4 | 0.20044
|  | 693.5 | 0.20293
|  | 693.6 | 0.20487
|  | 693.7 | 0.20626
|  | 693.8 | 0.2082
|  | 693.9 | 0.21015
|  | 694 | 0.19966
|  | 694.1 | 0.19642
|  | 694.2 | 0.19581
|  | 694.3 | 0.19192
|  | 694.4 | 0.18803
|  | 694.5 | 0.18541
|  | 694.6 | 0.18347
|  | 694.7 | 0.18203
|  | 694.8 | 0.17944
|  | 694.9 | 0.17685
|  | 695 | 0.18032
|  | 695.1 | 0.18097
|  | 695.2 | 0.18162
|  | 695.3 | 0.17889
|  | 695.4 | 0.17889
|  | 695.5 | 0.17708
|  | 695.6 | 0.17513
|  | 695.7 | 0.17319
|  | 695.8 | 0.16905
|  | 695.9 | 0.16645
|  | 696 | 0.17503
|  | 696.1 | 0.17764
|  | 696.2 | 0.18024
|  | 696.3 | 0.19112
|  | 696.4 | 0.19955
|  | 696.5 | 0.19333
|  | 696.6 | 0.19333
|  | 696.7 | 0.19333
|  | 696.8 | 0.18404
|  | 696.9 | 0.17689
|  | 697 | 0.16974
|  | 697.1 | 0.17072
|  | 697.2 | 0.16682
|  | 697.3 | 0.16342
|  | 697.4 | 0.16082
|  | 697.5 | 0.15823
|  | 697.6 | 0.16169
|  | 697.7 | 0.16104
|  | 697.8 | 0.16658
|  | 697.9 | 0.17048
|  | 698 | 0.17438
|  | 698.1 | 0.16621
|  | 698.2 | 0.16556
|  | 698.3 | 0.16151
|  | 698.4 | 0.15631
|  | 698.5 | 0.15112
|  | 698.6 | 0.1537
|  | 698.7 | 0.1524
|  | 698.8 | 0.15335
|  | 698.9 | 0.15465
|  | 699 | 0.15595
|  | 699.1 | 0.15304
|  | 699.2 | 0.15239
|  | 699.3 | 0.15174
|  | 699.4 | 0.16121
|  | 699.5 | 0.16381
|  | 699.6 | 0.15667
|  | 699.7 | 0.15667
|  | 699.8 | 0.15667
|  | 699.9 | 0.1585
|  | 700 | 0.15654
|  | 700.1 | 0.15708
|  | 700.2 | 0.15903
|  | 700.3 | 0.16098
|  | 700.4 | 0.16
|  | 700.5 | 0.16
|  | 700.6 | 0.15766
|  | 700.7 | 0.15635
|  | 700.8 | 0.15505
|  | 700.9 | 0.15339
|  | 701 | 0.15144
|  | 701.1 | 0.14948
|  | 701.2 | 0.15572
|  | 701.3 | 0.15702
|  | 701.4 | 0.15195
|  | 701.5 | 0.1513
|  | 701.6 | 0.15065
|  | 701.7 | 0.1455
|  | 701.8 | 0.14158
|  | 701.9 | 0.14839
|  | 702 | 0.15034
|  | 702.1 | 0.1523
|  | 702.2 | 0.14876
|  | 702.3 | 0.15006
|  | 702.4 | 0.15333
|  | 702.5 | 0.15333
|  | 702.6 | 0.15333
|  | 702.7 | 0.15461
|  | 702.8 | 0.15657
|  | 702.9 | 0.15549
|  | 703 | 0.15484
|  | 703.1 | 0.15418
|  | 703.2 | 0.15409
|  | 703.3 | 0.15344
|  | 703.4 | 0.15278
|  | 703.5 | 0.14806
|  | 703.6 | 0.1461
|  | 703.7 | 0.1486
|  | 703.8 | 0.14795
|  | 703.9 | 0.14729
|  | 704 | 0.14386
|  | 704.1 | 0.14321
|  | 704.2 | 0.148
|  | 704.3 | 0.14865
|  | 704.4 | 0.14931
|  | 704.5 | 0.14608
|  | 704.6 | 0.14674
|  | 704.7 | 0.15127
|  | 704.8 | 0.15193
|  | 704.9 | 0.15259
|  | 705 | 0.14621
|  | 705.1 | 0.14556
|  | 705.2 | 0.14869
|  | 705.3 | 0.14739
|  | 705.4 | 0.14608
|  | 705.5 | 0.15086
|  | 705.6 | 0.15414
|  | 705.7 | 0.15796
|  | 705.8 | 0.16124
|  | 705.9 | 0.16451
|  | 706 | 0.17013
|  | 706.1 | 0.17471
|  | 706.2 | 0.1793
|  | 706.3 | 0.18681
|  | 706.4 | 0.1927
|  | 706.5 | 0.20222
|  | 706.6 | 0.21007
|  | 706.7 | 0.21793
|  | 706.8 | 0.21455
|  | 706.9 | 0.21848
|  | 707 | 0.21091
|  | 707.1 | 0.20699
|  | 707.2 | 0.20306
|  | 707.3 | 0.19808
|  | 707.4 | 0.19283
|  | 707.5 | 0.18902
|  | 707.6 | 0.18508
|  | 707.7 | 0.18114
|  | 707.8 | 0.18352
|  | 707.9 | 0.18221
|  | 708 | 0.18111
|  | 708.1 | 0.18111
|  | 708.2 | 0.18111
|  | 708.3 | 0.1716
|  | 708.4 | 0.167
|  | 708.5 | 0.18031
|  | 708.6 | 0.18425
|  | 708.7 | 0.18819
|  | 708.8 | 0.20704
|  | 708.9 | 0.22213
|  | 709 | 0.23722
|  | 709.1 | 0.22816
|  | 709.2 | 0.23409
|  | 709.3 | 0.21317
|  | 709.4 | 0.20331
|  | 709.5 | 0.19344
|  | 709.6 | 0.17941
|  | 709.7 | 0.16563
|  | 709.8 | 0.15864
|  | 709.9 | 0.14945
|  | 710 | 0.14027
|  | 710.1 | 0.13995
|  | 710.2 | 0.13469
|  | 710.3 | 0.1325
|  | 710.4 | 0.12987
|  | 710.5 | 0.12724
|  | 710.6 | 0.12948
|  | 710.7 | 0.12883
|  | 710.8 | 0.13414
|  | 710.9 | 0.13743
|  | 711 | 0.14072
|  | 711.1 | 0.13445
|  | 711.2 | 0.13445
|  | 711.3 | 0.13766
|  | 711.4 | 0.137
|  | 711.5 | 0.13634
|  | 711.6 | 0.15177
|  | 711.7 | 0.15966
|  | 711.8 | 0.1666
|  | 711.9 | 0.17581
|  | 712 | 0.18501
|  | 712.1 | 0.18899
|  | 712.2 | 0.19492
|  | 712.3 | 0.19226
|  | 712.4 | 0.19291
|  | 712.5 | 0.19357
|  | 712.6 | 0.18794
|  | 712.7 | 0.18399
|  | 712.8 | 0.18004
|  | 712.9 | 0.16198
|  | 713 | 0.15276
|  | 713.1 | 0.14657
|  | 713.2 | 0.13668
|  | 713.3 | 0.1268
|  | 713.4 | 0.12682
|  | 713.5 | 0.12221
|  | 713.6 | 0.12117
|  | 713.7 | 0.11853
|  | 713.8 | 0.1159
|  | 713.9 | 0.12057
|  | 714 | 0.12123
|  | 714.1 | 0.11621
|  | 714.2 | 0.1149
|  | 714.3 | 0.11358
|  | 714.4 | 0.11724
|  | 714.5 | 0.11658
|  | 714.6 | 0.11333
|  | 714.7 | 0.11334
|  | 714.8 | 0.11334
|  | 714.9 | 0.11182
|  | 715 | 0.10984
|  | 715.1 | 0.10491
|  | 715.2 | 0.10227
|  | 715.3 | 0.099636
|  | 715.4 | 0.10397
|  | 715.5 | 0.10332
|  | 715.6 | 0.10248
|  | 715.7 | 0.1038
|  | 715.8 | 0.10512
|  | 715.9 | 0.1009
|  | 716 | 0.098921
|  | 716.1 | 0.10232
|  | 716.2 | 0.10297
|  | 716.3 | 0.10363
|  | 716.4 | 0.10152
|  | 716.5 | 0.10218
|  | 716.6 | 0.10562
|  | 716.7 | 0.10628
|  | 716.8 | 0.10694
|  | 716.9 | 0.10782
|  | 717 | 0.1098
|  | 717.1 | 0.10885
|  | 717.2 | 0.10819
|  | 717.3 | 0.10753
|  | 717.4 | 0.10302
|  | 717.5 | 0.10037
|  | 717.6 | 0.10221
|  | 717.7 | 0.10155
|  | 717.8 | 0.10089
|  | 717.9 | 0.10177
|  | 718 | 0.1031
|  | 718.1 | 0.10442
|  | 718.2 | 0.10157
|  | 718.3 | 0.10091
|  | 718.4 | 0.10364
|  | 718.5 | 0.1043
|  | 718.6 | 0.10496
|  | 718.7 | 0.10744
|  | 718.8 | 0.10942
|  | 718.9 | 0.10359
|  | 719 | 0.10161
|  | 719.1 | 0.099623
|  | 719.2 | 0.096973
|  | 719.3 | 0.093658
|  | 719.4 | 0.09693
|  | 719.5 | 0.097594
|  | 719.6 | 0.098258
|  | 719.7 | 0.092165
|  | 719.8 | 0.090842
|  | 719.9 | 0.095557
|  | 720 | 0.09556
|  | 720.1 | 0.095564
|  | 720.2 | 0.096143
|  | 720.3 | 0.098133
|  | 720.4 | 0.093541
|  | 720.5 | 0.090885
|  | 720.6 | 0.088229
|  | 720.7 | 0.088358
|  | 720.8 | 0.086371
|  | 720.9 | 0.091941
|  | 721 | 0.09459
|  | 721.1 | 0.097239
|  | 721.2 | 0.096038
|  | 721.3 | 0.09803
|  | 721.4 | 0.094978
|  | 721.5 | 0.092986
|  | 721.6 | 0.090994
|  | 721.7 | 0.091186
|  | 721.8 | 0.089858
|  | 721.9 | 0.086129
|  | 722 | 0.084137
|  | 722.1 | 0.082145
|  | 722.2 | 0.081819
|  | 722.3 | 0.079827
|  | 722.4 | 0.082554
|  | 722.5 | 0.083882
|  | 722.6 | 0.08521
|  | 722.7 | 0.082224
|  | 722.8 | 0.082225
|  | 722.9 | 0.081908
|  | 723 | 0.080578
|  | 723.1 | 0.079248
|  | 723.2 | 0.084788
|  | 723.3 | 0.08678
|  | 723.4 | 0.085853
|  | 723.5 | 0.087186
|  | 723.6 | 0.088518
|  | 723.7 | 0.085713
|  | 723.8 | 0.084382
|  | 723.9 | 0.084302
|  | 724 | 0.083637
|  | 724.1 | 0.082971
|  | 724.2 | 0.083333
|  | 724.3 | 0.083333
|  | 724.4 | 0.08458
|  | 724.5 | 0.085245
|  | 724.6 | 0.085909
|  | 724.7 | 0.092319
|  | 724.8 | 0.095647
|  | 724.9 | 0.089872
|  | 725 | 0.089211
|  | 725.1 | 0.088549
|  | 725.2 | 0.081876
|  | 725.3 | 0.076542
|  | 725.4 | 0.078762
|  | 725.5 | 0.078096
|  | 725.6 | 0.077429
|  | 725.7 | 0.077127
|  | 725.8 | 0.077793
|  | 725.9 | 0.078889
|  | 726 | 0.078889
|  | 726.1 | 0.078889
|  | 726.2 | 0.08092
|  | 726.3 | 0.082253
|  | 726.4 | 0.078636
|  | 726.5 | 0.077302
|  | 726.6 | 0.075969
|  | 726.7 | 0.076858
|  | 726.8 | 0.075524
|  | 726.9 | 0.074444
|  | 727 | 0.074444
|  | 727.1 | 0.074444
|  | 727.2 | 0.080269
|  | 727.3 | 0.082269
|  | 727.4 | 0.079016
|  | 727.5 | 0.079682
|  | 727.6 | 0.080349
|  | 727.7 | 0.075462
|  | 727.8 | 0.072121
|  | 727.9 | 0.074445
|  | 728 | 0.074447
|  | 728.1 | 0.074448
|  | 728.2 | 0.075385
|  | 728.3 | 0.076722
|  | 728.4 | 0.075417
|  | 728.5 | 0.074749
|  | 728.6 | 0.074081
|  | 728.7 | 0.075555
|  | 728.8 | 0.075554
|  | 728.9 | 0.078212
|  | 729 | 0.080217
|  | 729.1 | 0.082221
|  | 729.2 | 0.072514
|  | 729.3 | 0.069846
|  | 729.4 | 0.069076
|  | 729.5 | 0.06506
|  | 729.6 | 0.061044
|  | 729.7 | 0.072972
|  | 729.8 | 0.076988
|  | 729.9 | 0.075103
|  | 730 | 0.077771
|  | 730.1 | 0.08044
|  | 730.2 | 0.074654
|  | 730.3 | 0.071982
|  | 730.4 | 0.073159
|  | 730.5 | 0.07249
|  | 730.6 | 0.071821
|  | 730.7 | 0.071111
|  | 730.8 | 0.071111
|  | 730.9 | 0.070924
|  | 731 | 0.070254
|  | 731.1 | 0.069585
|  | 731.2 | 0.06672
|  | 731.3 | 0.065382
|  | 731.4 | 0.067777
|  | 731.5 | 0.067775
|  | 731.6 | 0.067774
|  | 731.7 | 0.068873
|  | 731.8 | 0.070213
|  | 731.9 | 0.071332
|  | 732 | 0.072001
|  | 732.1 | 0.072671
|  | 732.2 | 0.072784
|  | 732.3 | 0.073454
|  | 732.4 | 0.067942
|  | 732.5 | 0.065261
|  | 732.6 | 0.06258
|  | 732.7 | 0.058735
|  | 732.8 | 0.054712
|  | 732.9 | 0.063249
|  | 733 | 0.065928
|  | 733.1 | 0.068608
|  | 733.2 | 0.058282
|  | 733.3 | 0.057606
|  | 733.4 | 0.066941
|  | 733.5 | 0.067605
|  | 733.6 | 0.068269
|  | 733.7 | 0.064097
|  | 733.8 | 0.066113
|  | 733.9 | 0.06547
|  | 734 | 0.06279
|  | 734.1 | 0.060109
|  | 734.2 | 0.062869
|  | 734.3 | 0.06354
|  | 734.4 | 0.0619
|  | 734.5 | 0.061228
|  | 734.6 | 0.060556
|  | 734.7 | 0.058656
|  | 734.8 | 0.057312
|  | 734.9 | 0.058889
|  | 735 | 0.058889
|  | 735.1 | 0.05664
|  | 735.2 | 0.055968
|  | 735.3 | 0.055296
|  | 735.4 | 0.05629
|  | 735.5 | 0.055618
|  | 735.6 | 0.060269
|  | 735.7 | 0.063629
|  | 735.8 | 0.066989
|  | 735.9 | 0.065654
|  | 736 | 0.06767
|  | 736.1 | 0.065314
|  | 736.2 | 0.063297
|  | 736.3 | 0.061281
|  | 736.4 | 0.061355
|  | 736.5 | 0.060007
|  | 736.6 | 0.06367
|  | 736.7 | 0.065692
|  | 736.8 | 0.067713
|  | 736.9 | 0.066483
|  | 737 | 0.067827
|  | 737.1 | 0.065125
|  | 737.2 | 0.063105
|  | 737.3 | 0.061085
|  | 737.4 | 0.061231
|  | 737.5 | 0.059885
|  | 737.6 | 0.062577
|  | 737.7 | 0.063923
|  | 737.8 | 0.065268
|  | 737.9 | 0.062222
|  | 738 | 0.062222
|  | 738.1 | 0.081663
|  | 738.2 | 0.090425
|  | 738.3 | 0.099187
|  | 738.4 | 0.10186
|  | 738.5 | 0.11197
|  | 738.6 | 0.096392
|  | 738.7 | 0.089645
|  | 738.8 | 0.082897
|  | 738.9 | 0.067958
|  | 739 | 0.055812
|  | 739.1 | 0.062173
|  | 739.2 | 0.059474
|  | 739.3 | 0.056775
|  | 739.4 | 0.057953
|  | 739.5 | 0.059297
|  | 739.6 | 0.056801
|  | 739.7 | 0.05478
|  | 739.8 | 0.049955
|  | 739.9 | 0.046576
|  | 740 | 0.043196
|  | 740.1 | 0.048891
|  | 740.2 | 0.048895
|  | 740.3 | 0.04794
|  | 740.4 | 0.049964
|  | 740.5 | 0.051988
|  | 740.6 | 0.053745
|  | 740.7 | 0.05442
|  | 740.8 | 0.059456
|  | 740.9 | 0.063504
|  | 741 | 0.067553
|  | 741.1 | 0.057981
|  | 741.2 | 0.056634
|  | 741.3 | 0.059582
|  | 741.4 | 0.057557
|  | 741.5 | 0.055533
|  | 741.6 | 0.056554
|  | 741.7 | 0.057904
|  | 741.8 | 0.053705
|  | 741.9 | 0.051002
|  | 742 | 0.048298
|  | 742.1 | 0.054993
|  | 742.2 | 0.055671
|  | 742.3 | 0.056492
|  | 742.4 | 0.059195
|  | 742.5 | 0.061898
|  | 742.6 | 0.056584
|  | 742.7 | 0.055233
|  | 742.8 | 0.059173
|  | 742.9 | 0.059851
|  | 743 | 0.060528
|  | 743.1 | 0.054906
|  | 743.2 | 0.054226
|  | 743.3 | 0.058887
|  | 743.4 | 0.058883
|  | 743.5 | 0.055576
|  | 743.6 | 0.056253
|  | 743.7 | 0.056931
|  | 743.8 | 0.046524
|  | 743.9 | 0.040427
|  | 744 | 0.04311
|  | 744.1 | 0.041077
|  | 744.2 | 0.039045
|  | 744.3 | 0.041321
|  | 744.4 | 0.043354
|  | 744.5 | 0.045813
|  | 744.6 | 0.047168
|  | 744.7 | 0.048523
|  | 744.8 | 0.051978
|  | 744.9 | 0.054688
|  | 745 | 0.043164
|  | 745.1 | 0.038421
|  | 745.2 | 0.033679
|  | 745.3 | 0.043334
|  | 745.4 | 0.041969
|  | 745.5 | 0.043916
|  | 745.6 | 0.048665
|  | 745.7 | 0.053414
|  | 745.8 | 0.053442
|  | 745.9 | 0.054797
|  | 746 | 0.046576
|  | 746.1 | 0.043864
|  | 746.2 | 0.041152
|  | 746.3 | 0.048224
|  | 746.4 | 0.04755
|  | 746.5 | 0.043698
|  | 746.6 | 0.044381
|  | 746.7 | 0.048889
|  | 746.8 | 0.048893
|  | 746.9 | 0.048897
|  | 747 | 0.047515
|  | 747.1 | 0.048875
|  | 747.2 | 0.044977
|  | 747.3 | 0.041576
|  | 747.4 | 0.038175
|  | 747.5 | 0.048136
|  | 747.6 | 0.050177
|  | 747.7 | 0.049793
|  | 747.8 | 0.053184
|  | 747.9 | 0.056576
|  | 748 | 0.04724
|  | 748.1 | 0.043847
|  | 748.2 | 0.043465
|  | 748.3 | 0.040744
|  | 748.4 | 0.038023
|  | 748.5 | 0.03654
|  | 748.6 | 0.035179
|  | 748.7 | 0.037778
|  | 748.8 | 0.037778
|  | 748.9 | 0.038909
|  | 749 | 0.04095
|  | 749.1 | 0.042991
|  | 749.2 | 0.041111
|  | 749.3 | 0.041111
|  | 749.4 | 0.044599
|  | 749.5 | 0.04596
|  | 749.6 | 0.047322
|  | 749.7 | 0.046463
|  | 749.8 | 0.047825
|  | 749.9 | 0.045107
|  | 750 | 0.043066
|  | 750.1 | 0.041025
|  | 750.2 | 0.070603
|  | 750.3 | 0.081488
|  | 750.4 | 0.081307
|  | 750.5 | 0.092223
|  | 750.6 | 0.10314
|  | 750.7 | 0.07727
|  | 750.8 | 0.069086
|  | 750.9 | 0.065343
|  | 751 | 0.057159
|  | 751.1 | 0.048975
|  | 751.2 | 0.061348
|  | 751.3 | 0.062714
|  | 751.4 | 0.077036
|  | 751.5 | 0.085895
|  | 751.6 | 0.073128
|  | 751.7 | 0.071079
|  | 751.8 | 0.06903
|  | 751.9 | 0.061514
|  | 752 | 0.051975
|  | 752.1 | 0.048162
|  | 752.2 | 0.044754
|  | 752.3 | 0.041346
|  | 752.4 | 0.040118
|  | 752.5 | 0.037386
|  | 752.6 | 0.037552
|  | 752.7 | 0.036869
|  | 752.8 | 0.036186
|  | 752.9 | 0.034948
|  | 753 | 0.034265
|  | 753.1 | 0.032719
|  | 753.2 | 0.031352
|  | 753.3 | 0.029986
|  | 753.4 | 0.02862
|  | 753.5 | 0.027254
|  | 753.6 | 0.027388
|  | 753.7 | 0.026705
|  | 753.8 | 0.027867
|  | 753.9 | 0.02855
|  | 754 | 0.029233
|  | 754.1 | 0.024613
|  | 754.2 | 0.023247
|  | 754.3 | 0.02795
|  | 754.4 | 0.028632
|  | 754.5 | 0.029313
|  | 754.6 | 0.027782
|  | 754.7 | 0.029148
|  | 754.8 | 0.026999
|  | 754.9 | 0.02495
|  | 755 | 0.022901
|  | 755.1 | 0.030824
|  | 755.2 | 0.032875
|  | 755.3 | 0.027776
|  | 755.4 | 0.027772
|  | 755.5 | 0.029899
|  | 755.6 | 0.028528
|  | 755.7 | 0.027156
|  | 755.8 | 0.028649
|  | 755.9 | 0.030017
|  | 756 | 0.026249
|  | 756.1 | 0.024197
|  | 756.2 | 0.022145
|  | 756.3 | 0.029945
|  | 756.4 | 0.031317
|  | 756.5 | 0.024211
|  | 756.6 | 0.023525
|  | 756.7 | 0.02284
|  | 756.8 | 0.026541
|  | 756.9 | 0.025177
|  | 757 | 0.029722
|  | 757.1 | 0.03383
|  | 757.2 | 0.031084
|  | 757.3 | 0.030398
|  | 757.4 | 0.029712
|  | 757.5 | 0.025686
|  | 757.6 | 0.022257
|  | 757.7 | 0.021975
|  | 757.8 | 0.020604
|  | 757.9 | 0.019232
|  | 758 | 0.028128
|  | 758.1 | 0.031557
|  | 758.2 | 0.020233
|  | 758.3 | 0.01749
|  | 758.4 | 0.014746
|  | 758.5 | 0.021485
|  | 758.6 | 0.019416
|  | 758.7 | 0.014763
|  | 758.8 | 0.015445
|  | 758.9 | 0.019973
|  | 759 | 0.019287
|  | 759.1 | 0.018601
|  | 759.2 | 0.020562
|  | 759.3 | 0.023306
|  | 759.4 | 0.019622
|  | 759.5 | 0.017562
|  | 759.6 | 0.015501
|  | 759.7 | 0.021634
|  | 759.8 | 0.022324
|  | 759.9 | 0.023139
|  | 760 | 0.02589
|  | 760.1 | 0.028641
|  | 760.2 | 0.028401
|  | 760.3 | 0.029089
|  | 760.4 | 0.026335
|  | 760.5 | 0.025649
|  | 760.6 | 0.023168
|  | 760.7 | 0.020413
|  | 760.8 | 0.017658
|  | 760.9 | 0.019563
|  | 761 | 0.018877
|  | 761.1 | 0.017778
|  | 761.2 | 0.017779
|  | 761.3 | 0.01778
|  | 761.4 | 0.018889
|  | 761.5 | 0.018889
|  | 761.6 | 0.017778
|  | 761.7 | 0.017778
|  | 761.8 | 0.017778
|  | 761.9 | 0.018889
|  | 762 | 0.018889
|  | 762.1 | 0.016302
|  | 762.2 | 0.015613
|  | 762.3 | 0.016591
|  | 762.4 | 0.015902
|  | 762.5 | 0.015213
|  | 762.6 | 0.017617
|  | 762.7 | 0.018994
|  | 762.8 | 0.020372
|  | 762.9 | 0.021749
|  | 763 | 0.023127
|  | 763.1 | 0.026201
|  | 763.2 | 0.028267
|  | 763.3 | 0.087595
|  | 763.4 | 0.11794
|  | 763.5 | 0.13055
|  | 763.6 | 0.16366
|  | 763.7 | 0.19677
|  | 763.8 | 0.13551
|  | 763.9 | 0.11691
|  | 764 | 0.098875
|  | 764.1 | 0.068499
|  | 764.2 | 0.038123
|  | 764.3 | 0.041537
|  | 764.4 | 0.030497
|  | 764.5 | 0.032357
|  | 764.6 | 0.029594
|  | 764.7 | 0.02683
|  | 764.8 | 0.027777
|  | 764.9 | 0.027776
|  | 765 | 0.025944
|  | 765.1 | 0.024564
|  | 765.2 | 0.018432
|  | 765.3 | 0.014283
|  | 765.4 | 0.010134
|  | 765.5 | 0.014597
|  | 765.6 | 0.01322
|  | 765.7 | 0.01392
|  | 765.8 | 0.015992
|  | 765.9 | 0.018064
|  | 766 | 0.016667
|  | 766.1 | 0.016667
|  | 766.2 | 0.022384
|  | 766.3 | 0.02515
|  | 766.4 | 0.017674
|  | 766.5 | 0.015599
|  | 766.6 | 0.013525
|  | 766.7 | 0.0129
|  | 766.8 | 0.0087506
|  | 766.9 | 0.014129
|  | 767 | 0.017586
|  | 767.1 | 0.021044
|  | 767.2 | 0.014444
|  | 767.3 | 0.014444
|  | 767.4 | 0.017494
|  | 767.5 | 0.016803
|  | 767.6 | 0.016111
|  | 767.7 | 0.013333
|  | 767.8 | 0.013333
|  | 767.9 | 0.016665
|  | 768 | 0.016662
|  | 768.1 | 0.017054
|  | 768.2 | 0.019132
|  | 768.3 | 0.02121
|  | 768.4 | 0.018349
|  | 768.5 | 0.017658
|  | 768.6 | 0.016149
|  | 768.7 | 0.014763
|  | 768.8 | 0.013377
|  | 768.9 | 0.011994
|  | 769 | 0.010608
|  | 769.1 | 0.0092192
|  | 769.2 | 0.0078332
|  | 769.3 | 0.012548
|  | 769.4 | 0.01463
|  | 769.5 | 0.016711
|  | 769.6 | 0.016012
|  | 769.7 | 0.018094
|  | 769.8 | 0.014826
|  | 769.9 | 0.012743
|  | 770 | 0.01066
|  | 770.1 | 0.016214
|  | 770.2 | 0.016909
|  | 770.3 | 0.013335
|  | 770.4 | 0.013338
|  | 770.5 | 0.0082083
|  | 770.6 | 0.0033472
|  | 770.7 | -0.0015139
|  | 770.8 | 0.0023056
|  | 770.9 | 0.00091667
|  | 771 | 0.005625
|  | 771.1 | 0.0090972
|  | 771.2 | 0.012569
|  | 771.3 | 0.0066667
|  | 771.4 | 0.0066667
|  | 771.5 | 0.0025694
|  | 771.6 | -0.00090278
|  | 771.7 | 0.0035278
|  | 771.8 | 0.0049167
|  | 771.9 | 0.0063056
|  | 772 | 0.014111
|  | 772.1 | 0.019667
|  | 772.2 | 0.034653
|  | 772.3 | 0.045069
|  | 772.4 | 0.055486
|  | 772.5 | 0.063839
|  | 772.6 | 0.074977
|  | 772.7 | 0.056602
|  | 772.8 | 0.052419
|  | 772.9 | 0.049861
|  | 773 | 0.041528
|  | 773.1 | 0.033194
|  | 773.2 | 0.025808
|  | 773.3 | 0.020244
|  | 773.4 | 0.020734
|  | 773.5 | 0.016563
|  | 773.6 | 0.012391
|  | 773.7 | 0.021569
|  | 773.8 | 0.024352
|  | 773.9 | 0.013975
|  | 774 | 0.011187
|  | 774.1 | 0.0078029
|  | 774.2 | 0.0015374
|  | 774.3 | -0.004728
|  | 774.4 | 0.0060762
|  | 774.5 | 0.0081683
|  | 774.6 | 0.0063615
|  | 774.7 | 0.0084458
|  | 774.8 | 0.01053
|  | 774.9 | 0.0039017
|  | 775 | 0.0011181
|  | 775.1 | -0.00014412
|  | 775.2 | -0.0022362
|  | 775.3 | 0.0026616
|  | 775.4 | 0.0047536
|  | 775.5 | 0.0068457
|  | 775.6 | 0.0078429
|  | 775.7 | 0.010632
|  | 775.8 | 0.0057671
|  | 775.9 | 0.003675
|  | 776 | 0.0021106
|  | 776.1 | -0.00067875
|  | 776.2 | -0.0034682
|  | 776.3 | 0.0020177
|  | 776.4 | 0.0034124
|  | 776.5 | 0.0012924
|  | 776.6 | 0.0019898
|  | 776.7 | 0.0026871
|  | 776.8 | 0.0096057
|  | 776.9 | 0.011699
|  | 777 | 0.0041017
|  | 777.1 | 0.0033976
|  | 777.2 | 0.0029149
|  | 777.3 | -0.0012692
|  | 777.4 | -0.0054533
|  | 777.5 | 0.0047129
|  | 777.6 | 0.0082045
|  | 777.7 | 0.0089184
|  | 777.8 | 0.01241
|  | 777.9 | 0.015902
|  | 778 | 0.0080335
|  | 778.1 | 0.0059414
|  | 778.2 | 0.0055058
|  | 778.3 | 0.0034079
|  | 778.4 | 0.0045635
|  | 778.5 | 0.0052638
|  | 778.6 | 0.0059641
|  | 778.7 | 0.0027824
|  | 778.8 | 0.0020822
|  | 778.9 | 0.0016581
|  | 779 | 0.00025952
|  | 779.1 | 0.0022502
|  | 779.2 | 0.0036508
|  | 779.3 | 0.0050514
|  | 779.4 | 0.0037773
|  | 779.5 | 0.0044747
|  | 779.6 | 0.0016912
|  | 779.7 | -0.0004047
|  | 779.8 | -0.0025006
|  | 779.9 | 0.0011111
|  | 780 | 0.0011111
|  | 780.1 | 0.00034314
|  | 780.2 | 0.0010434
|  | 780.3 | -0.0049066
|  | 780.4 | -0.0091083
|  | 780.5 | -0.01331
|  | 780.6 | -0.0071779
|  | 780.7 | -0.0078782
|  | 780.8 | 0.0013399
|  | 780.9 | 0.0083427
|  | 781 | 0.015345
|  | 781.1 | -0.0020378
|  | 781.2 | -0.0041387
|  | 781.3 | 0.0020798
|  | 781.4 | -3.2817e-05
|  | 781.5 | 0.0021749
|  | 781.6 | 0.0070847
|  | 781.7 | 0.011995
|  | 781.8 | 0.012876
|  | 781.9 | 0.014977
|  | 782 | 0.0093511
|  | 782.1 | 0.0079505
|  | 782.2 | 0.0087066
|  | 782.3 | 0.0066038
|  | 782.4 | 0.0045011
|  | 782.5 | -0.00038049
|  | 782.6 | -0.0024813
|  | 782.7 | -0.0045034
|  | 782.8 | -0.0080137
|  | 782.9 | -0.011524
|  | 783 | -0.004211
|  | 783.1 | -0.0028055
|  | 783.2 | -0.0066649
|  | 783.3 | -0.006662
|  | 783.4 | -0.00081575
|  | 783.5 | 0.00059072
|  | 783.6 | 0.0019972
|  | 783.7 | -0.0027426
|  | 783.8 | -0.0020394
|  | 783.9 | -0.010725
|  | 784 | -0.017042
|  | 784.1 | -0.0051633
|  | 784.2 | -0.00096067
|  | 784.3 | 0.003242
|  | 784.4 | -0.0030239
|  | 784.5 | 0.00049226
|  | 784.6 | -0.0041913
|  | 784.7 | -0.0098172
|  | 784.8 | -0.015443
|  | 784.9 | -0.0070886
|  | 785 | -0.0063854
|  | 785.1 | -0.0069198
|  | 785.2 | -0.0055134
|  | 785.3 | -0.0082841
|  | 785.4 | -0.010394
|  | 785.5 | -0.012504
|  | 785.6 | -0.007166
|  | 785.7 | -0.0064627
|  | 785.8 | -0.011463
|  | 785.9 | -0.012166
|  | 786 | -0.012598
|  | 786.1 | -0.015419
|  | 786.2 | -0.01824
|  | 786.3 | -0.018318
|  | 786.4 | -0.019024
|  | 786.5 | -0.016385
|  | 786.6 | -0.015682
|  | 786.7 | -0.015509
|  | 786.8 | -0.014098
|  | 786.9 | -0.012688
|  | 787 | -0.0085896
|  | 787.1 | -0.0064739
|  | 787.2 | -0.0060249
|  | 787.3 | -0.0039093
|  | 787.4 | -0.0017936
|  | 787.5 | -0.015735
|  | 787.6 | -0.020672
|  | 787.7 | -0.014145
|  | 787.8 | -0.01556
|  | 787.9 | -0.012314
|  | 788 | -0.0073826
|  | 788.1 | -0.002451
|  | 788.2 | -0.014596
|  | 788.3 | -0.017411
|  | 788.4 | -0.0085499
|  | 788.5 | -0.0078437
|  | 788.6 | -0.013164
|  | 788.7 | -0.011751
|  | 788.8 | -0.010339
|  | 788.9 | -0.01709
|  | 789 | -0.022034
|  | 789.1 | -0.016106
|  | 789.2 | -0.014694
|  | 789.3 | -0.01662
|  | 789.4 | -0.015207
|  | 789.5 | -0.013795
|  | 789.6 | -0.018559
|  | 789.7 | -0.021384
|  | 789.8 | -0.016224
|  | 789.9 | -0.014812
|  | 790 | -0.013399
|  | 790.1 | 0.0025709
|  | 790.2 | 0.011065
|  | 790.3 | -0.0072505
|  | 790.4 | -0.010075
|  | 790.5 | -0.0081921
|  | 790.6 | -0.014548
|  | 790.7 | -0.020904
|  | 790.8 | -0.018267
|  | 790.9 | -0.017561
|  | 791 | -0.02
|  | 791.1 | -0.022124
|  | 791.2 | -0.016194
|  | 791.3 | -0.01336
|  | 791.4 | -0.010527
|  | 791.5 | -0.012728
|  | 791.6 | -0.01061
|  | 791.7 | -0.013185
|  | 791.8 | -0.015309
|  | 791.9 | -0.010906
|  | 792 | -0.00878
|  | 792.1 | -0.0066543
|  | 792.2 | -0.0073088
|  | 792.3 | -0.0051832
|  | 792.4 | -0.0086076
|  | 792.5 | -0.010735
|  | 792.6 | -0.011163
|  | 792.7 | -0.012577
|  | 792.8 | -0.013992
|  | 792.9 | -0.013333
|  | 793 | -0.013333
|  | 793.1 | -0.014678
|  | 793.2 | -0.015388
|  | 793.3 | -0.016097
|  | 793.4 | -0.014444
|  | 793.5 | -0.014444
|  | 793.6 | -0.010882
|  | 793.7 | -0.0087541
|  | 793.8 | -0.012421
|  | 793.9 | -0.01313
|  | 794 | -0.013839
|  | 794.1 | -0.013541
|  | 794.2 | -0.01496
|  | 794.3 | -0.0186
|  | 794.4 | -0.020019
|  | 794.5 | -0.020489
|  | 794.6 | -0.022617
|  | 794.7 | -0.024745
|  | 794.8 | 0.0061655
|  | 794.9 | 0.018222
|  | 795 | 0.023809
|  | 795.1 | 0.037284
|  | 795.2 | 0.031839
|  | 795.3 | 0.029712
|  | 795.4 | 0.027584
|  | 795.5 | 0.014005
|  | 795.6 | 0.0054769
|  | 795.7 | -0.00047071
|  | 795.8 | -0.0082872
|  | 795.9 | -0.0068652
|  | 796 | -0.0082837
|  | 796.1 | -0.0097021
|  | 796.2 | -0.009435
|  | 796.3 | -0.0087227
|  | 796.4 | -0.020636
|  | 796.5 | -0.027037
|  | 796.6 | -0.022443
|  | 796.7 | -0.024582
|  | 796.8 | -0.026721
|  | 796.9 | -0.022869
|  | 797 | -0.019317
|  | 797.1 | -0.02
|  | 797.2 | -0.02
|  | 797.3 | -0.018839
|  | 797.4 | -0.018127
|  | 797.5 | -0.017415
|  | 797.6 | -0.023782
|  | 797.7 | -0.025919
|  | 797.8 | -0.019723
|  | 797.9 | -0.019008
|  | 798 | -0.016451
|  | 798.1 | -0.011474
|  | 798.2 | -0.0064977
|  | 798.3 | -0.011111
|  | 798.4 | -0.011111
|  | 798.5 | -0.015726
|  | 798.6 | -0.019288
|  | 798.7 | -0.022293
|  | 798.8 | -0.025855
|  | 798.9 | -0.029416
|  | 799 | -0.020769
|  | 799.1 | -0.018632
|  | 799.2 | -0.020385
|  | 799.3 | -0.018248
|  | 799.4 | -0.016111
|  | 799.5 | -0.022849
|  | 799.6 | -0.025698
|  | 799.7 | -0.020639
|  | 799.8 | -0.019928
|  | 799.9 | -0.028712
|  | 800 | -0.031572
|  | 800.1 | -0.034432
|  | 800.2 | -0.028077
|  | 800.3 | -0.029501
|  | 800.4 | -0.01867
|  | 800.5 | -0.011532
|  | 800.6 | -0.0097556
|  | 800.7 | -0.005475
|  | 800.8 | -0.0011944
|  | 800.9 | 0.0052991
|  | 801 | 0.0095726
|  | 801.1 | -0.00091446
|  | 801.2 | -0.0023359
|  | 801.3 | 0.0015673
|  | 801.4 | -0.00056845
|  | 801.5 | -0.0027042
|  | 801.6 | 0.02875
|  | 801.7 | 0.042306
|  | 801.8 | 0.046671
|  | 801.9 | 0.058832
|  | 802 | 0.047574
|  | 802.1 | 0.043286
|  | 802.2 | 0.038998
|  | 802.3 | 0.016801
|  | 802.4 | 0.0032412
|  | 802.5 | 0.00071531
|  | 802.6 | -0.0085837
|  | 802.7 | -0.001557
|  | 802.8 | 0.00058894
|  | 802.9 | 0.0027349
|  | 803 | -0.0062566
|  | 803.1 | -0.0069719
|  | 803.2 | -0.011032
|  | 803.3 | -0.01604
|  | 803.4 | -0.0090844
|  | 803.5 | -0.0062232
|  | 803.6 | -0.0033619
|  | 803.7 | -0.011111
|  | 803.8 | -0.011111
|  | 803.9 | -0.0023891
|  | 804 | -0.00095851
|  | 804.1 | -0.0035002
|  | 804.2 | -0.00063901
|  | 804.3 | 0.0022222
|  | 804.4 | -0.0003195
|  | 804.5 | -0.0017501
|  | 804.6 | -0.00062809
|  | 804.7 | 8.9269e-05
|  | 804.8 | -0.0040794
|  | 804.9 | -0.0062315
|  | 805 | -0.0022079
|  | 805.1 | -0.0014926
|  | 805.2 | -0.0007773
|  | 805.3 | -0.0007294
|  | 805.4 | 0.0014237
|  | 805.5 | -0.003274
|  | 805.6 | -0.0061424
|  | 805.7 | -0.0067599
|  | 805.8 | -0.008912
|  | 805.9 | -0.011064
|  | 806 | -0.0034991
|  | 806.1 | -0.00062969
|  | 806.2 | 0.0014104
|  | 806.3 | 0.0049962
|  | 806.4 | -0.012977
|  | 806.5 | -0.023752
|  | 806.6 | -0.034528
|  | 806.7 | -0.042387
|  | 806.8 | -0.056
|  | 806.9 | -0.044344
|  | 807 | -0.041482
|  | 807.1 | -0.038242
|  | 807.2 | -0.031777
|  | 807.3 | -0.025311
|  | 807.4 | -0.032783
|  | 807.5 | -0.033501
|  | 807.6 | -0.035766
|  | 807.7 | -0.03864

## Lamp spectra interpolation
wavelength | raw blanked output (averaged) | new wavelength | forecast over 1, 1/2 nm increments | new wavelength | forecast over 1, 0.2 nm increments | new wavelength | forecast over 1, 0.5 nm increments | new wavelength | forecast over 1, 1 nm increments | 0.13703 | RLE is shown in column k
195.423 | 0 | 199.8 | 0 | 199.8 | 0 | 199.5 | 0 | 200 | 0.0082323 | 0.060078
195.796 | 0 | 199.9 | 0.0034274 | 200 | 0.0047715 | 200 | 0.0047388 | 201 | 0.01426 | 0.10407
196.169 | 0 | 200 | 0.0047715 | 200.2 | 0.009445 | 200.5 | 0.019958 | 202 | 0.010174 | 0.074251
196.542 | 0 | 200.1 | 0.0061156 | 200.4 | 0.018797 | 201 | 0.012189 | 203 | 0.0048449 | 0.035357
196.915 | 0 | 200.2 | 0.0074597 | 200.6 | 0.022123 | 201.5 | 0.010632 | 204 | 0.0072692 | 0.053049
197.287 | 0 | 200.3 | 0.01476 | 200.8 | 0.016663 | 202 | 0.012768 | 205 | 0.0031956 | 0.023321
197.66 | 0 | 200.4 | 0.018797 | 201 | 0.014314 | 202.5 | 0.0071235 | 206 | 0 | 0
198.032 | 0 | 200.5 | 0.022834 | 201.2 | 0.0055914 | 203 | 0.0018902 | 207 | 0.0055982 | 0.040854
198.404 | 0 | 200.6 | 0.026871 | 201.4 | 0.0078166 | 203.5 | 0.005521 | 208 | 0.012014 | 0.087679
198.777 | 0 | 200.7 | 0.016665 | 201.6 | 0.012961 | 204 | 0.0082461 | 209 | 0.011803 | 0.086138
199.149 | 0 | 200.8 | 0.016663 | 201.8 | 0.012319 | 204.5 | 0.0080405 | 210 | 0.01019 | 0.074363
199.521 | 0 | 200.9 | 0.01666 | 202 | 0.013342 | 205 | 0.0015464 | 211 | 0.015054 | 0.10986
199.893 | 0 | 201 | 0.016658 | 202.2 | 0.012641 | 205.5 | 0 | 212 | 0.010295 | 0.07513
200.265 | 0.01 | 201.1 | 0.0096237 | 202.4 | 0.0090435 | 206 | 0 | 213 | 0.0076485 | 0.055818
200.636 | 0.03 | 201.2 | 0.0055914 | 202.6 | 0.0052381 | 206.5 | 0 | 214 | 0.0029954 | 0.02186
201.008 | 0.01 | 201.3 | 0.0015591 | 202.8 | 0.0030982 | 207 | 0.0033232 | 215 | 0.01242 | 0.090642
201.38 | 0 | 201.4 | 0.010273 | 203 | 0.0015139 | 207.5 | 0.013471 | 216 | 0.013824 | 0.10089
201.751 | 0.02 | 201.5 | 0.011617 | 203.2 | 0.0010584 | 208 | 0.012607 | 217 | 0.0074197 | 0.054148
202.123 | 0.01 | 201.6 | 0.012961 | 203.4 | 0.0055371 | 208.5 | 0.0099652 | 218 | 0.0067931 | 0.049575
202.494 | 0.01 | 201.7 | 0.014306 | 203.6 | 0.0061494 | 209 | 0.014519 | 219 | 0.009092 | 0.066352
202.865 | 0 | 201.8 | 0.010654 | 203.8 | 0.0033357 | 209.5 | 0.010926 | 220 | 0.010905 | 0.07958
203.236 | 0 | 201.9 | 0.011998 | 204 | 0.0084199 | 210 | 0.0050925 | 221 | 0.0069407 | 0.050652
203.608 | 0.01 | 202 | 0.013342 | 204.2 | 0.012983 | 210.5 | 0.014551 | 222 | 0.005361 | 0.039123
203.979 | 0 | 202.1 | 0.014686 | 204.4 | 0.0092197 | 211 | 0.017276 | 223 | 0.010048 | 0.073332
204.349 | 0.02 | 202.2 | 0.012292 | 204.6 | 0.0066606 | 211.5 | 0.013336 | 224 | 0.0094895 | 0.069253
204.72 | 0 | 202.3 | 0.010946 | 204.8 | 0.0043278 | 212 | 0.0078516 | 225 | 0.0093902 | 0.068528
205.091 | 0 | 202.4 | 0.009599 | 205 | 0.00031147 | 212.5 | 0.0096974 | 226 | 0.010041 | 0.073276
205.462 | 0 | 202.5 | 0.0065858 | 205.2 | 0 | 213 | 0.0095182 | 227 | 0.010267 | 0.074927
205.832 | 0 | 202.6 | 0.0052381 | 205.4 | 0 | 213.5 | 0.0037301 | 228 | 0.0086278 | 0.062964
206.203 | 0 | 202.7 | 0.0038904 | 205.6 | 0 | 214 | 9.6357e-05 | 229 | 0.011119 | 0.081144
206.573 | 0 | 202.8 | 0.0025427 | 205.8 | 0 | 214.5 | 0.0051597 | 230 | 0.011719 | 0.085525
206.944 | 0 | 202.9 | 0.0028616 | 206 | 0 | 215 | 0.015821 | 231 | 0.0066492 | 0.048525
207.314 | 0.01 | 203 | 0.0015139 | 206.2 | 0 | 215.5 | 0.016281 | 232 | 0.0035504 | 0.02591
207.684 | 0.02 | 203.1 | 0.00016622 | 206.4 | 0 | 216 | 0.011895 | 233 | 0.0059312 | 0.043284
208.054 | 0.01 | 203.2 | -0.0011815 | 206.6 | 0 | 216.5 | 0.013298 | 234 | 0.0025986 | 0.018964
208.424 | 0.01 | 203.3 | 0.0041906 | 206.8 | 0 | 217 | 0.0038074 | 235 | 0.0073154 | 0.053386
208.794 | 0.01 | 203.4 | 0.0055371 | 207 | 0.0031785 | 217.5 | 0.0051542 | 236 | 0.010433 | 0.076141
209.164 | 0.02 | 203.5 | 0.0068836 | 207.2 | 0.0067911 | 218 | 0.0072705 | 237 | 0.014599 | 0.10654
209.533 | 0.01 | 203.6 | 0.0082301 | 207.4 | 0.01183 | 218.5 | 0.0079545 | 238 | 0.0090435 | 0.065998
209.903 | 0 | 203.7 | 0.0033344 | 207.6 | 0.015363 | 219 | 0.011106 | 239 | 0.0092193 | 0.067281
210.273 | 0.01 | 203.8 | 0.0033357 | 207.8 | 0.013333 | 219.5 | 0.0082157 | 240 | 0.012101 | 0.088307
210.642 | 0.02 | 203.9 | 0.0033369 | 208 | 0.013126 | 220 | 0.012876 | 241 | 0.010625 | 0.077542
211.012 | 0.01 | 204 | 0.010288 | 208.2 | 0.01136 | 220.5 | 0.011622 | 242 | 0.011111 | 0.081086
211.381 | 0.03 | 204.1 | 0.011635 | 208.4 | 0.0095556 | 221 | 0.0052537 | 243 | 0.006872 | 0.05015
211.75 | 0 | 204.2 | 0.012983 | 208.6 | 0.01 | 221.5 | 0.0039461 | 244 | 0.0036812 | 0.026865
212.119 | 0.01 | 204.3 | 0.014331 | 208.8 | 0.012727 | 222 | 0.00053753 | 245 | 0.0039811 | 0.029054
212.488 | 0.01 | 204.4 | 0.0066654 | 209 | 0.016117 | 222.5 | 0.011599 | 246 | 0.0071948 | 0.052506
212.857 | 0.01 | 204.5 | 0.006663 | 209.2 | 0.014712 | 223 | 0.011431 | 247 | 0.0048449 | 0.035357
213.226 | 0.01 | 204.6 | 0.0066606 | 209.4 | 0.013336 | 223.5 | 0.0071149 | 248 | 0.0033461 | 0.024419
213.595 | 0 | 204.7 | 0.0066582 | 209.6 | 0.0090075 | 224 | 0.013446 | 249 | 0.0013621 | 0.0099405
213.964 | 0 | 204.8 | 0.0045103 | 209.8 | 0.002783 | 224.5 | 0.0079076 | 250 | 0.0059198 | 0.043202
214.332 | 0 | 204.9 | 0.0018149 | 210 | 0.00447 | 225 | 0.011102 | 251 | 0.0058704 | 0.042841
214.701 | 0.01 | 205 | -0.0008805 | 210.2 | 0.0080244 | 225.5 | 0.009161 | 252 | 0.0068695 | 0.050133
215.069 | 0.02 | 205.1 | 0 | 210.4 | 0.013446 | 226 | 0.0074362 | 253 | 0.0051946 | 0.037909
215.437 | 0.02 | 205.2 | 0 | 210.6 | 0.016115 | 226.5 | 0.013525 | 254 | 0.0068336 | 0.049871
215.806 | 0.01 | 205.3 | 0 | 210.8 | 0.013331 | 227 | 0.010031 | 255 | 0.0031653 | 0.0231
216.174 | 0.01 | 205.4 | 0 | 211 | 0.015951 | 227.5 | 0.0072453 | 256 | 0 | 0
216.542 | 0.02 | 205.5 | 0 | 211.2 | 0.022545 | 228 | 0.0076215 | 257 | 0.0063932 | 0.046656
216.91 | 0 | 205.6 | 0 | 211.4 | 0.016231 | 228.5 | 0.011017 | 258 | 0.0122 | 0.089033
217.278 | 0 | 205.7 | 0 | 211.6 | 0.010366 | 229 | 0.0096327 | 259 | 0.010699 | 0.078077
217.646 | 0.01 | 205.8 | 0 | 211.8 | 0.010086 | 229.5 | 0.012708 | 260 | 0.0094846 | 0.069217
218.014 | 0 | 205.9 | 0 | 212 | 0.0065583 | 230 | 0.011815 | 261 | 0.014193 | 0.10358
218.381 | 0.02 | 206 | 0 | 212.2 | 0.0069106 | 230.5 | 0.010635 | 262 | 0.0096211 | 0.070213
218.749 | 0 | 206.1 | 0 | 212.4 | 0.0098645 | 231 | 0.0063723 | 263 | 0.0058793 | 0.042906
219.116 | 0.02 | 206.2 | 0 | 212.6 | 0.01 | 231.5 | 0.0029404 | 264 | 0.0054486 | 0.039763
219.484 | 0 | 206.3 | 0 | 212.8 | 0.01 | 232 | 0.0027884 | 265 | 0.01467 | 0.10706
219.851 | 0.01 | 206.4 | 0 | 213 | 0.01 | 232.5 | 0.0049225 | 266 | 0.01029 | 0.075097
220.218 | 0.02 | 206.5 | 0 | 213.2 | 0.0085547 | 233 | 0.01009 | 267 | 0.0053594 | 0.039112
220.586 | 0.01 | 206.6 | 0 | 213.4 | 0.0043089 | 233.5 | 0.0027812 | 268 | 0.0089037 | 0.064977
220.953 | 0 | 206.7 | 0 | 213.6 | 0.00271 | 234 | 0.00093814 | 269 | 0.010455 | 0.076299
221.32 | 0.01 | 206.8 | 0 | 213.8 | 0.00055556 | 234.5 | 0.0040764 | 270 | 0.010221 | 0.074591
221.687 | 0 | 206.9 | 0 | 214 | -0.00026649 | 235 | 0.009928 | 271 | 0.0031595 | 0.023058
222.053 | 0 | 207 | 0.0040932 | 214.2 | 0 | 235.5 | 0.0079417 | 272 | 0.0078296 | 0.057139
222.42 | 0 | 207.1 | 0.0054421 | 214.4 | 0.0032871 | 236 | 0.0083617 | 273 | 0.011268 | 0.082233
222.787 | 0.03 | 207.2 | 0.0067911 | 214.6 | 0.0069668 | 236.5 | 0.014997 | 274 | 0.012004 | 0.087604
223.153 | 0 | 207.3 | 0.00814 | 214.8 | 0.012143 | 237 | 0.014817 | 275 | 0.010471 | 0.076415
223.52 | 0.01 | 207.4 | 0.012324 | 215 | 0.016873 | 237.5 | 0.013984 | 276 | 0.010134 | 0.073958
223.886 | 0.01 | 207.5 | 0.015027 | 215.2 | 0.018447 | 238 | 0.0055958 | 277 | 0.0076026 | 0.055483
224.253 | 0.02 | 207.6 | 0.01773 | 215.4 | 0.018928 | 238.5 | 0.0075509 | 278 | 0.0093815 | 0.068464
224.619 | 0 | 207.7 | 0.013333 | 215.6 | 0.014459 | 239 | 0.010107 | 279 | 0.012638 | 0.092233
224.985 | 0.01 | 207.8 | 0.013333 | 215.8 | 0.012299 | 239.5 | 0.01 | 280 | 0.0083096 | 0.060642
225.351 | 0.02 | 207.9 | 0.013333 | 216 | 0.010695 | 240 | 0.012929 | 281 | 0.0037178 | 0.027132
225.717 | 0 | 208 | 0.013333 | 216.2 | 0.01269 | 240.5 | 0.013372 | 282 | 0.0065316 | 0.047666
226.083 | 0.01 | 208.1 | 0.012712 | 216.4 | 0.016404 | 241 | 0.007756 | 283 | 0.0036143 | 0.026377
226.449 | 0.01 | 208.2 | 0.01136 | 216.6 | 0.011609 | 241.5 | 0.010748 | 284 | 0.0073294 | 0.053488
226.814 | 0.02 | 208.3 | 0.010009 | 216.8 | 0.0064946 | 242 | 0.014531 | 285 | 0.010106 | 0.073752
227.18 | 0 | 208.4 | 0.0086577 | 217 | 0.0036202 | 242.5 | 0.0080538 | 286 | 0.014491 | 0.10575
227.546 | 0.01 | 208.5 | 0.01 | 217.2 | 0.0013074 | 243 | 0.0076834 | 287 | 0.0096108 | 0.070138
227.911 | 0.01 | 208.6 | 0.01 | 217.4 | 0.0049909 | 243.5 | 0.0048787 | 288 | 0.0094625 | 0.069055
228.276 | 0 | 208.7 | 0.01 | 217.6 | 0.0057971 | 244 | 0.0047682 | 289 | 0.012249 | 0.089388
228.642 | 0.02 | 208.8 | 0.013414 | 217.8 | 0.0033333 | 244.5 | 0.0013967 | 290 | 0.010826 | 0.079003
229.007 | 0 | 208.9 | 0.014766 | 218 | 0.0059466 | 245 | 0.0045254 | 291 | 0.011494 | 0.083883
229.372 | 0.02 | 209 | 0.016117 | 218.2 | 0.012532 | 245.5 | 0.0060213 | 292 | 0.0063996 | 0.046703
229.737 | 0.01 | 209.1 | 0.017468 | 218.4 | 0.0090734 | 246 | 0.0092823 | 293 | 0.0041374 | 0.030194
230.102 | 0.01 | 209.2 | 0.013334 | 218.6 | 0.0066613 | 246.5 | 0.0062809 | 294 | 0.0060714 | 0.044308
230.467 | 0.02 | 209.3 | 0.013335 | 218.8 | 0.011107 | 247 | 0.0055227 | 295 | 0.0076942 | 0.056151
230.832 | 0 | 209.4 | 0.013336 | 219 | 0.013327 | 247.5 | 0.0027311 | 296 | 0.0039477 | 0.028809
231.196 | 0.01 | 209.5 | 0.013337 | 219.2 | 0.0088838 | 248 | 0.0040259 | 297 | 0.0021253 | 0.01551
231.561 | 0 | 209.6 | 0.0081958 | 219.4 | 0.0076998 | 248.5 | 0.0032812 | 298 | 0.0023191 | 0.016924
231.926 | 0 | 209.7 | 0.0054894 | 219.6 | 0.0084151 | 249 | 0.0001157 | 299 | 0.0069407 | 0.050652
232.29 | 0.01 | 209.8 | 0.002783 | 219.8 | 0.008026 | 249.5 | 0.00068948 | 300 | 0.0055345 | 0.04039
232.654 | 0 | 209.9 | 7.6686e-05 | 220 | 0.01406 | 250 | 0.0069677 | 301 | 0.0074817 | 0.0546
233.019 | 0.02 | 210 | 0.0066667 | 220.2 | 0.016542 | 250.5 | 0.010102 | 302 | 0.0030081 | 0.021953
233.383 | 0 | 210.1 | 0.0066667 | 220.4 | 0.013331 | 251 | 0.0040504 | 303 | 0.0022088 | 0.01612
233.747 | 0 | 210.2 | 0.0066667 | 220.6 | 0.0099429 | 251.5 | 0.0034586 | 304 | 0.0022045 | 0.016088
234.111 | 0 | 210.3 | 0.01074 | 220.8 | 0.0041678 | 252 | 0.0088552 | 305 | 0.0040586 | 0.029619
234.475 | 0.01 | 210.4 | 0.013446 | 221 | 0.0049267 | 252.5 | 0.0082947 | 306 | 0.0053979 | 0.039393
234.839 | 0 | 210.5 | 0.016152 | 221.2 | 0.0066667 | 253 | 0.0021498 | 307 | 0.0086 | 0.062761
235.203 | 0.02 | 210.6 | 0.018859 | 221.4 | 0.0044444 | 253.5 | 0.0051394 | 308 | 0.0083615 | 0.061021
235.566 | 0 | 210.7 | 0.013333 | 221.6 | 0.0032727 | 254 | 0.0073625 | 309 | 0.0040752 | 0.02974
235.93 | 0.01 | 210.8 | 0.013331 | 221.8 | 0.0017865 | 254.5 | 0.007999 | 310 | 0.00178 | 0.01299
236.293 | 0.01 | 210.9 | 0.01333 | 222 | -0.00017389 | 255 | 0.001497 | 311 | 0.0023664 | 0.017269
236.657 | 0.02 | 211 | 0.013329 | 222.2 | 0 | 255.5 | 0 | 312 | 0.0035423 | 0.025851
237.02 | 0.01 | 211.1 | 0.021194 | 222.4 | 0.0044233 | 256 | 0 | 313 | 0.0060827 | 0.04439
237.384 | 0.02 | 211.2 | 0.022545 | 222.6 | 0.017357 | 256.5 | 0 | 314 | 0.0028111 | 0.020515
237.747 | 0.01 | 211.3 | 0.023896 | 222.8 | 0.013816 | 257 | 0.0050074 | 315 | 0.0023382 | 0.017064
238.11 | 0 | 211.4 | 0.013076 | 223 | 0.010008 | 257.5 | 0.014172 | 316 | 0.0026134 | 0.019072
238.473 | 0.01 | 211.5 | 0.011721 | 223.2 | 0.010469 | 258 | 0.011874 | 317 | 0.0036284 | 0.02648
238.836 | 0.01 | 211.6 | 0.010366 | 223.4 | 0.0066091 | 258.5 | 0.010554 | 318 | 0.0036156 | 0.026386
239.199 | 0.01 | 211.7 | 0.0090108 | 223.6 | 0.0069247 | 259 | 0.014823 | 319 | 0.0015687 | 0.011448
239.561 | 0.01 | 211.8 | 0.011978 | 223.8 | 0.011047 | 259.5 | 0.0067189 | 320 | 0.0023523 | 0.017166
239.924 | 0.01 | 211.9 | 0.0092683 | 224 | 0.014885 | 260 | 0.0077281 | 321 | 0.00070347 | 0.0051338
240.287 | 0.02 | 212 | 0.0065583 | 224.2 | 0.014406 | 260.5 | 0.014007 | 322 | 0.0050211 | 0.036643
240.649 | 0.01 | 212.1 | 0.0038482 | 224.4 | 0.0079927 | 261 | 0.018229 | 323 | 0.0051289 | 0.03743
241.012 | 0.01 | 212.2 | 0.0077642 | 224.6 | 0.0069306 | 261.5 | 0.010344 | 324 | 0.00061944 | 0.0045206
241.374 | 0 | 212.3 | 0.0091192 | 224.8 | 0.0075273 | 262 | 0.0085021 | 325 | 0.004298 | 0.031366
241.736 | 0.02 | 212.4 | 0.010474 | 225 | 0.0099044 | 262.5 | 0.010017 | 326 | 0.006517 | 0.04756
242.098 | 0.02 | 212.5 | 0.01 | 225.2 | 0.015874 | 263 | 0.0067593 | 327 | 0.0049996 | 0.036486
242.46 | 0 | 212.6 | 0.01 | 225.4 | 0.011967 | 263.5 | 0.00086132 | 328 | 0.0023661 | 0.017268
242.823 | 0.01 | 212.7 | 0.01 | 225.6 | 0.0065984 | 264 | 0.0023654 | 329 | 0.0043026 | 0.031399
243.184 | 0.01 | 212.8 | 0.01 | 225.8 | 0.0071995 | 264.5 | 0.013119 | 330 | 0.005766 | 0.042079
243.546 | 0 | 212.9 | 0.01 | 226 | 0.0068443 | 265 | 0.018486 | 331 | 0.0041478 | 0.03027
243.908 | 0.01 | 213 | 0.01 | 226.2 | 0.008265 | 265.5 | 0.012404 | 332 | 0.0017694 | 0.012913
244.27 | 0 | 213.1 | 0.01 | 226.4 | 0.011555 | 266 | 0.012877 | 333 | 0.0041132 | 0.030017
244.631 | 0 | 213.2 | 0.01 | 226.6 | 0.015403 | 266.5 | 0.0055908 | 334 | 0.0068411 | 0.049925
244.993 | 0.01 | 213.3 | 0.005664 | 226.8 | 0.014578 | 267 | 0.004076 | 335 | 0.0023786 | 0.017359
245.354 | 0 | 213.4 | 0.0043089 | 227 | 0.0074566 | 267.5 | 0.0064114 | 336 | 0.0051071 | 0.03727
245.716 | 0.01 | 213.5 | 0.0029539 | 227.2 | 0.0080581 | 268 | 0.0095589 | 337 | 0.0030915 | 0.022561
246.077 | 0.01 | 213.6 | 0.0032656 | 227.4 | 0.0069945 | 268.5 | 0.010741 | 338 | -1.5069e-05 | -0.00010997
246.438 | 0.01 | 213.7 | 0.0019106 | 227.6 | 0.0072725 | 269 | 0.0071993 | 339 | 0.0017034 | 0.012431
246.799 | 0 | 213.8 | 0.00055556 | 227.8 | 0.010148 | 269.5 | 0.013425 | 340 | 0.004429 | 0.032322
247.16 | 0.01 | 213.9 | -0.00079946 | 228 | 0.0070137 | 270 | 0.011429 | 341 | 0.0068932 | 0.050305
247.521 | 0 | 214 | 0 | 228.2 | 0.0057032 | 270.5 | 0.0058091 | 342 | 0.0057318 | 0.04183
247.882 | 0 | 214.1 | 0 | 228.4 | 0.011694 | 271 | 0.0037853 | 343 | 0.0050225 | 0.036653
248.243 | 0.01 | 214.2 | 0 | 228.6 | 0.011389 | 271.5 | -0.00011583 | 344 | 0.0039047 | 0.028496
248.603 | 0 | 214.3 | 0 | 228.8 | 0.0066706 | 272 | 0.012436 | 345 | 0.003764 | 0.027469
248.964 | 0 | 214.4 | 0.0042519 | 229 | 0.008894 | 272.5 | 0.011169 | 346 | 0.0065661 | 0.047918
249.324 | 0 | 214.5 | 0.0056093 | 229.2 | 0.013333 | 273 | 0.0083422 | 347 | 0.0040497 | 0.029554
249.685 | 0 | 214.6 | 0.0069668 | 229.4 | 0.011823 | 273.5 | 0.014293 | 348 | 0.0070208 | 0.051236
250.045 | 0.01 | 214.7 | 0.0083243 | 229.6 | 0.013123 | 274 | 0.0077423 | 349 | 0.0099853 | 0.072871
250.405 | 0.01 | 214.8 | 0.012696 | 229.8 | 0.012688 | 274.5 | 0.013977 | 350 | 0.013235 | 0.096584
250.765 | 0.01 | 214.9 | 0.015409 | 230 | 0.0097306 | 275 | 0.0091675 | 351 | 0.010748 | 0.078438
251.126 | 0 | 215 | 0.018123 | 230.2 | 0.013027 | 275.5 | 0.0082685 | 352 | 0.010122 | 0.073867
251.486 | 0 | 215.1 | 0.017088 | 230.4 | 0.014336 | 276 | 0.013754 | 353 | 0.012355 | 0.090164
251.845 | 0.01 | 215.2 | 0.018447 | 230.6 | 0.0081781 | 276.5 | 0.0083803 | 354 | 0.018737 | 0.13674
252.205 | 0.01 | 215.3 | 0.019805 | 230.8 | 0.0071026 | 277 | 0.0076712 | 355 | 0.021587 | 0.15754
252.565 | 0.01 | 215.4 | 0.021164 | 231 | 0.0076877 | 277.5 | 0.0067564 | 356 | 0.02294 | 0.16741
252.925 | 0 | 215.5 | 0.015816 | 231.2 | 0.0043265 | 278 | 0.0088027 | 357 | 0.024026 | 0.17534
253.284 | 0 | 215.6 | 0.014459 | 231.4 | 0.0033308 | 278.5 | 0.012585 | 358 | 0.025112 | 0.18327
253.644 | 0.01 | 215.7 | 0.013101 | 231.6 | 0.0025193 | 279 | 0.011807 | 359 | 0.031084 | 0.22684
254.003 | 0 | 215.8 | 0.011744 | 231.8 | 5.9361e-05 | 279.5 | 0.013523 | 360 | 0.034227 | 0.24978
254.362 | 0.02 | 215.9 | 0.012053 | 232 | 0.0029219 | 280 | 0.0074415 | 361 | 0.036754 | 0.26822
254.722 | 0 | 216 | 0.010695 | 232.2 | 0.0053839 | 280.5 | 0.0039646 | 362 | 0.043204 | 0.31529
255.081 | 0 | 216.1 | 0.0093379 | 232.4 | 0.0033333 | 281 | 0.0028423 | 363 | 0.060421 | 0.44094
255.44 | 0 | 216.2 | 0.013687 | 232.6 | 0.0057647 | 281.5 | 0.0043465 | 364 | 0.58959 | 4.3027
255.799 | 0 | 216.3 | 0.015045 | 232.8 | 0.012001 | 282 | 0.010272 | 365 | 1.1855 | 8.6517
256.158 | 0 | 216.4 | 0.016404 | 233 | 0.011597 | 282.5 | 0.0049765 | 366 | 0.46775 | 3.4135
256.517 | 0 | 216.5 | 0.017763 | 233.2 | 0.0066712 | 283 | 0.00074843 | 367 | 0.12402 | 0.9051
256.875 | 0 | 216.6 | 0.009212 | 233.4 | 0.0054419 | 283.5 | 0.0051181 | 368 | 0.064447 | 0.47032
257.234 | 0.01 | 216.7 | 0.0078533 | 233.6 | 0.00070513 | 284 | 0.0094368 | 369 | 0.067076 | 0.48951
257.593 | 0.02 | 216.8 | 0.0064946 | 233.8 | -0.00068071 | 284.5 | 0.0074332 | 370 | 0.072054 | 0.52584
257.951 | 0.01 | 216.9 | 0.0051359 | 234 | 0 | 285 | 0.0082081 | 371 | 0.075337 | 0.54979
258.31 | 0.01 | 217 | 0.004221 | 234.2 | 0.0034951 | 285.5 | 0.014677 | 372 | 0.083308 | 0.60796
258.668 | 0.01 | 217.1 | 0.0015036 | 234.4 | 0.005522 | 286 | 0.014815 | 373 | 0.089591 | 0.65381
259.026 | 0.02 | 217.2 | -0.0012138 | 234.6 | 0.0033333 | 286.5 | 0.013981 | 374 | 0.091125 | 0.66501
259.384 | 0.01 | 217.3 | 0.0036322 | 234.8 | 0.0058349 | 287 | 0.0065397 | 375 | 0.094603 | 0.69039
259.742 | 0 | 217.4 | 0.0049909 | 235 | 0.012212 | 287.5 | 0.0083113 | 376 | 0.099254 | 0.72434
260.1 | 0.01 | 217.5 | 0.0063496 | 235.2 | 0.011738 | 288 | 0.010076 | 377 | 0.1 | 0.72978
260.458 | 0.02 | 217.6 | 0.0077083 | 235.4 | 0.0066716 | 288.5 | 0.01 | 378 | 0.10538 | 0.76906
260.816 | 0.01 | 217.7 | 0.0033333 | 235.6 | 0.0081252 | 289 | 0.013993 | 379 | 0.11345 | 0.82792
261.174 | 0.03 | 217.8 | 0.0033333 | 235.8 | 0.0067903 | 289.5 | 0.012753 | 380 | 0.12084 | 0.88189
261.532 | 0 | 217.9 | 0.0033333 | 236 | 0.007354 | 290 | 0.0064384 | 381 | 0.12921 | 0.94297
261.889 | 0.01 | 218 | 0.0033333 | 236.2 | 0.010941 | 290.5 | 0.013286 | 382 | 0.13451 | 0.9816
262.247 | 0.01 | 218.1 | 0.011173 | 236.4 | 0.014801 | 291 | 0.013891 | 383 | 0.13703 | 1
262.604 | 0.01 | 218.2 | 0.012532 | 236.6 | 0.015688 | 291.5 | 0.0073061 | 384 | 0.14576 | 1.0637
262.961 | 0.01 | 218.3 | 0.01389 | 236.8 | 0.013335 | 292 | 0.0066413 | 385 | 0.15574 | 1.1365
263.319 | 0 | 218.4 | 0.0066662 | 237 | 0.014447 | 292.5 | 0.0052516 | 386 | 0.16193 | 1.1817
263.676 | 0 | 218.5 | 0.0066637 | 237.2 | 0.016669 | 293 | 0.0024292 | 387 | 0.17308 | 1.2631
264.033 | 0 | 218.6 | 0.0066613 | 237.4 | 0.014446 | 293.5 | 0.0047314 | 388 | 0.17889 | 1.3055
264.39 | 0.01 | 218.7 | 0.0066588 | 237.6 | 0.013336 | 294 | 0.005318 | 389 | 0.18864 | 1.3766
264.747 | 0.02 | 218.8 | 0.013332 | 237.8 | 0.0092208 | 294.5 | 0.0081647 | 390 | 0.20688 | 1.5097
265.104 | 0.02 | 218.9 | 0.01333 | 238 | 0.0030303 | 295 | 0.0089413 | 391 | 0.21241 | 1.5501
265.46 | 0.01 | 219 | 0.013327 | 238.2 | 0.0045363 | 295.5 | 0.0059766 | 392 | 0.21439 | 1.5646
265.817 | 0.01 | 219.1 | 0.013325 | 238.4 | 0.0067906 | 296 | 0.0038546 | 393 | 0.22706 | 1.657
266.174 | 0.02 | 219.2 | 0.0066646 | 238.6 | 0.008416 | 296.5 | 0.0020118 | 394 | 0.24071 | 1.7567
266.53 | 0 | 219.3 | 0.0066621 | 238.8 | 0.010321 | 297 | 0.0039188 | 395 | 0.24994 | 1.824
266.886 | 0 | 219.4 | 0.0066597 | 239 | 0.01 | 297.5 | 0.00044539 | 396 | 0.26201 | 1.9121
267.243 | 0.01 | 219.5 | 0.0097775 | 239.2 | 0.01 | 298 | 0.00040539 | 397 | 0.27539 | 2.0098
267.599 | 0 | 219.6 | 0.0084151 | 239.4 | 0.01 | 298.5 | 0.0061065 | 398 | 0.28883 | 2.1079
267.955 | 0.02 | 219.7 | 0.0070527 | 239.6 | 0.01 | 299 | 0.0097793 | 399 | 0.30037 | 2.192
268.311 | 0 | 219.8 | 0.0056903 | 239.8 | 0.01 | 299.5 | 0.0049363 | 400 | 0.31716 | 2.3145
268.667 | 0.02 | 219.9 | 0.011335 | 240 | 0.013379 | 300 | 0.0028204 | 401 | 0.3333 | 2.4324
269.023 | 0 | 220 | 0.01406 | 240.2 | 0.015409 | 300.5 | 0.0088469 | 402 | 0.35386 | 2.5824
269.379 | 0.01 | 220.1 | 0.016785 | 240.4 | 0.013335 | 301 | 0.0079893 | 403 | 0.54695 | 3.9915
269.735 | 0.02 | 220.2 | 0.01951 | 240.6 | 0.013103 | 301.5 | 0.0056088 | 404 | 2.9467 | 21.5045
270.091 | 0.01 | 220.3 | 0.013332 | 240.8 | 0.011256 | 302 | 0.0032638 | 405 | 2.8439 | 20.7542
270.446 | 0 | 220.4 | 0.013331 | 241 | 0.0079417 | 302.5 | 0.00015183 | 406 | 0.52078 | 3.8006
270.802 | 0.01 | 220.5 | 0.01333 | 241.2 | 0.0040702 | 303 | 0.0010377 | 407 | 0.5821 | 4.2481
271.157 | 0 | 220.6 | 0.00961 | 241.4 | 0.0082636 | 303.5 | 0.005437 | 408 | 0.63414 | 4.6278
271.512 | 0 | 220.7 | 0.0068889 | 241.6 | 0.013122 | 304 | 0.0012291 | 409 | 0.47088 | 3.4364
271.868 | 0 | 220.8 | 0.0041678 | 241.8 | 0.015823 | 304.5 | -5.2605e-05 | 410 | 0.47657 | 3.4779
272.223 | 0.03 | 220.9 | 0.0014467 | 242 | 0.017256 | 305 | 0.004379 | 411 | 0.49459 | 3.6094
272.578 | 0 | 221 | 0.0066667 | 242.2 | 0.010516 | 305.5 | 0.0078493 | 412 | 0.50329 | 3.6729
272.933 | 0.01 | 221.1 | 0.0066667 | 242.4 | 0.0073992 | 306 | 0.0022988 | 413 | 0.51942 | 3.7906
273.288 | 0.01 | 221.2 | 0.0066667 | 242.6 | 0.0080762 | 306.5 | 0.0060455 | 414 | 0.53612 | 3.9125
273.643 | 0.02 | 221.3 | 0.0066667 | 242.8 | 0.0065869 | 307 | 0.0097545 | 415 | 0.54792 | 3.9986
273.997 | 0 | 221.4 | 0.0033333 | 243 | 0.0091229 | 307.5 | 0.01 | 416 | 0.55327 | 4.0376
274.352 | 0.01 | 221.5 | 0.0033333 | 243.2 | 0.0073405 | 308 | 0.01 | 417 | 0.5668 | 4.1364
274.707 | 0.02 | 221.6 | 0.0033333 | 243.4 | 0.0036824 | 308.5 | 0.0050846 | 418 | 0.58225 | 4.2492
275.061 | 0 | 221.7 | 0.0031513 | 243.6 | 0.0052106 | 309 | 0.0027094 | 419 | 0.5957 | 4.3473
275.416 | 0.01 | 221.8 | 0.0017865 | 243.8 | 0.0066667 | 309.5 | 0.0044317 | 420 | 0.61416 | 4.482
275.77 | 0.01 | 221.9 | 0.0004216 | 244 | 0.0044444 | 310 | 0.00090818 | 421 | 0.62739 | 4.5786
276.124 | 0.02 | 222 | -0.00094328 | 244.2 | 0.0031934 | 310.5 | 0 | 422 | 0.63343 | 4.6227
276.478 | 0 | 222.1 | 0 | 244.4 | 0.0015298 | 311 | 0.0046457 | 423 | 0.651 | 4.7509
276.833 | 0.01 | 222.2 | 0 | 244.6 | 0.001064 | 311.5 | 0.0024534 | 424 | 0.66873 | 4.8803
277.187 | 0.01 | 222.3 | 0 | 244.8 | 0.0056673 | 312 | 0.00086491 | 425 | 0.68258 | 4.9813
277.541 | 0 | 222.4 | 0 | 245 | 0.0045731 | 312.5 | 0.0073084 | 426 | 0.69097 | 5.0425
277.894 | 0.02 | 222.5 | 0.01327 | 245.2 | 0.003336 | 313 | 0.0057989 | 427 | 0.69179 | 5.0486
278.248 | 0 | 222.6 | 0.017357 | 245.4 | 0.0055577 | 313.5 | 0.0051408 | 428 | 0.71781 | 5.2384
278.602 | 0.02 | 222.7 | 0.021444 | 245.6 | 0.0066698 | 314 | 0.0016979 | 429 | 0.7332 | 5.3508
278.955 | 0.01 | 222.8 | 0.01 | 245.8 | 0.0079074 | 314.5 | 0.0015946 | 430 | 0.76005 | 5.5467
279.309 | 0.01 | 222.9 | 0.010004 | 246 | 0.0099395 | 315 | 0.0047899 | 431 | 0.78816 | 5.7518
279.662 | 0.02 | 223 | 0.010008 | 246.2 | 0.01 | 315.5 | 0.00063003 | 432 | 0.79048 | 5.7688
280.016 | 0 | 223.1 | 0.010012 | 246.4 | 0.0086026 | 316 | 0.0031804 | 433 | 0.82765 | 6.04
280.369 | 0.01 | 223.2 | 0.012061 | 246.6 | 0.0044229 | 316.5 | 0.0040298 | 434 | 1.0939 | 7.9831
280.722 | 0 | 223.3 | 0.0093352 | 246.8 | 0.0054571 | 317 | 0.0028293 | 435 | 5.9032 | 43.0803
281.075 | 0 | 223.4 | 0.0066091 | 247 | 0.0066667 | 317.5 | 0.0040263 | 436 | 9.1354 | 66.6687
281.428 | 0.01 | 223.5 | 0.0038831 | 247.2 | 0.0044444 | 318 | 0.0028816 | 437 | 1.9132 | 13.9619
281.781 | 0 | 223.6 | 0.0077631 | 247.4 | 0.0033333 | 318.5 | 0.003939 | 438 | 0.89583 | 6.5376
282.134 | 0.02 | 223.7 | 0.009128 | 247.6 | 0.0021422 | 319 | 0.00076707 | 439 | 0.91907 | 6.7072
282.487 | 0 | 223.8 | 0.010493 | 247.8 | 0.0013019 | 319.5 | 0 | 440 | 0.94204 | 6.8748
282.84 | 0 | 223.9 | 0.01352 | 248 | 0.0049677 | 320 | 0.0049165 | 441 | 0.94781 | 6.9169
283.192 | 0 | 224 | 0.014885 | 248.2 | 0.0058082 | 320.5 | 0.0021403 | 442 | 0.95209 | 6.9482
283.545 | 0.01 | 224.1 | 0.01625 | 248.4 | 0.0033354 | 321 | -2.9845e-05 | 443 | 0.96446 | 7.0385
283.897 | 0 | 224.2 | 0.017614 | 248.6 | 0.0028893 | 321.5 | 0 | 444 | 0.97649 | 7.1262
284.25 | 0.02 | 224.3 | 0.0093551 | 248.8 | 0.0006069 | 322 | 0.0049806 | 445 | 0.98222 | 7.1681
284.602 | 0 | 224.4 | 0.0079927 | 249 | -0.00025981 | 322.5 | 0.010083 | 446 | 0.98292 | 7.1732
284.954 | 0.01 | 224.5 | 0.0066303 | 249.2 | 0 | 323 | 0.0051728 | 447 | 0.97999 | 7.1518
285.306 | 0.01 | 224.6 | 0.0052679 | 249.4 | 0 | 323.5 | 0.0001312 | 448 | 0.98807 | 7.2107
285.658 | 0.02 | 224.7 | 0.0088934 | 249.6 | 0.001182 | 324 | 0 | 449 | 0.99758 | 7.2802
286.01 | 0.01 | 224.8 | 0.0075273 | 249.8 | 0.0049322 | 324.5 | 0.0017271 | 450 | 1.0045 | 7.3308
286.362 | 0.02 | 224.9 | 0.0061612 | 250 | 0.0071513 | 325 | 0.0053311 | 451 | 1.0144 | 7.4026
286.714 | 0.01 | 225 | 0.01041 | 250.2 | 0.0088194 | 325.5 | 0.0058358 | 452 | 1.0239 | 7.472
287.066 | 0 | 225.1 | 0.013142 | 250.4 | 0.010602 | 326 | 0.0052047 | 453 | 1.0298 | 7.5153
287.417 | 0.01 | 225.2 | 0.015874 | 250.6 | 0.01 | 326.5 | 0.0085106 | 454 | 1.0364 | 7.5635
287.769 | 0.01 | 225.3 | 0.018607 | 250.8 | 0.0069946 | 327 | 0.0021889 | 455 | 1.0465 | 7.6369
288.12 | 0.01 | 225.4 | 0.0093306 | 251 | 0.0034104 | 327.5 | 0.0042995 | 456 | 1.054 | 7.692
288.472 | 0.01 | 225.5 | 0.0079645 | 251.2 | 0.0017463 | 328 | 0.0021881 | 457 | 1.0686 | 7.7987
288.823 | 0.01 | 225.6 | 0.0065984 | 251.4 | 0.0013245 | 328.5 | 0.00061079 | 458 | 1.074 | 7.8379
289.174 | 0.02 | 225.7 | 0.0052322 | 251.6 | 0.0049228 | 329 | 0.0072271 | 459 | 1.0795 | 7.878
289.526 | 0.01 | 225.8 | 0.0088661 | 251.8 | 0.0071476 | 329.5 | 0.0050699 | 460 | 1.0874 | 7.9355
289.877 | 0.01 | 225.9 | 0.0075 | 252 | 0.0088168 | 330 | 0.0069354 | 461 | 1.097 | 8.0057
290.228 | 0 | 226 | 0.0061339 | 252.2 | 0.010601 | 330.5 | 0.0052928 | 462 | 1.0846 | 7.915
290.579 | 0.02 | 226.1 | 0.0068989 | 252.4 | 0.01 | 331 | 0.0027694 | 463 | 1.1047 | 8.0619
290.929 | 0.02 | 226.2 | 0.008265 | 252.6 | 0.0069907 | 331.5 | 0.0043813 | 464 | 1.1143 | 8.1319
291.28 | 0 | 226.3 | 0.0096311 | 252.8 | 0.0034028 | 332 | 0.00092692 | 465 | 1.1225 | 8.192
291.631 | 0.01 | 226.4 | 0.010997 | 253 | 0.0017309 | 332.5 | 0 | 466 | 1.1304 | 8.2498
291.981 | 0.01 | 226.5 | 0.014035 | 253.2 | 0.0013157 | 333 | 0.0011578 | 467 | 1.1361 | 8.2908
292.332 | 0 | 226.6 | 0.015403 | 253.4 | 0.0049428 | 333.5 | 0.011182 | 468 | 1.1402 | 8.3207
292.682 | 0.01 | 226.7 | 0.01677 | 253.6 | 0.005798 | 334 | 0.0088452 | 469 | 1.1468 | 8.3688
293.033 | 0 | 226.8 | 0.018137 | 253.8 | 0.0033353 | 334.5 | 0.00049639 | 470 | 1.15 | 8.3926
293.383 | 0 | 226.9 | 0.0088265 | 254 | 0.0060085 | 335 | 0.0040103 | 471 | 1.1502 | 8.3942
293.733 | 0.01 | 227 | 0.0074566 | 254.2 | 0.012744 | 335.5 | 0.0026292 | 472 | 1.1564 | 8.4395
294.083 | 0 | 227.1 | 0.0060868 | 254.4 | 0.0091551 | 336 | 0.0047295 | 473 | 1.1645 | 8.4984
294.433 | 0.01 | 227.2 | 0.0097268 | 254.6 | 0.0066605 | 336.5 | 0.0079624 | 474 | 1.1671 | 8.517
294.783 | 0.01 | 227.3 | 0.0083607 | 254.8 | 0.0042828 | 337 | 0.0013573 | 475 | 1.1673 | 8.5188
295.133 | 0.01 | 227.4 | 0.0069945 | 255 | 0.00020821 | 337.5 | -4.5207e-05 | 476 | 1.1662 | 8.5107
295.483 | 0 | 227.5 | 0.0056284 | 255.2 | 0 | 338 | 0 | 477 | 1.1694 | 8.534
295.832 | 0.01 | 227.6 | 0.0074103 | 255.4 | 0 | 338.5 | 0 | 478 | 1.1732 | 8.5621
296.182 | 0 | 227.7 | 0.0087789 | 255.6 | 0 | 339 | 0.00045823 | 479 | 1.1723 | 8.5555
296.531 | 0 | 227.8 | 0.010148 | 255.8 | 0 | 339.5 | 0.004652 | 480 | 1.1719 | 8.5524
296.881 | 0.01 | 227.9 | 0.011516 | 256 | 0 | 340 | 0.0022499 | 481 | 1.1768 | 8.5884
297.23 | 0 | 228 | 0.0054475 | 256.2 | 0 | 340.5 | 0.006385 | 482 | 1.1817 | 8.624
297.579 | 0 | 228.1 | 0.0040776 | 256.4 | 0 | 341 | 0.0095674 | 483 | 1.185 | 8.6481
297.929 | 0 | 228.2 | 0.0027078 | 256.6 | 0 | 341.5 | 0.0047271 | 484 | 1.1892 | 8.6784
298.278 | 0 | 228.3 | 0.010324 | 256.8 | 0.0012258 | 342 | 0.0075828 | 485 | 1.1973 | 8.738
298.627 | 0.01 | 228.4 | 0.011694 | 257 | 0.0050729 | 342.5 | 0.0048855 | 486 | 1.2065 | 8.8048
298.976 | 0.01 | 228.5 | 0.013064 | 257.2 | 0.0087234 | 343 | 0.0048123 | 487 | 1.2066 | 8.8056
299.325 | 0.01 | 228.6 | 0.014434 | 257.4 | 0.014624 | 343.5 | 0.0053697 | 488 | 1.2016 | 8.7688
299.673 | 0 | 228.7 | 0.0066681 | 257.6 | 0.014693 | 344 | 0.0021181 | 489 | 1.1988 | 8.7484
300.022 | 0 | 228.8 | 0.0066706 | 257.8 | 0.013336 | 344.5 | 0.0042262 | 490 | 1.1949 | 8.72
300.371 | 0.01 | 228.9 | 0.0066731 | 258 | 0.012418 | 345 | 0.0022303 | 491 | 1.2104 | 8.8329
300.719 | 0.01 | 229 | 0.0066756 | 258.2 | 0.0098668 | 345.5 | 0.0048354 | 492 | 1.21 | 8.8306
301.068 | 0.01 | 229.1 | 0.013333 | 258.4 | 0.0094909 | 346 | 0.0096028 | 493 | 1.1781 | 8.5973
301.416 | 0 | 229.2 | 0.013333 | 258.6 | 0.01126 | 346.5 | 0.0052601 | 494 | 1.1715 | 8.549
301.764 | 0.01 | 229.3 | 0.013333 | 258.8 | 0.015177 | 347 | 0.0029173 | 495 | 1.16 | 8.4653
302.113 | 0 | 229.4 | 0.010384 | 259 | 0.015959 | 347.5 | 0.0039717 | 496 | 1.1582 | 8.4525
302.461 | 0 | 229.5 | 0.011753 | 259.2 | 0.013333 | 348 | 0.0071347 | 497 | 1.1532 | 8.4161
302.809 | 0 | 229.6 | 0.013123 | 259.4 | 0.0098821 | 348.5 | 0.0099559 | 498 | 1.1451 | 8.3569
303.157 | 0 | 229.7 | 0.014493 | 259.6 | 0.0039665 | 349 | 0.01 | 499 | 1.1359 | 8.2894
303.505 | 0.01 | 229.8 | 0.01247 | 259.8 | 0.0048355 | 349.5 | 0.01 | 500 | 1.1299 | 8.2458
303.853 | 0 | 229.9 | 0.0111 | 260 | 0.0066667 | 350 | 0.014058 | 501 | 1.1219 | 8.1871
304.2 | 0 | 230 | 0.0097306 | 260.2 | 0.011682 | 350.5 | 0.015646 | 502 | 1.1144 | 8.1326
304.548 | 0 | 230.1 | 0.0083607 | 260.4 | 0.015767 | 351 | 0.010981 | 503 | 1.1089 | 8.0927
304.895 | 0 | 230.2 | 0.014676 | 260.6 | 0.013333 | 351.5 | 0.0056171 | 504 | 1.1041 | 8.0574
305.243 | 0.01 | 230.3 | 0.016046 | 260.8 | 0.015947 | 352 | 0.01303 | 505 | 1.1019 | 8.0417
305.59 | 0.01 | 230.4 | 0.017416 | 261 | 0.02257 | 352.5 | 0.011718 | 506 | 1.0949 | 7.9901
305.938 | 0 | 230.5 | 0.0095479 | 261.2 | 0.01617 | 353 | 0.011156 | 507 | 1.0897 | 7.9523
306.285 | 0 | 230.6 | 0.0081781 | 261.4 | 0.010177 | 353.5 | 0.014191 | 508 | 1.0808 | 7.8873
306.632 | 0.01 | 230.7 | 0.0068082 | 261.6 | 0.0096069 | 354 | 0.01797 | 509 | 1.0799 | 7.8808
306.979 | 0.01 | 230.8 | 0.0054384 | 261.8 | 0.0070855 | 354.5 | 0.024049 | 510 | 1.0764 | 7.8551
307.326 | 0.01 | 230.9 | 0.0090614 | 262 | 0.0082137 | 355 | 0.020751 | 511 | 1.0685 | 7.7977
307.673 | 0.01 | 231 | 0.0076877 | 262.2 | 0.010207 | 355.5 | 0.019961 | 512 | 1.0666 | 7.784
308.02 | 0.01 | 231.1 | 0.0063141 | 262.4 | 0.01 | 356 | 0.021159 | 513 | 1.0605 | 7.7392
308.367 | 0.01 | 231.2 | 0.0033333 | 262.6 | 0.01 | 356.5 | 0.027699 | 514 | 1.0558 | 7.7049
308.713 | 0 | 231.3 | 0.003332 | 262.8 | 0.01 | 357 | 0.02422 | 515 | 1.0537 | 7.6898
309.06 | 0 | 231.4 | 0.0033308 | 263 | 0.0069507 | 357.5 | 0.020159 | 516 | 1.0456 | 7.6308
309.407 | 0.01 | 231.5 | 0.0033295 | 263.2 | 0.0033271 | 358 | 0.025362 | 517 | 1.042 | 7.6043
309.753 | 0 | 231.6 | 0.0027991 | 263.4 | 0.0016397 | 358.5 | 0.029817 | 518 | 1.04 | 7.5897
310.099 | 0 | 231.7 | 0.0014292 | 263.6 | 6.4264e-05 | 359 | 0.03 | 519 | 1.0394 | 7.5851
310.446 | 0 | 231.8 | 5.9361e-05 | 263.8 | 0 | 359.5 | 0.033434 | 520 | 1.0348 | 7.552
310.792 | 0 | 231.9 | -0.0013105 | 264 | 0.0014239 | 360 | 0.034985 | 521 | 1.0376 | 7.572
311.138 | 0.01 | 232 | 0.0043525 | 264.2 | 0.0056723 | 360.5 | 0.034261 | 522 | 1.0347 | 7.551
311.484 | 0 | 232.1 | 0.0057236 | 264.4 | 0.010145 | 361 | 0.036265 | 523 | 1.0295 | 7.5131
311.83 | 0 | 232.2 | 0.0070948 | 264.6 | 0.015882 | 361.5 | 0.039735 | 524 | 1.0344 | 7.5492
312.176 | 0 | 232.3 | 0.0033333 | 264.8 | 0.018301 | 362 | 0.044111 | 525 | 1.0313 | 7.5261
312.522 | 0.01 | 232.4 | 0.0033333 | 265 | 0.02021 | 362.5 | 0.045766 | 526 | 1.0317 | 7.5291
312.867 | 0.01 | 232.5 | 0.0033333 | 265.2 | 0.016947 | 363 | 0.051025 | 527 | 1.0311 | 7.5244
313.213 | 0 | 232.6 | 0.0033333 | 265.4 | 0.013068 | 363.5 | 0.084472 | 528 | 1.0356 | 7.5574
313.558 | 0.01 | 232.7 | 0.010627 | 265.6 | 0.011375 | 364 | 0.40216 | 529 | 1.0378 | 7.5736
313.904 | 0 | 232.8 | 0.012001 | 265.8 | 0.011014 | 364.5 | 1.2821 | 530 | 1.04 | 7.5899
314.249 | 0 | 232.9 | 0.013375 | 266 | 0.015896 | 365 | 1.4938 | 531 | 1.0386 | 7.5791
314.595 | 0 | 233 | 0.014748 | 266.2 | 0.01172 | 365.5 | 0.78062 | 532 | 1.0436 | 7.6157
314.94 | 0.01 | 233.1 | 0.0066687 | 266.4 | 0.0068301 | 366 | 0.4019 | 533 | 1.0542 | 7.693
315.285 | 0 | 233.2 | 0.0066712 | 266.6 | 0.0040071 | 366.5 | 0.22072 | 534 | 1.0671 | 7.7875
315.63 | 0 | 233.3 | 0.0066737 | 266.8 | 0.0014996 | 367 | 0.085369 | 535 | 1.0689 | 7.8003
315.975 | 0 | 233.4 | 0.0061996 | 267 | 0.0049283 | 367.5 | 0.065977 | 536 | 1.0718 | 7.8218
316.32 | 0.01 | 233.5 | 0.0034524 | 267.2 | 0.0058001 | 368 | 0.063755 | 537 | 1.0883 | 7.9424
316.665 | 0 | 233.6 | 0.00070513 | 267.4 | 0.0033354 | 368.5 | 0.063611 | 538 | 1.106 | 8.0716
317.01 | 0 | 233.7 | -0.0020421 | 267.6 | 0.0082564 | 369 | 0.066021 | 539 | 1.1273 | 8.227
317.354 | 0.01 | 233.8 | 0 | 267.8 | 0.012823 | 369.5 | 0.071597 | 540 | 1.1628 | 8.486
317.699 | 0 | 233.9 | 0 | 268 | 0.009187 | 370 | 0.071209 | 541 | 1.2255 | 8.9436
318.043 | 0 | 234 | 0 | 268.2 | 0.0066667 | 370.5 | 0.073357 | 542 | 1.306 | 9.5313
318.388 | 0.01 | 234.1 | 0 | 268.4 | 0.011111 | 371 | 0.07497 | 543 | 1.3243 | 9.6644
318.732 | 0 | 234.2 | 0.0045559 | 268.6 | 0.011111 | 371.5 | 0.077683 | 544 | 1.3441 | 9.8092
319.076 | 0 | 234.3 | 0.0059295 | 268.8 | 0.0066667 | 372 | 0.081272 | 545 | 2.4085 | 17.5769
319.42 | 0 | 234.4 | 0.0073031 | 269 | 0.0074173 | 372.5 | 0.090968 | 546 | 6.1436 | 44.8352
319.765 | 0 | 234.5 | 0.0033333 | 269.2 | 0.007514 | 373 | 0.089841 | 547 | 3.3903 | 24.742
320.109 | 0.01 | 234.6 | 0.0033333 | 269.4 | 0.010033 | 373.5 | 0.087963 | 548 | 1.3088 | 9.5513
320.453 | 0 | 234.7 | 0.0033333 | 269.6 | 0.016208 | 374 | 0.09134 | 549 | 1.298 | 9.4723
320.796 | 0 | 234.8 | 0.0033333 | 269.8 | 0.015228 | 374.5 | 0.094071 | 550 | 1.3039 | 9.5156
321.14 | 0 | 234.9 | 0.010838 | 270 | 0.012135 | 375 | 0.092014 | 551 | 1.3058 | 9.5296
321.484 | 0 | 235 | 0.012212 | 270.2 | 0.0069245 | 375.5 | 0.097723 | 552 | 1.3092 | 9.5544
321.828 | 0 | 235.1 | 0.013585 | 270.4 | 0.0040259 | 376 | 0.10004 | 553 | 1.3184 | 9.6217
322.171 | 0.01 | 235.2 | 0.014959 | 270.6 | 0.0066687 | 376.5 | 0.1 | 554 | 1.3229 | 9.6542
322.515 | 0.01 | 235.3 | 0.0066691 | 270.8 | 0.0055587 | 377 | 0.1 | 555 | 1.3244 | 9.6653
322.858 | 0.01 | 235.4 | 0.0066716 | 271 | 0.0033359 | 377.5 | 0.1 | 556 | 1.3299 | 9.7056
323.201 | 0 | 235.5 | 0.0066742 | 271.2 | 0.0024614 | 378 | 0.10617 | 557 | 1.3288 | 9.6971
323.544 | 0 | 235.6 | 0.0095375 | 271.4 | -8.9202e-05 | 378.5 | 0.10998 | 558 | 1.3349 | 9.7418
323.888 | 0 | 235.7 | 0.0081639 | 271.6 | -0.00049922 | 379 | 0.11203 | 559 | 1.3501 | 9.8529
324.231 | 0 | 235.8 | 0.0067903 | 271.8 | 0.0037879 | 379.5 | 0.11833 | 560 | 1.3605 | 9.929
324.574 | 0 | 235.9 | 0.0054167 | 272 | 0.015581 | 380 | 0.12009 | 561 | 1.3727 | 10.0173
324.917 | 0.01 | 236 | 0.0076346 | 272.2 | 0.017938 | 380.5 | 0.1241 | 562 | 1.3786 | 10.061
325.259 | 0 | 236.1 | 0.0090107 | 272.4 | 0.01 | 381 | 0.12931 | 563 | 1.3852 | 10.109
325.602 | 0.01 | 236.2 | 0.010387 | 272.6 | 0.01087 | 381.5 | 0.13422 | 564 | 1.3926 | 10.1627
325.945 | 0 | 236.3 | 0.013425 | 272.8 | 0.0070798 | 382 | 0.13532 | 565 | 1.4045 | 10.2495
326.287 | 0.01 | 236.4 | 0.014801 | 273 | 0.006964 | 382.5 | 0.13398 | 566 | 1.4143 | 10.3214
326.63 | 0.01 | 236.5 | 0.016177 | 273.2 | 0.010983 | 383 | 0.13714 | 567 | 1.4204 | 10.3661
326.972 | 0 | 236.6 | 0.017554 | 273.4 | 0.014911 | 383.5 | 0.13996 | 568 | 1.4301 | 10.4368
327.315 | 0 | 236.7 | 0.013334 | 273.6 | 0.014413 | 384 | 0.14666 | 569 | 1.4352 | 10.4735
327.657 | 0.01 | 236.8 | 0.013335 | 273.8 | 0.007784 | 384.5 | 0.15066 | 570 | 1.4387 | 10.4996
327.999 | 0 | 236.9 | 0.013336 | 274 | 0.0082973 | 385 | 0.15721 | 571 | 1.445 | 10.5451
328.341 | 0 | 237 | 0.013338 | 274.2 | 0.0071455 | 385.5 | 0.15934 | 572 | 1.4511 | 10.59
328.683 | 0 | 237.1 | 0.016668 | 274.4 | 0.010419 | 386 | 0.15775 | 573 | 1.4485 | 10.5706
329.025 | 0.01 | 237.2 | 0.016669 | 274.6 | 0.016986 | 386.5 | 0.16869 | 574 | 1.4543 | 10.6134
329.367 | 0.01 | 237.3 | 0.01667 | 274.8 | 0.011922 | 387 | 0.17439 | 575 | 1.4568 | 10.6314
329.709 | 0 | 237.4 | 0.013334 | 275 | 0.0075337 | 387.5 | 0.17616 | 576 | 1.6852 | 12.2979
330.05 | 0.01 | 237.5 | 0.013335 | 275.2 | 0.008047 | 388 | 0.17968 | 577 | 2.2767 | 16.6147
330.392 | 0.01 | 237.6 | 0.013336 | 275.4 | 0.0065751 | 388.5 | 0.18083 | 578 | 1.8667 | 13.6231
330.734 | 0 | 237.7 | 0.013337 | 275.6 | 0.0092678 | 389 | 0.18813 | 579 | 2.1987 | 16.0458
331.075 | 0 | 237.8 | 0.0085399 | 275.8 | 0.013202 | 389.5 | 0.19695 | 580 | 1.7828 | 13.0105
331.417 | 0.01 | 237.9 | 0.0057851 | 276 | 0.016582 | 390 | 0.20828 | 581 | 1.4755 | 10.7676
331.758 | 0 | 238 | 0.0030303 | 276.2 | 0.011478 | 390.5 | 0.2154 | 582 | 1.4625 | 10.673
332.099 | 0 | 238.1 | 0.00027548 | 276.4 | 0.0077702 | 391 | 0.21182 | 583 | 1.4667 | 10.7035
332.44 | 0 | 238.2 | 0.0066667 | 276.6 | 0.0082864 | 391.5 | 0.21 | 584 | 1.4611 | 10.663
332.781 | 0 | 238.3 | 0.0066667 | 276.8 | 0.0066547 | 392 | 0.21359 | 585 | 1.4462 | 10.554
333.122 | 0 | 238.4 | 0.0066667 | 277 | 0.0090279 | 392.5 | 0.21958 | 586 | 1.4361 | 10.4805
333.463 | 0.01 | 238.5 | 0.0070386 | 277.2 | 0.0073309 | 393 | 0.22501 | 587 | 1.4519 | 10.5955
333.804 | 0.02 | 238.6 | 0.008416 | 277.4 | 0.0036582 | 393.5 | 0.23657 | 588 | 1.428 | 10.4215
334.145 | 0 | 238.7 | 0.0097934 | 277.6 | 0.0084448 | 394 | 0.24048 | 589 | 1.3807 | 10.0762
334.485 | 0 | 238.8 | 0.011171 | 277.8 | 0.01086 | 394.5 | 0.24509 | 590 | 1.3707 | 10.0028
334.826 | 0 | 238.9 | 0.01 | 278 | 0.0066638 | 395 | 0.24713 | 591 | 1.3598 | 9.9235
335.167 | 0.01 | 239 | 0.01 | 278.2 | 0.0088843 | 395.5 | 0.25758 | 592 | 1.3583 | 9.9124
335.507 | 0 | 239.1 | 0.01 | 278.4 | 0.013333 | 396 | 0.26113 | 593 | 1.3571 | 9.9037
335.847 | 0 | 239.2 | 0.01 | 278.6 | 0.012686 | 396.5 | 0.26731 | 594 | 1.3299 | 9.7053
336.188 | 0.01 | 239.3 | 0.01 | 278.8 | 0.012809 | 397 | 0.27546 | 595 | 1.2899 | 9.4137
336.528 | 0.01 | 239.4 | 0.01 | 279 | 0.012739 | 397.5 | 0.28341 | 596 | 1.2656 | 9.2361
336.868 | 0 | 239.5 | 0.01 | 279.2 | 0.0098743 | 398 | 0.28923 | 597 | 1.2446 | 9.0827
337.208 | 0 | 239.6 | 0.01 | 279.4 | 0.013041 | 398.5 | 0.29386 | 598 | 1.2345 | 9.0092
337.548 | 0 | 239.7 | 0.01 | 279.6 | 0.014319 | 399 | 0.29912 | 599 | 1.2346 | 9.0101
337.888 | 0 | 239.8 | 0.01 | 279.8 | 0.0080501 | 399.5 | 0.30813 | 600 | 1.2146 | 8.864
338.228 | 0 | 239.9 | 0.01 | 280 | 0.0068854 | 400 | 0.31863 | 601 | 1.1822 | 8.6277
338.567 | 0 | 240 | 0.01438 | 280.2 | 0.0073891 | 400.5 | 0.3247 | 602 | 1.1496 | 8.3896
338.907 | 0 | 240.1 | 0.015758 | 280.4 | 0.0042131 | 401 | 0.33188 | 603 | 1.1349 | 8.2823
339.246 | 0 | 240.2 | 0.017135 | 280.6 | 0.0033333 | 401.5 | 0.34333 | 604 | 1.1163 | 8.1465
339.586 | 0.01 | 240.3 | 0.013334 | 280.8 | 0.0021246 | 402 | 0.35208 | 605 | 1.0991 | 8.0212
339.925 | 0 | 240.4 | 0.013335 | 281 | 0.0012984 | 402.5 | 0.36617 | 606 | 1.0804 | 7.8844
340.265 | 0 | 240.5 | 0.013336 | 281.2 | 0.0051039 | 403 | 0.41985 | 607 | 1.0649 | 7.7718
340.604 | 0.01 | 240.6 | 0.013337 | 281.4 | 0.0059301 | 403.5 | 0.85484 | 608 | 1.061 | 7.7428
340.943 | 0.01 | 240.7 | 0.012635 | 281.6 | 0.0033333 | 404 | 2.9028 | 609 | 1.0667 | 7.7846
341.282 | 0.01 | 240.8 | 0.011256 | 281.8 | 0.0084293 | 404.5 | 5.0825 | 610 | 1.1633 | 8.4893
341.621 | 0 | 240.9 | 0.0098775 | 282 | 0.013102 | 405 | 2.7538 | 611 | 1.5183 | 11.0803
341.96 | 0.01 | 241 | 0.0084988 | 282.2 | 0.0092839 | 405.5 | 0.69536 | 612 | 1.4526 | 10.6007
342.299 | 0.01 | 241.1 | 0.0054488 | 282.4 | 0.0065439 | 406 | 0.42463 | 613 | 1.1533 | 8.4163
342.638 | 0 | 241.2 | 0.0040702 | 282.6 | 0.0034655 | 406.5 | 0.44236 | 614 | 1.0791 | 7.8753
342.976 | 0.01 | 241.3 | 0.0026915 | 282.8 | -0.00052251 | 407 | 0.55501 | 615 | 1.0253 | 7.4822
343.315 | 0 | 241.4 | 0.010359 | 283 | 0 | 407.5 | 0.74894 | 616 | 0.94797 | 6.9181
343.654 | 0.01 | 241.5 | 0.01174 | 283.2 | 0.0027678 | 408 | 0.66954 | 617 | 0.90196 | 6.5824
343.992 | 0 | 241.6 | 0.013122 | 283.4 | 0.0062804 | 408.5 | 0.48394 | 618 | 0.86281 | 6.2966
344.33 | 0 | 241.7 | 0.014503 | 283.6 | 0.0047897 | 409 | 0.46121 | 619 | 0.84135 | 6.14
344.669 | 0.01 | 241.8 | 0.015101 | 283.8 | 0.00557 | 409.5 | 0.4675 | 620 | 0.83474 | 6.0917
345.007 | 0 | 241.9 | 0.017864 | 284 | 0.011458 | 410 | 0.4724 | 621 | 0.81406 | 5.9409
345.345 | 0 | 242 | 0.020626 | 284.2 | 0.011282 | 410.5 | 0.48982 | 622 | 0.79377 | 5.7928
345.683 | 0.01 | 242.1 | 0.013278 | 284.4 | 0.0066707 | 411 | 0.49541 | 623 | 0.78061 | 5.6967
346.021 | 0.01 | 242.2 | 0.010516 | 284.6 | 0.0073191 | 411.5 | 0.49854 | 624 | 0.76305 | 5.5686
346.359 | 0.01 | 242.3 | 0.0077532 | 284.8 | 0.0071875 | 412 | 0.50243 | 625 | 0.73596 | 5.3709
346.697 | 0 | 242.4 | 0.0049908 | 285 | 0.0072759 | 412.5 | 0.50889 | 626 | 0.72201 | 5.2691
347.035 | 0 | 242.5 | 0.0094536 | 285.2 | 0.010161 | 413 | 0.5196 | 627 | 0.70741 | 5.1626
347.372 | 0.01 | 242.6 | 0.0080762 | 285.4 | 0.014113 | 413.5 | 0.52978 | 628 | 0.68604 | 5.0066
347.71 | 0 | 242.7 | 0.0066988 | 285.6 | 0.015644 | 414 | 0.53758 | 629 | 0.67036 | 4.8922
348.047 | 0.01 | 242.8 | 0.0053214 | 285.8 | 0.013333 | 414.5 | 0.54101 | 630 | 0.67346 | 4.9148
348.385 | 0.01 | 242.9 | 0.0077404 | 286 | 0.014444 | 415 | 0.54829 | 631 | 0.68405 | 4.9921
348.722 | 0.01 | 243 | 0.0091229 | 286.2 | 0.016667 | 415.5 | 0.55445 | 632 | 0.63344 | 4.6227
349.059 | 0.01 | 243.1 | 0.010505 | 286.4 | 0.014444 | 416 | 0.55172 | 633 | 0.58672 | 4.2818
349.397 | 0.01 | 243.2 | 0.0064499 | 286.6 | 0.013333 | 416.5 | 0.55363 | 634 | 0.56824 | 4.1469
349.734 | 0.01 | 243.3 | 0.0050661 | 286.8 | 0.0085354 | 417 | 0.5675 | 635 | 0.54917 | 4.0077
350.071 | 0.02 | 243.4 | 0.0036824 | 287 | 0.004419 | 417.5 | 0.57926 | 636 | 0.53414 | 3.898
350.408 | 0.01 | 243.5 | 0.0022986 | 287.2 | 0.0066649 | 418 | 0.58386 | 637 | 0.523 | 3.8168
350.744 | 0.02 | 243.6 | 0.0066667 | 287.4 | 0.0070559 | 418.5 | 0.58365 | 638 | 0.50884 | 3.7135
351.081 | 0.01 | 243.7 | 0.0066667 | 287.6 | 0.0092638 | 419 | 0.59572 | 639 | 0.49537 | 3.6151
351.418 | 0 | 243.8 | 0.0066667 | 287.8 | 0.010229 | 419.5 | 0.60773 | 640 | 0.4783 | 3.4905
351.755 | 0.01 | 243.9 | 0.0066667 | 288 | 0.01 | 420 | 0.6121 | 641 | 0.46465 | 3.3909
352.091 | 0.02 | 244 | 0.0033333 | 288.2 | 0.01 | 420.5 | 0.62264 | 642 | 0.45712 | 3.3359
352.428 | 0.01 | 244.1 | 0.0033333 | 288.4 | 0.01 | 421 | 0.62953 | 643 | 0.44614 | 3.2558
352.764 | 0.01 | 244.2 | 0.0033333 | 288.6 | 0.01 | 421.5 | 0.63 | 644 | 0.43358 | 3.1642
353.1 | 0.01 | 244.3 | 0.0029136 | 288.8 | 0.011477 | 422 | 0.63 | 645 | 0.42105 | 3.0727
353.437 | 0.02 | 244.4 | 0.0015298 | 289 | 0.015855 | 422.5 | 0.6403 | 646 | 0.41251 | 3.0105
353.773 | 0.01 | 244.5 | 0.00014607 | 289.2 | 0.014648 | 423 | 0.65245 | 647 | 0.39759 | 2.9015
354.109 | 0.02 | 244.6 | -0.0012377 | 289.4 | 0.01333 | 423.5 | 0.66026 | 648 | 0.38933 | 2.8413
354.445 | 0.03 | 244.7 | 0.0042835 | 289.6 | 0.012152 | 424 | 0.67112 | 649 | 0.38613 | 2.8179
354.781 | 0.02 | 244.8 | 0.0056673 | 289.8 | 0.0088735 | 424.5 | 0.67482 | 650 | 0.38506 | 2.8101
355.117 | 0.02 | 244.9 | 0.007051 | 290 | 0.0049145 | 425 | 0.68302 | 651 | 0.37632 | 2.7463
355.453 | 0.02 | 245 | 0.0033334 | 290.2 | 0.0055271 | 425.5 | 0.6899 | 652 | 0.36082 | 2.6332
355.788 | 0.02 | 245.1 | 0.0033347 | 290.4 | 0.01245 | 426 | 0.69386 | 653 | 0.35144 | 2.5648
356.124 | 0.02 | 245.2 | 0.003336 | 290.6 | 0.014871 | 426.5 | 0.68914 | 654 | 0.33594 | 2.4517
356.459 | 0.03 | 245.3 | 0.0033373 | 290.8 | 0.019651 | 427 | 0.68296 | 655 | 0.32094 | 2.3422
356.795 | 0.03 | 245.4 | 0.0066672 | 291 | 0.014095 | 427.5 | 0.70327 | 656 | 0.31163 | 2.2742
357.13 | 0.02 | 245.5 | 0.0066685 | 291.2 | 0.0079281 | 428 | 0.72436 | 657 | 0.31 | 2.2623
357.466 | 0.02 | 245.6 | 0.0066698 | 291.4 | 0.0082906 | 428.5 | 0.72579 | 658 | 0.30281 | 2.2099
357.801 | 0.02 | 245.7 | 0.0066711 | 291.6 | 0.0066546 | 429 | 0.72947 | 659 | 0.29488 | 2.152
358.136 | 0.03 | 245.8 | 0.0078336 | 291.8 | 0.0090834 | 429.5 | 0.74434 | 660 | 0.28391 | 2.0719
358.471 | 0.03 | 245.9 | 0.0092174 | 292 | 0.0072946 | 430 | 0.76004 | 661 | 0.27392 | 1.999
358.806 | 0.03 | 246 | 0.010601 | 292.2 | 0.0035458 | 430.5 | 0.77575 | 662 | 0.27767 | 2.0264
359.141 | 0.03 | 246.1 | 0.01 | 292.4 | 0.0051496 | 431 | 0.79106 | 663 | 0.27185 | 1.9839
359.476 | 0.03 | 246.2 | 0.01 | 292.6 | 0.0055535 | 431.5 | 0.79766 | 664 | 0.25975 | 1.8956
359.811 | 0.04 | 246.3 | 0.01 | 292.8 | 0.0033317 | 432 | 0.78774 | 665 | 0.25117 | 1.833
360.145 | 0.03 | 246.4 | 0.01 | 293 | 0.0030106 | 432.5 | 0.78604 | 666 | 0.24099 | 1.7587
360.48 | 0.04 | 246.5 | 0.0058079 | 293.2 | 0.00094513 | 433 | 0.82321 | 667 | 0.23989 | 1.7507
360.814 | 0.03 | 246.6 | 0.0044229 | 293.4 | 0.0026996 | 433.5 | 0.87369 | 668 | 0.23248 | 1.6966
361.149 | 0.04 | 246.7 | 0.0030379 | 293.6 | 0.0064333 | 434 | 0.96343 | 669 | 0.22254 | 1.624
361.483 | 0.04 | 246.8 | 0.0066667 | 293.8 | 0.0048429 | 434.5 | 1.4446 | 670 | 0.21802 | 1.591
361.817 | 0.04 | 246.9 | 0.0066667 | 294 | 0.0044444 | 435 | 4.6899 | 671 | 0.21687 | 1.5827
362.152 | 0.05 | 247 | 0.0066667 | 294.2 | 0.0066667 | 435.5 | 11.575 | 672 | 0.20947 | 1.5287
362.486 | 0.04 | 247.1 | 0.0066667 | 294.4 | 0.0069857 | 436 | 12.0735 | 673 | 0.20294 | 1.481
362.82 | 0.05 | 247.2 | 0.0033333 | 294.6 | 0.0090524 | 436.5 | 3.7578 | 674 | 0.19818 | 1.4462
363.154 | 0.05 | 247.3 | 0.0033333 | 294.8 | 0.01016 | 437 | 1.0942 | 675 | 0.19606 | 1.4308
363.488 | 0.07 | 247.4 | 0.0033333 | 295 | 0.01 | 437.5 | 0.88745 | 676 | 0.19344 | 1.4117
363.822 | 0.13 | 247.5 | 0.0033333 | 295.2 | 0.0066635 | 438 | 0.89394 | 677 | 0.18594 | 1.357
364.155 | 0.49 | 247.6 | 0.0022392 | 295.4 | 0.0045999 | 438.5 | 0.90609 | 678 | 0.177 | 1.2917
364.489 | 1.37 | 247.7 | 0.00085411 | 295.6 | 0.0066651 | 439 | 0.91899 | 679 | 0.17547 | 1.2805
364.823 | 1.86 | 247.8 | -0.00053093 | 295.8 | 0.0055528 | 439.5 | 0.93212 | 680 | 0.17384 | 1.2687
365.156 | 1.37 | 247.9 | 0.0035826 | 296 | 0.003331 | 440 | 0.94416 | 681 | 0.16824 | 1.2278
365.489 | 0.63 | 248 | 0.0049677 | 296.2 | 0.0026801 | 440.5 | 0.94984 | 682 | 0.16414 | 1.1978
365.823 | 0.43 | 248.1 | 0.0063527 | 296.4 | 0.00020834 | 441 | 0.94795 | 683 | 0.15837 | 1.1558
366.156 | 0.39 | 248.2 | 0.0077378 | 296.6 | 0.0029469 | 441.5 | 0.94563 | 684 | 0.15263 | 1.1139
366.489 | 0.22 | 248.3 | 0.0033341 | 296.8 | 0.0054199 | 442 | 0.95179 | 685 | 0.14911 | 1.0882
366.822 | 0.08 | 248.4 | 0.0033354 | 297 | 0.003335 | 442.5 | 0.95885 | 686 | 0.14534 | 1.0607
367.156 | 0.07 | 248.5 | 0.0033366 | 297.2 | 0.0030015 | 443 | 0.96319 | 687 | 0.15466 | 1.1287
367.488 | 0.06 | 248.6 | 0.0033379 | 297.4 | 0.0008978 | 443.5 | 0.97134 | 688 | 0.15623 | 1.1402
367.821 | 0.07 | 248.7 | 0.0019932 | 297.6 | -0.00017829 | 444 | 0.98243 | 689 | 0.14694 | 1.0723
368.154 | 0.06 | 248.8 | 0.0006069 | 297.8 | 0 | 444.5 | 0.97569 | 690 | 0.15237 | 1.112
368.487 | 0.06 | 248.9 | -0.00077942 | 298 | 0 | 445 | 0.98309 | 691 | 0.16 | 1.1677
368.82 | 0.07 | 249 | 0 | 298.2 | 0.0012162 | 445.5 | 0.98789 | 692 | 0.12857 | 0.9383
369.152 | 0.06 | 249.1 | 0 | 298.4 | 0.0050812 | 446 | 0.98091 | 693 | 0.13202 | 0.96348
369.485 | 0.08 | 249.2 | 0 | 298.6 | 0.007391 | 446.5 | 0.97997 | 694 | 0.13119 | 0.95738
369.817 | 0.07 | 249.3 | 0 | 298.8 | 0.0091452 | 447 | 0.98 | 695 | 0.12181 | 0.88894
370.149 | 0.07 | 249.4 | 0 | 299 | 0.010193 | 447.5 | 0.98 | 696 | 0.12964 | 0.94608
370.482 | 0.07 | 249.5 | 0 | 299.2 | 0.01 | 448 | 0.98943 | 697 | 0.12893 | 0.94087
370.814 | 0.08 | 249.6 | 0 | 299.4 | 0.0065796 | 448.5 | 0.99478 | 698 | 0.11526 | 0.84114
371.146 | 0.07 | 249.7 | 0.0035459 | 299.6 | 0.0032738 | 449 | 0.99799 | 699 | 0.11018 | 0.80407
371.478 | 0.08 | 249.8 | 0.0049322 | 299.8 | 0.0015169 | 449.5 | 0.99998 | 700 | 0.10873 | 0.79349
371.81 | 0.08 | 249.9 | 0.0063185 | 300 | 0.0010608 | 450 | 1.0037 | 701 | 0.10637 | 0.77626
372.142 | 0.08 | 250 | 0.0077049 | 300.2 | 0.0058835 | 450.5 | 1.0099 | 702 | 0.10768 | 0.78581
372.474 | 0.09 | 250.1 | 0.0074306 | 300.4 | 0.0076423 | 451 | 1.0139 | 703 | 0.10528 | 0.76828
372.805 | 0.1 | 250.2 | 0.0088194 | 300.6 | 0.0099585 | 451.5 | 1.0193 | 704 | 0.10448 | 0.76246
373.137 | 0.08 | 250.3 | 0.010208 | 300.8 | 0.010465 | 452 | 1.0238 | 705 | 0.10945 | 0.79872
373.469 | 0.09 | 250.4 | 0.011597 | 301 | 0.0087343 | 452.5 | 1.0285 | 706 | 0.13332 | 0.97295
373.8 | 0.09 | 250.5 | 0.01 | 301.2 | 0.004769 | 453 | 1.0301 | 707 | 0.14745 | 1.0761
374.132 | 0.09 | 250.6 | 0.01 | 301.4 | 0.0039675 | 453.5 | 1.0308 | 708 | 0.13006 | 0.94913
374.463 | 0.1 | 250.7 | 0.01 | 301.6 | 0.0066667 | 454 | 1.0376 | 709 | 0.13574 | 0.99064
374.794 | 0.09 | 250.8 | 0.0061856 | 301.8 | 0.0044437 | 454.5 | 1.0409 | 710 | 0.11344 | 0.82788
375.125 | 0.09 | 250.9 | 0.004798 | 302 | 0.0033301 | 455 | 1.0484 | 711 | 0.13878 | 1.0128
375.456 | 0.1 | 251 | 0.0034104 | 302.2 | 0.0020176 | 455.5 | 1.05 | 712 | 0.15875 | 1.1585
375.787 | 0.1 | 251.1 | 0.0020228 | 302.4 | -4.8908e-05 | 456 | 1.051 | 713 | 0.1101 | 0.80351
376.118 | 0.1 | 251.2 | 0.0023019 | 302.6 | 0 | 456.5 | 1.061 | 714 | 0.09246 | 0.67476
376.449 | 0.1 | 251.3 | 0.00091428 | 302.8 | 0 | 457 | 1.0706 | 715 | 0.082298 | 0.60059
376.78 | 0.1 | 251.4 | -0.00047332 | 303 | 0 | 457.5 | 1.0743 | 716 | 0.080145 | 0.58488
377.111 | 0.1 | 251.5 | 0.0035326 | 303.2 | 0.003113 | 458 | 1.0731 | 717 | 0.082139 | 0.59943
377.441 | 0.1 | 251.6 | 0.0049228 | 303.4 | 0.0068247 | 458.5 | 1.0747 | 718 | 0.07325 | 0.53457
377.772 | 0.1 | 251.7 | 0.0063129 | 303.6 | 0.0049761 | 459 | 1.0807 | 719 | 0.075185 | 0.54869
378.102 | 0.11 | 251.8 | 0.0077031 | 303.8 | 0.0031062 | 459.5 | 1.0831 | 720 | 0.079991 | 0.58376
378.433 | 0.11 | 251.9 | 0.0074266 | 304 | 0.0012124 | 460 | 1.0841 | 721 | 0.076695 | 0.5597
378.763 | 0.11 | 252 | 0.0088168 | 304.2 | -0.00063126 | 460.5 | 1.0949 | 722 | 0.073176 | 0.53402
379.093 | 0.11 | 252.1 | 0.010207 | 304.4 | 0 | 461 | 1.0963 | 723 | 0.063187 | 0.46113
379.424 | 0.12 | 252.2 | 0.011597 | 304.6 | 0 | 461.5 | 1.0998 | 724 | 0.066503 | 0.48533
379.754 | 0.12 | 252.3 | 0.01 | 304.8 | 0.0011335 | 462 | 1.0613 | 725 | 0.067297 | 0.49112
380.084 | 0.12 | 252.4 | 0.01 | 305 | 0.00484 | 462.5 | 1.0927 | 726 | 0.065436 | 0.47754
380.414 | 0.12 | 252.5 | 0.01 | 305.2 | 0.0071636 | 463 | 1.1115 | 727 | 0.064969 | 0.47413
380.744 | 0.13 | 252.6 | 0.0061806 | 305.4 | 0.0089315 | 463.5 | 1.11 | 728 | 0.065934 | 0.48118
381.073 | 0.13 | 252.7 | 0.0047917 | 305.6 | 0.0073289 | 464 | 1.1134 | 729 | 0.066399 | 0.48457
381.403 | 0.13 | 252.8 | 0.0034028 | 305.8 | 0.0036484 | 464.5 | 1.1195 | 730 | 0.057344 | 0.41849
381.733 | 0.14 | 252.9 | 0.0020139 | 306 | 0.0018805 | 465 | 1.121 | 731 | 0.05863 | 0.42787
382.062 | 0.13 | 253 | 0.0022851 | 306.2 | 0.0013676 | 465.5 | 1.1271 | 732 | 0.060869 | 0.44421
382.392 | 0.14 | 253.1 | 0.00089363 | 306.4 | 0.0049904 | 466 | 1.1302 | 733 | 0.059947 | 0.43748
382.721 | 0.13 | 253.2 | -0.00049783 | 306.6 | 0.0073167 | 466.5 | 1.1341 | 734 | 0.054835 | 0.40017
383.051 | 0.14 | 253.3 | 0.0035513 | 306.8 | 0.0090874 | 467 | 1.1353 | 735 | 0.055023 | 0.40155
383.38 | 0.14 | 253.4 | 0.0049428 | 307 | 0.010176 | 467.5 | 1.1388 | 736 | 0.059834 | 0.43666
383.709 | 0.14 | 253.5 | 0.0063343 | 307.2 | 0.01 | 468 | 1.14 | 737 | 0.056826 | 0.41471
384.038 | 0.15 | 253.6 | 0.0077257 | 307.4 | 0.01 | 468.5 | 1.1416 | 738 | 0.075425 | 0.55044
384.367 | 0.15 | 253.7 | 0.0033341 | 307.6 | 0.01 | 469 | 1.1486 | 739 | 0.07614 | 0.55566
384.696 | 0.15 | 253.8 | 0.0033353 | 307.8 | 0.01 | 469.5 | 1.15 | 740 | 0.053363 | 0.38943
385.025 | 0.16 | 253.9 | 0.0033366 | 308 | 0.01 | 470 | 1.15 | 741 | 0.05457 | 0.39824
385.354 | 0.16 | 254 | 0.0033379 | 308.2 | 0.01 | 470.5 | 1.15 | 742 | 0.054333 | 0.39651
385.682 | 0.16 | 254.1 | 0.011351 | 308.4 | 0.0069765 | 471 | 1.15 | 743 | 0.053327 | 0.38917
386.011 | 0.15 | 254.2 | 0.012744 | 308.6 | 0.0033013 | 471.5 | 1.1507 | 744 | 0.050782 | 0.3706
386.34 | 0.17 | 254.3 | 0.014136 | 308.8 | 0.0015278 | 472 | 1.1575 | 745 | 0.049162 | 0.35877
386.668 | 0.17 | 254.4 | 0.0066657 | 309 | 0.0012498 | 472.5 | 1.1611 | 746 | 0.054228 | 0.39575
386.997 | 0.18 | 254.5 | 0.0066631 | 309.2 | 0.0053506 | 473 | 1.1652 | 747 | 0.05107 | 0.3727
387.325 | 0.17 | 254.6 | 0.0066605 | 309.4 | 0.0061195 | 473.5 | 1.1672 | 748 | 0.04634 | 0.33818
387.653 | 0.18 | 254.7 | 0.006658 | 309.6 | 0.003336 | 474 | 1.166 | 749 | 0.045158 | 0.32956
387.981 | 0.18 | 254.8 | 0.0044867 | 309.8 | 0.0024002 | 474.5 | 1.1679 | 750 | 0.059114 | 0.43141
388.309 | 0.18 | 254.9 | 0.0017038 | 310 | 0.00032434 | 475 | 1.1679 | 751 | 0.083031 | 0.60595
388.637 | 0.18 | 255 | -0.0010792 | 310.2 | 0 | 475.5 | 1.1661 | 752 | 0.069532 | 0.50743
388.965 | 0.19 | 255.1 | 0 | 310.4 | 0 | 476 | 1.1664 | 753 | 0.049265 | 0.35952
389.293 | 0.19 | 255.2 | 0 | 310.6 | 0 | 476.5 | 1.1661 | 754 | 0.048425 | 0.3534
389.621 | 0.2 | 255.3 | 0 | 310.8 | 0.002781 | 477 | 1.17 | 755 | 0.048062 | 0.35075
389.949 | 0.21 | 255.4 | 0 | 311 | 0.0063391 | 477.5 | 1.172 | 756 | 0.041861 | 0.30549
390.276 | 0.21 | 255.5 | 0 | 311.2 | 0.004817 | 478 | 1.1737 | 757 | 0.039652 | 0.28937
390.604 | 0.22 | 255.6 | 0 | 311.4 | 0.0032563 | 478.5 | 1.174 | 758 | 0.037904 | 0.27662
390.932 | 0.21 | 255.7 | 0 | 311.6 | 0.001657 | 479 | 1.173 | 759 | 0.040118 | 0.29278
391.259 | 0.21 | 255.8 | 0 | 311.8 | -0.0003404 | 479.5 | 1.1701 | 760 | 0.037203 | 0.2715
391.586 | 0.21 | 255.9 | 0 | 312 | 0 | 480 | 1.1704 | 761 | 0.035521 | 0.25923
391.914 | 0.21 | 256 | 0 | 312.2 | 0.0029351 | 480.5 | 1.1753 | 762 | 0.040817 | 0.29787
392.241 | 0.22 | 256.1 | 0 | 312.4 | 0.0065703 | 481 | 1.1753 | 763 | 0.087998 | 0.64219
392.568 | 0.22 | 256.2 | 0 | 312.6 | 0.008355 | 481.5 | 1.1799 | 764 | 0.11148 | 0.81354
392.895 | 0.22 | 256.3 | 0 | 312.8 | 0.008713 | 482 | 1.1812 | 765 | 0.044227 | 0.32276
393.222 | 0.23 | 256.4 | 0 | 313 | 0.0047458 | 482.5 | 1.1841 | 766 | 0.036407 | 0.26569
393.549 | 0.24 | 256.5 | 0 | 313.2 | 0.0039378 | 483 | 1.1848 | 767 | 0.033091 | 0.24149
393.876 | 0.24 | 256.6 | 0 | 313.4 | 0.0066641 | 483.5 | 1.1861 | 768 | 0.034354 | 0.25071
394.202 | 0.24 | 256.7 | 0 | 313.6 | 0.0044423 | 484 | 1.1897 | 769 | 0.031484 | 0.22976
394.529 | 0.25 | 256.8 | 0 | 313.8 | 0.00333 | 484.5 | 1.1917 | 770 | 0.033391 | 0.24368
394.855 | 0.24 | 256.9 | 0.0036775 | 314 | 0.0019193 | 485 | 1.1956 | 771 | 0.032744 | 0.23896
395.182 | 0.25 | 257 | 0.0050729 | 314.2 | -0.00015551 | 485.5 | 1.2047 | 772 | 0.059179 | 0.43188
395.508 | 0.26 | 257.1 | 0.0064682 | 314.4 | 0 | 486 | 1.2091 | 773 | 0.06691 | 0.4883
395.835 | 0.26 | 257.2 | 0.0078636 | 314.6 | 0.0027558 | 486.5 | 1.2058 | 774 | 0.03326 | 0.24273
396.161 | 0.26 | 257.3 | 0.011838 | 314.8 | 0.0063034 | 487 | 1.2095 | 775 | 0.026429 | 0.19287
396.487 | 0.27 | 257.4 | 0.014624 | 315 | 0.0048055 | 487.5 | 1.2046 | 776 | 0.03073 | 0.22426
396.813 | 0.27 | 257.5 | 0.017409 | 315.2 | 0.0032609 | 488 | 1.2001 | 777 | 0.029652 | 0.21639
397.139 | 0.28 | 257.6 | 0.013333 | 315.4 | 0.0016667 | 488.5 | 1.2 | 778 | 0.023774 | 0.17349
397.465 | 0.28 | 257.7 | 0.013335 | 315.6 | -0.00033816 | 489 | 1.2 | 779 | 0.027803 | 0.2029
397.791 | 0.29 | 257.8 | 0.013336 | 315.8 | 0 | 489.5 | 1.1963 | 780 | 0.025642 | 0.18713
398.117 | 0.29 | 257.9 | 0.013337 | 316 | 0.0029469 | 490 | 1.1931 | 781 | 0.025511 | 0.18617
398.443 | 0.29 | 258 | 0.012655 | 316.2 | 0.0065942 | 490.5 | 1.1953 | 782 | 0.026511 | 0.19347
398.768 | 0.3 | 258.1 | 0.011261 | 316.4 | 0.0049034 | 491 | 1.2088 | 783 | 0.024048 | 0.1755
399.094 | 0.3 | 258.2 | 0.0098668 | 316.6 | 0.0031643 | 491.5 | 1.2269 | 784 | 0.028505 | 0.20802
399.42 | 0.3 | 258.3 | 0.0084727 | 316.8 | 0.0013768 | 492 | 1.2185 | 785 | 0.024674 | 0.18007
399.745 | 0.32 | 258.4 | 0.01 | 317 | 0.0010165 | 492.5 | 1.1846 | 786 | 0.025628 | 0.18703
400.07 | 0.32 | 258.5 | 0.01 | 317.2 | 0.0060945 | 493 | 1.1753 | 787 | 0.027887 | 0.20351
400.396 | 0.32 | 258.6 | 0.01 | 317.4 | 0.0047364 | 493.5 | 1.1742 | 788 | 0.025819 | 0.18842
400.721 | 0.33 | 258.7 | 0.01378 | 317.6 | 0.003325 | 494 | 1.1748 | 789 | 0.010017 | 0.073101
401.046 | 0.33 | 258.8 | 0.015177 | 317.8 | 0.0018619 | 494.5 | 1.1653 | 790 | 0.017191 | 0.12546
401.371 | 0.34 | 258.9 | 0.016574 | 318 | 0.0011746 | 495 | 1.1595 | 791 | 0.019529 | 0.14252
401.696 | 0.35 | 259 | 0.01797 | 318.2 | 0.0056083 | 495.5 | 1.1551 | 792 | 0.022992 | 0.16779
402.021 | 0.35 | 259.1 | 0.013333 | 318.4 | 0.0045763 | 496 | 1.1618 | 793 | 0.017485 | 0.1276
402.346 | 0.36 | 259.2 | 0.013333 | 318.6 | 0.0033363 | 496.5 | 1.1578 | 794 | 0.017477 | 0.12755
402.67 | 0.37 | 259.3 | 0.013333 | 318.8 | 0.0021914 | 497 | 1.1527 | 795 | 0.040085 | 0.29253
402.995 | 0.4 | 259.4 | 0.0095531 | 319 | 0.00010982 | 497.5 | 1.1492 | 796 | 0.032235 | 0.23524
403.32 | 0.49 | 259.5 | 0.0067598 | 319.2 | 0 | 498 | 1.1451 | 797 | 0.026178 | 0.19104
403.644 | 0.96 | 259.6 | 0.0039665 | 319.4 | 0 | 498.5 | 1.1411 | 798 | 0.023585 | 0.17212
403.969 | 2.44 | 259.7 | 0.0011732 | 319.6 | 0 | 499 | 1.1325 | 799 | 0.024294 | 0.17729
404.293 | 4.93 | 259.8 | 0.0066667 | 319.8 | 0.0030475 | 499.5 | 1.134 | 800 | 0.029383 | 0.21443
404.617 | 5.86 | 259.9 | 0.0066667 | 320 | 0.0067473 | 500 | 1.127 | 801 | 0.041791 | 0.30498
404.942 | 2.96 | 260 | 0.0066667 | 320.2 | 0.0049549 | 500.5 | 1.1287 | 802 | 0.054692 | 0.39914
405.266 | 0.73 | 260.1 | 0.0066667 | 320.4 | 0.0031036 | 501 | 1.1237 | 803 | 0.02469 | 0.18019
405.59 | 0.45 | 260.2 | 0.012793 | 320.6 | 0.0011877 | 501.5 | 1.1131 | 804 | 0.021166 | 0.15447
405.914 | 0.42 | 260.3 | 0.015587 | 320.8 | -8.9535e-05 | 502 | 1.1174 | 805 | 0.021523 | 0.15707
406.238 | 0.43 | 260.4 | 0.01838 | 321 | 0 | 502.5 | 1.1127 | 806 | 0.02031 | 0.14822
406.562 | 0.43 | 260.5 | 0.013333 | 321.2 | 0 | 503 | 1.1062 | 807 | 0.011282 | 0.082331
406.885 | 0.5 | 260.6 | 0.013333 | 321.4 | 0 | 503.5 | 1.1079
407.209 | 0.62 | 260.7 | 0.013333 | 321.6 | 0 | 504 | 1.1027
407.533 | 0.81 | 260.8 | 0.013333 | 321.8 | 0.0014619 | 504.5 | 1.1017
407.856 | 0.79 | 260.9 | 0.021173 | 322 | 0.0058406 | 505 | 1.1037
408.18 | 0.55 | 261 | 0.02257 | 322.2 | 0.0076393 | 505.5 | 1.1004
408.503 | 0.47 | 261.1 | 0.023966 | 322.4 | 0.0099935 | 506 | 1.0943
408.826 | 0.46 | 261.2 | 0.01297 | 322.6 | 0.010483 | 506.5 | 1.0899
409.15 | 0.46 | 261.3 | 0.011574 | 322.8 | 0.0086848 | 507 | 1.0944
409.473 | 0.47 | 261.4 | 0.010177 | 323 | 0.0045967 | 507.5 | 1.0847
409.796 | 0.47 | 261.5 | 0.0087803 | 323.2 | 0.0022368 | 508 | 1.0798
410.119 | 0.47 | 261.6 | 0.01142 | 323.4 | 0.00043246 | 508.5 | 1.0778
410.442 | 0.49 | 261.7 | 0.0086203 | 323.6 | -0.00034176 | 509 | 1.0776
410.765 | 0.5 | 261.8 | 0.0058205 | 323.8 | 0 | 509.5 | 1.0843
411.087 | 0.49 | 261.9 | 0.0068158 | 324 | 0 | 510 | 1.076
411.41 | 0.5 | 262 | 0.0082137 | 324.2 | 0 | 510.5 | 1.0688
411.733 | 0.5 | 262.1 | 0.0096117 | 324.4 | 0 | 511 | 1.0711
412.055 | 0.5 | 262.2 | 0.01101 | 324.6 | 0.0029608 | 511.5 | 1.0656
412.378 | 0.51 | 262.3 | 0.01 | 324.8 | 0.0066278 | 512 | 1.0658
412.7 | 0.51 | 262.4 | 0.01 | 325 | 0.0049187 | 512.5 | 1.0685
413.022 | 0.52 | 262.5 | 0.01 | 325.2 | 0.0044468 | 513 | 1.058
413.345 | 0.53 | 262.6 | 0.01 | 325.4 | 0.0066687 | 513.5 | 1.0549
413.667 | 0.53 | 262.7 | 0.01 | 325.6 | 0.0055583 | 514 | 1.0548
413.989 | 0.54 | 262.8 | 0.01 | 325.8 | 0.0033333 | 514.5 | 1.0577
414.311 | 0.54 | 262.9 | 0.01 | 326 | 0.0055546 | 515 | 1.0541
414.633 | 0.54 | 263 | 0.0061256 | 326.2 | 0.0067263 | 515.5 | 1.0494
414.955 | 0.55 | 263.1 | 0.0047264 | 326.4 | 0.0083106 | 516 | 1.0429
415.277 | 0.55 | 263.2 | 0.0033271 | 326.6 | 0.0088797 | 516.5 | 1.0446
415.598 | 0.56 | 263.3 | 0.0019279 | 326.8 | 0.0041813 | 517 | 1.0414
415.92 | 0.55 | 263.4 | 0.0021953 | 327 | 0.0023741 | 517.5 | 1.04
416.242 | 0.55 | 263.5 | 0.00079602 | 327.2 | 1.1352e-05 | 518 | 1.04
416.563 | 0.55 | 263.6 | -0.00060323 | 327.4 | 0.0030561 | 518.5 | 1.04
416.885 | 0.57 | 263.7 | 0 | 327.6 | 0.0056226 | 519 | 1.04
417.206 | 0.57 | 263.8 | 0 | 327.8 | 0.0033333 | 519.5 | 1.0381
417.527 | 0.58 | 263.9 | 0 | 328 | 0.0028363 | 520 | 1.0323
417.848 | 0.59 | 264 | 0 | 328.2 | 0.00039474 | 520.5 | 1.0341
418.17 | 0.58 | 264.1 | 0.0042717 | 328.4 | -0.00035575 | 521 | 1.0357
418.491 | 0.58 | 264.2 | 0.0056723 | 328.6 | 0.001194 | 521.5 | 1.043
418.812 | 0.59 | 264.3 | 0.0070728 | 328.8 | 0.0050439 | 522 | 1.0359
419.132 | 0.6 | 264.4 | 0.01028 | 329 | 0.0074123 | 522.5 | 1.0251
419.453 | 0.61 | 264.5 | 0.013081 | 329.2 | 0.0092251 | 523 | 1.0291
419.774 | 0.61 | 264.6 | 0.015882 | 329.4 | 0.0071979 | 523.5 | 1.0343
420.095 | 0.61 | 264.7 | 0.018683 | 329.6 | 0.0032602 | 524 | 1.0353
420.415 | 0.62 | 264.8 | 0.017409 | 329.8 | 0.0050425 | 524.5 | 1.0337
420.736 | 0.63 | 264.9 | 0.01881 | 330 | 0.0069066 | 525 | 1.0302
421.056 | 0.63 | 265 | 0.02021 | 330.2 | 0.0088569 | 525.5 | 1.03
421.377 | 0.63 | 265.1 | 0.021611 | 330.4 | 0.0073192 | 526 | 1.0338
421.697 | 0.63 | 265.2 | 0.015316 | 330.6 | 0.0036257 | 526.5 | 1.0313
422.017 | 0.63 | 265.3 | 0.013914 | 330.8 | 0.0018074 | 527 | 1.03
422.337 | 0.63 | 265.4 | 0.012512 | 331 | 0.0013411 | 527.5 | 1.0319
422.657 | 0.65 | 265.5 | 0.012777 | 331.2 | 0.0051595 | 528 | 1.0387
422.977 | 0.65 | 265.6 | 0.011375 | 331.4 | 0.006016 | 528.5 | 1.0361
423.297 | 0.66 | 265.7 | 0.0099735 | 331.6 | 0.003336 | 529 | 1.0372
423.617 | 0.66 | 265.8 | 0.0085717 | 331.8 | 0.0024354 | 529.5 | 1.0401
423.937 | 0.67 | 265.9 | 0.014496 | 332 | 0.00034539 | 530 | 1.04
424.257 | 0.68 | 266 | 0.015896 | 332.2 | 0 | 530.5 | 1.04
424.576 | 0.67 | 266.1 | 0.017297 | 332.4 | 0 | 531 | 1.0387
424.896 | 0.68 | 266.2 | 0.0096312 | 332.6 | 0 | 531.5 | 1.0369
425.215 | 0.69 | 266.3 | 0.0082306 | 332.8 | 0 | 532 | 1.0465
425.535 | 0.69 | 266.4 | 0.0068301 | 333 | 0 | 532.5 | 1.0472
425.854 | 0.69 | 266.5 | 0.0054295 | 333.2 | 0.0034734 | 533 | 1.0534
426.173 | 0.7 | 266.6 | 0.0047004 | 333.4 | 0.008146 | 533.5 | 1.0619
426.493 | 0.69 | 266.7 | 0.0018914 | 333.6 | 0.014018 | 534 | 1.0694
426.812 | 0.68 | 266.8 | -0.0009176 | 333.8 | 0.015142 | 534.5 | 1.07
427.131 | 0.68 | 266.9 | 0.0035251 | 334 | 0.0071261 | 535 | 1.0695
427.45 | 0.7 | 267 | 0.0049283 | 334.2 | 0.0042677 | 535.5 | 1.0671
427.769 | 0.72 | 267.1 | 0.0063315 | 334.4 | 0.00042225 | 536 | 1.0729
428.087 | 0.73 | 267.2 | 0.0077346 | 334.6 | 0 | 536.5 | 1.0754
428.406 | 0.73 | 267.3 | 0.0033341 | 334.8 | 0.0014728 | 537 | 1.0902
428.725 | 0.72 | 267.4 | 0.0033354 | 335 | 0.0058847 | 537.5 | 1.0994
429.043 | 0.73 | 267.5 | 0.0033367 | 335.2 | 0.0046733 | 538 | 1.105
429.362 | 0.74 | 267.6 | 0.010014 | 335.4 | 0.0033367 | 538.5 | 1.1137
429.68 | 0.75 | 267.7 | 0.011419 | 335.6 | 0.001933 | 539 | 1.1262
429.999 | 0.76 | 267.8 | 0.012823 | 335.8 | 0.0012089 | 539.5 | 1.1421
430.317 | 0.77 | 267.9 | 0.014228 | 336 | 0.0055762 | 540 | 1.163
430.635 | 0.78 | 268 | 0.0066667 | 336.2 | 0.0074034 | 540.5 | 1.1834
430.953 | 0.79 | 268.1 | 0.0066667 | 336.4 | 0.0097862 | 541 | 1.2202
431.271 | 0.8 | 268.2 | 0.0066667 | 336.6 | 0.0070001 | 541.5 | 1.273
431.589 | 0.8 | 268.3 | 0.0066667 | 336.8 | 0.0032222 | 542 | 1.3175
431.907 | 0.79 | 268.4 | 0.013333 | 337 | 0.0013922 | 542.5 | 1.3276
432.225 | 0.78 | 268.5 | 0.013333 | 337.2 | -0.00054248 | 543 | 1.321
432.543 | 0.78 | 268.6 | 0.013333 | 337.4 | 0 | 543.5 | 1.3242
432.861 | 0.81 | 268.7 | 0.0066667 | 337.6 | 0 | 544 | 1.3378
433.178 | 0.84 | 268.8 | 0.0066667 | 337.8 | 0 | 544.5 | 1.3703
433.496 | 0.87 | 268.9 | 0.0066667 | 338 | 0 | 545 | 1.8152
433.813 | 0.91 | 269 | 0.0066667 | 338.2 | 0 | 545.5 | 4.04
434.131 | 0.98 | 269.1 | 0.0089185 | 338.4 | 0 | 546 | 7.6004
434.448 | 1.14 | 269.2 | 0.007514 | 338.6 | 0 | 546.5 | 6.7905
434.765 | 2.03 | 269.3 | 0.0061096 | 338.8 | 0 | 547 | 2.0511
435.082 | 5.07 | 269.4 | 0.01059 | 339 | 0 | 547.5 | 1.3294
435.399 | 9.27 | 269.5 | 0.013399 | 339.2 | 0.0013747 | 548 | 1.3008
435.717 | 15.51 | 269.6 | 0.016208 | 339.4 | 0.0055976 | 548.5 | 1.2962
436.033 | 14.47 | 269.7 | 0.019017 | 339.6 | 0.0045799 | 549 | 1.2997
436.35 | 4.18 | 269.8 | 0.013333 | 339.8 | 0.0033364 | 549.5 | 1.2979
436.667 | 1.19 | 269.9 | 0.013333 | 340 | 0.0021114 | 550 | 1.3061
436.984 | 0.94 | 270 | 0.013333 | 340.2 | 0.0013019 | 550.5 | 1.3076
437.301 | 0.89 | 270.1 | 0.0097375 | 340.4 | 0.0053255 | 551 | 1.3032
437.617 | 0.89 | 270.2 | 0.0069245 | 340.6 | 0.0077165 | 551.5 | 1.3066
437.934 | 0.89 | 270.3 | 0.0041116 | 340.8 | 0.0095575 | 552 | 1.3085
438.25 | 0.9 | 270.4 | 0.0012986 | 341 | 0.010344 | 552.5 | 1.3126
438.566 | 0.91 | 270.5 | 0.0066674 | 341.2 | 0.0088004 | 553 | 1.3182
438.883 | 0.91 | 270.6 | 0.0066687 | 341.4 | 0.0049263 | 553.5 | 1.3246
439.199 | 0.93 | 270.7 | 0.00667 | 341.6 | 0.0040315 | 554 | 1.3238
439.515 | 0.93 | 270.8 | 0.0066713 | 341.8 | 0.0066667 | 554.5 | 1.3203
439.831 | 0.94 | 270.9 | 0.0033346 | 342 | 0.0075516 | 555 | 1.3266
440.147 | 0.95 | 271 | 0.0033359 | 342.2 | 0.00853 | 555.5 | 1.3264
440.463 | 0.95 | 271.1 | 0.0033373 | 342.4 | 0.005177 | 556 | 1.333
440.779 | 0.95 | 271.2 | 0.0027277 | 342.6 | 0.0041983 | 556.5 | 1.3304
441.095 | 0.95 | 271.3 | 0.0013192 | 342.8 | 0.0066643 | 557 | 1.3261
441.41 | 0.94 | 271.4 | -8.9202e-05 | 343 | 0.0044425 | 557.5 | 1.3299
441.726 | 0.95 | 271.5 | -0.0014977 | 343.2 | 0.0033301 | 558 | 1.3351
442.042 | 0.95 | 271.6 | 0 | 343.4 | 0.005554 | 558.5 | 1.3397
442.357 | 0.96 | 271.7 | 0 | 343.6 | 0.0055558 | 559 | 1.3549
442.672 | 0.96 | 271.8 | 0 | 343.8 | 0.0033355 | 559.5 | 1.3557
442.988 | 0.96 | 271.9 | 0.011364 | 344 | 0.0027625 | 560 | 1.3587
443.303 | 0.97 | 272 | 0.015581 | 344.2 | 0.00025641 | 560.5 | 1.3672
443.618 | 0.97 | 272.1 | 0.019798 | 344.4 | 0.0029936 | 561 | 1.3745
443.933 | 0.99 | 272.2 | 0.024016 | 344.6 | 0.0054977 | 561.5 | 1.3762
444.248 | 0.98 | 272.3 | 0.01 | 344.8 | 0.0033352 | 562 | 1.3797
444.563 | 0.97 | 272.4 | 0.01 | 345 | 0.0028775 | 562.5 | 1.38
444.878 | 0.98 | 272.5 | 0.01 | 345.2 | 0.0004783 | 563 | 1.3856
445.193 | 0.99 | 272.6 | 0.012714 | 345.4 | 0.0029241 | 563.5 | 1.3901
445.508 | 0.99 | 272.7 | 0.0098967 | 345.6 | 0.00655 | 564 | 1.3928
445.822 | 0.98 | 272.8 | 0.0070798 | 345.8 | 0.0083974 | 564.5 | 1.3948
446.137 | 0.98 | 272.9 | 0.0042629 | 346 | 0.010411 | 565 | 1.4044
446.452 | 0.98 | 273 | 0.0076103 | 346.2 | 0.01 | 565.5 | 1.4142
446.766 | 0.98 | 273.1 | 0.0090188 | 346.4 | 0.0068803 | 566 | 1.4163
447.08 | 0.98 | 273.2 | 0.010427 | 346.6 | 0.0036571 | 566.5 | 1.4124
447.395 | 0.98 | 273.3 | 0.013502 | 346.8 | 0.0018097 | 567 | 1.4231
447.709 | 0.98 | 273.4 | 0.014911 | 347 | 0.0011607 | 567.5 | 1.4258
448.023 | 0.99 | 273.5 | 0.016319 | 347.2 | 0.0057815 | 568 | 1.4299
448.337 | 1 | 273.6 | 0.017728 | 347.4 | 0.0046422 | 568.5 | 1.4347
448.651 | 0.99 | 273.7 | 0.0091925 | 347.6 | 0.00333 | 569 | 1.4323
448.965 | 1 | 273.8 | 0.007784 | 347.8 | 0.0055526 | 569.5 | 1.4385
449.279 | 1 | 273.9 | 0.0063756 | 348 | 0.0069243 | 570 | 1.4363
449.593 | 1 | 274 | 0.0099624 | 348.2 | 0.0089273 | 570.5 | 1.4414
449.906 | 1 | 274.1 | 0.008554 | 348.4 | 0.010136 | 571 | 1.4477
450.22 | 1.01 | 274.2 | 0.0071455 | 348.6 | 0.01 | 571.5 | 1.4458
450.534 | 1.01 | 274.3 | 0.0057371 | 348.8 | 0.01 | 572 | 1.4561
450.847 | 1.01 | 274.4 | 0.011352 | 349 | 0.01 | 572.5 | 1.4514
451.161 | 1.02 | 274.5 | 0.014169 | 349.2 | 0.01 | 573 | 1.4423
451.474 | 1.01 | 274.6 | 0.016986 | 349.4 | 0.01 | 573.5 | 1.4517
451.787 | 1.03 | 274.7 | 0.019803 | 349.6 | 0.01 | 574 | 1.4576
452.1 | 1.02 | 274.8 | 0.0086854 | 349.8 | 0.01337 | 574.5 | 1.4537
452.413 | 1.03 | 274.9 | 0.007277 | 350 | 0.01547 | 575 | 1.4518
452.727 | 1.03 | 275 | 0.0058686 | 350.2 | 0.013333 | 575.5 | 1.4649
453.04 | 1.03 | 275.1 | 0.0094554 | 350.4 | 0.014444 | 576 | 1.5594
453.352 | 1.03 | 275.2 | 0.008047 | 350.6 | 0.016664 | 576.5 | 2.0311
453.665 | 1.03 | 275.3 | 0.0066385 | 350.8 | 0.014442 | 577 | 2.6628
453.978 | 1.04 | 275.4 | 0.0052301 | 351 | 0.012032 | 577.5 | 2.136
454.291 | 1.04 | 275.5 | 0.0078567 | 351.2 | 0.0064688 | 578 | 1.601
454.603 | 1.04 | 275.6 | 0.0092678 | 351.4 | 0.0035674 | 578.5 | 1.8632
454.916 | 1.05 | 275.7 | 0.010679 | 351.6 | 0.0066667 | 579 | 2.4711
455.228 | 1.05 | 275.8 | 0.013757 | 351.8 | 0.010778 | 579.5 | 2.2618
455.541 | 1.05 | 275.9 | 0.015169 | 352 | 0.014981 | 580 | 1.5936
455.853 | 1.05 | 276 | 0.016582 | 352.2 | 0.013332 | 580.5 | 1.493
456.165 | 1.05 | 276.1 | 0.017994 | 352.4 | 0.012972 | 581 | 1.4717
456.477 | 1.06 | 276.2 | 0.0089266 | 352.6 | 0.010771 | 581.5 | 1.4617
456.79 | 1.07 | 276.3 | 0.0075141 | 352.8 | 0.0097616 | 582 | 1.4621
457.102 | 1.07 | 276.4 | 0.0061017 | 353 | 0.01 | 582.5 | 1.4637
457.414 | 1.08 | 276.5 | 0.0096948 | 353.2 | 0.013706 | 583 | 1.4672
457.725 | 1.07 | 276.6 | 0.0082864 | 353.4 | 0.015808 | 583.5 | 1.4691
458.037 | 1.07 | 276.7 | 0.0068779 | 353.6 | 0.013336 | 584 | 1.4616
458.349 | 1.08 | 276.8 | 0.0054695 | 353.8 | 0.015557 | 584.5 | 1.4527
458.661 | 1.07 | 276.9 | 0.0076168 | 354 | 0.016667 | 585 | 1.4495
458.972 | 1.08 | 277 | 0.0090279 | 354.2 | 0.021687 | 585.5 | 1.4364
459.284 | 1.09 | 277.1 | 0.010439 | 354.4 | 0.025893 | 586 | 1.4337
459.595 | 1.08 | 277.2 | 0.0064831 | 354.6 | 0.023333 | 586.5 | 1.4382
459.907 | 1.08 | 277.3 | 0.0050706 | 354.8 | 0.022649 | 587 | 1.4544
460.218 | 1.09 | 277.4 | 0.0036582 | 355 | 0.020074 | 587.5 | 1.4631
460.529 | 1.1 | 277.5 | 0.0022458 | 355.2 | 0.019529 | 588 | 1.431
460.84 | 1.09 | 277.6 | 0.010838 | 355.4 | 0.02 | 588.5 | 1.39
461.151 | 1.1 | 277.7 | 0.01225 | 355.6 | 0.02 | 589 | 1.3749
461.462 | 1.1 | 277.8 | 0.013663 | 355.8 | 0.02 | 589.5 | 1.3772
461.773 | 1.1 | 277.9 | 0.0066665 | 356 | 0.02 | 590 | 1.3732
462.084 | 1 | 278 | 0.0066638 | 356.2 | 0.023477 | 590.5 | 1.3616
462.395 | 1.1 | 278.1 | 0.0066612 | 356.4 | 0.026894 | 591 | 1.3575
462.706 | 1.11 | 278.2 | 0.0066585 | 356.6 | 0.028762 | 591.5 | 1.3603
463.016 | 1.11 | 278.3 | 0.013333 | 356.8 | 0.027312 | 592 | 1.3581
463.327 | 1.11 | 278.4 | 0.013333 | 357 | 0.023608 | 592.5 | 1.3564
463.638 | 1.11 | 278.5 | 0.013333 | 357.2 | 0.02174 | 593 | 1.362
463.948 | 1.11 | 278.6 | 0.013333 | 357.4 | 0.020041 | 593.5 | 1.3528
464.258 | 1.12 | 278.7 | 0.011393 | 357.6 | 0.02 | 594 | 1.3299
464.569 | 1.12 | 278.8 | 0.012809 | 357.8 | 0.021604 | 594.5 | 1.307
464.879 | 1.12 | 278.9 | 0.014226 | 358 | 0.026303 | 595 | 1.2869
465.189 | 1.12 | 279 | 0.012702 | 358.2 | 0.028177 | 595.5 | 1.276
465.499 | 1.13 | 279.1 | 0.011288 | 358.4 | 0.029907 | 596 | 1.2669
465.809 | 1.13 | 279.2 | 0.0098743 | 358.6 | 0.03 | 596.5 | 1.2539
466.119 | 1.13 | 279.3 | 0.0084606 | 358.8 | 0.03 | 597 | 1.2445
466.429 | 1.13 | 279.4 | 0.014625 | 359 | 0.03 | 597.5 | 1.2354
466.739 | 1.14 | 279.5 | 0.016038 | 359.2 | 0.03 | 598 | 1.2308
467.048 | 1.13 | 279.6 | 0.017452 | 359.4 | 0.031231 | 598.5 | 1.2374
467.358 | 1.14 | 279.7 | 0.0094665 | 359.6 | 0.035184 | 599 | 1.2338
467.668 | 1.14 | 279.8 | 0.0080501 | 359.8 | 0.03606 | 599.5 | 1.2327
467.977 | 1.14 | 279.9 | 0.0066336 | 360 | 0.033336 | 600 | 1.2109
468.286 | 1.14 | 280 | 0.0052172 | 360.2 | 0.035558 | 600.5 | 1.2003
468.596 | 1.14 | 280.1 | 0.0088055 | 360.4 | 0.035558 | 601 | 1.1836
468.905 | 1.15 | 280.2 | 0.0073891 | 360.6 | 0.033335 | 601.5 | 1.1628
469.214 | 1.15 | 280.3 | 0.0059726 | 360.8 | 0.034448 | 602 | 1.146
469.523 | 1.15 | 280.4 | 0.0033333 | 361 | 0.036669 | 602.5 | 1.14
469.832 | 1.15 | 280.5 | 0.0033333 | 361.2 | 0.037678 | 603 | 1.139
470.141 | 1.15 | 280.6 | 0.0033333 | 361.4 | 0.039785 | 603.5 | 1.1257
470.45 | 1.15 | 280.7 | 0.0033333 | 361.6 | 0.04 | 604 | 1.1141
470.759 | 1.15 | 280.8 | 0.0022285 | 361.8 | 0.041523 | 604.5 | 1.1092
471.068 | 1.15 | 280.9 | 0.00081209 | 362 | 0.046065 | 605 | 1.0979
471.377 | 1.15 | 281 | -0.00060434 | 362.2 | 0.044743 | 605.5 | 1.0903
471.685 | 1.15 | 281.1 | 0.0036874 | 362.4 | 0.044446 | 606 | 1.0804
471.994 | 1.16 | 281.2 | 0.0051039 | 362.6 | 0.046667 | 606.5 | 1.0704
472.302 | 1.16 | 281.3 | 0.0065203 | 362.8 | 0.047066 | 607 | 1.0619
472.611 | 1.16 | 281.4 | 0.0079367 | 363 | 0.049361 | 607.5 | 1.0625
472.919 | 1.17 | 281.5 | 0.0033333 | 363.2 | 0.056647 | 608 | 1.0604
473.227 | 1.16 | 281.6 | 0.0033333 | 363.4 | 0.069947 | 608.5 | 1.06
473.536 | 1.17 | 281.7 | 0.0033333 | 363.6 | 0.096747 | 609 | 1.0608
473.844 | 1.17 | 281.8 | 0.010269 | 363.8 | 0.16958 | 609.5 | 1.0793
474.152 | 1.16 | 281.9 | 0.011686 | 364 | 0.34225 | 610 | 1.1287
474.46 | 1.17 | 282 | 0.013102 | 364.2 | 0.69464 | 610.5 | 1.2818
474.768 | 1.17 | 282.1 | 0.014518 | 364.4 | 1.1044 | 611 | 1.5717
475.075 | 1.17 | 282.2 | 0.0066667 | 364.6 | 1.4676 | 611.5 | 1.7015
475.383 | 1.16 | 282.3 | 0.0066667 | 364.8 | 1.6947 | 612 | 1.4365
475.691 | 1.17 | 282.4 | 0.0066667 | 365 | 1.5335 | 612.5 | 1.2198
475.999 | 1.17 | 282.5 | 0.0062984 | 365.2 | 1.2532 | 613 | 1.1422
476.306 | 1.16 | 282.6 | 0.0034655 | 365.4 | 0.88391 | 613.5 | 1.0978
476.614 | 1.17 | 282.7 | 0.00063267 | 365.6 | 0.65408 | 614 | 1.0779
476.921 | 1.17 | 282.8 | -0.0022002 | 365.8 | 0.447 | 614.5 | 1.0617
477.228 | 1.17 | 282.9 | 0 | 366 | 0.4195 | 615 | 1.029
477.536 | 1.17 | 283 | 0 | 366.2 | 0.33919 | 615.5 | 0.98503
477.843 | 1.18 | 283.1 | 0 | 366.4 | 0.26529 | 616 | 0.94331
478.15 | 1.17 | 283.2 | 0.0034421 | 366.6 | 0.17833 | 616.5 | 0.91557
478.457 | 1.17 | 283.3 | 0.0048612 | 366.8 | 0.10763 | 617 | 0.90529
478.764 | 1.18 | 283.4 | 0.0062804 | 367 | 0.083396 | 617.5 | 0.88503
479.071 | 1.17 | 283.5 | 0.0076995 | 367.2 | 0.065077 | 618 | 0.86127
479.378 | 1.17 | 283.6 | 0.0033341 | 367.4 | 0.064992 | 618.5 | 0.84214
479.684 | 1.17 | 283.7 | 0.0033354 | 367.6 | 0.066668 | 619 | 0.84366
479.991 | 1.17 | 283.8 | 0.0033368 | 367.8 | 0.065558 | 619.5 | 0.83826
480.298 | 1.17 | 283.9 | 0.010038 | 368 | 0.063333 | 620 | 0.83652
480.604 | 1.18 | 284 | 0.011458 | 368.2 | 0.062372 | 620.5 | 0.82942
480.911 | 1.17 | 284.1 | 0.012879 | 368.4 | 0.061436 | 621 | 0.81223
481.217 | 1.18 | 284.2 | 0.014299 | 368.6 | 0.06503 | 621.5 | 0.80053
481.523 | 1.18 | 284.3 | 0.006668 | 368.8 | 0.065966 | 622 | 0.79505
481.83 | 1.18 | 284.4 | 0.0066707 | 369 | 0.063336 | 622.5 | 0.78575
482.136 | 1.18 | 284.5 | 0.0066734 | 369.2 | 0.06876 | 623 | 0.78119
482.442 | 1.19 | 284.6 | 0.0066761 | 369.4 | 0.072062 | 623.5 | 0.77489
482.748 | 1.18 | 284.7 | 0.008608 | 369.6 | 0.071737 | 624 | 0.76353
483.054 | 1.19 | 284.8 | 0.0071875 | 369.8 | 0.073358 | 624.5 | 0.75075
483.36 | 1.18 | 284.9 | 0.005767 | 370 | 0.070577 | 625 | 0.73185
483.666 | 1.19 | 285 | 0.0073201 | 370.2 | 0.06969 | 625.5 | 0.72527
483.971 | 1.19 | 285.1 | 0.0087405 | 370.4 | 0.071203 | 626 | 0.72056
484.277 | 1.19 | 285.2 | 0.010161 | 370.6 | 0.075112 | 626.5 | 0.7202
484.583 | 1.19 | 285.3 | 0.011581 | 370.8 | 0.076022 | 627 | 0.70746
484.888 | 1.2 | 285.4 | 0.014669 | 371 | 0.073333 | 627.5 | 0.69458
485.194 | 1.19 | 285.5 | 0.016089 | 371.2 | 0.075556 | 628 | 0.68498
485.499 | 1.21 | 285.6 | 0.017509 | 371.4 | 0.076777 | 628.5 | 0.67856
485.805 | 1.21 | 285.7 | 0.013333 | 371.6 | 0.078504 | 629 | 0.6679
486.11 | 1.21 | 285.8 | 0.013333 | 371.8 | 0.080509 | 629.5 | 0.66463
486.415 | 1.2 | 285.9 | 0.013333 | 372 | 0.08 | 630 | 0.67212
486.72 | 1.21 | 286 | 0.013333 | 372.2 | 0.083307 | 630.5 | 0.68364
487.025 | 1.21 | 286.1 | 0.016667 | 372.4 | 0.087909 | 631 | 0.69741
487.33 | 1.21 | 286.2 | 0.016667 | 372.6 | 0.093811 | 631.5 | 0.67111
487.635 | 1.2 | 286.3 | 0.016667 | 372.8 | 0.095081 | 632 | 0.62969
487.94 | 1.2 | 286.4 | 0.013333 | 373 | 0.087059 | 632.5 | 0.59951
488.245 | 1.2 | 286.5 | 0.013333 | 373.2 | 0.087382 | 633 | 0.58305
488.549 | 1.2 | 286.6 | 0.013333 | 373.4 | 0.086908 | 633.5 | 0.57759
488.854 | 1.2 | 286.7 | 0.013333 | 373.6 | 0.088649 | 634 | 0.569
489.159 | 1.2 | 286.8 | 0.0075568 | 373.8 | 0.090608 | 634.5 | 0.55812
489.463 | 1.2 | 286.9 | 0.0047159 | 374 | 0.09 | 635 | 0.54868
489.767 | 1.19 | 287 | 0.001875 | 374.2 | 0.093412 | 635.5 | 0.54069
490.072 | 1.19 | 287.1 | 0.0066662 | 374.4 | 0.095528 | 636 | 0.53504
490.376 | 1.2 | 287.2 | 0.0066649 | 374.6 | 0.093333 | 636.5 | 0.52668
490.68 | 1.19 | 287.3 | 0.0066635 | 374.8 | 0.092769 | 637 | 0.52468
490.984 | 1.21 | 287.4 | 0.0066622 | 375 | 0.090222 | 637.5 | 0.51765
491.288 | 1.22 | 287.5 | 0.007842 | 375.2 | 0.093051 | 638 | 0.5082
491.592 | 1.23 | 287.6 | 0.0092638 | 375.4 | 0.096932 | 638.5 | 0.50068
491.896 | 1.24 | 287.7 | 0.010686 | 375.6 | 0.098842 | 639 | 0.49873
492.2 | 1.2 | 287.8 | 0.01 | 375.8 | 0.10012 | 639.5 | 0.4867
492.504 | 1.18 | 287.9 | 0.01 | 376 | 0.1 | 640 | 0.47956
492.808 | 1.18 | 288 | 0.01 | 376.2 | 0.1 | 640.5 | 0.46863
493.111 | 1.17 | 288.1 | 0.01 | 376.4 | 0.1 | 641 | 0.46257
493.415 | 1.18 | 288.2 | 0.01 | 376.6 | 0.1 | 641.5 | 0.46274
493.718 | 1.17 | 288.3 | 0.01 | 376.8 | 0.1 | 642 | 0.45769
494.022 | 1.18 | 288.4 | 0.01 | 377 | 0.1 | 642.5 | 0.45091
494.325 | 1.17 | 288.5 | 0.01 | 377.2 | 0.1 | 643 | 0.4489
494.628 | 1.16 | 288.6 | 0.01 | 377.4 | 0.1 | 643.5 | 0.4386
494.931 | 1.16 | 288.7 | 0.01 | 377.6 | 0.1 | 644 | 0.43713
495.235 | 1.16 | 288.8 | 0.01 | 377.8 | 0.10301 | 644.5 | 0.42502
495.538 | 1.15 | 288.9 | 0.01443 | 378 | 0.10679 | 645 | 0.41994
495.841 | 1.16 | 289 | 0.015855 | 378.2 | 0.1087 | 645.5 | 0.41819
496.144 | 1.17 | 289.1 | 0.017279 | 378.4 | 0.11027 | 646 | 0.41104
496.446 | 1.15 | 289.2 | 0.013333 | 378.6 | 0.11 | 646.5 | 0.40832
496.749 | 1.16 | 289.3 | 0.013332 | 378.8 | 0.11 | 647 | 0.39445
497.052 | 1.15 | 289.4 | 0.01333 | 379 | 0.11114 | 647.5 | 0.38999
497.355 | 1.15 | 289.5 | 0.013329 | 379.2 | 0.11495 | 648 | 0.38802
497.657 | 1.15 | 289.6 | 0.012275 | 379.4 | 0.11742 | 648.5 | 0.38998
497.96 | 1.14 | 289.7 | 0.010852 | 379.6 | 0.11934 | 649 | 0.38757
498.262 | 1.15 | 289.8 | 0.0094292 | 379.8 | 0.12028 | 649.5 | 0.38085
498.564 | 1.14 | 289.9 | 0.006339 | 380 | 0.12 | 650 | 0.38524
498.867 | 1.13 | 290 | 0.0049145 | 380.2 | 0.12 | 650.5 | 0.38908
499.169 | 1.13 | 290.1 | 0.00349 | 380.4 | 0.12155 | 651 | 0.3773
499.471 | 1.14 | 290.2 | 0.0020655 | 380.6 | 0.12615 | 651.5 | 0.36258
499.773 | 1.13 | 290.3 | 0.011026 | 380.8 | 0.12808 | 652 | 0.35988
500.075 | 1.12 | 290.4 | 0.01245 | 381 | 0.12987 | 652.5 | 0.36
500.377 | 1.13 | 290.5 | 0.013875 | 381.2 | 0.13 | 653 | 0.35076
500.679 | 1.13 | 290.6 | 0.013942 | 381.4 | 0.1316 | 653.5 | 0.34356
500.981 | 1.13 | 290.7 | 0.016797 | 381.6 | 0.13632 | 654 | 0.3349
501.282 | 1.11 | 290.8 | 0.019651 | 381.8 | 0.13483 | 654.5 | 0.32937
501.584 | 1.11 | 290.9 | 0.022506 | 382 | 0.13445 | 655 | 0.31919
501.886 | 1.12 | 291 | 0.011316 | 382.2 | 0.13667 | 655.5 | 0.31427
502.187 | 1.12 | 291.1 | 0.0084618 | 382.4 | 0.13445 | 656 | 0.31062
502.489 | 1.11 | 291.2 | 0.0056074 | 382.6 | 0.13334 | 656.5 | 0.31
502.79 | 1.11 | 291.3 | 0.0097151 | 382.8 | 0.13556 | 657 | 0.31
503.091 | 1.1 | 291.4 | 0.0082906 | 383 | 0.13692 | 657.5 | 0.31
503.392 | 1.11 | 291.5 | 0.0068661 | 383.2 | 0.13893 | 658 | 0.3005
503.694 | 1.11 | 291.6 | 0.0054416 | 383.4 | 0.14015 | 658.5 | 0.29793
503.995 | 1.1 | 291.7 | 0.0076562 | 383.6 | 0.14 | 659 | 0.2927
504.296 | 1.1 | 291.8 | 0.0090834 | 383.8 | 0.14365 | 659.5 | 0.29401
504.597 | 1.1 | 291.9 | 0.010511 | 384 | 0.1472 | 660 | 0.28283
504.898 | 1.11 | 292 | 0.0064003 | 384.2 | 0.14913 | 660.5 | 0.27489
505.198 | 1.1 | 292.1 | 0.004973 | 384.4 | 0.15022 | 661 | 0.27394
505.499 | 1.1 | 292.2 | 0.0035458 | 384.6 | 0.15113 | 661.5 | 0.27295
505.8 | 1.1 | 292.3 | 0.0021186 | 384.8 | 0.15491 | 662 | 0.28006
506.1 | 1.09 | 292.4 | 0.0066657 | 385 | 0.1574 | 662.5 | 0.28
506.401 | 1.09 | 292.5 | 0.0066644 | 385.2 | 0.15933 | 663 | 0.27132
506.701 | 1.09 | 292.6 | 0.006663 | 385.4 | 0.16028 | 663.5 | 0.26424
507.002 | 1.1 | 292.7 | 0.0033331 | 385.6 | 0.1588 | 664 | 0.25711
507.302 | 1.09 | 292.8 | 0.0033317 | 385.8 | 0.15487 | 664.5 | 0.25791
507.602 | 1.08 | 292.9 | 0.0033304 | 386 | 0.15551 | 665 | 0.25264
507.903 | 1.08 | 293 | 0.003329 | 386.2 | 0.16287 | 665.5 | 0.24295
508.203 | 1.08 | 293.1 | 0.0023723 | 386.4 | 0.16593 | 666 | 0.24001
508.503 | 1.08 | 293.2 | 0.00094513 | 386.6 | 0.1711 | 666.5 | 0.24
508.803 | 1.07 | 293.3 | -0.00048208 | 386.8 | 0.17534 | 667 | 0.24133
509.103 | 1.08 | 293.4 | 0.0035762 | 387 | 0.17451 | 667.5 | 0.23835
509.402 | 1.09 | 293.5 | 0.0050048 | 387.2 | 0.17334 | 668 | 0.22978
509.702 | 1.08 | 293.6 | 0.0064333 | 387.4 | 0.17556 | 668.5 | 0.22931
510.002 | 1.07 | 293.7 | 0.0078619 | 387.6 | 0.17691 | 669 | 0.22204
510.301 | 1.08 | 293.8 | 0.0033333 | 387.8 | 0.17891 | 669.5 | 0.21626
510.601 | 1.06 | 293.9 | 0.0033333 | 388 | 0.18014 | 670 | 0.21553
510.9 | 1.07 | 294 | 0.0033333 | 388.2 | 0.18 | 670.5 | 0.22226
511.2 | 1.08 | 294.1 | 0.0066667 | 388.4 | 0.18 | 671 | 0.21453
511.499 | 1.06 | 294.2 | 0.0066667 | 388.6 | 0.18143 | 671.5 | 0.21383
511.798 | 1.06 | 294.3 | 0.0066667 | 388.8 | 0.18582 | 672 | 0.20663
512.098 | 1.07 | 294.4 | 0.0066667 | 389 | 0.18776 | 672.5 | 0.20795
512.397 | 1.07 | 294.5 | 0.0076238 | 389.2 | 0.1908 | 673 | 0.20163
512.696 | 1.07 | 294.6 | 0.0090524 | 389.4 | 0.19496 | 673.5 | 0.19924
512.995 | 1.05 | 294.7 | 0.010481 | 389.6 | 0.19897 | 674 | 0.19561
513.294 | 1.06 | 294.8 | 0.01 | 389.8 | 0.20546 | 674.5 | 0.19968
513.593 | 1.05 | 294.9 | 0.01 | 390 | 0.20831 | 675 | 0.19339
513.891 | 1.06 | 295 | 0.01 | 390.2 | 0.21106 | 675.5 | 0.19512
514.19 | 1.05 | 295.1 | 0.01 | 390.4 | 0.21522 | 676 | 0.19655
514.489 | 1.06 | 295.2 | 0.0057095 | 390.6 | 0.21612 | 676.5 | 0.18866
514.787 | 1.06 | 295.3 | 0.004281 | 390.8 | 0.21333 | 677 | 0.18881
515.086 | 1.05 | 295.4 | 0.0028524 | 391 | 0.21213 | 677.5 | 0.18035
515.384 | 1.05 | 295.5 | 0.0066664 | 391.2 | 0.21 | 678 | 0.17937
515.683 | 1.05 | 295.6 | 0.0066651 | 391.4 | 0.21 | 678.5 | 0.17127
515.981 | 1.04 | 295.7 | 0.0066637 | 391.6 | 0.21 | 679 | 0.17856
516.279 | 1.04 | 295.8 | 0.0066623 | 391.8 | 0.21 | 679.5 | 0.17657
516.577 | 1.05 | 295.9 | 0.0033324 | 392 | 0.21361 | 680 | 0.17461
516.875 | 1.04 | 296 | 0.003331 | 392.2 | 0.21715 | 680.5 | 0.17034
517.173 | 1.04 | 296.1 | 0.0033297 | 392.4 | 0.2191 | 681 | 0.16951
517.471 | 1.04 | 296.2 | 0.0030709 | 392.6 | 0.22021 | 681.5 | 0.16487
517.769 | 1.04 | 296.3 | 0.0016396 | 392.8 | 0.22114 | 682 | 0.16481
518.067 | 1.04 | 296.4 | 0.00020834 | 393 | 0.22494 | 682.5 | 0.16273
518.365 | 1.04 | 296.5 | -0.001223 | 393.2 | 0.22895 | 683 | 0.15986
518.662 | 1.04 | 296.6 | 0.0043162 | 393.4 | 0.23544 | 683.5 | 0.15252
518.96 | 1.04 | 296.7 | 0.0057475 | 393.6 | 0.23831 | 684 | 0.15088
519.258 | 1.04 | 296.8 | 0.0071787 | 393.8 | 0.23983 | 684.5 | 0.15449
519.555 | 1.04 | 296.9 | 0.0033336 | 394 | 0.24 | 685 | 0.1479
519.852 | 1.03 | 297 | 0.003335 | 394.2 | 0.24161 | 685.5 | 0.14494
520.15 | 1.03 | 297.1 | 0.0033363 | 394.4 | 0.24636 | 686 | 0.14392
520.447 | 1.04 | 297.2 | 0.0033377 | 394.6 | 0.24485 | 686.5 | 0.14716
520.744 | 1.03 | 297.3 | 0.0023305 | 394.8 | 0.24445 | 687 | 0.1568
521.041 | 1.03 | 297.4 | 0.0008978 | 395 | 0.24667 | 687.5 | 0.16001
521.338 | 1.05 | 297.5 | -0.00053486 | 395.2 | 0.25029 | 688 | 0.15389
521.635 | 1.04 | 297.6 | 0 | 395.4 | 0.25669 | 688.5 | 0.15481
521.932 | 1.04 | 297.7 | 0 | 395.6 | 0.25914 | 689 | 0.14592
522.229 | 1.03 | 297.8 | 0 | 395.8 | 0.26024 | 689.5 | 0.1401
522.526 | 1.02 | 297.9 | 0 | 396 | 0.26 | 690 | 0.14
522.823 | 1.03 | 298 | 0 | 396.2 | 0.26313 | 690.5 | 0.17702
523.119 | 1.03 | 298.1 | 0 | 396.4 | 0.26644 | 691 | 0.17155
523.416 | 1.03 | 298.2 | 0 | 396.6 | 0.2684 | 691.5 | 0.13144
523.712 | 1.04 | 298.3 | 0.0036485 | 396.8 | 0.27202 | 692 | 0.12869
524.009 | 1.03 | 298.4 | 0.0050812 | 397 | 0.2762 | 692.5 | 0.12558
524.305 | 1.04 | 298.5 | 0.0065138 | 397.2 | 0.27816 | 693 | 0.1339
524.601 | 1.03 | 298.6 | 0.0079465 | 397.4 | 0.28123 | 693.5 | 0.13659
524.898 | 1.03 | 298.7 | 0.0077125 | 397.6 | 0.2854 | 694 | 0.13165
525.194 | 1.03 | 298.8 | 0.0091452 | 397.8 | 0.28736 | 694.5 | 0.12532
525.49 | 1.03 | 298.9 | 0.010578 | 398 | 0.28987 | 695 | 0.11664
525.786 | 1.03 | 299 | 0.01 | 398.2 | 0.29047 | 695.5 | 0.12347
526.082 | 1.04 | 299.1 | 0.01 | 398.4 | 0.2914 | 696 | 0.12499
526.378 | 1.03 | 299.2 | 0.01 | 398.6 | 0.29575 | 696.5 | 0.14046
526.673 | 1.03 | 299.3 | 0.01 | 398.8 | 0.29771 | 697 | 0.12855
526.969 | 1.03 | 299.4 | 0.0055864 | 399 | 0.29964 | 697.5 | 0.11777
527.265 | 1.03 | 299.5 | 0.0041523 | 399.2 | 0.3 | 698 | 0.11747
527.56 | 1.03 | 299.6 | 0.0027183 | 399.4 | 0.30304 | 698.5 | 0.11054
527.856 | 1.04 | 299.7 | 0.0029509 | 399.6 | 0.3122 | 699 | 0.11
528.151 | 1.04 | 299.8 | 0.0015169 | 399.8 | 0.31613 | 699.5 | 0.11
528.447 | 1.04 | 299.9 | 8.2858e-05 | 400 | 0.31976 | 700 | 0.11
528.742 | 1.03 | 300 | -0.0013512 | 400.2 | 0.32 | 700.5 | 0.10619
529.037 | 1.04 | 300.1 | 0.0044508 | 400.4 | 0.32278 | 701 | 0.10631
529.333 | 1.04 | 300.2 | 0.0058835 | 400.6 | 0.32647 | 701.5 | 0.10661
529.628 | 1.04 | 300.3 | 0.0073161 | 400.8 | 0.32844 | 702 | 0.10749
529.923 | 1.04 | 300.4 | 0.0070877 | 401 | 0.33151 | 702.5 | 0.10894
530.218 | 1.04 | 300.5 | 0.0085231 | 401.2 | 0.3357 | 703 | 0.10283
530.513 | 1.04 | 300.6 | 0.0099585 | 401.4 | 0.3407 | 703.5 | 0.10406
530.808 | 1.04 | 300.7 | 0.011394 | 401.6 | 0.34591 | 704 | 0.10112
531.102 | 1.04 | 300.8 | 0.01 | 401.8 | 0.34827 | 704.5 | 0.10826
531.397 | 1.03 | 300.9 | 0.01 | 402 | 0.3519 | 705 | 0.10725
531.692 | 1.04 | 301 | 0.01 | 402.2 | 0.35609 | 705.5 | 0.11283
531.986 | 1.05 | 301.1 | 0.006203 | 402.4 | 0.36135 | 706 | 0.12942
532.281 | 1.05 | 301.2 | 0.004769 | 402.6 | 0.37036 | 706.5 | 0.15772
532.575 | 1.04 | 301.3 | 0.0033349 | 402.8 | 0.38466 | 707 | 0.15188
532.87 | 1.06 | 301.4 | 0.0019009 | 403 | 0.41704 | 707.5 | 0.13276
533.164 | 1.05 | 301.5 | 0.0066667 | 403.2 | 0.45785 | 708 | 0.12752
533.458 | 1.06 | 301.6 | 0.0066667 | 403.4 | 0.64483 | 708.5 | 0.1299
533.752 | 1.07 | 301.7 | 0.0066667 | 403.6 | 1.0316 | 709 | 0.14938
534.046 | 1.07 | 301.8 | 0.0033328 | 403.8 | 1.7645 | 709.5 | 0.12796
534.34 | 1.07 | 301.9 | 0.0033315 | 404 | 2.8711 | 710 | 0.10253
534.634 | 1.07 | 302 | 0.0033301 | 404.2 | 4.0728 | 710.5 | 0.10984
534.928 | 1.07 | 302.1 | 0.0033287 | 404.4 | 4.9747 | 711 | 0.10931
535.222 | 1.07 | 302.2 | 0.0020797 | 404.6 | 5.2883 | 711.5 | 0.1972
535.516 | 1.06 | 302.3 | 0.00064434 | 404.8 | 4.0283 | 712 | 0.14998
535.809 | 1.08 | 302.4 | -0.00079106 | 405 | 2.7928 | 712.5 | 0.12906
536.103 | 1.07 | 302.5 | 0 | 405.2 | 1.4404 | 713 | 0.10205
536.397 | 1.07 | 302.6 | 0 | 405.4 | 0.86096 | 713.5 | 0.0992
536.69 | 1.08 | 302.7 | 0 | 405.6 | 0.49429 | 714 | 0.091727
536.983 | 1.09 | 302.8 | 0 | 405.8 | 0.43287 | 714.5 | 0.086453
537.277 | 1.1 | 302.9 | 0 | 406 | 0.41443 | 715 | 0.084374
537.57 | 1.1 | 303 | 0 | 406.2 | 0.42657 | 715.5 | 0.076066
537.863 | 1.1 | 303.1 | 0 | 406.4 | 0.42917 | 716 | 0.078807
538.156 | 1.11 | 303.2 | 0.0039511 | 406.6 | 0.45216 | 716.5 | 0.085563
538.449 | 1.11 | 303.3 | 0.0053879 | 406.8 | 0.48946 | 717 | 0.084572
538.742 | 1.12 | 303.4 | 0.0068247 | 407 | 0.55034 | 717.5 | 0.076281
539.035 | 1.13 | 303.5 | 0.0082615 | 407.2 | 0.62522 | 718 | 0.073422
539.328 | 1.13 | 303.6 | 0.0033333 | 407.4 | 0.73471 | 718.5 | 0.070049
539.621 | 1.15 | 303.7 | 0.0033333 | 407.6 | 0.77475 | 719 | 0.075534
539.914 | 1.16 | 303.8 | 0.0033333 | 407.8 | 0.76446 | 719.5 | 0.079972
540.206 | 1.17 | 303.9 | 0.002652 | 408 | 0.65891 | 720 | 0.08
540.499 | 1.18 | 304 | 0.0012124 | 408.2 | 0.58526 | 720.5 | 0.08
540.792 | 1.2 | 304.1 | -0.00022712 | 408.4 | 0.49433 | 721 | 0.074642
541.084 | 1.23 | 304.2 | -0.0016667 | 408.6 | 0.46352 | 721.5 | 0.075443
541.376 | 1.25 | 304.3 | 0 | 408.8 | 0.46001 | 722 | 0.075828
541.669 | 1.3 | 304.4 | 0 | 409 | 0.46065 | 722.5 | 0.068257
541.961 | 1.32 | 304.5 | 0 | 409.2 | 0.46296 | 723 | 0.060378
542.253 | 1.33 | 304.6 | 0 | 409.4 | 0.46665 | 723.5 | 0.060926
542.545 | 1.33 | 304.7 | 0 | 409.6 | 0.46863 | 724 | 0.068553
542.837 | 1.32 | 304.8 | 0 | 409.8 | 0.47006 | 724.5 | 0.07003
543.129 | 1.32 | 304.9 | 0.0034005 | 410 | 0.47 | 725 | 0.06873
543.421 | 1.32 | 305 | 0.00484 | 410.2 | 0.47715 | 725.5 | 0.06313
543.713 | 1.33 | 305.1 | 0.0062796 | 410.4 | 0.48567 | 726 | 0.065291
544.005 | 1.34 | 305.2 | 0.0077191 | 410.6 | 0.494 | 726.5 | 0.067888
544.296 | 1.34 | 305.3 | 0.007492 | 410.8 | 0.49511 | 727 | 0.064445
544.588 | 1.37 | 305.4 | 0.0089315 | 411 | 0.49445 | 727.5 | 0.062573
544.88 | 1.44 | 305.5 | 0.010371 | 411.2 | 0.49667 | 728 | 0.066786
545.171 | 1.87 | 305.6 | 0.0065275 | 411.4 | 0.49713 | 728.5 | 0.068445
545.463 | 3.61 | 305.7 | 0.005088 | 411.6 | 0.49961 | 729 | 0.069554
545.754 | 5.82 | 305.8 | 0.0036484 | 411.8 | 0.50039 | 729.5 | 0.061199
546.045 | 7.75 | 305.9 | 0.0022089 | 412 | 0.50134 | 730 | 0.058701
546.337 | 10.37 | 306 | 0.002436 | 412.2 | 0.50558 | 730.5 | 0.052132
546.628 | 4.59 | 306.1 | 0.00099648 | 412.4 | 0.50757 | 731 | 0.062682
546.919 | 1.76 | 306.2 | -0.00044305 | 412.6 | 0.51012 | 731.5 | 0.061076
547.21 | 1.41 | 306.3 | 0.0035495 | 412.8 | 0.51433 | 732 | 0.064076
547.501 | 1.33 | 306.4 | 0.0049904 | 413 | 0.51895 | 732.5 | 0.057454
547.792 | 1.3 | 306.5 | 0.0064313 | 413.2 | 0.52551 | 733 | 0.062278
548.083 | 1.3 | 306.6 | 0.0078722 | 413.4 | 0.5284 | 733.5 | 0.060109
548.373 | 1.29 | 306.7 | 0.0076465 | 413.6 | 0.53118 | 734 | 0.054394
548.664 | 1.3 | 306.8 | 0.0090874 | 413.8 | 0.5354 | 734.5 | 0.050001
548.955 | 1.3 | 306.9 | 0.010528 | 414 | 0.53739 | 735 | 0.055355
549.245 | 1.3 | 307 | 0.01 | 414.2 | 0.53994 | 735.5 | 0.059714
549.536 | 1.3 | 307.1 | 0.01 | 414.4 | 0.5405 | 736 | 0.066257
549.826 | 1.29 | 307.2 | 0.01 | 414.6 | 0.54146 | 736.5 | 0.05353
550.116 | 1.32 | 307.3 | 0.01 | 414.8 | 0.54593 | 737 | 0.058844
550.407 | 1.31 | 307.4 | 0.01 | 415 | 0.54792 | 737.5 | 0.058104
550.697 | 1.3 | 307.5 | 0.01 | 415.2 | 0.55103 | 738 | 0.072381
550.987 | 1.3 | 307.6 | 0.01 | 415.4 | 0.55525 | 738.5 | 0.09579
551.277 | 1.31 | 307.7 | 0.01 | 415.6 | 0.55449 | 739 | 0.070114
551.567 | 1.3 | 307.8 | 0.01 | 415.8 | 0.55333 | 739.5 | 0.062517
551.857 | 1.32 | 307.9 | 0.01 | 416 | 0.55199 | 740 | 0.048056
552.147 | 1.3 | 308 | 0.01 | 416.2 | 0.54984 | 740.5 | 0.049515
552.437 | 1.31 | 308.1 | 0.01 | 416.4 | 0.55 | 741 | 0.056522
552.727 | 1.32 | 308.2 | 0.01 | 416.6 | 0.55624 | 741.5 | 0.057675
553.016 | 1.31 | 308.3 | 0.01 | 416.8 | 0.56292 | 742 | 0.052938
553.306 | 1.33 | 308.4 | 0.0061859 | 417 | 0.56692 | 742.5 | 0.052387
553.595 | 1.32 | 308.5 | 0.0047436 | 417.2 | 0.57266 | 743 | 0.050427
553.885 | 1.33 | 308.6 | 0.0033013 | 417.4 | 0.57636 | 743.5 | 0.057166
554.174 | 1.32 | 308.7 | 0.001859 | 417.6 | 0.58186 | 744 | 0.049408
554.464 | 1.32 | 308.8 | 0.0020833 | 417.8 | 0.58574 | 744.5 | 0.045772
554.753 | 1.32 | 308.9 | 0.00064103 | 418 | 0.58333 | 745 | 0.04769
555.042 | 1.33 | 309 | -0.00080128 | 418.2 | 0.5825 | 745.5 | 0.054023
555.331 | 1.33 | 309.1 | 0.0039097 | 418.4 | 0.58151 | 746 | 0.052189
555.62 | 1.32 | 309.2 | 0.0053506 | 418.6 | 0.58503 | 746.5 | 0.056472
555.909 | 1.34 | 309.3 | 0.0067915 | 418.8 | 0.58916 | 747 | 0.049658
556.198 | 1.33 | 309.4 | 0.0082325 | 419 | 0.59588 | 747.5 | 0.04708
556.487 | 1.33 | 309.5 | 0.0033346 | 419.2 | 0.60211 | 748 | 0.046027
556.776 | 1.33 | 309.6 | 0.003336 | 419.4 | 0.60699 | 748.5 | 0.045913
557.065 | 1.32 | 309.7 | 0.0033374 | 419.6 | 0.60896 | 749 | 0.045957
557.353 | 1.33 | 309.8 | 0.0026541 | 419.8 | 0.61017 | 749.5 | 0.043605
557.642 | 1.33 | 309.9 | 0.0012091 | 420 | 0.61114 | 750 | 0.052087
557.931 | 1.34 | 310 | -0.00023603 | 420.2 | 0.61498 | 750.5 | 0.081652
558.219 | 1.33 | 310.1 | 0 | 420.4 | 0.61909 | 751 | 0.072273
558.507 | 1.34 | 310.2 | 0 | 420.6 | 0.62576 | 751.5 | 0.095169
558.796 | 1.35 | 310.3 | 0 | 420.8 | 0.62859 | 752 | 0.066764
559.084 | 1.36 | 310.4 | 0 | 421 | 0.63001 | 752.5 | 0.046662
559.372 | 1.36 | 310.5 | 0 | 421.2 | 0.63 | 753 | 0.048984
559.66 | 1.35 | 310.6 | 0 | 421.4 | 0.63 | 753.5 | 0.052149
559.948 | 1.36 | 310.7 | 0 | 421.6 | 0.63 | 754 | 0.048259
560.236 | 1.36 | 310.8 | 0.0034489 | 421.8 | 0.63 | 754.5 | 0.044868
560.524 | 1.37 | 310.9 | 0.004894 | 422 | 0.63 | 755 | 0.047655
560.812 | 1.37 | 311 | 0.0063391 | 422.2 | 0.63 | 755.5 | 0.051663
561.1 | 1.38 | 311.1 | 0.0077842 | 422.4 | 0.6368 | 756 | 0.039771
561.388 | 1.37 | 311.2 | 0.0033333 | 422.6 | 0.64377 | 756.5 | 0.034148
561.675 | 1.38 | 311.3 | 0.0033333 | 422.8 | 0.6478 | 757 | 0.040451
561.963 | 1.38 | 311.4 | 0.0033333 | 423 | 0.65329 | 757.5 | 0.044355
562.251 | 1.38 | 311.5 | 0.0031021 | 423.2 | 0.65626 | 758 | 0.036239
562.538 | 1.38 | 311.6 | 0.001657 | 423.4 | 0.65828 | 758.5 | 0.033118
562.825 | 1.38 | 311.7 | 0.00021195 | 423.6 | 0.66196 | 759 | 0.049131
563.113 | 1.39 | 311.8 | -0.0012331 | 423.8 | 0.66619 | 759.5 | 0.038106
563.4 | 1.39 | 311.9 | 0 | 424 | 0.67161 | 760 | 0.038937
563.687 | 1.39 | 312 | 0 | 424.2 | 0.67555 | 760.5 | 0.034565
563.974 | 1.39 | 312.1 | 0 | 424.4 | 0.67334 | 761 | 0.033285
564.261 | 1.4 | 312.2 | 0.0036802 | 424.6 | 0.67556 | 761.5 | 0.038714
564.548 | 1.39 | 312.3 | 0.0051252 | 424.8 | 0.67782 | 762 | 0.049852
564.835 | 1.4 | 312.4 | 0.0065703 | 425 | 0.68327 | 762.5 | 0.033885
565.122 | 1.41 | 312.5 | 0.0080154 | 425.2 | 0.68797 | 763 | 0.058924
565.409 | 1.41 | 312.6 | 0.0078008 | 425.4 | 0.68956 | 763.5 | 0.17119
565.696 | 1.42 | 312.7 | 0.0092487 | 425.6 | 0.69037 | 764 | 0.11116
565.983 | 1.42 | 312.8 | 0.010697 | 425.8 | 0.69135 | 764.5 | 0.052093
566.269 | 1.41 | 312.9 | 0.0061937 | 426 | 0.69562 | 765 | 0.04231
566.556 | 1.41 | 313 | 0.0047458 | 426.2 | 0.69462 | 765.5 | 0.038278
566.842 | 1.42 | 313.1 | 0.0032979 | 426.4 | 0.69214 | 766 | 0.041317
567.129 | 1.43 | 313.2 | 0.0018501 | 426.6 | 0.68664 | 766.5 | 0.029627
567.415 | 1.42 | 313.3 | 0.0066654 | 426.8 | 0.68195 | 767 | 0.031739
567.701 | 1.43 | 313.4 | 0.0066641 | 427 | 0.68039 | 767.5 | 0.037906
567.987 | 1.43 | 313.5 | 0.0066627 | 427.2 | 0.68654 | 768 | 0.036492
568.273 | 1.43 | 313.6 | 0.0033328 | 427.4 | 0.69673 | 768.5 | 0.028664
568.56 | 1.44 | 313.7 | 0.0033314 | 427.6 | 0.7094 | 769 | 0.030661
568.846 | 1.43 | 313.8 | 0.00333 | 427.8 | 0.71889 | 769.5 | 0.035127
569.132 | 1.43 | 313.9 | 0.0033286 | 428 | 0.72576 | 770 | 0.029877
569.417 | 1.44 | 314 | 0.0019385 | 428.2 | 0.72843 | 770.5 | 0.035169
569.703 | 1.44 | 314.1 | 0.00049067 | 428.4 | 0.72892 | 771 | 0.025532
569.989 | 1.43 | 314.2 | -0.0009572 | 428.6 | 0.72363 | 771.5 | 0.03753
570.275 | 1.44 | 314.3 | 0 | 428.8 | 0.72513 | 772 | 0.046778
570.56 | 1.44 | 314.4 | 0 | 429 | 0.72837 | 772.5 | 0.09323
570.846 | 1.45 | 314.5 | 0 | 429.2 | 0.73492 | 773 | 0.071158
571.131 | 1.45 | 314.6 | 0.0034105 | 429.4 | 0.7412 | 773.5 | 0.036343
571.417 | 1.44 | 314.7 | 0.004857 | 429.6 | 0.74748 | 774 | 0.031516
571.702 | 1.45 | 314.8 | 0.0063034 | 429.8 | 0.75376 | 774.5 | 0.031921
571.987 | 1.46 | 314.9 | 0.0077499 | 430 | 0.76004 | 775 | 0.02492
572.272 | 1.46 | 315 | 0.0033333 | 430.2 | 0.76632 | 775.5 | 0.022445
572.558 | 1.45 | 315.1 | 0.0033333 | 430.4 | 0.77261 | 776 | 0.030205
572.843 | 1.44 | 315.2 | 0.0033333 | 430.6 | 0.7789 | 776.5 | 0.039541
573.128 | 1.44 | 315.3 | 0.0031159 | 430.8 | 0.78519 | 777 | 0.027786
573.413 | 1.45 | 315.4 | 0.0016667 | 431 | 0.79148 | 777.5 | 0.021628
573.698 | 1.46 | 315.5 | 0.00021739 | 431.2 | 0.7965 | 778 | 0.02
573.982 | 1.45 | 315.6 | -0.0012319 | 431.4 | 0.79869 | 778.5 | 0.029693
574.267 | 1.47 | 315.7 | 0 | 431.6 | 0.79723 | 779 | 0.028994
574.552 | 1.45 | 315.8 | 0 | 431.8 | 0.79335 | 779.5 | 0.024723
574.836 | 1.44 | 315.9 | 0 | 432 | 0.78759 | 780 | 0.026874
575.121 | 1.46 | 316 | 0.0036957 | 432.2 | 0.78229 | 780.5 | 0.025328
575.405 | 1.46 | 316.1 | 0.0051449 | 432.4 | 0.78058 | 781 | 0.025092
575.69 | 1.47 | 316.2 | 0.0065942 | 432.6 | 0.7897 | 781.5 | 0.026112
575.974 | 1.48 | 316.3 | 0.0080435 | 432.8 | 0.80441 | 782 | 0.027682
576.259 | 1.67 | 316.4 | 0.0033333 | 433 | 0.82317 | 782.5 | 0.025739
576.543 | 2.11 | 316.5 | 0.0033333 | 433.2 | 0.84205 | 783 | 0.019999
576.827 | 2.44 | 316.6 | 0.0033333 | 433.4 | 0.86208 | 783.5 | 0.026406
577.111 | 3.02 | 316.7 | 0.0028261 | 433.6 | 0.88483 | 784 | 0.029903
577.395 | 2.43 | 316.8 | 0.0013768 | 433.8 | 0.91258 | 784.5 | 0.029206
577.679 | 1.62 | 316.9 | -7.2464e-05 | 434 | 0.95234 | 785 | 0.022654
577.963 | 1.52 | 317 | -0.0015217 | 434.2 | 1.0254 | 785.5 | 0.022162
578.247 | 1.57 | 317.1 | 0.0046438 | 434.4 | 1.2161 | 786 | 0.0245
578.53 | 1.87 | 317.2 | 0.0060945 | 434.6 | 1.6351 | 786.5 | 0.030222
578.814 | 2.26 | 317.3 | 0.0075451 | 434.8 | 2.7826 | 787 | 0.020287
579.098 | 2.67 | 317.4 | 0.0033327 | 435 | 4.483 | 787.5 | 0.033151
579.381 | 2.8 | 317.5 | 0.0033313 | 435.2 | 6.8042 | 788 | 0.034278
579.665 | 1.78 | 317.6 | 0.0033299 | 435.4 | 9.8374 | 788.5 | 0.010029
579.948 | 1.54 | 317.7 | 0.003314 | 435.6 | 13.2495 | 789 | 0.010022
580.232 | 1.49 | 317.8 | 0.0018619 | 435.8 | 14.4188 | 789.5 | 0.01
580.515 | 1.5 | 317.9 | 0.00040981 | 436 | 13.3994 | 790 | 0.022908
580.798 | 1.48 | 318 | -0.0010423 | 436.2 | 8.4022 | 790.5 | 0.018667
581.081 | 1.47 | 318.1 | 0.0041562 | 436.4 | 5.2163 | 791 | 0.014074
581.364 | 1.45 | 318.2 | 0.0056083 | 436.6 | 2.2609 | 791.5 | 0.025846
581.648 | 1.47 | 318.3 | 0.0070603 | 436.8 | 1.4236 | 792 | 0.01903
581.931 | 1.46 | 318.4 | 0.0033335 | 437 | 0.95449 | 792.5 | 0.024099
582.213 | 1.46 | 318.5 | 0.0033349 | 437.2 | 0.90446 | 793 | 0.015606
582.496 | 1.47 | 318.6 | 0.0033363 | 437.4 | 0.88229 | 793.5 | 0.012749
582.779 | 1.46 | 318.7 | 0.0033377 | 437.6 | 0.88797 | 794 | 0.022902
583.062 | 1.47 | 318.8 | 0.002345 | 437.8 | 0.89 | 794.5 | 0.016781
583.344 | 1.47 | 318.9 | 0.00089147 | 438 | 0.89345 | 795 | 0.062012
583.627 | 1.47 | 319 | -0.00056202 | 438.2 | 0.89836 | 795.5 | 0.041463
583.909 | 1.46 | 319.1 | 0 | 438.4 | 0.90475 | 796 | 0.023634
584.192 | 1.46 | 319.2 | 0 | 438.6 | 0.90796 | 796.5 | 0.031608
584.474 | 1.46 | 319.3 | 0 | 438.8 | 0.91212 | 797 | 0.020166
584.757 | 1.44 | 319.4 | 0 | 439 | 0.92037 | 797.5 | 0.026759
585.039 | 1.46 | 319.5 | 0 | 439.2 | 0.92447 | 798 | 0.024361
585.321 | 1.44 | 319.6 | 0 | 439.4 | 0.92969 | 798.5 | 0.019635
585.603 | 1.43 | 319.7 | 0 | 439.6 | 0.9346 | 799 | 0.021111
585.885 | 1.44 | 319.8 | 0.0038459 | 439.8 | 0.93876 | 799.5 | 0.032135
586.167 | 1.43 | 319.9 | 0.0052966 | 440 | 0.94535 | 800 | 0.035468
586.449 | 1.43 | 320 | 0.0067473 | 440.2 | 0.94837 | 800.5 | 0.020545
586.731 | 1.45 | 320.1 | 0.0081979 | 440.4 | 0.94992 | 801 | 0.049233
587.013 | 1.45 | 320.2 | 0.0033333 | 440.6 | 0.95 | 801.5 | 0.055594
587.295 | 1.47 | 320.3 | 0.0033333 | 440.8 | 0.95 | 802 | 0.083741
587.576 | 1.47 | 320.4 | 0.0033333 | 441 | 0.94886 | 802.5 | 0.024743
587.858 | 1.44 | 320.5 | 0.002644 | 441.2 | 0.945 | 803 | 0.02602
588.139 | 1.42 | 320.6 | 0.0011877 | 441.4 | 0.94397 | 803.5 | 0.023309
588.421 | 1.4 | 320.7 | -0.00026861 | 441.6 | 0.94667 | 804 | 0.020822
588.702 | 1.37 | 320.8 | 0 | 441.8 | 0.94798 | 804.5 | 0.019368
588.984 | 1.38 | 320.9 | 0 | 442 | 0.95156 | 805 | 0.025813
589.265 | 1.37 | 321 | 0 | 442.2 | 0.95584 | 805.5 | 0.019388
589.546 | 1.38 | 321.1 | 0 | 442.4 | 0.9579 | 806 | 0.023385
589.827 | 1.38 | 321.2 | 0 | 442.6 | 0.95982 | 806.5 | 0.018158
590.108 | 1.37 | 321.3 | 0 | 442.8 | 0.96 | 807 | 0.013065
590.389 | 1.36 | 321.4 | 0 | 443 | 0.96288 | 807.5 | 0.0026221
590.67 | 1.36 | 321.5 | 0 | 443.2 | 0.9667
590.951 | 1.36 | 321.6 | 0 | 443.4 | 0.96876
591.232 | 1.35 | 321.7 | 0 | 443.6 | 0.97348
591.513 | 1.37 | 321.8 | 0 | 443.8 | 0.98244
591.794 | 1.35 | 321.9 | 0.0043857 | 444 | 0.98311
592.074 | 1.36 | 322 | 0.0058406 | 444.2 | 0.98175
592.355 | 1.36 | 322.1 | 0.0072955 | 444.4 | 0.97517
592.635 | 1.35 | 322.2 | 0.0070837 | 444.6 | 0.97511
592.916 | 1.37 | 322.3 | 0.0085386 | 444.8 | 0.97801
593.196 | 1.36 | 322.4 | 0.0099935 | 445 | 0.98387
593.476 | 1.35 | 322.5 | 0.011448 | 445.2 | 0.9874
593.757 | 1.35 | 322.6 | 0.01 | 445.4 | 0.98995
594.037 | 1.32 | 322.7 | 0.01 | 445.6 | 0.98678
594.317 | 1.32 | 322.8 | 0.01 | 445.8 | 0.98258
594.597 | 1.3 | 322.9 | 0.0060544 | 446 | 0.98051
594.877 | 1.29 | 323 | 0.0045967 | 446.2 | 0.97964
595.157 | 1.28 | 323.1 | 0.003139 | 446.4 | 0.98
595.437 | 1.28 | 323.2 | 0.0016812 | 446.6 | 0.98
595.716 | 1.27 | 323.3 | 0.0018902 | 446.8 | 0.98
595.996 | 1.27 | 323.4 | 0.00043246 | 447 | 0.98
596.276 | 1.26 | 323.5 | -0.0010253 | 447.2 | 0.98
596.555 | 1.25 | 323.6 | 0 | 447.4 | 0.98
596.835 | 1.25 | 323.7 | 0 | 447.6 | 0.98
597.114 | 1.24 | 323.8 | 0 | 447.8 | 0.98372
597.394 | 1.24 | 323.9 | 0 | 448 | 0.98893
597.673 | 1.23 | 324 | 0 | 448.2 | 0.99564
597.952 | 1.23 | 324.1 | 0 | 448.4 | 0.99516
598.232 | 1.23 | 324.2 | 0 | 448.6 | 0.99444
598.511 | 1.24 | 324.3 | 0 | 448.8 | 0.99667
598.79 | 1.24 | 324.4 | 0 | 449 | 0.99757
599.069 | 1.23 | 324.5 | 0 | 449.2 | 0.99974
599.348 | 1.23 | 324.6 | 0.0037123 | 449.4 | 1
599.627 | 1.24 | 324.7 | 0.0051701 | 449.6 | 1
599.905 | 1.21 | 324.8 | 0.0066278 | 449.8 | 1
600.184 | 1.2 | 324.9 | 0.0080855 | 450 | 1.0038
600.463 | 1.2 | 325 | 0.0033345 | 450.2 | 1.0075
600.742 | 1.2 | 325.1 | 0.0033359 | 450.4 | 1.0095
601.02 | 1.18 | 325.2 | 0.0033374 | 450.6 | 1.0104
601.299 | 1.17 | 325.3 | 0.0066672 | 450.8 | 1.0114
601.577 | 1.16 | 325.4 | 0.0066687 | 451 | 1.0158
601.855 | 1.15 | 325.5 | 0.0066701 | 451.2 | 1.0147
602.134 | 1.14 | 325.6 | 0.0066715 | 451.4 | 1.0157
602.412 | 1.14 | 325.7 | 0.0033333 | 451.6 | 1.022
602.69 | 1.14 | 325.8 | 0.0033333 | 451.8 | 1.0219
602.968 | 1.14 | 325.9 | 0.0033333 | 452 | 1.0234
603.246 | 1.14 | 326 | 0.0066659 | 452.2 | 1.0261
603.524 | 1.12 | 326.1 | 0.0066645 | 452.4 | 1.0271
603.802 | 1.12 | 326.2 | 0.006663 | 452.6 | 1.0296
604.08 | 1.11 | 326.3 | 0.0068515 | 452.8 | 1.0304
604.358 | 1.11 | 326.4 | 0.0083106 | 453 | 1.03
604.636 | 1.11 | 326.5 | 0.0097698 | 453.2 | 1.03
604.913 | 1.1 | 326.6 | 0.011229 | 453.4 | 1.03
605.191 | 1.09 | 326.7 | 0.0056404 | 453.6 | 1.0313
605.468 | 1.09 | 326.8 | 0.0041813 | 453.8 | 1.0355
605.746 | 1.09 | 326.9 | 0.0027221 | 454 | 1.0376
606.023 | 1.08 | 327 | 0.0029296 | 454.2 | 1.0396
606.301 | 1.07 | 327.1 | 0.0014705 | 454.4 | 1.04
606.578 | 1.07 | 327.2 | 1.1352e-05 | 454.6 | 1.0416
606.855 | 1.07 | 327.3 | -0.0014478 | 454.8 | 1.0465
607.132 | 1.05 | 327.4 | 0.0045785 | 455 | 1.0486
607.409 | 1.07 | 327.5 | 0.0060376 | 455.2 | 1.0503
607.686 | 1.06 | 327.6 | 0.0074968 | 455.4 | 1.05
607.963 | 1.06 | 327.7 | 0.0033333 | 455.6 | 1.05
608.24 | 1.06 | 327.8 | 0.0033333 | 455.8 | 1.05
608.517 | 1.06 | 327.9 | 0.0033333 | 456 | 1.05
608.794 | 1.06 | 328 | 0.0033187 | 456.2 | 1.0531
609.071 | 1.05 | 328.1 | 0.0018567 | 456.4 | 1.0578
609.347 | 1.08 | 328.2 | 0.00039474 | 456.6 | 1.0639
609.624 | 1.08 | 328.3 | -0.0010673 | 456.8 | 1.0675
609.9 | 1.1 | 328.4 | 0 | 457 | 1.07
610.177 | 1.16 | 328.5 | 0 | 457.2 | 1.0743
610.453 | 1.24 | 328.6 | 0 | 457.4 | 1.076
610.729 | 1.39 | 328.7 | 0.0035819 | 457.6 | 1.0733
611.006 | 1.6 | 328.8 | 0.0050439 | 457.8 | 1.072
611.282 | 1.75 | 328.9 | 0.0065058 | 458 | 1.0713
611.558 | 1.75 | 329 | 0.0079678 | 458.2 | 1.0759
611.834 | 1.55 | 329.1 | 0.0077632 | 458.4 | 1.0747
612.11 | 1.34 | 329.2 | 0.0092251 | 458.6 | 1.0744
612.386 | 1.23 | 329.3 | 0.010687 | 458.8 | 1.0767
612.662 | 1.18 | 329.4 | 0.0061842 | 459 | 1.0805
612.938 | 1.14 | 329.5 | 0.0047222 | 459.2 | 1.0849
613.213 | 1.13 | 329.6 | 0.0032602 | 459.4 | 1.0833
613.489 | 1.09 | 329.7 | 0.0017982 | 459.6 | 1.0827
613.765 | 1.08 | 329.8 | 0.0066654 | 459.8 | 1.08
614.04 | 1.08 | 329.9 | 0.0066639 | 460 | 1.0832
614.316 | 1.07 | 330 | 0.0066625 | 460.2 | 1.089
614.591 | 1.06 | 330.1 | 0.0073935 | 460.4 | 1.0959
614.867 | 1.04 | 330.2 | 0.0088569 | 460.6 | 1.0952
615.142 | 1.02 | 330.3 | 0.01032 | 460.8 | 1.0944
615.417 | 0.99 | 330.4 | 0.0065497 | 461 | 1.0967
615.692 | 0.97 | 330.5 | 0.0050877 | 461.2 | 1.0977
615.967 | 0.94 | 330.6 | 0.0036257 | 461.4 | 1.0999
616.242 | 0.93 | 330.7 | 0.0021637 | 461.6 | 1.1
616.517 | 0.91 | 330.8 | 0.0023617 | 461.8 | 1.0695
616.792 | 0.91 | 330.9 | 0.00089681 | 462 | 1.0477
617.067 | 0.91 | 331 | -0.00056803 | 462.2 | 1.0667
617.342 | 0.89 | 331.1 | 0.0036947 | 462.4 | 1.0754
617.617 | 0.88 | 331.2 | 0.0051595 | 462.6 | 1.1063
617.891 | 0.87 | 331.3 | 0.0066243 | 462.8 | 1.114
618.166 | 0.85 | 331.4 | 0.0080892 | 463 | 1.1104
618.441 | 0.84 | 331.5 | 0.0033345 | 463.2 | 1.11
618.715 | 0.84 | 331.6 | 0.003336 | 463.4 | 1.11
618.99 | 0.85 | 331.7 | 0.0033374 | 463.6 | 1.11
619.264 | 0.84 | 331.8 | 0.0027175 | 463.8 | 1.11
619.538 | 0.84 | 331.9 | 0.0012512 | 464 | 1.1133
619.812 | 0.83 | 332 | -0.00021505 | 464.2 | 1.1168
620.087 | 0.84 | 332.1 | 0 | 464.4 | 1.1189
620.361 | 0.84 | 332.2 | 0 | 464.6 | 1.1202
620.635 | 0.82 | 332.3 | 0 | 464.8 | 1.12
620.909 | 0.82 | 332.4 | 0 | 465 | 1.12
621.183 | 0.8 | 332.5 | 0 | 465.2 | 1.1229
621.456 | 0.8 | 332.6 | 0 | 465.4 | 1.1262
621.73 | 0.8 | 332.7 | 0 | 465.6 | 1.1283
622.004 | 0.8 | 332.8 | 0 | 465.8 | 1.1305
622.278 | 0.78 | 332.9 | 0 | 466 | 1.13
622.551 | 0.79 | 333 | 0 | 466.2 | 1.13
622.825 | 0.78 | 333.1 | 0 | 466.4 | 1.1315
623.098 | 0.78 | 333.2 | 0.004477 | 466.6 | 1.1361
623.372 | 0.78 | 333.3 | 0.0059433 | 466.8 | 1.1348
623.645 | 0.77 | 333.4 | 0.0074096 | 467 | 1.1344
623.918 | 0.77 | 333.5 | 0.011085 | 467.2 | 1.1367
624.192 | 0.75 | 333.6 | 0.014018 | 467.4 | 1.1377
624.465 | 0.76 | 333.7 | 0.01695 | 467.6 | 1.1398
624.738 | 0.74 | 333.8 | 0.019883 | 467.8 | 1.14
625.011 | 0.73 | 333.9 | 0.0085924 | 468 | 1.14
625.284 | 0.72 | 334 | 0.0071261 | 468.2 | 1.14
625.557 | 0.73 | 334.1 | 0.0056598 | 468.4 | 1.14
625.83 | 0.72 | 334.2 | 0.0050408 | 468.6 | 1.1428
626.102 | 0.71 | 334.3 | 0.0021025 | 468.8 | 1.1466
626.375 | 0.75 | 334.4 | -0.00083577 | 469 | 1.1488
626.648 | 0.7 | 334.5 | 0 | 469.2 | 1.1504
626.92 | 0.71 | 334.6 | 0 | 469.4 | 1.15
627.193 | 0.71 | 334.7 | 0 | 469.6 | 1.15
627.465 | 0.7 | 334.8 | 0 | 469.8 | 1.15
627.738 | 0.68 | 334.9 | 0.0044184 | 470 | 1.15
628.01 | 0.69 | 335 | 0.0058847 | 470.2 | 1.15
628.282 | 0.68 | 335.1 | 0.0073509 | 470.4 | 1.15
628.555 | 0.68 | 335.2 | 0.0033338 | 470.6 | 1.15
628.827 | 0.67 | 335.3 | 0.0033353 | 470.8 | 1.15
629.099 | 0.66 | 335.4 | 0.0033367 | 471 | 1.15
629.371 | 0.67 | 335.5 | 0.0033381 | 471.2 | 1.15
629.643 | 0.66 | 335.6 | 0.0019657 | 471.4 | 1.15
629.915 | 0.67 | 335.7 | 0.0004951 | 471.6 | 1.1512
630.187 | 0.68 | 335.8 | -0.00097549 | 471.8 | 1.1552
630.458 | 0.68 | 335.9 | 0.0041071 | 472 | 1.1573
630.73 | 0.69 | 336 | 0.0055762 | 472.2 | 1.16
631.002 | 0.71 | 336.1 | 0.0070454 | 472.4 | 1.1605
631.273 | 0.69 | 336.2 | 0.0068479 | 472.6 | 1.1616
631.545 | 0.67 | 336.3 | 0.008317 | 472.8 | 1.1664
631.816 | 0.64 | 336.4 | 0.0097862 | 473 | 1.1649
632.088 | 0.62 | 336.5 | 0.011255 | 473.2 | 1.1644
632.359 | 0.61 | 336.6 | 0.0056078 | 473.4 | 1.1667
632.63 | 0.59 | 336.7 | 0.0041373 | 473.6 | 1.1679
632.902 | 0.58 | 336.8 | 0.0026667 | 473.8 | 1.1687
633.173 | 0.59 | 336.9 | 0.0028627 | 474 | 1.1641
633.444 | 0.57 | 337 | 0.0013922 | 474.2 | 1.1653
633.715 | 0.58 | 337.1 | -7.8431e-05 | 474.4 | 1.1669
633.986 | 0.57 | 337.2 | -0.001549 | 474.6 | 1.1689
634.257 | 0.56 | 337.3 | 0 | 474.8 | 1.1702
634.528 | 0.56 | 337.4 | 0 | 475 | 1.1688
634.798 | 0.55 | 337.5 | 0 | 475.2 | 1.1646
635.069 | 0.55 | 337.6 | 0 | 475.4 | 1.1654
635.34 | 0.54 | 337.7 | 0 | 475.6 | 1.1667
635.61 | 0.54 | 337.8 | 0 | 475.8 | 1.1684
635.881 | 0.54 | 337.9 | 0 | 476 | 1.1672
636.151 | 0.53 | 338 | 0 | 476.2 | 1.1634
636.422 | 0.53 | 338.1 | 0 | 476.4 | 1.165
636.692 | 0.52 | 338.2 | 0 | 476.6 | 1.1671
636.962 | 0.53 | 338.3 | 0 | 476.8 | 1.1697
637.233 | 0.52 | 338.4 | 0 | 477 | 1.1704
637.503 | 0.52 | 338.5 | 0 | 477.2 | 1.17
637.773 | 0.51 | 338.6 | 0 | 477.4 | 1.17
638.043 | 0.51 | 338.7 | 0 | 477.6 | 1.1735
638.313 | 0.5 | 338.8 | 0 | 477.8 | 1.1757
638.583 | 0.5 | 338.9 | 0 | 478 | 1.1733
638.853 | 0.5 | 339 | 0 | 478.2 | 1.1722
639.122 | 0.5 | 339.1 | 0 | 478.4 | 1.1714
639.392 | 0.49 | 339.2 | 0 | 478.6 | 1.1757
639.662 | 0.48 | 339.3 | 0.0041241 | 478.8 | 1.1747
639.931 | 0.48 | 339.4 | 0.0055976 | 479 | 1.1732
640.201 | 0.48 | 339.5 | 0.0070711 | 479.2 | 1.1712
640.47 | 0.47 | 339.6 | 0.0033335 | 479.4 | 1.1699
640.74 | 0.46 | 339.7 | 0.003335 | 479.6 | 1.17
641.009 | 0.46 | 339.8 | 0.0033364 | 479.8 | 1.17
641.279 | 0.47 | 339.9 | 0.0033379 | 480 | 1.17
641.548 | 0.46 | 340 | 0.0022342 | 480.2 | 1.1711
641.817 | 0.46 | 340.1 | 0.00076219 | 480.4 | 1.175
642.086 | 0.46 | 340.2 | -0.00070984 | 480.6 | 1.1761
642.355 | 0.45 | 340.3 | 0.0038535 | 480.8 | 1.1733
642.624 | 0.45 | 340.4 | 0.0053255 | 481 | 1.1756
642.893 | 0.45 | 340.5 | 0.0067975 | 481.2 | 1.1771
643.162 | 0.45 | 340.6 | 0.0082695 | 481.4 | 1.1797
643.431 | 0.44 | 340.7 | 0.0080826 | 481.6 | 1.1804
643.699 | 0.43 | 340.8 | 0.0095575 | 481.8 | 1.18
643.968 | 0.44 | 340.9 | 0.011032 | 482 | 1.18
644.237 | 0.44 | 341 | 0.01 | 482.2 | 1.1835
644.505 | 0.42 | 341.1 | 0.01 | 482.4 | 1.1857
644.774 | 0.42 | 341.2 | 0.01 | 482.6 | 1.1833
645.042 | 0.42 | 341.3 | 0.0064012 | 482.8 | 1.1856
645.311 | 0.42 | 341.4 | 0.0049263 | 483 | 1.1856
645.579 | 0.42 | 341.5 | 0.0034513 | 483.2 | 1.1833
645.847 | 0.41 | 341.6 | 0.0019764 | 483.4 | 1.1856
646.115 | 0.41 | 341.7 | 0.0066667 | 483.6 | 1.1869
646.384 | 0.41 | 341.8 | 0.0066667 | 483.8 | 1.1889
646.652 | 0.41 | 341.9 | 0.0066667 | 484 | 1.1902
646.92 | 0.39 | 342 | 0.0072566 | 484.2 | 1.19
647.188 | 0.39 | 342.1 | 0.0087316 | 484.4 | 1.19
647.455 | 0.4 | 342.2 | 0.010206 | 484.6 | 1.193
647.723 | 0.38 | 342.3 | 0.0066519 | 484.8 | 1.1952
647.991 | 0.39 | 342.4 | 0.005177 | 485 | 1.1933
648.259 | 0.39 | 342.5 | 0.0037021 | 485.2 | 1.1984
648.526 | 0.39 | 342.6 | 0.0022271 | 485.4 | 1.2028
648.794 | 0.39 | 342.7 | 0.0066658 | 485.6 | 1.2066
649.061 | 0.39 | 342.8 | 0.0066643 | 485.8 | 1.211
649.329 | 0.38 | 342.9 | 0.0066629 | 486 | 1.21
649.596 | 0.38 | 343 | 0.003333 | 486.2 | 1.2062
649.864 | 0.38 | 343.1 | 0.0033315 | 486.4 | 1.204
650.131 | 0.39 | 343.2 | 0.0033301 | 486.6 | 1.2067
650.398 | 0.39 | 343.3 | 0.0033286 | 486.8 | 1.2081
650.665 | 0.39 | 343.4 | 0.0066667 | 487 | 1.2103
650.932 | 0.38 | 343.5 | 0.0066667 | 487.2 | 1.21
651.199 | 0.37 | 343.6 | 0.0066667 | 487.4 | 1.2065
651.466 | 0.36 | 343.7 | 0.003334 | 487.6 | 1.2028
651.733 | 0.36 | 343.8 | 0.0033355 | 487.8 | 1.2006
652 | 0.36 | 343.9 | 0.0033369 | 488 | 1.1997
652.267 | 0.36 | 344 | 0.003215 | 488.2 | 1.2
652.534 | 0.36 | 344.1 | 0.0017357 | 488.4 | 1.2
652.8 | 0.36 | 344.2 | 0.00025641 | 488.6 | 1.2
653.067 | 0.34 | 344.3 | -0.0012229 | 488.8 | 1.2
653.333 | 0.35 | 344.4 | 0.0043629 | 489 | 1.2
653.6 | 0.34 | 344.5 | 0.0058407 | 489.2 | 1.2
653.866 | 0.34 | 344.6 | 0.0073186 | 489.4 | 1.1987
654.133 | 0.33 | 344.7 | 0.0033338 | 489.6 | 1.1944
654.399 | 0.33 | 344.8 | 0.0033352 | 489.8 | 1.1922
654.665 | 0.33 | 344.9 | 0.0033367 | 490 | 1.1915
654.931 | 0.32 | 345 | 0.0033382 | 490.2 | 1.1954
655.197 | 0.31 | 345.1 | 0.0019576 | 490.4 | 1.1946
655.463 | 0.32 | 345.2 | 0.0004783 | 490.6 | 1.1957
655.729 | 0.31 | 345.3 | -0.001001 | 490.8 | 1.202
655.995 | 0.31 | 345.4 | 0.0041469 | 491 | 1.2078
656.261 | 0.31 | 345.5 | 0.0056262 | 491.2 | 1.2167
656.527 | 0.31 | 345.6 | 0.0071055 | 491.4 | 1.2237
656.793 | 0.31 | 345.7 | 0.0069181 | 491.6 | 1.2303
657.059 | 0.31 | 345.8 | 0.0083974 | 491.8 | 1.2312
657.324 | 0.31 | 345.9 | 0.0098767 | 492 | 1.2182
657.59 | 0.31 | 346 | 0.011356 | 492.2 | 1.2061
657.855 | 0.31 | 346.1 | 0.01 | 492.4 | 1.1869
658.121 | 0.29 | 346.2 | 0.01 | 492.6 | 1.1803
658.386 | 0.3 | 346.3 | 0.01 | 492.8 | 1.1774
658.651 | 0.3 | 346.4 | 0.0060602 | 493 | 1.1735
658.917 | 0.29 | 346.5 | 0.0045809 | 493.2 | 1.1751
659.182 | 0.29 | 346.6 | 0.0031016 | 493.4 | 1.1756
659.447 | 0.3 | 346.7 | 0.003289 | 493.6 | 1.1733
659.712 | 0.29 | 346.8 | 0.0018097 | 493.8 | 1.1756
659.977 | 0.28 | 346.9 | 0.00033037 | 494 | 1.1756
660.242 | 0.28 | 347 | -0.0011489 | 494.2 | 1.1733
660.507 | 0.27 | 347.1 | 0.0043008 | 494.4 | 1.1684
660.772 | 0.28 | 347.2 | 0.0057815 | 494.6 | 1.1624
661.037 | 0.28 | 347.3 | 0.0072623 | 494.8 | 1.1605
661.301 | 0.26 | 347.4 | 0.0033329 | 495 | 1.1596
661.566 | 0.28 | 347.5 | 0.0033315 | 495.2 | 1.1585
661.831 | 0.28 | 347.6 | 0.00333 | 495.4 | 1.1539
662.095 | 0.28 | 347.7 | 0.0033285 | 495.6 | 1.1552
662.36 | 0.28 | 347.8 | 0.0066653 | 495.8 | 1.1584
662.624 | 0.28 | 347.9 | 0.0066639 | 496 | 1.1652
662.888 | 0.28 | 348 | 0.0066624 | 496.2 | 1.1617
663.153 | 0.26 | 348.1 | 0.0074465 | 496.4 | 1.1574
663.417 | 0.27 | 348.2 | 0.0089273 | 496.6 | 1.1575
663.681 | 0.26 | 348.3 | 0.010408 | 496.8 | 1.1542
663.945 | 0.26 | 348.4 | 0.01 | 497 | 1.1531
664.209 | 0.25 | 348.5 | 0.01 | 497.2 | 1.1509
664.473 | 0.26 | 348.6 | 0.01 | 497.4 | 1.1497
664.737 | 0.26 | 348.7 | 0.01 | 497.6 | 1.1487
665.001 | 0.25 | 348.8 | 0.01 | 497.8 | 1.1443
665.265 | 0.25 | 348.9 | 0.01 | 498 | 1.1453
665.529 | 0.24 | 349 | 0.01 | 498.2 | 1.1456
665.792 | 0.24 | 349.1 | 0.01 | 498.4 | 1.1433
666.056 | 0.24 | 349.2 | 0.01 | 498.6 | 1.1392
666.319 | 0.24 | 349.3 | 0.01 | 498.8 | 1.1335
666.583 | 0.24 | 349.4 | 0.01 | 499 | 1.1311
666.846 | 0.24 | 349.5 | 0.01 | 499.2 | 1.1329
667.11 | 0.24 | 349.6 | 0.01 | 499.4 | 1.1353
667.373 | 0.25 | 349.7 | 0.01 | 499.6 | 1.1333
667.636 | 0.23 | 349.8 | 0.014313 | 499.8 | 1.1294
667.9 | 0.23 | 349.9 | 0.015796 | 500 | 1.125
668.163 | 0.23 | 350 | 0.01728 | 500.2 | 1.1267
668.426 | 0.23 | 350.1 | 0.013333 | 500.4 | 1.1275
668.689 | 0.23 | 350.2 | 0.013333 | 500.6 | 1.1297
668.952 | 0.22 | 350.3 | 0.013333 | 500.8 | 1.13
669.215 | 0.22 | 350.4 | 0.013333 | 501 | 1.124
669.478 | 0.22 | 350.5 | 0.016665 | 501.2 | 1.1172
669.74 | 0.21 | 350.6 | 0.016664 | 501.4 | 1.1128
670.003 | 0.22 | 350.7 | 0.016662 | 501.6 | 1.1128
670.266 | 0.21 | 350.8 | 0.013333 | 501.8 | 1.1164
670.528 | 0.23 | 350.9 | 0.013331 | 502 | 1.1186
670.791 | 0.22 | 351 | 0.01333 | 502.2 | 1.1172
671.054 | 0.21 | 351.1 | 0.0094362 | 502.4 | 1.1137
671.316 | 0.21 | 351.2 | 0.0064688 | 502.6 | 1.1115
671.578 | 0.22 | 351.3 | 0.0035015 | 502.8 | 1.1071
671.841 | 0.21 | 351.4 | 0.00053412 | 503 | 1.1049
672.103 | 0.2 | 351.5 | 0.0066667 | 503.2 | 1.1067
672.365 | 0.21 | 351.6 | 0.0066667 | 503.4 | 1.1073
672.627 | 0.21 | 351.7 | 0.0066667 | 503.6 | 1.1084
672.889 | 0.2 | 351.8 | 0.011347 | 503.8 | 1.1049
673.151 | 0.2 | 351.9 | 0.014319 | 504 | 1.1027
673.413 | 0.2 | 352 | 0.017291 | 504.2 | 1.1005
673.675 | 0.2 | 352.1 | 0.013333 | 504.4 | 1.1
673.937 | 0.19 | 352.2 | 0.013332 | 504.6 | 1.1028
674.199 | 0.2 | 352.3 | 0.01333 | 504.8 | 1.105
674.461 | 0.2 | 352.4 | 0.013329 | 505 | 1.1033
674.722 | 0.2 | 352.5 | 0.012258 | 505.2 | 1.1028
674.984 | 0.19 | 352.6 | 0.010771 | 505.4 | 1.1005
675.245 | 0.19 | 352.7 | 0.0092848 | 505.6 | 1.1
675.507 | 0.2 | 352.8 | 0.01 | 505.8 | 1.0983
675.768 | 0.19 | 352.9 | 0.01 | 506 | 1.0933
676.03 | 0.2 | 353 | 0.01 | 506.2 | 1.0911
676.291 | 0.2 | 353.1 | 0.01 | 506.4 | 1.0895
676.552 | 0.18 | 353.2 | 0.014815 | 506.6 | 1.09
676.813 | 0.19 | 353.3 | 0.016302 | 506.8 | 1.0939
677.074 | 0.19 | 353.4 | 0.017788 | 507 | 1.0961
677.336 | 0.19 | 353.5 | 0.013334 | 507.2 | 1.0933
677.597 | 0.16 | 353.6 | 0.013336 | 507.4 | 1.0878
677.857 | 0.2 | 353.7 | 0.013337 | 507.6 | 1.0817
678.118 | 0.17 | 353.8 | 0.016667 | 507.8 | 1.08
678.379 | 0.17 | 353.9 | 0.016667 | 508 | 1.0795
678.64 | 0.17 | 354 | 0.016667 | 508.2 | 1.08
678.901 | 0.18 | 354.1 | 0.016667 | 508.4 | 1.08
679.161 | 0.18 | 354.2 | 0.022708 | 508.6 | 1.0761
679.422 | 0.18 | 354.3 | 0.025685 | 508.8 | 1.0739
679.682 | 0.17 | 354.4 | 0.028661 | 509 | 1.0767
679.943 | 0.18 | 354.5 | 0.023333 | 509.2 | 1.0822
680.203 | 0.17 | 354.6 | 0.023333 | 509.4 | 1.0866
680.464 | 0.17 | 354.7 | 0.023333 | 509.6 | 1.0833
680.724 | 0.17 | 354.8 | 0.023051 | 509.8 | 1.0778
680.984 | 0.17 | 354.9 | 0.021563 | 510 | 1.0734
681.244 | 0.17 | 355 | 0.020074 | 510.2 | 1.0767
681.505 | 0.16 | 355.1 | 0.018586 | 510.4 | 1.0706
681.765 | 0.17 | 355.2 | 0.02 | 510.6 | 1.0667
682.025 | 0.16 | 355.3 | 0.02 | 510.8 | 1.0667
682.285 | 0.17 | 355.4 | 0.02 | 511 | 1.0717
682.544 | 0.16 | 355.5 | 0.02 | 511.2 | 1.075
682.804 | 0.16 | 355.6 | 0.02 | 511.4 | 1.0672
683.064 | 0.16 | 355.7 | 0.02 | 511.6 | 1.0633
683.324 | 0.16 | 355.8 | 0.02 | 511.8 | 1.0628
683.583 | 0.15 | 355.9 | 0.02 | 512 | 1.0661
683.843 | 0.14 | 356 | 0.02 | 512.2 | 1.0684
684.102 | 0.16 | 356.1 | 0.02 | 512.4 | 1.07
684.362 | 0.15 | 356.2 | 0.02447 | 512.6 | 1.0677
684.621 | 0.16 | 356.3 | 0.02596 | 512.8 | 1.0599
684.881 | 0.14 | 356.4 | 0.02745 | 513 | 1.0582
685.14 | 0.15 | 356.5 | 0.027272 | 513.2 | 1.056
685.399 | 0.15 | 356.6 | 0.028762 | 513.4 | 1.0533
685.658 | 0.14 | 356.7 | 0.030252 | 513.6 | 1.0556
685.917 | 0.14 | 356.8 | 0.026587 | 513.8 | 1.0556
686.177 | 0.15 | 356.9 | 0.025098 | 514 | 1.0533
686.436 | 0.15 | 357 | 0.023608 | 514.2 | 1.0556
686.694 | 0.14 | 357.1 | 0.022119 | 514.4 | 1.0567
686.953 | 0.16 | 357.2 | 0.022296 | 514.6 | 1.0585
687.212 | 0.16 | 357.3 | 0.020806 | 514.8 | 1.0571
687.471 | 0.17 | 357.4 | 0.019316 | 515 | 1.0537
687.73 | 0.15 | 357.5 | 0.02 | 515.2 | 1.0514
687.988 | 0.15 | 357.6 | 0.02 | 515.4 | 1.0499
688.247 | 0.16 | 357.7 | 0.02 | 515.6 | 1.0488
688.505 | 0.15 | 357.8 | 0.02 | 515.8 | 1.0447
688.764 | 0.16 | 357.9 | 0.024811 | 516 | 1.0425
689.022 | 0.14 | 358 | 0.026303 | 516.2 | 1.0416
689.281 | 0.14 | 358.1 | 0.027796 | 516.4 | 1.0454
689.539 | 0.14 | 358.2 | 0.027622 | 516.6 | 1.0446
689.797 | 0.14 | 358.3 | 0.029114 | 516.8 | 1.0432
690.055 | 0.14 | 358.4 | 0.030607 | 517 | 1.0412
690.314 | 0.14 | 358.5 | 0.03 | 517.2 | 1.0399
690.572 | 0.2 | 358.6 | 0.03 | 517.4 | 1.04
690.83 | 0.21 | 358.7 | 0.03 | 517.6 | 1.04
691.088 | 0.15 | 358.8 | 0.03 | 517.8 | 1.04
691.345 | 0.13 | 358.9 | 0.03 | 518 | 1.04
691.603 | 0.13 | 359 | 0.03 | 518.2 | 1.04
691.861 | 0.13 | 359.1 | 0.03 | 518.4 | 1.04
692.119 | 0.13 | 359.2 | 0.03 | 518.6 | 1.04
692.376 | 0.12 | 359.3 | 0.03 | 518.8 | 1.04
692.634 | 0.13 | 359.4 | 0.03 | 519 | 1.04
692.891 | 0.12 | 359.5 | 0.033692 | 519.2 | 1.04
693.149 | 0.15 | 359.6 | 0.035184 | 519.4 | 1.04
693.406 | 0.13 | 359.7 | 0.036677 | 519.6 | 1.0367
693.664 | 0.14 | 359.8 | 0.038169 | 519.8 | 1.0331
693.921 | 0.13 | 359.9 | 0.033335 | 520 | 1.0309
694.178 | 0.13 | 360 | 0.033336 | 520.2 | 1.0331
694.435 | 0.13 | 360.1 | 0.033338 | 520.4 | 1.0356
694.693 | 0.12 | 360.2 | 0.036667 | 520.6 | 1.0333
694.95 | 0.11 | 360.3 | 0.036669 | 520.8 | 1.0321
695.207 | 0.12 | 360.4 | 0.03667 | 521 | 1.0328
695.464 | 0.13 | 360.5 | 0.033334 | 521.2 | 1.042
695.721 | 0.12 | 360.6 | 0.033335 | 521.4 | 1.0431
695.977 | 0.12 | 360.7 | 0.033337 | 521.6 | 1.0431
696.234 | 0.13 | 360.8 | 0.033338 | 521.8 | 1.0406
696.491 | 0.15 | 360.9 | 0.036668 | 522 | 1.0361
696.748 | 0.14 | 361 | 0.036669 | 522.2 | 1.0312
697.004 | 0.13 | 361.1 | 0.036671 | 522.4 | 1.0242
697.261 | 0.11 | 361.2 | 0.037434 | 522.6 | 1.0247
697.517 | 0.12 | 361.3 | 0.03893 | 522.8 | 1.0271
697.774 | 0.12 | 361.4 | 0.040425 | 523 | 1.0297
698.03 | 0.12 | 361.5 | 0.04 | 523.2 | 1.0304
698.286 | 0.11 | 361.6 | 0.04 | 523.4 | 1.0316
698.543 | 0.11 | 361.7 | 0.04 | 523.6 | 1.0364
698.799 | 0.11 | 361.8 | 0.04 | 523.8 | 1.0349
699.055 | 0.11 | 361.9 | 0.04457 | 524 | 1.0344
699.311 | 0.11 | 362 | 0.046065 | 524.2 | 1.0367
699.567 | 0.11 | 362.1 | 0.047561 | 524.4 | 1.0344
699.823 | 0.11 | 362.2 | 0.043334 | 524.6 | 1.0328
700.079 | 0.11 | 362.3 | 0.043336 | 524.8 | 1.0306
700.335 | 0.11 | 362.4 | 0.043337 | 525 | 1.03
700.591 | 0.1 | 362.5 | 0.046667 | 525.2 | 1.03
700.846 | 0.11 | 362.6 | 0.046667 | 525.4 | 1.03
701.102 | 0.11 | 362.7 | 0.046667 | 525.6 | 1.03
701.358 | 0.09 | 362.8 | 0.046667 | 525.8 | 1.0329
701.613 | 0.12 | 362.9 | 0.047864 | 526 | 1.0352
701.869 | 0.1 | 363 | 0.049361 | 526.2 | 1.0333
702.124 | 0.11 | 363.1 | 0.050858 | 526.4 | 1.0325
702.38 | 0.11 | 363.2 | 0.058044 | 526.6 | 1.0303
702.635 | 0.11 | 363.3 | 0.061038 | 526.8 | 1.03
702.89 | 0.1 | 363.4 | 0.064032 | 527 | 1.03
703.145 | 0.1 | 363.5 | 0.08477 | 527.2 | 1.03
703.401 | 0.11 | 363.6 | 0.096747 | 527.4 | 1.03
703.656 | 0.1 | 363.7 | 0.10872 | 527.6 | 1.0332
703.911 | 0.1 | 363.8 | 0.1207 | 527.8 | 1.0368
704.166 | 0.1 | 363.9 | 0.27931 | 528 | 1.0391
704.421 | 0.11 | 364 | 0.34225 | 528.2 | 1.0403
704.675 | 0.11 | 364.1 | 0.4052 | 528.4 | 1.0386
704.93 | 0.11 | 364.2 | 0.74639 | 528.6 | 1.0341
705.185 | 0.1 | 364.3 | 0.93234 | 528.8 | 1.0352
705.44 | 0.11 | 364.4 | 1.1183 | 529 | 1.037
705.694 | 0.12 | 364.5 | 1.2626 | 529.2 | 1.0394
705.949 | 0.13 | 364.6 | 1.4676 | 529.4 | 1.0404
706.203 | 0.13 | 364.7 | 1.6727 | 529.6 | 1.04
706.458 | 0.16 | 364.8 | 1.8778 | 529.8 | 1.04
706.712 | 0.17 | 364.9 | 1.5334 | 530 | 1.04
706.967 | 0.15 | 365 | 1.5335 | 530.2 | 1.04
707.221 | 0.14 | 365.1 | 1.5335 | 530.4 | 1.04
707.475 | 0.14 | 365.2 | 1.2054 | 530.6 | 1.04
707.729 | 0.12 | 365.3 | 1.0207 | 530.8 | 1.04
707.984 | 0.13 | 365.4 | 0.83604 | 531 | 1.04
708.238 | 0.13 | 365.5 | 0.79497 | 531.2 | 1.0361
708.492 | 0.12 | 365.6 | 0.65408 | 531.4 | 1.0355
708.746 | 0.14 | 365.7 | 0.51319 | 531.6 | 1.0379
709 | 0.16 | 365.8 | 0.3723 | 531.8 | 1.0437
709.253 | 0.15 | 365.9 | 0.4555 | 532 | 1.0475
709.507 | 0.13 | 366 | 0.4195 | 532.2 | 1.0484
709.761 | 0.1 | 366.1 | 0.38351 | 532.4 | 1.0446
710.015 | 0.1 | 366.2 | 0.33279 | 532.6 | 1.0485
710.268 | 0.1 | 366.3 | 0.30126 | 532.8 | 1.0522
710.522 | 0.12 | 366.4 | 0.26973 | 533 | 1.0522
710.775 | 0.1 | 366.5 | 0.22488 | 533.2 | 1.0558
711.029 | 0.1 | 366.6 | 0.17833 | 533.4 | 1.0583
711.282 | 0.1 | 366.7 | 0.13179 | 533.6 | 1.0648
711.535 | 0.29 | 366.8 | 0.08524 | 533.8 | 1.0683
711.789 | 0.13 | 366.9 | 0.10587 | 534 | 1.07
712.042 | 0.14 | 367 | 0.083396 | 534.2 | 1.07
712.295 | 0.14 | 367.1 | 0.060917 | 534.4 | 1.07
712.548 | 0.13 | 367.2 | 0.068659 | 534.6 | 1.07
712.801 | 0.11 | 367.3 | 0.065656 | 534.8 | 1.07
713.054 | 0.09 | 367.4 | 0.062653 | 535 | 1.07
713.307 | 0.1 | 367.5 | 0.066667 | 535.2 | 1.0684
713.56 | 0.1 | 367.6 | 0.066668 | 535.4 | 1.0636
713.813 | 0.1 | 367.7 | 0.06667 | 535.6 | 1.0688
714.066 | 0.09 | 367.8 | 0.066671 | 535.8 | 1.0732
714.318 | 0.08 | 367.9 | 0.063333 | 536 | 1.0732
714.571 | 0.09 | 368 | 0.063333 | 536.2 | 1.0722
714.823 | 0.09 | 368.1 | 0.063333 | 536.4 | 1.0728
715.076 | 0.08 | 368.2 | 0.062643 | 536.6 | 1.0774
715.328 | 0.08 | 368.3 | 0.061141 | 536.8 | 1.0838
715.581 | 0.07 | 368.4 | 0.05964 | 537 | 1.0906
715.833 | 0.08 | 368.5 | 0.063529 | 537.2 | 1.0961
716.086 | 0.08 | 368.6 | 0.06503 | 537.4 | 1.0988
716.338 | 0.08 | 368.7 | 0.066532 | 537.6 | 1.1002
716.59 | 0.09 | 368.8 | 0.068033 | 537.8 | 1.1013
716.842 | 0.09 | 368.9 | 0.063335 | 538 | 1.1057
717.094 | 0.08 | 369 | 0.063336 | 538.2 | 1.108
717.346 | 0.08 | 369.1 | 0.063338 | 538.4 | 1.1114
717.598 | 0.07 | 369.2 | 0.070718 | 538.6 | 1.1159
717.85 | 0.08 | 369.3 | 0.072224 | 538.8 | 1.1217
718.102 | 0.07 | 369.4 | 0.07373 | 539 | 1.1273
718.354 | 0.07 | 369.5 | 0.070231 | 539.2 | 1.1295
718.605 | 0.07 | 369.6 | 0.071737 | 539.4 | 1.1376
718.857 | 0.07 | 369.7 | 0.073243 | 539.6 | 1.1464
719.109 | 0.08 | 369.8 | 0.074749 | 539.8 | 1.1558
719.36 | 0.08 | 369.9 | 0.072083 | 540 | 1.1634
719.612 | 0.08 | 370 | 0.070577 | 540.2 | 1.1698
719.863 | 0.08 | 370.1 | 0.069071 | 540.4 | 1.1777
720.115 | 0.08 | 370.2 | 0.07 | 540.6 | 1.1885
720.366 | 0.08 | 370.3 | 0.07 | 540.8 | 1.2034
720.617 | 0.08 | 370.4 | 0.07 | 541 | 1.2206
720.869 | 0.08 | 370.5 | 0.073609 | 541.2 | 1.2366
721.12 | 0.07 | 370.6 | 0.075112 | 541.4 | 1.2609
721.371 | 0.07 | 370.7 | 0.076615 | 541.6 | 1.2851
721.622 | 0.08 | 370.8 | 0.078118 | 541.8 | 1.3057
721.873 | 0.08 | 370.9 | 0.073333 | 542 | 1.3201
722.124 | 0.07 | 371 | 0.073333 | 542.2 | 1.3267
722.375 | 0.08 | 371.1 | 0.073333 | 542.4 | 1.3292
722.626 | 0.06 | 371.2 | 0.076667 | 542.6 | 1.3269
722.876 | 0.06 | 371.3 | 0.076667 | 542.8 | 1.3229
723.127 | 0.06 | 371.4 | 0.076667 | 543 | 1.3205
723.378 | 0.06 | 371.5 | 0.076998 | 543.2 | 1.3196
723.628 | 0.06 | 371.6 | 0.078504 | 543.4 | 1.3216
723.879 | 0.07 | 371.7 | 0.08001 | 543.6 | 1.3264
724.129 | 0.07 | 371.8 | 0.081516 | 543.8 | 1.3325
724.38 | 0.07 | 371.9 | 0.08 | 544 | 1.3382
724.63 | 0.07 | 372 | 0.08 | 544.2 | 1.3428
724.881 | 0.07 | 372.1 | 0.08 | 544.4 | 1.3553
725.131 | 0.07 | 372.2 | 0.084207 | 544.6 | 1.3828
725.381 | 0.06 | 372.3 | 0.085713 | 544.8 | 1.4665
725.631 | 0.06 | 372.4 | 0.087219 | 545 | 1.6632
725.881 | 0.08 | 372.5 | 0.090794 | 545.2 | 2.316
726.131 | 0.05 | 372.6 | 0.093811 | 545.4 | 3.3211
726.381 | 0.08 | 372.7 | 0.096828 | 545.6 | 4.6971
726.631 | 0.06 | 372.8 | 0.099844 | 545.8 | 6.0646
726.881 | 0.08 | 372.9 | 0.08857 | 546 | 7.5495
727.131 | 0.05 | 373 | 0.087059 | 546.2 | 9.1872
727.381 | 0.07 | 373.1 | 0.085549 | 546.4 | 7.9603
727.631 | 0.06 | 373.2 | 0.089051 | 546.6 | 5.7797
727.88 | 0.06 | 373.3 | 0.087545 | 546.8 | 3.0288
728.13 | 0.08 | 373.4 | 0.086039 | 547 | 1.7637
728.379 | 0.05 | 373.5 | 0.08714 | 547.2 | 1.3608
728.629 | 0.08 | 373.6 | 0.088649 | 547.4 | 1.3596
728.878 | 0.07 | 373.7 | 0.090158 | 547.6 | 1.3076
729.128 | 0.07 | 373.8 | 0.091667 | 547.8 | 1.3077
729.377 | 0.06 | 373.9 | 0.09 | 548 | 1.3
729.626 | 0.06 | 374 | 0.09 | 548.2 | 1.2946
729.875 | 0.06 | 374.1 | 0.09 | 548.4 | 1.2954
730.125 | 0.06 | 374.2 | 0.094363 | 548.6 | 1.2969
730.374 | 0.05 | 374.3 | 0.095871 | 548.8 | 1.299
730.623 | 0.05 | 374.4 | 0.097379 | 549 | 1.3002
730.872 | 0.06 | 374.5 | 0.093333 | 549.2 | 1.3
731.121 | 0.07 | 374.6 | 0.093333 | 549.4 | 1.3
731.37 | 0.06 | 374.7 | 0.093333 | 549.6 | 1.2965
731.618 | 0.06 | 374.8 | 0.093243 | 549.8 | 1.2973
731.867 | 0.06 | 374.9 | 0.091732 | 550 | 1.3093
732.116 | 0.07 | 375 | 0.090222 | 550.2 | 1.3118
732.365 | 0.06 | 375.1 | 0.088711 | 550.4 | 1.3121
732.613 | 0.05 | 375.2 | 0.094466 | 550.6 | 1.3045
732.862 | 0.07 | 375.3 | 0.095977 | 550.8 | 1.3016
733.11 | 0.06 | 375.4 | 0.097487 | 551 | 1.3029
733.359 | 0.06 | 375.5 | 0.097331 | 551.2 | 1.3052
733.607 | 0.06 | 375.6 | 0.098842 | 551.4 | 1.3033
733.855 | 0.06 | 375.7 | 0.10035 | 551.6 | 1.3087
734.104 | 0.05 | 375.8 | 0.1 | 551.8 | 1.311
734.352 | 0.05 | 375.9 | 0.1 | 552 | 1.3067
734.6 | 0.05 | 376 | 0.1 | 552.2 | 1.3077
734.848 | 0.05 | 376.1 | 0.1 | 552.4 | 1.3084
735.096 | 0.06 | 376.2 | 0.1 | 552.6 | 1.3156
735.344 | 0.06 | 376.3 | 0.1 | 552.8 | 1.3152
735.592 | 0.05 | 376.4 | 0.1 | 553 | 1.316
735.84 | 0.08 | 376.5 | 0.1 | 553.2 | 1.3232
736.088 | 0.07 | 376.6 | 0.1 | 553.4 | 1.3233
736.336 | 0.05 | 376.7 | 0.1 | 553.6 | 1.3256
736.583 | 0.05 | 376.8 | 0.1 | 553.8 | 1.3256
736.831 | 0.06 | 376.9 | 0.1 | 554 | 1.3233
737.079 | 0.06 | 377 | 0.1 | 554.2 | 1.3225
737.326 | 0.06 | 377.1 | 0.1 | 554.4 | 1.3202
737.574 | 0.06 | 377.2 | 0.1 | 554.6 | 1.32
737.821 | 0.05 | 377.3 | 0.1 | 554.8 | 1.3233
738.068 | 0.07 | 377.4 | 0.1 | 555 | 1.3271
738.316 | 0.11 | 377.5 | 0.1 | 555.2 | 1.3294
738.563 | 0.1 | 377.6 | 0.1 | 555.4 | 1.3268
738.81 | 0.07 | 377.7 | 0.1 | 555.6 | 1.3257
739.057 | 0.07 | 377.8 | 0.10376 | 555.8 | 1.3331
739.304 | 0.06 | 377.9 | 0.10527 | 556 | 1.3332
739.552 | 0.07 | 378 | 0.10679 | 556.2 | 1.3327
739.799 | 0.05 | 378.1 | 0.1083 | 556.4 | 1.3305
740.045 | 0.04 | 378.2 | 0.10814 | 556.6 | 1.33
740.292 | 0.05 | 378.3 | 0.10966 | 556.8 | 1.3269
740.539 | 0.05 | 378.4 | 0.11117 | 557 | 1.3247
740.786 | 0.05 | 378.5 | 0.11 | 557.2 | 1.3267
741.033 | 0.06 | 378.6 | 0.11 | 557.4 | 1.3278
741.279 | 0.06 | 378.7 | 0.11 | 557.6 | 1.3315
741.526 | 0.06 | 378.8 | 0.11 | 557.8 | 1.3361
741.773 | 0.05 | 378.9 | 0.11 | 558 | 1.3348
742.019 | 0.05 | 379 | 0.11 | 558.2 | 1.3344
742.265 | 0.06 | 379.1 | 0.11343 | 558.4 | 1.3367
742.512 | 0.05 | 379.2 | 0.11495 | 558.6 | 1.3422
742.758 | 0.05 | 379.3 | 0.11646 | 558.8 | 1.3501
743.004 | 0.05 | 379.4 | 0.11798 | 559 | 1.3559
743.251 | 0.05 | 379.5 | 0.11782 | 559.2 | 1.3587
743.497 | 0.06 | 379.6 | 0.11934 | 559.4 | 1.357
743.743 | 0.06 | 379.7 | 0.12085 | 559.6 | 1.3546
743.989 | 0.04 | 379.8 | 0.12 | 559.8 | 1.3567
744.235 | 0.05 | 379.9 | 0.12 | 560 | 1.3578
744.481 | 0.05 | 380 | 0.12 | 560.2 | 1.3616
744.727 | 0.04 | 380.1 | 0.12 | 560.4 | 1.3662
744.973 | 0.05 | 380.2 | 0.12 | 560.6 | 1.3685
745.219 | 0.05 | 380.3 | 0.12 | 560.8 | 1.372
745.464 | 0.05 | 380.4 | 0.12 | 561 | 1.3766
745.71 | 0.06 | 380.5 | 0.12464 | 561.2 | 1.375
745.956 | 0.05 | 380.6 | 0.12615 | 561.4 | 1.3756
746.201 | 0.05 | 380.7 | 0.12767 | 561.6 | 1.3768
746.447 | 0.05 | 380.8 | 0.12752 | 561.8 | 1.3788
746.692 | 0.07 | 380.9 | 0.12904 | 562 | 1.3802
746.938 | 0.05 | 381 | 0.13056 | 562.2 | 1.38
747.183 | 0.04 | 381.1 | 0.13 | 562.4 | 1.38
747.428 | 0.04 | 381.2 | 0.13 | 562.6 | 1.38
747.673 | 0.06 | 381.3 | 0.13 | 562.8 | 1.3815
747.919 | 0.04 | 381.4 | 0.13 | 563 | 1.3864
748.164 | 0.05 | 381.5 | 0.1348 | 563.2 | 1.3887
748.409 | 0.04 | 381.6 | 0.13632 | 563.4 | 1.3905
748.654 | 0.05 | 381.7 | 0.13783 | 563.6 | 1.39
748.899 | 0.05 | 381.8 | 0.13333 | 563.8 | 1.39
749.144 | 0.04 | 381.9 | 0.13334 | 564 | 1.3931
749.389 | 0.05 | 382 | 0.13334 | 564.2 | 1.3954
749.633 | 0.04 | 382.1 | 0.13667 | 564.4 | 1.3933
749.878 | 0.04 | 382.2 | 0.13667 | 564.6 | 1.3956
750.123 | 0.05 | 382.3 | 0.13667 | 564.8 | 1.3985
750.368 | 0.1 | 382.4 | 0.13333 | 565 | 1.4057
750.612 | 0.08 | 382.5 | 0.13333 | 565.2 | 1.409
750.857 | 0.06 | 382.6 | 0.13334 | 565.4 | 1.4121
751.101 | 0.07 | 382.7 | 0.13334 | 565.6 | 1.4161
751.346 | 0.1 | 382.8 | 0.13667 | 565.8 | 1.4185
751.59 | 0.11 | 382.9 | 0.13667 | 566 | 1.4171
751.834 | 0.07 | 383 | 0.13667 | 566.2 | 1.4134
752.079 | 0.06 | 383.1 | 0.13742 | 566.4 | 1.4111
752.323 | 0.05 | 383.2 | 0.13893 | 566.6 | 1.4131
752.567 | 0.04 | 383.3 | 0.14045 | 566.8 | 1.4185
752.811 | 0.05 | 383.4 | 0.14 | 567 | 1.4255
753.055 | 0.05 | 383.5 | 0.14 | 567.2 | 1.4252
753.299 | 0.05 | 383.6 | 0.14 | 567.4 | 1.4244
753.543 | 0.05 | 383.7 | 0.14 | 567.6 | 1.4267
753.787 | 0.06 | 383.8 | 0.14472 | 567.8 | 1.4284
754.031 | 0.04 | 383.9 | 0.14624 | 568 | 1.43
754.275 | 0.05 | 384 | 0.14776 | 568.2 | 1.4313
754.518 | 0.04 | 384.1 | 0.14761 | 568.4 | 1.4355
754.762 | 0.05 | 384.2 | 0.14913 | 568.6 | 1.4347
755.006 | 0.04 | 384.3 | 0.15065 | 568.8 | 1.433
755.249 | 0.06 | 384.4 | 0.15 | 569 | 1.4306
755.493 | 0.05 | 384.5 | 0.15 | 569.2 | 1.4332
755.736 | 0.05 | 384.6 | 0.15 | 569.4 | 1.4375
755.98 | 0.04 | 384.7 | 0.15339 | 569.6 | 1.4399
756.223 | 0.03 | 384.8 | 0.15491 | 569.8 | 1.4366
756.466 | 0.04 | 384.9 | 0.15643 | 570 | 1.4355
756.709 | 0.03 | 385 | 0.15795 | 570.2 | 1.4368
756.953 | 0.04 | 385.1 | 0.15781 | 570.4 | 1.4389
757.196 | 0.05 | 385.2 | 0.15933 | 570.6 | 1.4435
757.439 | 0.04 | 385.3 | 0.16085 | 570.8 | 1.447
757.682 | 0.05 | 385.4 | 0.16 | 571 | 1.4494
757.925 | 0.03 | 385.5 | 0.16 | 571.2 | 1.4468
758.168 | 0.04 | 385.6 | 0.16 | 571.4 | 1.4441
758.411 | 0.03 | 385.7 | 0.1564 | 571.6 | 1.4467
758.653 | 0.03 | 385.8 | 0.15487 | 571.8 | 1.4523
758.896 | 0.05 | 385.9 | 0.15335 | 572 | 1.4575
759.139 | 0.06 | 386 | 0.15183 | 572.2 | 1.4584
759.382 | 0.04 | 386.1 | 0.16135 | 572.4 | 1.4544
759.624 | 0.03 | 386.2 | 0.16287 | 572.6 | 1.4487
759.867 | 0.04 | 386.3 | 0.16439 | 572.8 | 1.443
760.109 | 0.04 | 386.4 | 0.16517 | 573 | 1.4406
760.352 | 0.04 | 386.5 | 0.16822 | 573.2 | 1.4433
760.594 | 0.03 | 386.6 | 0.17126 | 573.4 | 1.4492
760.836 | 0.03 | 386.7 | 0.17382 | 573.6 | 1.4543
761.079 | 0.04 | 386.8 | 0.17534 | 573.8 | 1.4533
761.321 | 0.03 | 386.9 | 0.17686 | 574 | 1.4586
761.563 | 0.04 | 387 | 0.17333 | 574.2 | 1.4609
761.805 | 0.05 | 387.1 | 0.17333 | 574.4 | 1.4567
762.047 | 0.06 | 387.2 | 0.17334 | 574.6 | 1.451
762.289 | 0.04 | 387.3 | 0.17334 | 574.8 | 1.4456
762.531 | 0.03 | 387.4 | 0.17667 | 575 | 1.4529
762.773 | 0.03 | 387.5 | 0.17667 | 575.2 | 1.4568
763.015 | 0.05 | 387.6 | 0.17667 | 575.4 | 1.4626
763.257 | 0.08 | 387.7 | 0.17738 | 575.6 | 1.4674
763.498 | 0.23 | 387.8 | 0.17891 | 575.8 | 1.4739
763.74 | 0.19 | 387.9 | 0.18043 | 576 | 1.5369
763.982 | 0.1 | 388 | 0.18 | 576.2 | 1.6675
764.223 | 0.05 | 388.1 | 0.18 | 576.4 | 1.9098
764.465 | 0.04 | 388.2 | 0.18 | 576.6 | 2.1524
764.706 | 0.06 | 388.3 | 0.18 | 576.8 | 2.4494
764.948 | 0.03 | 388.4 | 0.18 | 577 | 2.8005
765.189 | 0.05 | 388.5 | 0.18 | 577.2 | 2.7386
765.43 | 0.03 | 388.6 | 0.18 | 577.4 | 2.3563
765.672 | 0.04 | 388.7 | 0.18429 | 577.6 | 1.9241
765.913 | 0.05 | 388.8 | 0.18582 | 577.8 | 1.6628
766.154 | 0.03 | 388.9 | 0.18734 | 578 | 1.5424
766.395 | 0.04 | 389 | 0.1872 | 578.2 | 1.5978
766.636 | 0.02 | 389.1 | 0.18872 | 578.4 | 1.7479
766.877 | 0.03 | 389.2 | 0.19025 | 578.6 | 1.967
767.118 | 0.03 | 389.3 | 0.19344 | 578.8 | 2.2408
767.359 | 0.05 | 389.4 | 0.19496 | 579 | 2.4984
767.6 | 0.03 | 389.5 | 0.19649 | 579.2 | 2.6742
767.841 | 0.04 | 389.6 | 0.19801 | 579.4 | 2.4623
768.081 | 0.04 | 389.7 | 0.20241 | 579.6 | 2.0883
768.322 | 0.03 | 389.8 | 0.20546 | 579.8 | 1.7392
768.563 | 0.03 | 389.9 | 0.20851 | 580 | 1.5399
768.803 | 0.02 | 390 | 0.20745 | 580.2 | 1.5019
769.044 | 0.04 | 390.1 | 0.20898 | 580.4 | 1.4981
769.284 | 0.03 | 390.2 | 0.21051 | 580.6 | 1.4888
769.525 | 0.04 | 390.3 | 0.21369 | 580.8 | 1.4826
769.765 | 0.03 | 390.4 | 0.21522 | 581 | 1.4721
770.005 | 0.02 | 390.5 | 0.21675 | 581.2 | 1.4604
770.246 | 0.04 | 390.6 | 0.21828 | 581.4 | 1.4606
770.486 | 0.05 | 390.7 | 0.21333 | 581.6 | 1.4625
770.726 | 0.02 | 390.8 | 0.21333 | 581.8 | 1.4627
770.966 | 0.02 | 390.9 | 0.21333 | 582 | 1.4623
771.206 | 0.03 | 391 | 0.21229 | 582.2 | 1.4613
771.446 | 0.04 | 391.1 | 0.21076 | 582.4 | 1.4649
771.686 | 0.04 | 391.2 | 0.20923 | 582.6 | 1.4633
771.926 | 0.04 | 391.3 | 0.21 | 582.8 | 1.4656
772.166 | 0.04 | 391.4 | 0.21 | 583 | 1.4669
772.406 | 0.1 | 391.5 | 0.21 | 583.2 | 1.4691
772.645 | 0.11 | 391.6 | 0.21 | 583.4 | 1.4703
772.885 | 0.08 | 391.7 | 0.21 | 583.6 | 1.4685
773.125 | 0.06 | 391.8 | 0.21 | 583.8 | 1.4636
773.364 | 0.04 | 391.9 | 0.21 | 584 | 1.4612
773.604 | 0.03 | 392 | 0.21465 | 584.2 | 1.46
773.843 | 0.03 | 392.1 | 0.21618 | 584.4 | 1.4575
774.083 | 0.02 | 392.2 | 0.2177 | 584.6 | 1.4489
774.322 | 0.05 | 392.3 | 0.21757 | 584.8 | 1.4507
774.561 | 0.02 | 392.4 | 0.2191 | 585 | 1.4511
774.801 | 0.03 | 392.5 | 0.22063 | 585.2 | 1.4467
775.04 | 0.02 | 392.6 | 0.22 | 585.4 | 1.4399
775.279 | 0.03 | 392.7 | 0.22 | 585.6 | 1.433
775.518 | 0.02 | 392.8 | 0.22 | 585.8 | 1.4356
775.757 | 0.02 | 392.9 | 0.22341 | 586 | 1.4333
775.996 | 0.03 | 393 | 0.22494 | 586.2 | 1.4324
776.235 | 0.04 | 393.1 | 0.22647 | 586.4 | 1.4329
776.474 | 0.04 | 393.2 | 0.228 | 586.6 | 1.442
776.713 | 0.04 | 393.3 | 0.23239 | 586.8 | 1.4469
776.951 | 0.02 | 393.4 | 0.23544 | 587 | 1.454
777.19 | 0.03 | 393.5 | 0.2385 | 587.2 | 1.4622
777.429 | 0.02 | 393.6 | 0.23745 | 587.4 | 1.4671
777.667 | 0.02 | 393.7 | 0.23898 | 587.6 | 1.4609
777.906 | 0.02 | 393.8 | 0.2405 | 587.8 | 1.447
778.145 | 0.02 | 393.9 | 0.24 | 588 | 1.4307
778.383 | 0.02 | 394 | 0.24 | 588.2 | 1.4154
778.621 | 0.04 | 394.1 | 0.24 | 588.4 | 1.3999
778.86 | 0.03 | 394.2 | 0.24 | 588.6 | 1.3807
779.098 | 0.03 | 394.3 | 0.24483 | 588.8 | 1.376
779.336 | 0.02 | 394.4 | 0.24636 | 589 | 1.3743
779.575 | 0.03 | 394.5 | 0.24789 | 589.2 | 1.3744
779.813 | 0.02 | 394.6 | 0.24333 | 589.4 | 1.3767
780.051 | 0.03 | 394.7 | 0.24334 | 589.6 | 1.3779
780.289 | 0.03 | 394.8 | 0.24334 | 589.8 | 1.3787
780.527 | 0.03 | 394.9 | 0.24667 | 590 | 1.3736
780.765 | 0.01 | 395 | 0.24667 | 590.2 | 1.3672
781.003 | 0.03 | 395.1 | 0.24667 | 590.4 | 1.3626
781.241 | 0.03 | 395.2 | 0.25056 | 590.6 | 1.3603
781.478 | 0.03 | 395.3 | 0.25362 | 590.8 | 1.36
781.716 | 0.02 | 395.4 | 0.25669 | 591 | 1.3566
781.954 | 0.03 | 395.5 | 0.25975 | 591.2 | 1.3558
782.191 | 0.03 | 395.6 | 0.25807 | 591.4 | 1.363
782.429 | 0.03 | 395.7 | 0.2596 | 591.6 | 1.3594
782.667 | 0.02 | 395.8 | 0.26113 | 591.8 | 1.3582
782.904 | 0.02 | 395.9 | 0.26 | 592 | 1.3572
783.142 | 0.02 | 396 | 0.26 | 592.2 | 1.3589
783.379 | 0.02 | 396.1 | 0.26 | 592.4 | 1.3569
783.616 | 0.03 | 396.2 | 0.26393 | 592.6 | 1.3558
783.853 | 0.04 | 396.3 | 0.26547 | 592.8 | 1.3629
784.091 | 0.02 | 396.4 | 0.267 | 593 | 1.3632
784.328 | 0.03 | 396.5 | 0.26687 | 593.2 | 1.3598
784.565 | 0.03 | 396.6 | 0.2684 | 593.4 | 1.354
784.802 | 0.03 | 396.7 | 0.26993 | 593.6 | 1.3511
785.039 | 0.01 | 396.8 | 0.27147 | 593.8 | 1.3398
785.276 | 0.03 | 396.9 | 0.27467 | 594 | 1.3287
785.513 | 0.02 | 397 | 0.2762 | 594.2 | 1.3213
785.75 | 0.02 | 397.1 | 0.27774 | 594.4 | 1.311
785.987 | 0.02 | 397.2 | 0.2776 | 594.6 | 1.3026
786.223 | 0.03 | 397.3 | 0.27914 | 594.8 | 1.2932
786.46 | 0.05 | 397.4 | 0.28067 | 595 | 1.2856
786.697 | 0.01 | 397.5 | 0.28387 | 595.2 | 1.2818
786.933 | 0.02 | 397.6 | 0.2854 | 595.4 | 1.2784
787.17 | 0.02 | 397.7 | 0.28694 | 595.6 | 1.2737
787.406 | 0.03 | 397.8 | 0.2868 | 595.8 | 1.2713
787.643 | 0.04 | 397.9 | 0.28834 | 596 | 1.2672
787.879 | 0.04 | 398 | 0.28987 | 596.2 | 1.2623
788.116 | 0.04 | 398.1 | 0.29141 | 596.4 | 1.2556
788.352 | 0 | 398.2 | 0.29 | 596.6 | 1.2518
788.588 | 0.01 | 398.3 | 0.29 | 596.8 | 1.2484
788.824 | 0.01 | 398.4 | 0.29 | 597 | 1.2437
789.06 | 0.01 | 398.5 | 0.29421 | 597.2 | 1.2412
789.297 | 0.01 | 398.6 | 0.29575 | 597.4 | 1.2371
789.533 | 0.01 | 398.7 | 0.29728 | 597.6 | 1.2335
789.769 | 0.01 | 398.8 | 0.29715 | 597.8 | 1.2311
790.004 | 0.03 | 398.9 | 0.29869 | 598 | 1.2298
790.24 | 0.03 | 399 | 0.30022 | 598.2 | 1.2315
790.476 | 0.01 | 399.1 | 0.3 | 598.4 | 1.2363
790.712 | 0.02 | 399.2 | 0.3 | 598.6 | 1.2388
790.948 | 0.01 | 399.3 | 0.3 | 598.8 | 1.2371
791.183 | 0.01 | 399.4 | 0.3 | 599 | 1.2335
791.419 | 0.03 | 399.5 | 0.30913 | 599.2 | 1.231
791.655 | 0.03 | 399.6 | 0.3122 | 599.4 | 1.2332
791.89 | 0.02 | 399.7 | 0.31527 | 599.6 | 1.2326
792.126 | 0.01 | 399.8 | 0.31503 | 599.8 | 1.2205
792.361 | 0.03 | 399.9 | 0.3181 | 600 | 1.2098
792.596 | 0.03 | 400 | 0.32118 | 600.2 | 1.2023
792.832 | 0.01 | 400.1 | 0.32 | 600.4 | 1.2002
793.067 | 0.02 | 400.2 | 0.32 | 600.6 | 1.2
793.302 | 0.01 | 400.3 | 0.32 | 600.8 | 1.193
793.537 | 0.01 | 400.4 | 0.3234 | 601 | 1.1836
793.772 | 0.02 | 400.5 | 0.32494 | 601.2 | 1.1742
794.007 | 0.03 | 400.6 | 0.32647 | 601.4 | 1.1664
794.242 | 0.02 | 400.7 | 0.32801 | 601.6 | 1.1592
794.477 | 0.01 | 400.8 | 0.32788 | 601.8 | 1.152
794.712 | 0.02 | 400.9 | 0.32942 | 602 | 1.1448
794.947 | 0.08 | 401 | 0.33096 | 602.2 | 1.1412
795.182 | 0.09 | 401.1 | 0.33416 | 602.4 | 1.1396
795.417 | 0.03 | 401.2 | 0.3357 | 602.6 | 1.14
795.651 | 0.04 | 401.3 | 0.33724 | 602.8 | 1.14
795.886 | 0.01 | 401.4 | 0.34089 | 603 | 1.14
796.121 | 0.02 | 401.5 | 0.34397 | 603.2 | 1.1371
796.355 | 0.05 | 401.6 | 0.34705 | 603.4 | 1.1278
796.59 | 0.02 | 401.7 | 0.34673 | 603.6 | 1.1228
796.824 | 0.03 | 401.8 | 0.34827 | 603.8 | 1.1173
797.059 | 0.01 | 401.9 | 0.34981 | 604 | 1.1137
797.293 | 0.02 | 402 | 0.35134 | 604.2 | 1.1112
797.527 | 0.03 | 402.1 | 0.35455 | 604.4 | 1.1098
797.761 | 0.03 | 402.2 | 0.35609 | 604.6 | 1.1085
797.996 | 0.03 | 402.3 | 0.35763 | 604.8 | 1.1037
798.23 | 0.01 | 402.4 | 0.36167 | 605 | 1.0974
798.464 | 0.03 | 402.5 | 0.36476 | 605.2 | 1.0926
798.698 | 0.01 | 402.6 | 0.36784 | 605.4 | 1.0903
798.932 | 0.03 | 402.7 | 0.3785 | 605.6 | 1.09
799.166 | 0.01 | 402.8 | 0.38466 | 605.8 | 1.0865
799.4 | 0.03 | 402.9 | 0.39083 | 606 | 1.0811
799.634 | 0.04 | 403 | 0.42092 | 606.2 | 1.0736
799.867 | 0.04 | 403.1 | 0.43938 | 606.4 | 1.0704
800.101 | 0.04 | 403.2 | 0.45785 | 606.6 | 1.0699
800.335 | 0.01 | 403.3 | 0.47631 | 606.8 | 1.0672
800.568 | 0.02 | 403.4 | 0.68596 | 607 | 1.0581
800.802 | 0.03 | 403.5 | 0.77222 | 607.2 | 1.0604
801.036 | 0.08 | 403.6 | 0.85847 | 607.4 | 1.0628
801.269 | 0.04 | 403.7 | 1.464 | 607.6 | 1.0627
801.503 | 0.04 | 403.8 | 1.7645 | 607.8 | 1.0613
801.736 | 0.1 | 403.9 | 2.0651 | 608 | 1.0598
801.969 | 0.12 | 404 | 2.9683 | 608.2 | 1.06
802.203 | 0.05 | 404.1 | 3.5799 | 608.4 | 1.06
802.436 | 0.01 | 404.2 | 4.1916 | 608.6 | 1.06
802.669 | 0.02 | 404.3 | 4.4469 | 608.8 | 1.0571
802.902 | 0.03 | 404.4 | 4.9747 | 609 | 1.0574
803.135 | 0.03 | 404.5 | 5.5025 | 609.2 | 1.068
803.368 | 0.01 | 404.6 | 6.0303 | 609.4 | 1.0742
803.601 | 0.03 | 404.7 | 4.3322 | 609.6 | 1.0838
803.834 | 0.03 | 404.8 | 4.0283 | 609.8 | 1.093
804.067 | 0.01 | 404.9 | 3.7245 | 610 | 1.1222
804.3 | 0.02 | 405 | 2.7222 | 610.2 | 1.1709
804.533 | 0.02 | 405.1 | 1.9317 | 610.4 | 1.2347
804.765 | 0.02 | 405.2 | 1.1412 | 610.6 | 1.3246
804.998 | 0.03 | 405.3 | 1.2483 | 610.8 | 1.4478
805.231 | 0.03 | 405.4 | 0.86096 | 611 | 1.5829
805.463 | 0.01 | 405.5 | 0.47361 | 611.2 | 1.6843
805.696 | 0.02 | 405.6 | 0.52855 | 611.4 | 1.7321
805.928 | 0.03 | 405.7 | 0.48071 | 611.6 | 1.6864
806.161 | 0.01 | 405.8 | 0.43287 | 611.8 | 1.5751
806.393 | 0.04 | 405.9 | 0.38503 | 612 | 1.4234
806.625 | 0 | 406 | 0.43068 | 612.2 | 1.3111
806.858 | 0.03 | 406.1 | 0.42759 | 612.4 | 1.242
807.09 | 0.01 | 406.2 | 0.42451 | 612.6 | 1.194
807.322 | 0 | 406.3 | 0.42762 | 612.8 | 1.1608
807.554 | 0 | 406.4 | 0.42917 | 613 | 1.1414
807.786 | 0.01 | 406.5 | 0.43071 | 613.2 | 1.1245
|  | 406.6 | 0.45748 | 613.4 | 1.1047
|  | 406.7 | 0.46829 | 613.6 | 1.0899
|  | 406.8 | 0.47911 | 613.8 | 1.0815
|  | 406.9 | 0.52097 | 614 | 1.0785
|  | 407 | 0.55034 | 614.2 | 1.0738
|  | 407.1 | 0.57971 | 614.4 | 1.0674
|  | 407.2 | 0.60909 | 614.6 | 1.0568
|  | 407.3 | 0.68687 | 614.8 | 1.0445
|  | 407.4 | 0.73471 | 615 | 1.0303
|  | 407.5 | 0.78255 | 615.2 | 1.0123
|  | 407.6 | 0.7577 | 615.4 | 0.99377
|  | 407.7 | 0.78399 | 615.6 | 0.97614
|  | 407.8 | 0.81029 | 615.8 | 0.95685
|  | 407.9 | 0.69911 | 616 | 0.94301
|  | 408 | 0.65891 | 616.2 | 0.93007
|  | 408.1 | 0.61871 | 616.4 | 0.91805
|  | 408.2 | 0.59327 | 616.6 | 0.91208
|  | 408.3 | 0.5438 | 616.8 | 0.91
|  | 408.4 | 0.49433 | 617 | 0.90738
|  | 408.5 | 0.44486 | 617.2 | 0.8985
|  | 408.6 | 0.47982 | 617.4 | 0.88992
|  | 408.7 | 0.46589 | 617.6 | 0.88031
|  | 408.8 | 0.45196 | 617.8 | 0.87216
|  | 408.9 | 0.4622 | 618 | 0.86073
|  | 409 | 0.46065 | 618.2 | 0.85092
|  | 409.1 | 0.45911 | 618.4 | 0.84295
|  | 409.2 | 0.46411 | 618.6 | 0.84043
|  | 409.3 | 0.46566 | 618.8 | 0.84339
|  | 409.4 | 0.4672 | 619 | 0.84446
|  | 409.5 | 0.46708 | 619.2 | 0.84312
|  | 409.6 | 0.46863 | 619.4 | 0.84085
|  | 409.7 | 0.47018 | 619.6 | 0.83609
|  | 409.8 | 0.47 | 619.8 | 0.83409
|  | 409.9 | 0.47 | 620 | 0.83675
|  | 410 | 0.47 | 620.2 | 0.83873
|  | 410.1 | 0.47 | 620.4 | 0.83358
|  | 410.2 | 0.47917 | 620.6 | 0.82572
|  | 410.3 | 0.48227 | 620.8 | 0.82064
|  | 410.4 | 0.48537 | 621 | 0.81112
|  | 410.5 | 0.48936 | 621.2 | 0.80492
|  | 410.6 | 0.494 | 621.4 | 0.80036
|  | 410.7 | 0.49865 | 621.6 | 0.8
|  | 410.8 | 0.49333 | 621.8 | 0.8
|  | 410.9 | 0.49334 | 622 | 0.79661
|  | 411 | 0.49334 | 622.2 | 0.78853
|  | 411.1 | 0.49667 | 622.4 | 0.78776
|  | 411.2 | 0.49667 | 622.6 | 0.7842
|  | 411.3 | 0.49667 | 622.8 | 0.78287
|  | 411.4 | 0.49667 | 623 | 0.78069
|  | 411.5 | 0.49806 | 623.2 | 0.78
|  | 411.6 | 0.49961 | 623.4 | 0.77682
|  | 411.7 | 0.50116 | 623.6 | 0.77305
|  | 411.8 | 0.5 | 623.8 | 0.77049
|  | 411.9 | 0.5 | 624 | 0.7619
|  | 412 | 0.5 | 624.2 | 0.75818
|  | 412.1 | 0.50403 | 624.4 | 0.75452
|  | 412.2 | 0.50558 | 624.6 | 0.74753
|  | 412.3 | 0.50713 | 624.8 | 0.74002
|  | 412.4 | 0.50701 | 625 | 0.73004
|  | 412.5 | 0.50856 | 625.2 | 0.72549
|  | 412.6 | 0.51012 | 625.4 | 0.72667
|  | 412.7 | 0.51167 | 625.6 | 0.72444
|  | 412.8 | 0.51489 | 625.8 | 0.72136
|  | 412.9 | 0.51644 | 626 | 0.71375
|  | 413 | 0.51799 | 626.2 | 0.72656
|  | 413.1 | 0.52241 | 626.4 | 0.72494
|  | 413.2 | 0.52551 | 626.6 | 0.71658
|  | 413.3 | 0.52861 | 626.8 | 0.70881
|  | 413.4 | 0.52752 | 627 | 0.70652
|  | 413.5 | 0.52908 | 627.2 | 0.70706
|  | 413.6 | 0.53063 | 627.4 | 0.70077
|  | 413.7 | 0.53385 | 627.6 | 0.68925
|  | 413.8 | 0.5354 | 627.8 | 0.68654
|  | 413.9 | 0.53695 | 628 | 0.68518
|  | 414 | 0.53684 | 628.2 | 0.68323
|  | 414.1 | 0.53839 | 628.4 | 0.68118
|  | 414.2 | 0.53994 | 628.6 | 0.67639
|  | 414.3 | 0.5415 | 628.8 | 0.67116
|  | 414.4 | 0.54 | 629 | 0.66587
|  | 414.5 | 0.54 | 629.2 | 0.66667
|  | 414.6 | 0.54 | 629.4 | 0.66444
|  | 414.7 | 0.54437 | 629.6 | 0.66444
|  | 414.8 | 0.54593 | 629.8 | 0.66667
|  | 414.9 | 0.54748 | 630 | 0.6722
|  | 415 | 0.54737 | 630.2 | 0.67749
|  | 415.1 | 0.54892 | 630.4 | 0.68115
|  | 415.2 | 0.55047 | 630.6 | 0.68594
|  | 415.3 | 0.5537 | 630.8 | 0.6959
|  | 415.4 | 0.55525 | 631 | 0.70253
|  | 415.5 | 0.5568 | 631.2 | 0.69379
|  | 415.6 | 0.55333 | 631.4 | 0.68067
|  | 415.7 | 0.55333 | 631.6 | 0.66241
|  | 415.8 | 0.55333 | 631.8 | 0.64372
|  | 415.9 | 0.55333 | 632 | 0.62824
|  | 416 | 0.55209 | 632.2 | 0.61713
|  | 416.1 | 0.55054 | 632.4 | 0.60495
|  | 416.2 | 0.54899 | 632.6 | 0.59389
|  | 416.3 | 0.55 | 632.8 | 0.58396
|  | 416.4 | 0.55 | 633 | 0.58392
|  | 416.5 | 0.55 | 633.2 | 0.58127
|  | 416.6 | 0.55781 | 633.4 | 0.57748
|  | 416.7 | 0.56092 | 633.6 | 0.57712
|  | 416.8 | 0.56403 | 633.8 | 0.57398
|  | 416.9 | 0.56381 | 634 | 0.56954
|  | 417 | 0.56692 | 634.2 | 0.56348
|  | 417.1 | 0.57003 | 634.4 | 0.56069
|  | 417.2 | 0.57315 | 634.6 | 0.55589
|  | 417.3 | 0.5748 | 634.8 | 0.55275
|  | 417.4 | 0.57636 | 635 | 0.54905
|  | 417.5 | 0.57791 | 635.2 | 0.54425
|  | 417.6 | 0.58227 | 635.4 | 0.54166
|  | 417.7 | 0.58539 | 635.6 | 0.53963
|  | 417.8 | 0.5885 | 635.8 | 0.53877
|  | 417.9 | 0.58333 | 636 | 0.53446
|  | 418 | 0.58333 | 636.2 | 0.53188
|  | 418.1 | 0.58333 | 636.4 | 0.52818
|  | 418.2 | 0.58286 | 636.6 | 0.52509
|  | 418.3 | 0.58131 | 636.8 | 0.52667
|  | 418.4 | 0.57975 | 637 | 0.52444
|  | 418.5 | 0.58347 | 637.2 | 0.52292
|  | 418.6 | 0.58503 | 637.4 | 0.52024
|  | 418.7 | 0.58659 | 637.6 | 0.51543
|  | 418.8 | 0.58815 | 637.8 | 0.51228
|  | 418.9 | 0.59276 | 638 | 0.50857
|  | 419 | 0.59588 | 638.2 | 0.50376
|  | 419.1 | 0.599 | 638.4 | 0.50117
|  | 419.2 | 0.60211 | 638.6 | 0.49996
|  | 419.3 | 0.60523 | 638.8 | 0.5
|  | 419.4 | 0.60835 | 639 | 0.5
|  | 419.5 | 0.6074 | 639.2 | 0.4962
|  | 419.6 | 0.60896 | 639.4 | 0.48969
|  | 419.7 | 0.61051 | 639.6 | 0.48364
|  | 419.8 | 0.61 | 639.8 | 0.48077
|  | 419.9 | 0.61 | 640 | 0.47964
|  | 420 | 0.61 | 640.2 | 0.47827
|  | 420.1 | 0.61342 | 640.4 | 0.47223
|  | 420.2 | 0.61498 | 640.6 | 0.46519
|  | 420.3 | 0.61653 | 640.8 | 0.46135
|  | 420.4 | 0.61809 | 641 | 0.46129
|  | 420.5 | 0.62264 | 641.2 | 0.46507
|  | 420.6 | 0.62576 | 641.4 | 0.46334
|  | 420.7 | 0.62888 | 641.6 | 0.46207
|  | 420.8 | 0.62767 | 641.8 | 0.45972
|  | 420.9 | 0.62923 | 642 | 0.4588
|  | 421 | 0.63079 | 642.2 | 0.45455
|  | 421.1 | 0.63 | 642.4 | 0.45194
|  | 421.2 | 0.63 | 642.6 | 0.44981
|  | 421.3 | 0.63 | 642.8 | 0.45
|  | 421.4 | 0.63 | 643 | 0.45
|  | 421.5 | 0.63 | 643.2 | 0.44669
|  | 421.6 | 0.63 | 643.4 | 0.44125
|  | 421.7 | 0.63 | 643.6 | 0.43593
|  | 421.8 | 0.63 | 643.8 | 0.43667
|  | 421.9 | 0.63 | 644 | 0.43768
|  | 422 | 0.63 | 644.2 | 0.43703
|  | 422.1 | 0.63 | 644.4 | 0.42725
|  | 422.2 | 0.63 | 644.6 | 0.42203
|  | 422.3 | 0.63 | 644.8 | 0.41981
|  | 422.4 | 0.63864 | 645 | 0.42
|  | 422.5 | 0.64176 | 645.2 | 0.42
|  | 422.6 | 0.64489 | 645.4 | 0.42
|  | 422.7 | 0.64468 | 645.6 | 0.41689
|  | 422.8 | 0.6478 | 645.8 | 0.4131
|  | 422.9 | 0.65093 | 646 | 0.41048
|  | 423 | 0.65369 | 646.2 | 0.40954
|  | 423.1 | 0.65526 | 646.4 | 0.41
|  | 423.2 | 0.65682 | 646.6 | 0.40718
|  | 423.3 | 0.65671 | 646.8 | 0.39781
|  | 423.4 | 0.65828 | 647 | 0.39257
|  | 423.5 | 0.65984 | 647.2 | 0.39298
|  | 423.6 | 0.6614 | 647.4 | 0.39396
|  | 423.7 | 0.66463 | 647.6 | 0.38729
|  | 423.8 | 0.66619 | 647.8 | 0.38689
|  | 423.9 | 0.66776 | 648 | 0.38741
|  | 424 | 0.67197 | 648.2 | 0.38976
|  | 424.1 | 0.67509 | 648.4 | 0.39
|  | 424.2 | 0.67822 | 648.6 | 0.39
|  | 424.3 | 0.67333 | 648.8 | 0.39
|  | 424.4 | 0.67334 | 649 | 0.38865
|  | 424.5 | 0.67334 | 649.2 | 0.38407
|  | 424.6 | 0.67667 | 649.4 | 0.38144
|  | 424.7 | 0.67667 | 649.6 | 0.38004
|  | 424.8 | 0.67667 | 649.8 | 0.38134
|  | 424.9 | 0.68014 | 650 | 0.38588
|  | 425 | 0.68327 | 650.2 | 0.38851
|  | 425.1 | 0.6864 | 650.4 | 0.38994
|  | 425.2 | 0.68953 | 650.6 | 0.38867
|  | 425.3 | 0.68799 | 650.8 | 0.38414
|  | 425.4 | 0.68956 | 651 | 0.37781
|  | 425.5 | 0.69112 | 651.2 | 0.36996
|  | 425.6 | 0.69 | 651.4 | 0.3638
|  | 425.7 | 0.69 | 651.6 | 0.36082
|  | 425.8 | 0.69 | 651.8 | 0.35965
|  | 425.9 | 0.69405 | 652 | 0.36
|  | 426 | 0.69562 | 652.2 | 0.36
|  | 426.1 | 0.69719 | 652.4 | 0.36
|  | 426.2 | 0.69333 | 652.6 | 0.36
|  | 426.3 | 0.69333 | 652.8 | 0.35653
|  | 426.4 | 0.69333 | 653 | 0.34827
|  | 426.5 | 0.68977 | 653.2 | 0.34749
|  | 426.6 | 0.68664 | 653.4 | 0.34409
|  | 426.7 | 0.68351 | 653.6 | 0.3427
|  | 426.8 | 0.68038 | 653.8 | 0.33902
|  | 426.9 | 0.68195 | 654 | 0.33416
|  | 427 | 0.68039 | 654.2 | 0.33151
|  | 427.1 | 0.67882 | 654.4 | 0.33006
|  | 427.2 | 0.68883 | 654.6 | 0.32867
|  | 427.3 | 0.69196 | 654.8 | 0.32413
|  | 427.4 | 0.6951 | 655 | 0.31777
|  | 427.5 | 0.70313 | 655.2 | 0.31566
|  | 427.6 | 0.7094 | 655.4 | 0.31556
|  | 427.7 | 0.71567 | 655.6 | 0.31333
|  | 427.8 | 0.71814 | 655.8 | 0.31182
|  | 427.9 | 0.72285 | 656 | 0.31004
|  | 428 | 0.72756 | 656.2 | 0.31
|  | 428.1 | 0.72687 | 656.4 | 0.31
|  | 428.2 | 0.72843 | 656.6 | 0.31
|  | 428.3 | 0.73 | 656.8 | 0.31
|  | 428.4 | 0.73157 | 657 | 0.31
|  | 428.5 | 0.72519 | 657.2 | 0.31
|  | 428.6 | 0.72363 | 657.4 | 0.31
|  | 428.7 | 0.72206 | 657.6 | 0.31
|  | 428.8 | 0.72667 | 657.8 | 0.30722
|  | 428.9 | 0.72666 | 658 | 0.29788
|  | 429 | 0.72666 | 658.2 | 0.29641
|  | 429.1 | 0.73178 | 658.4 | 0.29745
|  | 429.2 | 0.73492 | 658.6 | 0.29842
|  | 429.3 | 0.73806 | 658.8 | 0.29387
|  | 429.4 | 0.7412 | 659 | 0.29121
|  | 429.5 | 0.74434 | 659.2 | 0.29304
|  | 429.6 | 0.74748 | 659.4 | 0.29545
|  | 429.7 | 0.75062 | 659.6 | 0.29333
|  | 429.8 | 0.75376 | 659.8 | 0.28764
|  | 429.9 | 0.7569 | 660 | 0.28227
|  | 430 | 0.76004 | 660.2 | 0.27857
|  | 430.1 | 0.76318 | 660.4 | 0.27369
|  | 430.2 | 0.76632 | 660.6 | 0.27504
|  | 430.3 | 0.76946 | 660.8 | 0.27765
|  | 430.4 | 0.77261 | 661 | 0.277
|  | 430.5 | 0.77575 | 661.2 | 0.26716
|  | 430.6 | 0.7789 | 661.4 | 0.27002
|  | 430.7 | 0.78204 | 661.6 | 0.27545
|  | 430.8 | 0.78519 | 661.8 | 0.28018
|  | 430.9 | 0.78833 | 662 | 0.28
|  | 431 | 0.79148 | 662.2 | 0.28
|  | 431.1 | 0.79462 | 662.4 | 0.28
|  | 431.2 | 0.79777 | 662.6 | 0.28
|  | 431.3 | 0.79712 | 662.8 | 0.27763
|  | 431.4 | 0.79869 | 663 | 0.26911
|  | 431.5 | 0.80027 | 663.2 | 0.26721
|  | 431.6 | 0.79649 | 663.4 | 0.26529
|  | 431.7 | 0.79492 | 663.6 | 0.26321
|  | 431.8 | 0.79335 | 663.8 | 0.26108
|  | 431.9 | 0.79178 | 664 | 0.25618
|  | 432 | 0.78708 | 664.2 | 0.25408
|  | 432.1 | 0.78393 | 664.4 | 0.25684
|  | 432.2 | 0.78079 | 664.6 | 0.25907
|  | 432.3 | 0.78215 | 664.8 | 0.25667
|  | 432.4 | 0.78058 | 665 | 0.25224
|  | 432.5 | 0.77901 | 665.2 | 0.24901
|  | 432.6 | 0.79269 | 665.4 | 0.24411
|  | 432.7 | 0.79741 | 665.6 | 0.24142
|  | 432.8 | 0.80212 | 665.8 | 0.24003
|  | 432.9 | 0.81372 | 666 | 0.24
|  | 433 | 0.82317 | 666.2 | 0.24
|  | 433.1 | 0.83261 | 666.4 | 0.24
|  | 433.2 | 0.84205 | 666.6 | 0.24
|  | 433.3 | 0.8515 | 666.8 | 0.24
|  | 433.4 | 0.86094 | 667 | 0.24
|  | 433.5 | 0.87381 | 667.2 | 0.244
|  | 433.6 | 0.88483 | 667.4 | 0.24134
|  | 433.7 | 0.89586 | 667.6 | 0.23584
|  | 433.8 | 0.90688 | 667.8 | 0.23046
|  | 433.9 | 0.93502 | 668 | 0.22889
|  | 434 | 0.95234 | 668.2 | 0.23
|  | 434.1 | 0.96967 | 668.4 | 0.23
|  | 434.2 | 1.0351 | 668.6 | 0.22882
|  | 434.3 | 1.0713 | 668.8 | 0.22456
|  | 434.4 | 1.1075 | 669 | 0.22187
|  | 434.5 | 1.4695 | 669.2 | 0.21971
|  | 434.6 | 1.6351 | 669.4 | 0.21875
|  | 434.7 | 1.8007 | 669.6 | 0.21434
|  | 434.8 | 2.9636 | 669.8 | 0.21526
|  | 434.9 | 3.5835 | 670 | 0.21556
|  | 435 | 4.2034 | 670.2 | 0.21577
|  | 435.1 | 5.6622 | 670.4 | 0.22255
|  | 435.2 | 6.8042 | 670.6 | 0.22303
|  | 435.3 | 7.9461 | 670.8 | 0.21959
|  | 435.4 | 9.961 | 671 | 0.21345
|  | 435.5 | 11.6052 | 671.2 | 0.21054
|  | 435.6 | 13.2495 | 671.4 | 0.21347
|  | 435.7 | 14.8937 | 671.6 | 0.2145
|  | 435.8 | 13.7706 | 671.8 | 0.21147
|  | 435.9 | 14.592 | 672 | 0.20393
|  | 436 | 15.4134 | 672.2 | 0.20448
|  | 436.1 | 10.1929 | 672.4 | 0.20775
|  | 436.2 | 8.4022 | 672.6 | 0.20856
|  | 436.3 | 6.6116 | 672.8 | 0.20392
|  | 436.4 | 5.566 | 673 | 0.20122
|  | 436.5 | 3.4714 | 673.2 | 0.19977
|  | 436.6 | 1.3767 | 673.4 | 0.2
|  | 436.7 | 1.9347 | 673.6 | 0.19873
|  | 436.8 | 1.4236 | 673.8 | 0.19428
|  | 436.9 | 0.91261 | 674 | 0.19524
|  | 437 | 0.9991 | 674.2 | 0.19732
|  | 437.1 | 0.95178 | 674.4 | 0.1997
|  | 437.2 | 0.90446 | 674.6 | 0.2
|  | 437.3 | 0.85714 | 674.8 | 0.19615
|  | 437.4 | 0.89882 | 675 | 0.19247
|  | 437.5 | 0.89091 | 675.2 | 0.19156
|  | 437.6 | 0.88301 | 675.4 | 0.19629
|  | 437.7 | 0.89 | 675.6 | 0.19496
|  | 437.8 | 0.89 | 675.8 | 0.19556
|  | 437.9 | 0.89 | 676 | 0.19712
|  | 438 | 0.89438 | 676.2 | 0.19698
|  | 438.1 | 0.89596 | 676.4 | 0.18916
|  | 438.2 | 0.89754 | 676.6 | 0.18719
|  | 438.3 | 0.90158 | 676.8 | 0.18692
|  | 438.4 | 0.90475 | 677 | 0.18953
|  | 438.5 | 0.90791 | 677.2 | 0.19
|  | 438.6 | 0.9072 | 677.4 | 0.17896
|  | 438.7 | 0.90878 | 677.6 | 0.17976
|  | 438.8 | 0.91036 | 677.8 | 0.18334
|  | 438.9 | 0.91721 | 678 | 0.17939
|  | 439 | 0.92037 | 678.2 | 0.17538
|  | 439.1 | 0.92353 | 678.4 | 0.16985
|  | 439.2 | 0.92336 | 678.6 | 0.17149
|  | 439.3 | 0.92653 | 678.8 | 0.1764
|  | 439.4 | 0.92969 | 679 | 0.17913
|  | 439.5 | 0.93286 | 679.2 | 0.18017
|  | 439.6 | 0.93468 | 679.4 | 0.17839
|  | 439.7 | 0.93626 | 679.6 | 0.17503
|  | 439.8 | 0.93784 | 679.8 | 0.17667
|  | 439.9 | 0.94218 | 680 | 0.17445
|  | 440 | 0.94535 | 680.2 | 0.17272
|  | 440.1 | 0.94851 | 680.4 | 0.17035
|  | 440.2 | 0.94751 | 680.6 | 0.17
|  | 440.3 | 0.94909 | 680.8 | 0.17
|  | 440.4 | 0.95067 | 681 | 0.17
|  | 440.5 | 0.95 | 681.2 | 0.16853
|  | 440.6 | 0.95 | 681.4 | 0.16368
|  | 440.7 | 0.95 | 681.6 | 0.16503
|  | 440.8 | 0.95 | 681.8 | 0.16444
|  | 440.9 | 0.95 | 682 | 0.16444
|  | 441 | 0.95 | 682.2 | 0.16556
|  | 441.1 | 0.94658 | 682.4 | 0.16334
|  | 441.2 | 0.945 | 682.6 | 0.16198
|  | 441.3 | 0.94341 | 682.8 | 0.15958
|  | 441.4 | 0.94183 | 683 | 0.16
|  | 441.5 | 0.94667 | 683.2 | 0.16
|  | 441.6 | 0.94667 | 683.4 | 0.15616
|  | 441.7 | 0.94667 | 683.6 | 0.14938
|  | 441.8 | 0.94784 | 683.8 | 0.14609
|  | 441.9 | 0.94942 | 684 | 0.15303
|  | 442 | 0.951 | 684.2 | 0.15354
|  | 442.1 | 0.95426 | 684.4 | 0.15571
|  | 442.2 | 0.95584 | 684.6 | 0.15394
|  | 442.3 | 0.95743 | 684.8 | 0.14822
|  | 442.4 | 0.95735 | 685 | 0.1477
|  | 442.5 | 0.95894 | 685.2 | 0.14778
|  | 442.6 | 0.96052 | 685.4 | 0.14704
|  | 442.7 | 0.96 | 685.6 | 0.14334
|  | 442.8 | 0.96 | 685.8 | 0.14059
|  | 442.9 | 0.96 | 686 | 0.14348
|  | 443 | 0.96353 | 686.2 | 0.14767
|  | 443.1 | 0.96511 | 686.4 | 0.14848
|  | 443.2 | 0.9667 | 686.6 | 0.14634
|  | 443.3 | 0.96828 | 686.8 | 0.15205
|  | 443.4 | 0.96821 | 687 | 0.15605
|  | 443.5 | 0.96979 | 687.2 | 0.1623
|  | 443.6 | 0.97138 | 687.4 | 0.16381
|  | 443.7 | 0.97927 | 687.6 | 0.15751
|  | 443.8 | 0.98244 | 687.8 | 0.1532
|  | 443.9 | 0.98562 | 688 | 0.15304
|  | 444 | 0.98106 | 688.2 | 0.15542
|  | 444.1 | 0.98265 | 688.4 | 0.15334
|  | 444.2 | 0.98424 | 688.6 | 0.15556
|  | 444.3 | 0.97835 | 688.8 | 0.15111
|  | 444.4 | 0.97517 | 689 | 0.14549
|  | 444.5 | 0.972 | 689.2 | 0.14115
|  | 444.6 | 0.97667 | 689.4 | 0.14
|  | 444.7 | 0.97667 | 689.6 | 0.14
|  | 444.8 | 0.97667 | 689.8 | 0.14
|  | 444.9 | 0.9807 | 690 | 0.14
|  | 445 | 0.98387 | 690.2 | 0.14
|  | 445.1 | 0.98705 | 690.4 | 0.16387
|  | 445.2 | 0.98678 | 690.6 | 0.18981
|  | 445.3 | 0.98837 | 690.8 | 0.19828
|  | 445.4 | 0.98995 | 691 | 0.1705
|  | 445.5 | 0.99154 | 691.2 | 0.14588
|  | 445.6 | 0.9852 | 691.4 | 0.13185
|  | 445.7 | 0.98361 | 691.6 | 0.12915
|  | 445.8 | 0.98202 | 691.8 | 0.13
|  | 445.9 | 0.9821 | 692 | 0.13
|  | 446 | 0.98051 | 692.2 | 0.12608
|  | 446.1 | 0.97892 | 692.4 | 0.12549
|  | 446.2 | 0.98 | 692.6 | 0.12556
|  | 446.3 | 0.98 | 692.8 | 0.12678
|  | 446.4 | 0.98 | 693 | 0.13756
|  | 446.5 | 0.98 | 693.2 | 0.13735
|  | 446.6 | 0.98 | 693.4 | 0.13757
|  | 446.7 | 0.98 | 693.6 | 0.13592
|  | 446.8 | 0.98 | 693.8 | 0.13334
|  | 446.9 | 0.98 | 694 | 0.13166
|  | 447 | 0.98 | 694.2 | 0.12995
|  | 447.1 | 0.98 | 694.4 | 0.12847
|  | 447.2 | 0.98 | 694.6 | 0.12286
|  | 447.3 | 0.98 | 694.8 | 0.11583
|  | 447.4 | 0.98 | 695 | 0.11509
|  | 447.5 | 0.98 | 695.2 | 0.11898
|  | 447.6 | 0.98 | 695.4 | 0.12482
|  | 447.7 | 0.98 | 695.6 | 0.12333
|  | 447.8 | 0.98478 | 695.8 | 0.12165
|  | 447.9 | 0.98637 | 696 | 0.12311
|  | 448 | 0.98797 | 696.2 | 0.1302
|  | 448.1 | 0.99245 | 696.4 | 0.14013
|  | 448.2 | 0.99564 | 696.6 | 0.14212
|  | 448.3 | 0.99882 | 696.8 | 0.1387
|  | 448.4 | 0.99333 | 697 | 0.12843
|  | 448.5 | 0.99333 | 697.2 | 0.11851
|  | 448.6 | 0.99333 | 697.4 | 0.11728
|  | 448.7 | 0.99667 | 697.6 | 0.11794
|  | 448.8 | 0.99667 | 697.8 | 0.12008
|  | 448.9 | 0.99667 | 698 | 0.11843
|  | 449 | 0.99722 | 698.2 | 0.1139
|  | 449.1 | 0.99882 | 698.4 | 0.11112
|  | 449.2 | 1.0004 | 698.6 | 0.10972
|  | 449.3 | 1 | 698.8 | 0.11
|  | 449.4 | 1 | 699 | 0.11
|  | 449.5 | 1 | 699.2 | 0.11
|  | 449.6 | 1 | 699.4 | 0.11
|  | 449.7 | 1 | 699.6 | 0.11
|  | 449.8 | 1 | 699.8 | 0.11
|  | 449.9 | 1 | 700 | 0.11
|  | 450 | 1.0048 | 700.2 | 0.11
|  | 450.1 | 1.0064 | 700.4 | 0.10628
|  | 450.2 | 1.008 | 700.6 | 0.10559
|  | 450.3 | 1.0079 | 700.8 | 0.10701
|  | 450.4 | 1.0095 | 701 | 0.10967
|  | 450.5 | 1.0111 | 701.2 | 0.10224
|  | 450.6 | 1.01 | 701.4 | 0.10418
|  | 450.7 | 1.01 | 701.6 | 0.10862
|  | 450.8 | 1.01 | 701.8 | 0.10713
|  | 450.9 | 1.0142 | 702 | 0.10742
|  | 451 | 1.0158 | 702.2 | 0.1079
|  | 451.1 | 1.0174 | 702.4 | 0.11003
|  | 451.2 | 1.0133 | 702.6 | 0.10846
|  | 451.3 | 1.0133 | 702.8 | 0.10399
|  | 451.4 | 1.0133 | 703 | 0.10118
|  | 451.5 | 1.0204 | 703.2 | 0.10333
|  | 451.6 | 1.022 | 703.4 | 0.10601
|  | 451.7 | 1.0236 | 703.6 | 0.10305
|  | 451.8 | 1.0202 | 703.8 | 0.10051
|  | 451.9 | 1.0218 | 704 | 0.099516
|  | 452 | 1.0234 | 704.2 | 0.10332
|  | 452.1 | 1.025 | 704.4 | 0.10737
|  | 452.2 | 1.0267 | 704.6 | 0.10947
|  | 452.3 | 1.0267 | 704.8 | 0.11
|  | 452.4 | 1.0267 | 705 | 0.10621
|  | 452.5 | 1.028 | 705.2 | 0.10556
|  | 452.6 | 1.0296 | 705.4 | 0.10857
|  | 452.7 | 1.0312 | 705.6 | 0.1163
|  | 452.8 | 1.03 | 705.8 | 0.12415
|  | 452.9 | 1.03 | 706 | 0.12847
|  | 453 | 1.03 | 706.2 | 0.13565
|  | 453.1 | 1.03 | 706.4 | 0.15132
|  | 453.2 | 1.03 | 706.6 | 0.16452
|  | 453.3 | 1.03 | 706.8 | 0.16232
|  | 453.4 | 1.03 | 707 | 0.15105
|  | 453.5 | 1.03 | 707.2 | 0.14228
|  | 453.6 | 1.03 | 707.4 | 0.13798
|  | 453.7 | 1.0339 | 707.6 | 0.12841
|  | 453.8 | 1.0355 | 707.8 | 0.12658
|  | 453.9 | 1.0371 | 708 | 0.12753
|  | 454 | 1.037 | 708.2 | 0.12844
|  | 454.1 | 1.0386 | 708.4 | 0.12636
|  | 454.2 | 1.0402 | 708.6 | 0.13213
|  | 454.3 | 1.04 | 708.8 | 0.14349
|  | 454.4 | 1.04 | 709 | 0.1547
|  | 454.5 | 1.04 | 709.2 | 0.14995
|  | 454.6 | 1.04 | 709.4 | 0.13799
|  | 454.7 | 1.0449 | 709.6 | 0.11908
|  | 454.8 | 1.0465 | 709.8 | 0.10572
|  | 454.9 | 1.0481 | 710 | 0.099226
|  | 455 | 1.048 | 710.2 | 0.10264
|  | 455.1 | 1.0496 | 710.4 | 0.11186
|  | 455.2 | 1.0512 | 710.6 | 0.10972
|  | 455.3 | 1.05 | 710.8 | 0.10471
|  | 455.4 | 1.05 | 711 | 0.099854
|  | 455.5 | 1.05 | 711.2 | 0.12336
|  | 455.6 | 1.05 | 711.4 | 0.20764
|  | 455.7 | 1.05 | 711.6 | 0.20177
|  | 455.8 | 1.05 | 711.8 | 0.17334
|  | 455.9 | 1.05 | 712 | 0.13854
|  | 456 | 1.05 | 712.2 | 0.13806
|  | 456.1 | 1.05 | 712.4 | 0.13459
|  | 456.2 | 1.0539 | 712.6 | 0.12462
|  | 456.3 | 1.0555 | 712.8 | 0.11052
|  | 456.4 | 1.0571 | 713 | 0.098511
|  | 456.5 | 1.0607 | 713.2 | 0.097115
|  | 456.6 | 1.0639 | 713.4 | 0.098041
|  | 456.7 | 1.0671 | 713.6 | 0.10016
|  | 456.8 | 1.0668 | 713.8 | 0.098316
|  | 456.9 | 1.0684 | 714 | 0.092186
|  | 457 | 1.07 | 714.2 | 0.08468
|  | 457.1 | 1.0716 | 714.4 | 0.084687
|  | 457.2 | 1.0749 | 714.6 | 0.087716
|  | 457.3 | 1.0765 | 714.8 | 0.088529
|  | 457.4 | 1.0781 | 715 | 0.083722
|  | 457.5 | 1.0733 | 715.2 | 0.08087
|  | 457.6 | 1.0733 | 715.4 | 0.0758
|  | 457.7 | 1.0733 | 715.6 | 0.075532
|  | 457.8 | 1.0721 | 715.8 | 0.077103
|  | 457.9 | 1.0705 | 716 | 0.079317
|  | 458 | 1.0689 | 716.2 | 0.08
|  | 458.1 | 1.0743 | 716.4 | 0.083704
|  | 458.2 | 1.0759 | 716.6 | 0.087421
|  | 458.3 | 1.0775 | 716.8 | 0.088399
|  | 458.4 | 1.0733 | 717 | 0.084087
|  | 458.5 | 1.0733 | 717.2 | 0.08123
|  | 458.6 | 1.0733 | 717.4 | 0.076151
|  | 458.7 | 1.0767 | 717.6 | 0.075648
|  | 458.8 | 1.0767 | 717.8 | 0.075556
|  | 458.9 | 1.0767 | 718 | 0.073333
|  | 459 | 1.0809 | 718.2 | 0.071376
|  | 459.1 | 1.0841 | 718.4 | 0.069802
|  | 459.2 | 1.0873 | 718.6 | 0.07
|  | 459.3 | 1.0833 | 718.8 | 0.071396
|  | 459.4 | 1.0833 | 719 | 0.076171
|  | 459.5 | 1.0833 | 719.2 | 0.079037
|  | 459.6 | 1.0833 | 719.4 | 0.080158
|  | 459.7 | 1.0817 | 719.6 | 0.08
|  | 459.8 | 1.08 | 719.8 | 0.08
|  | 459.9 | 1.0784 | 720 | 0.08
|  | 460 | 1.0848 | 720.2 | 0.08
|  | 460.1 | 1.0864 | 720.4 | 0.08
|  | 460.2 | 1.088 | 720.6 | 0.08
|  | 460.3 | 1.0926 | 720.8 | 0.078681
|  | 460.4 | 1.0959 | 721 | 0.074057
|  | 460.5 | 1.0991 | 721.2 | 0.071186
|  | 460.6 | 1.0933 | 721.4 | 0.073187
|  | 460.7 | 1.0933 | 721.6 | 0.07734
|  | 460.8 | 1.0933 | 721.8 | 0.078187
|  | 460.9 | 1.0967 | 722 | 0.074137
|  | 461 | 1.0967 | 722.2 | 0.075159
|  | 461.1 | 1.0967 | 722.4 | 0.071226
|  | 461.2 | 1.0975 | 722.6 | 0.065575
|  | 461.3 | 1.0991 | 722.8 | 0.061133
|  | 461.4 | 1.1007 | 723 | 0.06
|  | 461.5 | 1.1 | 723.2 | 0.06
|  | 461.6 | 1.1 | 723.4 | 0.06
|  | 461.7 | 1.1 | 723.6 | 0.061588
|  | 461.8 | 1.0623 | 723.8 | 0.066207
|  | 461.9 | 1.0462 | 724 | 0.06909
|  | 462 | 1.0302 | 724.2 | 0.070362
|  | 462.1 | 1.0667 | 724.4 | 0.07
|  | 462.2 | 1.0667 | 724.6 | 0.07
|  | 462.3 | 1.0667 | 724.8 | 0.07
|  | 462.4 | 1.0709 | 725 | 0.07
|  | 462.5 | 1.0886 | 725.2 | 0.066191
|  | 462.6 | 1.1063 | 725.4 | 0.062398
|  | 462.7 | 1.1239 | 725.6 | 0.063111
|  | 462.8 | 1.1082 | 725.8 | 0.068602
|  | 462.9 | 1.1098 | 726 | 0.060953
|  | 463 | 1.1114 | 726.2 | 0.066318
|  | 463.1 | 1.11 | 726.4 | 0.066476
|  | 463.2 | 1.11 | 726.6 | 0.06892
|  | 463.3 | 1.11 | 726.8 | 0.069873
|  | 463.4 | 1.11 | 727 | 0.060953
|  | 463.5 | 1.11 | 727.2 | 0.062509
|  | 463.6 | 1.11 | 727.4 | 0.062016
|  | 463.7 | 1.11 | 727.6 | 0.062901
|  | 463.8 | 1.11 | 727.8 | 0.063112
|  | 463.9 | 1.11 | 728 | 0.071466
|  | 464 | 1.1142 | 728.2 | 0.065779
|  | 464.1 | 1.1158 | 728.4 | 0.066648
|  | 464.2 | 1.1174 | 728.6 | 0.069852
|  | 464.3 | 1.1173 | 728.8 | 0.071991
|  | 464.4 | 1.1189 | 729 | 0.070897
|  | 464.5 | 1.1206 | 729.2 | 0.065776
|  | 464.6 | 1.12 | 729.4 | 0.062317
|  | 464.7 | 1.12 | 729.6 | 0.059906
|  | 464.8 | 1.12 | 729.8 | 0.06
|  | 464.9 | 1.12 | 730 | 0.06
|  | 465 | 1.12 | 730.2 | 0.056104
|  | 465.1 | 1.12 | 730.4 | 0.052257
|  | 465.2 | 1.1235 | 730.6 | 0.051493
|  | 465.3 | 1.1251 | 730.8 | 0.057631
|  | 465.4 | 1.1267 | 731 | 0.065141
|  | 465.5 | 1.1267 | 731.2 | 0.065274
|  | 465.6 | 1.1283 | 731.4 | 0.062255
|  | 465.7 | 1.1299 | 731.6 | 0.059802
|  | 465.8 | 1.1315 | 731.8 | 0.061332
|  | 465.9 | 1.13 | 732 | 0.066004
|  | 466 | 1.13 | 732.2 | 0.064893
|  | 466.1 | 1.13 | 732.4 | 0.058822
|  | 466.2 | 1.13 | 732.6 | 0.05561
|  | 466.3 | 1.13 | 732.8 | 0.062095
|  | 466.4 | 1.13 | 733 | 0.062789
|  | 466.5 | 1.1345 | 733.2 | 0.061952
|  | 466.6 | 1.1361 | 733.4 | 0.05984
|  | 466.7 | 1.1377 | 733.6 | 0.06
|  | 466.8 | 1.1333 | 733.8 | 0.058589
|  | 466.9 | 1.1333 | 734 | 0.053754
|  | 467 | 1.1333 | 734.2 | 0.050838
|  | 467.1 | 1.1367 | 734.4 | 0.049793
|  | 467.2 | 1.1367 | 734.6 | 0.05
|  | 467.3 | 1.1367 | 734.8 | 0.051461
|  | 467.4 | 1.1373 | 735 | 0.055842
|  | 467.5 | 1.139 | 735.2 | 0.058763
|  | 467.6 | 1.1406 | 735.4 | 0.056613
|  | 467.7 | 1.14 | 735.6 | 0.061622
|  | 467.8 | 1.14 | 735.8 | 0.069498
|  | 467.9 | 1.14 | 736 | 0.069382
|  | 468 | 1.14 | 736.2 | 0.059892
|  | 468.1 | 1.14 | 736.4 | 0.052644
|  | 468.2 | 1.14 | 736.6 | 0.053128
|  | 468.3 | 1.14 | 736.8 | 0.057154
|  | 468.4 | 1.14 | 737 | 0.059377
|  | 468.5 | 1.14 | 737.2 | 0.06
|  | 468.6 | 1.1434 | 737.4 | 0.06
|  | 468.7 | 1.145 | 737.6 | 0.05675
|  | 468.8 | 1.1466 | 737.8 | 0.055938
|  | 468.9 | 1.1482 | 738 | 0.068576
|  | 469 | 1.1482 | 738.2 | 0.09263
|  | 469.1 | 1.1498 | 738.4 | 0.10257
|  | 469.2 | 1.1514 | 738.6 | 0.092365
|  | 469.3 | 1.15 | 738.8 | 0.076973
|  | 469.4 | 1.15 | 739 | 0.069597
|  | 469.5 | 1.15 | 739.2 | 0.063772
|  | 469.6 | 1.15 | 739.4 | 0.06503
|  | 469.7 | 1.15 | 739.6 | 0.060902
|  | 469.8 | 1.15 | 739.8 | 0.052476
|  | 469.9 | 1.15 | 740 | 0.044971
|  | 470 | 1.15 | 740.2 | 0.046723
|  | 470.1 | 1.15 | 740.4 | 0.048853
|  | 470.2 | 1.15 | 740.6 | 0.050292
|  | 470.3 | 1.15 | 740.8 | 0.053086
|  | 470.4 | 1.15 | 741 | 0.057113
|  | 470.5 | 1.15 | 741.2 | 0.059366
|  | 470.6 | 1.15 | 741.4 | 0.06
|  | 470.7 | 1.15 | 741.6 | 0.056104
|  | 470.8 | 1.15 | 741.8 | 0.052224
|  | 470.9 | 1.15 | 742 | 0.051482
|  | 471 | 1.15 | 742.2 | 0.055108
|  | 471.1 | 1.15 | 742.4 | 0.05333
|  | 471.2 | 1.15 | 742.6 | 0.051459
|  | 471.3 | 1.15 | 742.8 | 0.049837
|  | 471.4 | 1.15 | 743 | 0.05
|  | 471.5 | 1.15 | 743.2 | 0.051444
|  | 471.6 | 1.15 | 743.4 | 0.055807
|  | 471.7 | 1.1536 | 743.6 | 0.05876
|  | 471.8 | 1.1552 | 743.8 | 0.05292
|  | 471.9 | 1.1568 | 744 | 0.048157
|  | 472 | 1.1568 | 744.2 | 0.047148
|  | 472.1 | 1.1584 | 744.4 | 0.048096
|  | 472.2 | 1.16 | 744.6 | 0.044248
|  | 472.3 | 1.1616 | 744.8 | 0.045183
|  | 472.4 | 1.16 | 745 | 0.04771
|  | 472.5 | 1.16 | 745.2 | 0.050176
|  | 472.6 | 1.16 | 745.4 | 0.051353
|  | 472.7 | 1.1648 | 745.6 | 0.056098
|  | 472.8 | 1.1664 | 745.8 | 0.054934
|  | 472.9 | 1.168 | 746 | 0.052052
|  | 473 | 1.1633 | 746.2 | 0.049582
|  | 473.1 | 1.1633 | 746.4 | 0.052946
|  | 473.2 | 1.1633 | 746.6 | 0.059471
|  | 473.3 | 1.1667 | 746.8 | 0.056661
|  | 473.4 | 1.1667 | 747 | 0.049864
|  | 473.5 | 1.1667 | 747.2 | 0.042448
|  | 473.6 | 1.1677 | 747.4 | 0.043152
|  | 473.7 | 1.1693 | 747.6 | 0.049986
|  | 473.8 | 1.171 | 747.8 | 0.04666
|  | 473.9 | 1.1658 | 748 | 0.047098
|  | 474 | 1.1641 | 748.2 | 0.044322
|  | 474.1 | 1.1625 | 748.4 | 0.044444
|  | 474.2 | 1.1667 | 748.6 | 0.04698
|  | 474.3 | 1.1667 | 748.8 | 0.047966
|  | 474.4 | 1.1667 | 749 | 0.044605
|  | 474.5 | 1.1673 | 749.2 | 0.045299
|  | 474.6 | 1.1689 | 749.4 | 0.044446
|  | 474.7 | 1.1706 | 749.6 | 0.042882
|  | 474.8 | 1.17 | 749.8 | 0.041894
|  | 474.9 | 1.17 | 750 | 0.045823
|  | 475 | 1.17 | 750.2 | 0.068544
|  | 475.1 | 1.1663 | 750.4 | 0.082819
|  | 475.2 | 1.1646 | 750.6 | 0.082857
|  | 475.3 | 1.163 | 750.8 | 0.068863
|  | 475.4 | 1.1667 | 751 | 0.067063
|  | 475.5 | 1.1667 | 751.2 | 0.080892
|  | 475.6 | 1.1667 | 751.4 | 0.098888
|  | 475.7 | 1.1668 | 751.6 | 0.095084
|  | 475.8 | 1.1684 | 751.8 | 0.080094
|  | 475.9 | 1.1701 | 752 | 0.065161
|  | 476 | 1.1666 | 752.2 | 0.055037
|  | 476.1 | 1.165 | 752.4 | 0.046846
|  | 476.2 | 1.1634 | 752.6 | 0.04536
|  | 476.3 | 1.1618 | 752.8 | 0.047275
|  | 476.4 | 1.1667 | 753 | 0.049677
|  | 476.5 | 1.1667 | 753.2 | 0.05
|  | 476.6 | 1.1667 | 753.4 | 0.05
|  | 476.7 | 1.1681 | 753.6 | 0.053684
|  | 476.8 | 1.1697 | 753.8 | 0.051323
|  | 476.9 | 1.1713 | 754 | 0.047302
|  | 477 | 1.17 | 754.2 | 0.046152
|  | 477.1 | 1.17 | 754.4 | 0.043337
|  | 477.2 | 1.17 | 754.6 | 0.04556
|  | 477.3 | 1.17 | 754.8 | 0.044446
|  | 477.4 | 1.17 | 755 | 0.0462
|  | 477.5 | 1.17 | 755.2 | 0.052318
|  | 477.6 | 1.1744 | 755.4 | 0.052436
|  | 477.7 | 1.176 | 755.6 | 0.051128
|  | 477.8 | 1.1776 | 755.8 | 0.045912
|  | 477.9 | 1.1733 | 756 | 0.039176
|  | 478 | 1.1733 | 756.2 | 0.034225
|  | 478.1 | 1.1733 | 756.4 | 0.035556
|  | 478.2 | 1.1725 | 756.6 | 0.033333
|  | 478.3 | 1.1709 | 756.8 | 0.035558
|  | 478.4 | 1.1693 | 757 | 0.041556
|  | 478.5 | 1.174 | 757.2 | 0.044239
|  | 478.6 | 1.1757 | 757.4 | 0.044444
|  | 478.7 | 1.1773 | 757.6 | 0.044321
|  | 478.8 | 1.1733 | 757.8 | 0.037572
|  | 478.9 | 1.1733 | 758 | 0.03679
|  | 479 | 1.1733 | 758.2 | 0.034355
|  | 479.1 | 1.1729 | 758.4 | 0.032719
|  | 479.2 | 1.1712 | 758.6 | 0.03317
|  | 479.3 | 1.1696 | 758.8 | 0.042742
|  | 479.4 | 1.17 | 759 | 0.053086
|  | 479.5 | 1.17 | 759.2 | 0.051564
|  | 479.6 | 1.17 | 759.4 | 0.041633
|  | 479.7 | 1.17 | 759.6 | 0.034169
|  | 479.8 | 1.17 | 759.8 | 0.036898
|  | 479.9 | 1.17 | 760 | 0.039418
|  | 480 | 1.17 | 760.2 | 0.040494
|  | 480.1 | 1.17 | 760.4 | 0.036427
|  | 480.2 | 1.17 | 760.6 | 0.032654
|  | 480.3 | 1.1734 | 760.8 | 0.031622
|  | 480.4 | 1.175 | 761 | 0.034897
|  | 480.5 | 1.1766 | 761.2 | 0.033337
|  | 480.6 | 1.1783 | 761.4 | 0.035558
|  | 480.7 | 1.1733 | 761.6 | 0.041286
|  | 480.8 | 1.1733 | 761.8 | 0.049793
|  | 480.9 | 1.1733 | 762 | 0.053629
|  | 481 | 1.1767 | 762.2 | 0.046132
|  | 481.1 | 1.1767 | 762.4 | 0.036453
|  | 481.2 | 1.1767 | 762.6 | 0.030668
|  | 481.3 | 1.178 | 762.8 | 0.036513
|  | 481.4 | 1.1797 | 763 | 0.050025
|  | 481.5 | 1.1813 | 763.2 | 0.090234
|  | 481.6 | 1.18 | 763.4 | 0.15886
|  | 481.7 | 1.18 | 763.6 | 0.18979
|  | 481.8 | 1.18 | 763.8 | 0.16671
|  | 481.9 | 1.18 | 764 | 0.1058
|  | 482 | 1.18 | 764.2 | 0.060959
|  | 482.1 | 1.18 | 764.4 | 0.048647
|  | 482.2 | 1.1844 | 764.6 | 0.052796
|  | 482.3 | 1.186 | 764.8 | 0.045186
|  | 482.4 | 1.1876 | 765 | 0.042797
|  | 482.5 | 1.1833 | 765.2 | 0.038945
|  | 482.6 | 1.1833 | 765.4 | 0.037298
|  | 482.7 | 1.1833 | 765.6 | 0.038743
|  | 482.8 | 1.1867 | 765.8 | 0.045314
|  | 482.9 | 1.1867 | 766 | 0.041257
|  | 483 | 1.1867 | 766.2 | 0.037379
|  | 483.1 | 1.1833 | 766.4 | 0.031563
|  | 483.2 | 1.1833 | 766.6 | 0.027414
|  | 483.3 | 1.1833 | 766.8 | 0.027471
|  | 483.4 | 1.1867 | 767 | 0.029219
|  | 483.5 | 1.1867 | 767.2 | 0.038527
|  | 483.6 | 1.1867 | 767.4 | 0.039184
|  | 483.7 | 1.1872 | 767.6 | 0.037086
|  | 483.8 | 1.1889 | 767.8 | 0.037226
|  | 483.9 | 1.1905 | 768 | 0.038054
|  | 484 | 1.19 | 768.2 | 0.034198
|  | 484.1 | 1.19 | 768.4 | 0.031158
|  | 484.2 | 1.19 | 768.6 | 0.026448
|  | 484.3 | 1.19 | 768.8 | 0.025854
|  | 484.4 | 1.19 | 769 | 0.032428
|  | 484.5 | 1.19 | 769.2 | 0.033699
|  | 484.6 | 1.1936 | 769.4 | 0.03667
|  | 484.7 | 1.1953 | 769.6 | 0.034449
|  | 484.8 | 1.1969 | 769.8 | 0.028752
|  | 484.9 | 1.1933 | 770 | 0.025519
|  | 485 | 1.1933 | 770.2 | 0.035361
|  | 485.1 | 1.1933 | 770.4 | 0.040811
|  | 485.2 | 1.2001 | 770.6 | 0.031917
|  | 485.3 | 1.2017 | 770.8 | 0.024083
|  | 485.4 | 1.2034 | 771 | 0.023097
|  | 485.5 | 1.2034 | 771.2 | 0.029417
|  | 485.6 | 1.2066 | 771.4 | 0.036597
|  | 485.7 | 1.2099 | 771.6 | 0.039222
|  | 485.8 | 1.2132 | 771.8 | 0.04
|  | 485.9 | 1.21 | 772 | 0.04
|  | 486 | 1.21 | 772.2 | 0.060333
|  | 486.1 | 1.21 | 772.4 | 0.087709
|  | 486.2 | 1.2052 | 772.6 | 0.10108
|  | 486.3 | 1.2036 | 772.8 | 0.088784
|  | 486.4 | 1.2019 | 773 | 0.071354
|  | 486.5 | 1.2067 | 773.2 | 0.053335
|  | 486.6 | 1.2067 | 773.4 | 0.040432
|  | 486.7 | 1.2067 | 773.6 | 0.031579
|  | 486.8 | 1.208 | 773.8 | 0.028678
|  | 486.9 | 1.2096 | 774 | 0.027644
|  | 487 | 1.2113 | 774.2 | 0.038226
|  | 487.1 | 1.21 | 774.4 | 0.034132
|  | 487.2 | 1.21 | 774.6 | 0.029758
|  | 487.3 | 1.21 | 774.8 | 0.024757
|  | 487.4 | 1.2055 | 775 | 0.024447
|  | 487.5 | 1.2039 | 775.2 | 0.025556
|  | 487.6 | 1.2022 | 775.4 | 0.023333
|  | 487.7 | 1.2023 | 775.6 | 0.021492
|  | 487.8 | 1.2006 | 775.8 | 0.023361
|  | 487.9 | 1.199 | 776 | 0.030281
|  | 488 | 1.2 | 776.2 | 0.036971
|  | 488.1 | 1.2 | 776.4 | 0.039382
|  | 488.2 | 1.2 | 776.6 | 0.04
|  | 488.3 | 1.2 | 776.8 | 0.031719
|  | 488.4 | 1.2 | 777 | 0.027119
|  | 488.5 | 1.2 | 777.2 | 0.024519
|  | 488.6 | 1.2 | 777.4 | 0.022834
|  | 488.7 | 1.2 | 777.6 | 0.020525
|  | 488.8 | 1.2 | 777.8 | 0.02
|  | 488.9 | 1.2 | 778 | 0.02
|  | 489 | 1.2 | 778.2 | 0.02
|  | 489.1 | 1.2 | 778.4 | 0.026321
|  | 489.2 | 1.2 | 778.6 | 0.033004
|  | 489.3 | 1.2 | 778.8 | 0.032624
|  | 489.4 | 1.2 | 779 | 0.029834
|  | 489.5 | 1.1961 | 779.2 | 0.024524
|  | 489.6 | 1.1944 | 779.4 | 0.025254
|  | 489.7 | 1.1928 | 779.6 | 0.024448
|  | 489.8 | 1.1928 | 779.8 | 0.024448
|  | 489.9 | 1.1912 | 780 | 0.02701
|  | 490 | 1.1895 | 780.2 | 0.029164
|  | 490.1 | 1.1938 | 780.4 | 0.03
|  | 490.2 | 1.1954 | 780.6 | 0.02211
|  | 490.3 | 1.1971 | 780.8 | 0.02091
|  | 490.4 | 1.1933 | 781 | 0.024692
|  | 490.5 | 1.1933 | 781.2 | 0.029673
|  | 490.6 | 1.1933 | 781.4 | 0.028737
|  | 490.7 | 1.2003 | 781.6 | 0.024103
|  | 490.8 | 1.202 | 781.8 | 0.02511
|  | 490.9 | 1.2036 | 782 | 0.02802
|  | 491 | 1.2075 | 782.2 | 0.029917
|  | 491.1 | 1.2124 | 782.4 | 0.028392
|  | 491.2 | 1.2173 | 782.6 | 0.023627
|  | 491.3 | 1.2204 | 782.8 | 0.020524
|  | 491.4 | 1.2237 | 783 | 0.019473
|  | 491.5 | 1.227 | 783.2 | 0.02
|  | 491.6 | 1.2303 | 783.4 | 0.023221
|  | 491.7 | 1.2336 | 783.6 | 0.029142
|  | 491.8 | 1.2368 | 783.8 | 0.033441
|  | 491.9 | 1.2231 | 784 | 0.028575
|  | 492 | 1.2182 | 784.2 | 0.027693
|  | 492.1 | 1.2133 | 784.4 | 0.028022
|  | 492.2 | 1.2083 | 784.6 | 0.030098
|  | 492.3 | 1.1968 | 784.8 | 0.026399
|  | 492.4 | 1.1869 | 785 | 0.01917
|  | 492.5 | 1.1771 | 785.2 | 0.022391
|  | 492.6 | 1.1835 | 785.4 | 0.022616
|  | 492.7 | 1.1802 | 785.6 | 0.021871
|  | 492.8 | 1.1769 | 785.8 | 0.019796
|  | 492.9 | 1.1751 | 786 | 0.023114
|  | 493 | 1.1735 | 786.2 | 0.030588
|  | 493.1 | 1.1719 | 786.4 | 0.037017
|  | 493.2 | 1.1767 | 786.6 | 0.026287
|  | 493.3 | 1.1767 | 786.8 | 0.020105
|  | 493.4 | 1.1767 | 787 | 0.017339
|  | 493.5 | 1.1733 | 787.2 | 0.023416
|  | 493.6 | 1.1733 | 787.4 | 0.029416
|  | 493.7 | 1.1733 | 787.6 | 0.036676
|  | 493.8 | 1.1767 | 787.8 | 0.039292
|  | 493.9 | 1.1767 | 788 | 0.04
|  | 494 | 1.1767 | 788.2 | 0.023543
|  | 494.1 | 1.1733 | 788.4 | 0.010655
|  | 494.2 | 1.1733 | 788.6 | 0.0077401
|  | 494.3 | 1.1733 | 788.8 | 0.010066
|  | 494.4 | 1.1675 | 789 | 0.01
|  | 494.5 | 1.1642 | 789.2 | 0.01
|  | 494.6 | 1.1609 | 789.4 | 0.01
|  | 494.7 | 1.1621 | 789.6 | 0.01
|  | 494.8 | 1.1605 | 789.8 | 0.016745
|  | 494.9 | 1.1588 | 790 | 0.025372
|  | 495 | 1.16 | 790.2 | 0.026607
|  | 495.1 | 1.16 | 790.4 | 0.018945
|  | 495.2 | 1.16 | 790.6 | 0.017373
|  | 495.3 | 1.1556 | 790.8 | 0.013974
|  | 495.4 | 1.1539 | 791 | 0.011884
|  | 495.5 | 1.1523 | 791.2 | 0.016365
|  | 495.6 | 1.1567 | 791.4 | 0.024754
|  | 495.7 | 1.1567 | 791.6 | 0.027824
|  | 495.8 | 1.1567 | 791.8 | 0.022959
|  | 495.9 | 1.1619 | 792 | 0.015343
|  | 496 | 1.1652 | 792.2 | 0.018788
|  | 496.1 | 1.1685 | 792.4 | 0.025978
|  | 496.2 | 1.1591 | 792.6 | 0.023785
|  | 496.3 | 1.1574 | 792.8 | 0.017384
|  | 496.4 | 1.1558 | 793 | 0.016099
|  | 496.5 | 1.1591 | 793.2 | 0.013333
|  | 496.6 | 1.1575 | 793.4 | 0.011234
|  | 496.7 | 1.1558 | 793.6 | 0.013532
|  | 496.8 | 1.1533 | 793.8 | 0.021147
|  | 496.9 | 1.1533 | 794 | 0.026161
|  | 497 | 1.1533 | 794.2 | 0.0214
|  | 497.1 | 1.1525 | 794.4 | 0.015825
|  | 497.2 | 1.1509 | 794.6 | 0.016667
|  | 497.3 | 1.1492 | 794.8 | 0.043702
|  | 497.4 | 1.15 | 795 | 0.074005
|  | 497.5 | 1.15 | 795.2 | 0.068329
|  | 497.6 | 1.15 | 795.4 | 0.047343
|  | 497.7 | 1.146 | 795.6 | 0.034266
|  | 497.8 | 1.1443 | 795.8 | 0.022548
|  | 497.9 | 1.1427 | 796 | 0.018482
|  | 498 | 1.1467 | 796.2 | 0.029871
|  | 498.1 | 1.1467 | 796.4 | 0.03398
|  | 498.2 | 1.1467 | 796.6 | 0.0305
|  | 498.3 | 1.1433 | 796.8 | 0.023785
|  | 498.4 | 1.1433 | 797 | 0.017915
|  | 498.5 | 1.1433 | 797.2 | 0.018799
|  | 498.6 | 1.1388 | 797.4 | 0.024573
|  | 498.7 | 1.1355 | 797.6 | 0.029145
|  | 498.8 | 1.1322 | 797.8 | 0.030121
|  | 498.9 | 1.1328 | 798 | 0.024012
|  | 499 | 1.1311 | 798.2 | 0.018949
|  | 499.1 | 1.1295 | 798.4 | 0.021111
|  | 499.2 | 1.1338 | 798.6 | 0.018889
|  | 499.3 | 1.1355 | 798.8 | 0.023333
|  | 499.4 | 1.1372 | 799 | 0.018889
|  | 499.5 | 1.1333 | 799.2 | 0.021111
|  | 499.6 | 1.1333 | 799.4 | 0.026581
|  | 499.7 | 1.1333 | 799.6 | 0.036884
|  | 499.8 | 1.1291 | 799.8 | 0.03944
|  | 499.9 | 1.1258 | 800 | 0.04
|  | 500 | 1.1225 | 800.2 | 0.026966
|  | 500.1 | 1.1267 | 800.4 | 0.018006
|  | 500.2 | 1.1267 | 800.6 | 0.021079
|  | 500.3 | 1.1267 | 800.8 | 0.037153
|  | 500.4 | 1.127 | 801 | 0.058667
|  | 500.5 | 1.1287 | 801.2 | 0.051879
|  | 500.6 | 1.1304 | 801.4 | 0.042149
|  | 500.7 | 1.13 | 801.6 | 0.063807
|  | 500.8 | 1.13 | 801.8 | 0.09927
|  | 500.9 | 1.13 | 802 | 0.092509
|  | 501 | 1.1227 | 802.2 | 0.059444
|  | 501.1 | 1.1194 | 802.4 | 0.024375
|  | 501.2 | 1.1161 | 802.6 | 0.019995
|  | 501.3 | 1.1161 | 802.8 | 0.025622
|  | 501.4 | 1.1128 | 803 | 0.029866
|  | 501.5 | 1.1095 | 803.2 | 0.02257
|  | 501.6 | 1.1136 | 803.4 | 0.020973
|  | 501.7 | 1.1153 | 803.6 | 0.02475
|  | 501.8 | 1.1169 | 803.8 | 0.026652
|  | 501.9 | 1.1169 | 804 | 0.018667
|  | 502 | 1.1186 | 804.2 | 0.017146
|  | 502.1 | 1.1202 | 804.4 | 0.018257
|  | 502.2 | 1.1165 | 804.6 | 0.02032
|  | 502.3 | 1.1148 | 804.8 | 0.023437
|  | 502.4 | 1.1131 | 805 | 0.027266
|  | 502.5 | 1.1131 | 805.2 | 0.026737
|  | 502.6 | 1.1115 | 805.4 | 0.018541
|  | 502.7 | 1.1098 | 805.6 | 0.018822
|  | 502.8 | 1.1065 | 805.8 | 0.024487
|  | 502.9 | 1.1048 | 806 | 0.021181
|  | 503 | 1.1032 | 806.2 | 0.024487
|  | 503.1 | 1.1067 | 806.4 | 0.020176
|  | 503.2 | 1.1067 | 806.6 | 0.016101
|  | 503.3 | 1.1067 | 806.8 | 0.018526
|  | 503.4 | 1.1068 | 807 | 0.014446
|  | 503.5 | 1.1085 | 807.2 | 0.0062213
|  | 503.6 | 1.1101 | 807.4 | 0.00030172
|  | 503.7 | 1.1066 | 807.6 | 0.0034339
|  | 503.8 | 1.1049
|  | 503.9 | 1.1032
|  | 504 | 1.1033
|  | 504.1 | 1.1016
|  | 504.2 | 1.0999
|  | 504.3 | 1.1
|  | 504.4 | 1.1
|  | 504.5 | 1.1
|  | 504.6 | 1.1034
|  | 504.7 | 1.105
|  | 504.8 | 1.1067
|  | 504.9 | 1.1033
|  | 505 | 1.1033
|  | 505.1 | 1.1033
|  | 505.2 | 1.1033
|  | 505.3 | 1.1016
|  | 505.4 | 1.1
|  | 505.5 | 1.1
|  | 505.6 | 1.1
|  | 505.7 | 1.1
|  | 505.8 | 1.1
|  | 505.9 | 1.095
|  | 506 | 1.0933
|  | 506.1 | 1.0917
|  | 506.2 | 1.0917
|  | 506.3 | 1.09
|  | 506.4 | 1.0883
|  | 506.5 | 1.09
|  | 506.6 | 1.09
|  | 506.7 | 1.09
|  | 506.8 | 1.095
|  | 506.9 | 1.0966
|  | 507 | 1.0983
|  | 507.1 | 1.0933
|  | 507.2 | 1.0933
|  | 507.3 | 1.0933
|  | 507.4 | 1.0867
|  | 507.5 | 1.0834
|  | 507.6 | 1.0801
|  | 507.7 | 1.0817
|  | 507.8 | 1.08
|  | 507.9 | 1.0784
|  | 508 | 1.08
|  | 508.1 | 1.08
|  | 508.2 | 1.08
|  | 508.3 | 1.08
|  | 508.4 | 1.08
|  | 508.5 | 1.08
|  | 508.6 | 1.0751
|  | 508.7 | 1.0734
|  | 508.8 | 1.0717
|  | 508.9 | 1.0767
|  | 509 | 1.0767
|  | 509.1 | 1.0767
|  | 509.2 | 1.0832
|  | 509.3 | 1.0866
|  | 509.4 | 1.0899
|  | 509.5 | 1.0833
|  | 509.6 | 1.0833
|  | 509.7 | 1.0833
|  | 509.8 | 1.0767
|  | 509.9 | 1.0734
|  | 510 | 1.0701
|  | 510.1 | 1.0767
|  | 510.2 | 1.0767
|  | 510.3 | 1.0767
|  | 510.4 | 1.0684
|  | 510.5 | 1.0667
|  | 510.6 | 1.065
|  | 510.7 | 1.0683
|  | 510.8 | 1.0667
|  | 510.9 | 1.065
|  | 511 | 1.0733
|  | 511.1 | 1.0767
|  | 511.2 | 1.08
|  | 511.3 | 1.0683
|  | 511.4 | 1.0667
|  | 511.5 | 1.0666
|  | 511.6 | 1.0633
|  | 511.7 | 1.0599
|  | 511.8 | 1.0634
|  | 511.9 | 1.065
|  | 512 | 1.0667
|  | 512.1 | 1.0667
|  | 512.2 | 1.0684
|  | 512.3 | 1.07
|  | 512.4 | 1.07
|  | 512.5 | 1.07
|  | 512.6 | 1.07
|  | 512.7 | 1.0632
|  | 512.8 | 1.0599
|  | 512.9 | 1.0565
|  | 513 | 1.0599
|  | 513.1 | 1.0582
|  | 513.2 | 1.0566
|  | 513.3 | 1.0533
|  | 513.4 | 1.0533
|  | 513.5 | 1.0533
|  | 513.6 | 1.0567
|  | 513.7 | 1.0567
|  | 513.8 | 1.0567
|  | 513.9 | 1.0533
|  | 514 | 1.0533
|  | 514.1 | 1.0533
|  | 514.2 | 1.0567
|  | 514.3 | 1.0567
|  | 514.4 | 1.0567
|  | 514.5 | 1.0569
|  | 514.6 | 1.0585
|  | 514.7 | 1.0602
|  | 514.8 | 1.0565
|  | 514.9 | 1.0548
|  | 515 | 1.0531
|  | 515.1 | 1.0531
|  | 515.2 | 1.0514
|  | 515.3 | 1.0497
|  | 515.4 | 1.05
|  | 515.5 | 1.05
|  | 515.6 | 1.05
|  | 515.7 | 1.0464
|  | 515.8 | 1.0447
|  | 515.9 | 1.043
|  | 516 | 1.043
|  | 516.1 | 1.0413
|  | 516.2 | 1.0397
|  | 516.3 | 1.0437
|  | 516.4 | 1.0454
|  | 516.5 | 1.047
|  | 516.6 | 1.0433
|  | 516.7 | 1.0433
|  | 516.8 | 1.0433
|  | 516.9 | 1.0429
|  | 517 | 1.0412
|  | 517.1 | 1.0396
|  | 517.2 | 1.04
|  | 517.3 | 1.04
|  | 517.4 | 1.04
|  | 517.5 | 1.04
|  | 517.6 | 1.04
|  | 517.7 | 1.04
|  | 517.8 | 1.04
|  | 517.9 | 1.04
|  | 518 | 1.04
|  | 518.1 | 1.04
|  | 518.2 | 1.04
|  | 518.3 | 1.04
|  | 518.4 | 1.04
|  | 518.5 | 1.04
|  | 518.6 | 1.04
|  | 518.7 | 1.04
|  | 518.8 | 1.04
|  | 518.9 | 1.04
|  | 519 | 1.04
|  | 519.1 | 1.04
|  | 519.2 | 1.04
|  | 519.3 | 1.04
|  | 519.4 | 1.04
|  | 519.5 | 1.04
|  | 519.6 | 1.0359
|  | 519.7 | 1.0342
|  | 519.8 | 1.0325
|  | 519.9 | 1.0325
|  | 520 | 1.0309
|  | 520.1 | 1.0292
|  | 520.2 | 1.0342
|  | 520.3 | 1.0359
|  | 520.4 | 1.0375
|  | 520.5 | 1.0333
|  | 520.6 | 1.0333
|  | 520.7 | 1.0333
|  | 520.8 | 1.0324
|  | 520.9 | 1.0307
|  | 521 | 1.029
|  | 521.1 | 1.0387
|  | 521.2 | 1.042
|  | 521.3 | 1.0454
|  | 521.4 | 1.041
|  | 521.5 | 1.0427
|  | 521.6 | 1.0444
|  | 521.7 | 1.0422
|  | 521.8 | 1.0406
|  | 521.9 | 1.0389
|  | 522 | 1.0355
|  | 522.1 | 1.0338
|  | 522.2 | 1.0322
|  | 522.3 | 1.0276
|  | 522.4 | 1.0242
|  | 522.5 | 1.0209
|  | 522.6 | 1.0267
|  | 522.7 | 1.0267
|  | 522.8 | 1.0267
|  | 522.9 | 1.028
|  | 523 | 1.0297
|  | 523.1 | 1.0313
|  | 523.2 | 1.03
|  | 523.3 | 1.03
|  | 523.4 | 1.03
|  | 523.5 | 1.0348
|  | 523.6 | 1.0364
|  | 523.7 | 1.0381
|  | 523.8 | 1.0333
|  | 523.9 | 1.0333
|  | 524 | 1.0333
|  | 524.1 | 1.0367
|  | 524.2 | 1.0367
|  | 524.3 | 1.0367
|  | 524.4 | 1.0333
|  | 524.5 | 1.0333
|  | 524.6 | 1.0333
|  | 524.7 | 1.0317
|  | 524.8 | 1.03
|  | 524.9 | 1.03
|  | 525 | 1.03
|  | 525.1 | 1.03
|  | 525.2 | 1.03
|  | 525.3 | 1.03
|  | 525.4 | 1.03
|  | 525.5 | 1.03
|  | 525.6 | 1.03
|  | 525.7 | 1.03
|  | 525.8 | 1.0336
|  | 525.9 | 1.0353
|  | 526 | 1.0369
|  | 526.1 | 1.0333
|  | 526.2 | 1.0333
|  | 526.3 | 1.0333
|  | 526.4 | 1.033
|  | 526.5 | 1.0313
|  | 526.6 | 1.0296
|  | 526.7 | 1.03
|  | 526.8 | 1.03
|  | 526.9 | 1.03
|  | 527 | 1.03
|  | 527.1 | 1.03
|  | 527.2 | 1.03
|  | 527.3 | 1.03
|  | 527.4 | 1.03
|  | 527.5 | 1.03
|  | 527.6 | 1.034
|  | 527.7 | 1.0357
|  | 527.8 | 1.0374
|  | 527.9 | 1.0374
|  | 528 | 1.0391
|  | 528.1 | 1.0408
|  | 528.2 | 1.04
|  | 528.3 | 1.04
|  | 528.4 | 1.04
|  | 528.5 | 1.0358
|  | 528.6 | 1.0341
|  | 528.7 | 1.0324
|  | 528.8 | 1.0367
|  | 528.9 | 1.0367
|  | 529 | 1.0367
|  | 529.1 | 1.0377
|  | 529.2 | 1.0394
|  | 529.3 | 1.0411
|  | 529.4 | 1.04
|  | 529.5 | 1.04
|  | 529.6 | 1.04
|  | 529.7 | 1.04
|  | 529.8 | 1.04
|  | 529.9 | 1.04
|  | 530 | 1.04
|  | 530.1 | 1.04
|  | 530.2 | 1.04
|  | 530.3 | 1.04
|  | 530.4 | 1.04
|  | 530.5 | 1.04
|  | 530.6 | 1.04
|  | 530.7 | 1.04
|  | 530.8 | 1.04
|  | 530.9 | 1.04
|  | 531 | 1.04
|  | 531.1 | 1.04
|  | 531.2 | 1.035
|  | 531.3 | 1.0333
|  | 531.4 | 1.0367
|  | 531.5 | 1.0367
|  | 531.6 | 1.0367
|  | 531.7 | 1.0403
|  | 531.8 | 1.0437
|  | 531.9 | 1.0471
|  | 532 | 1.0469
|  | 532.1 | 1.0486
|  | 532.2 | 1.0503
|  | 532.3 | 1.0463
|  | 532.4 | 1.0446
|  | 532.5 | 1.0429
|  | 532.6 | 1.0504
|  | 532.7 | 1.0521
|  | 532.8 | 1.0538
|  | 532.9 | 1.0505
|  | 533 | 1.0522
|  | 533.1 | 1.0539
|  | 533.2 | 1.0567
|  | 533.3 | 1.0567
|  | 533.4 | 1.0567
|  | 533.5 | 1.0614
|  | 533.6 | 1.0648
|  | 533.7 | 1.0682
|  | 533.8 | 1.0675
|  | 533.9 | 1.0692
|  | 534 | 1.0709
|  | 534.1 | 1.07
|  | 534.2 | 1.07
|  | 534.3 | 1.07
|  | 534.4 | 1.07
|  | 534.5 | 1.07
|  | 534.6 | 1.07
|  | 534.7 | 1.07
|  | 534.8 | 1.07
|  | 534.9 | 1.07
|  | 535 | 1.07
|  | 535.1 | 1.07
|  | 535.2 | 1.07
|  | 535.3 | 1.0653
|  | 535.4 | 1.0636
|  | 535.5 | 1.0619
|  | 535.6 | 1.0714
|  | 535.7 | 1.0731
|  | 535.8 | 1.0748
|  | 535.9 | 1.0715
|  | 536 | 1.0732
|  | 536.1 | 1.0749
|  | 536.2 | 1.0717
|  | 536.3 | 1.07
|  | 536.4 | 1.0734
|  | 536.5 | 1.0751
|  | 536.6 | 1.0768
|  | 536.7 | 1.0803
|  | 536.8 | 1.0838
|  | 536.9 | 1.0872
|  | 537 | 1.0906
|  | 537.1 | 1.094
|  | 537.2 | 1.0974
|  | 537.3 | 1.0971
|  | 537.4 | 1.0988
|  | 537.5 | 1.1005
|  | 537.6 | 1.1
|  | 537.7 | 1.1
|  | 537.8 | 1.1
|  | 537.9 | 1.104
|  | 538 | 1.1057
|  | 538.1 | 1.1074
|  | 538.2 | 1.1074
|  | 538.3 | 1.1091
|  | 538.4 | 1.1108
|  | 538.5 | 1.1142
|  | 538.6 | 1.1159
|  | 538.7 | 1.1176
|  | 538.8 | 1.122
|  | 538.9 | 1.1254
|  | 539 | 1.1288
|  | 539.1 | 1.1278
|  | 539.2 | 1.1295
|  | 539.3 | 1.1312
|  | 539.4 | 1.1391
|  | 539.5 | 1.1425
|  | 539.6 | 1.1459
|  | 539.7 | 1.1507
|  | 539.8 | 1.1558
|  | 539.9 | 1.1609
|  | 540 | 1.163
|  | 540.1 | 1.1664
|  | 540.2 | 1.1698
|  | 540.3 | 1.1732
|  | 540.4 | 1.1766
|  | 540.5 | 1.1834
|  | 540.6 | 1.1885
|  | 540.7 | 1.1936
|  | 540.8 | 1.204
|  | 540.9 | 1.2126
|  | 541 | 1.2211
|  | 541.1 | 1.228
|  | 541.2 | 1.2366
|  | 541.3 | 1.2452
|  | 541.4 | 1.2628
|  | 541.5 | 1.2748
|  | 541.6 | 1.2868
|  | 541.7 | 1.2938
|  | 541.8 | 1.3057
|  | 541.9 | 1.3177
|  | 542 | 1.3187
|  | 542.1 | 1.3238
|  | 542.2 | 1.3289
|  | 542.3 | 1.3275
|  | 542.4 | 1.3292
|  | 542.5 | 1.3309
|  | 542.6 | 1.3257
|  | 542.7 | 1.324
|  | 542.8 | 1.3223
|  | 542.9 | 1.3223
|  | 543 | 1.3205
|  | 543.1 | 1.3188
|  | 543.2 | 1.32
|  | 543.3 | 1.32
|  | 543.4 | 1.32
|  | 543.5 | 1.3247
|  | 543.6 | 1.3264
|  | 543.7 | 1.3281
|  | 543.8 | 1.333
|  | 543.9 | 1.3364
|  | 544 | 1.3398
|  | 544.1 | 1.3383
|  | 544.2 | 1.34
|  | 544.3 | 1.3502
|  | 544.4 | 1.3553
|  | 544.5 | 1.3605
|  | 544.6 | 1.3854
|  | 544.7 | 1.4025
|  | 544.8 | 1.4196
|  | 544.9 | 1.5774
|  | 545 | 1.6632
|  | 545.1 | 1.7489
|  | 545.2 | 2.4134
|  | 545.3 | 2.7857
|  | 545.4 | 3.1581
|  | 545.5 | 4.0196
|  | 545.6 | 4.6971
|  | 545.7 | 5.3746
|  | 545.8 | 6.0539
|  | 545.9 | 6.7652
|  | 546 | 7.4766
|  | 546.1 | 8.4067
|  | 546.2 | 9.1872
|  | 546.3 | 9.9677
|  | 546.4 | 7.2272
|  | 546.5 | 6.686
|  | 546.6 | 6.1448
|  | 546.7 | 4.5082
|  | 546.8 | 3.0288
|  | 546.9 | 1.5494
|  | 547 | 2.1441
|  | 547.1 | 1.5977
|  | 547.2 | 1.0513
|  | 547.3 | 1.4335
|  | 547.4 | 1.3596
|  | 547.5 | 1.2857
|  | 547.6 | 1.328
|  | 547.7 | 1.3091
|  | 547.8 | 1.3096
|  | 547.9 | 1.3044
|  | 548 | 1.2993
|  | 548.1 | 1.2964
|  | 548.2 | 1.2946
|  | 548.3 | 1.2929
|  | 548.4 | 1.2967
|  | 548.5 | 1.2967
|  | 548.6 | 1.2967
|  | 548.7 | 1.2973
|  | 548.8 | 1.299
|  | 548.9 | 1.3007
|  | 549 | 1.3
|  | 549.1 | 1.3
|  | 549.2 | 1.3
|  | 549.3 | 1.3
|  | 549.4 | 1.3
|  | 549.5 | 1.3
|  | 549.6 | 1.2956
|  | 549.7 | 1.2938
|  | 549.8 | 1.2921
|  | 549.9 | 1.3059
|  | 550 | 1.3093
|  | 550.1 | 1.3128
|  | 550.2 | 1.3095
|  | 550.3 | 1.313
|  | 550.4 | 1.3164
|  | 550.5 | 1.3068
|  | 550.6 | 1.3033
|  | 550.7 | 1.3033
|  | 550.8 | 1.3016
|  | 550.9 | 1.2998
|  | 551 | 1.3036
|  | 551.1 | 1.3053
|  | 551.2 | 1.307
|  | 551.3 | 1.3033
|  | 551.4 | 1.3033
|  | 551.5 | 1.3033
|  | 551.6 | 1.3106
|  | 551.7 | 1.3123
|  | 551.8 | 1.314
|  | 551.9 | 1.3067
|  | 552 | 1.3067
|  | 552.1 | 1.3067
|  | 552.2 | 1.3091
|  | 552.3 | 1.3074
|  | 552.4 | 1.3056
|  | 552.5 | 1.3122
|  | 552.6 | 1.3156
|  | 552.7 | 1.3191
|  | 552.8 | 1.3133
|  | 552.9 | 1.3133
|  | 553 | 1.3133
|  | 553.1 | 1.3214
|  | 553.2 | 1.3232
|  | 553.3 | 1.3249
|  | 553.4 | 1.3216
|  | 553.5 | 1.3234
|  | 553.6 | 1.3267
|  | 553.7 | 1.3267
|  | 553.8 | 1.3267
|  | 553.9 | 1.3233
|  | 554 | 1.3233
|  | 554.1 | 1.3233
|  | 554.2 | 1.3229
|  | 554.3 | 1.3212
|  | 554.4 | 1.3194
|  | 554.5 | 1.32
|  | 554.6 | 1.32
|  | 554.7 | 1.32
|  | 554.8 | 1.3241
|  | 554.9 | 1.3259
|  | 555 | 1.3276
|  | 555.1 | 1.3277
|  | 555.2 | 1.3294
|  | 555.3 | 1.3311
|  | 555.4 | 1.3255
|  | 555.5 | 1.3237
|  | 555.6 | 1.322
|  | 555.7 | 1.3314
|  | 555.8 | 1.3331
|  | 555.9 | 1.3348
|  | 556 | 1.3316
|  | 556.1 | 1.3333
|  | 556.2 | 1.3333
|  | 556.3 | 1.3316
|  | 556.4 | 1.3298
|  | 556.5 | 1.33
|  | 556.6 | 1.33
|  | 556.7 | 1.33
|  | 556.8 | 1.3263
|  | 556.9 | 1.3245
|  | 557 | 1.3228
|  | 557.1 | 1.3267
|  | 557.2 | 1.3267
|  | 557.3 | 1.3267
|  | 557.4 | 1.3275
|  | 557.5 | 1.3292
|  | 557.6 | 1.3309
|  | 557.7 | 1.3343
|  | 557.8 | 1.3361
|  | 557.9 | 1.3378
|  | 558 | 1.3333
|  | 558.1 | 1.3333
|  | 558.2 | 1.3333
|  | 558.3 | 1.3367
|  | 558.4 | 1.3367
|  | 558.5 | 1.3367
|  | 558.6 | 1.3432
|  | 558.7 | 1.3467
|  | 558.8 | 1.3502
|  | 558.9 | 1.3536
|  | 559 | 1.3571
|  | 559.1 | 1.3569
|  | 559.2 | 1.3587
|  | 559.3 | 1.3604
|  | 559.4 | 1.3562
|  | 559.5 | 1.3544
|  | 559.6 | 1.3527
|  | 559.7 | 1.3567
|  | 559.8 | 1.3567
|  | 559.9 | 1.3567
|  | 560 | 1.3576
|  | 560.1 | 1.3593
|  | 560.2 | 1.361
|  | 560.3 | 1.3644
|  | 560.4 | 1.3662
|  | 560.5 | 1.3679
|  | 560.6 | 1.368
|  | 560.7 | 1.3697
|  | 560.8 | 1.3715
|  | 560.9 | 1.3749
|  | 561 | 1.3766
|  | 561.1 | 1.3783
|  | 561.2 | 1.3733
|  | 561.3 | 1.3733
|  | 561.4 | 1.3767
|  | 561.5 | 1.3767
|  | 561.6 | 1.3767
|  | 561.7 | 1.3771
|  | 561.8 | 1.3788
|  | 561.9 | 1.3806
|  | 562 | 1.38
|  | 562.1 | 1.38
|  | 562.2 | 1.38
|  | 562.3 | 1.38
|  | 562.4 | 1.38
|  | 562.5 | 1.38
|  | 562.6 | 1.38
|  | 562.7 | 1.38
|  | 562.8 | 1.38
|  | 562.9 | 1.3846
|  | 563 | 1.3864
|  | 563.1 | 1.3881
|  | 563.2 | 1.3882
|  | 563.3 | 1.3899
|  | 563.4 | 1.3917
|  | 563.5 | 1.39
|  | 563.6 | 1.39
|  | 563.7 | 1.39
|  | 563.8 | 1.39
|  | 563.9 | 1.39
|  | 564 | 1.3938
|  | 564.1 | 1.3955
|  | 564.2 | 1.3973
|  | 564.3 | 1.3933
|  | 564.4 | 1.3933
|  | 564.5 | 1.3933
|  | 564.6 | 1.3967
|  | 564.7 | 1.3967
|  | 564.8 | 1.3967
|  | 564.9 | 1.4023
|  | 565 | 1.4057
|  | 565.1 | 1.4092
|  | 565.2 | 1.408
|  | 565.3 | 1.4098
|  | 565.4 | 1.4115
|  | 565.5 | 1.4149
|  | 565.6 | 1.4167
|  | 565.7 | 1.4167
|  | 565.8 | 1.4185
|  | 565.9 | 1.4202
|  | 566 | 1.4164
|  | 566.1 | 1.4146
|  | 566.2 | 1.4129
|  | 566.3 | 1.4128
|  | 566.4 | 1.4111
|  | 566.5 | 1.4093
|  | 566.6 | 1.4141
|  | 566.7 | 1.4159
|  | 566.8 | 1.4176
|  | 566.9 | 1.422
|  | 567 | 1.4255
|  | 567.1 | 1.429
|  | 567.2 | 1.4233
|  | 567.3 | 1.4233
|  | 567.4 | 1.4233
|  | 567.5 | 1.4267
|  | 567.6 | 1.4267
|  | 567.7 | 1.4267
|  | 567.8 | 1.4284
|  | 567.9 | 1.4301
|  | 568 | 1.43
|  | 568.1 | 1.43
|  | 568.2 | 1.43
|  | 568.3 | 1.4338
|  | 568.4 | 1.4355
|  | 568.5 | 1.4373
|  | 568.6 | 1.4333
|  | 568.7 | 1.4333
|  | 568.8 | 1.4333
|  | 568.9 | 1.4324
|  | 569 | 1.4306
|  | 569.1 | 1.4289
|  | 569.2 | 1.4345
|  | 569.3 | 1.4363
|  | 569.4 | 1.438
|  | 569.5 | 1.4381
|  | 569.6 | 1.4399
|  | 569.7 | 1.4416
|  | 569.8 | 1.435
|  | 569.9 | 1.4332
|  | 570 | 1.4367
|  | 570.1 | 1.4367
|  | 570.2 | 1.4367
|  | 570.3 | 1.4371
|  | 570.4 | 1.4389
|  | 570.5 | 1.4406
|  | 570.6 | 1.444
|  | 570.7 | 1.4458
|  | 570.8 | 1.4475
|  | 570.9 | 1.4476
|  | 571 | 1.4494
|  | 571.1 | 1.4511
|  | 571.2 | 1.4455
|  | 571.3 | 1.4437
|  | 571.4 | 1.442
|  | 571.5 | 1.4467
|  | 571.6 | 1.4467
|  | 571.7 | 1.4467
|  | 571.8 | 1.4534
|  | 571.9 | 1.4569
|  | 572 | 1.4569
|  | 572.1 | 1.4586
|  | 572.2 | 1.4604
|  | 572.3 | 1.4562
|  | 572.4 | 1.4544
|  | 572.5 | 1.4527
|  | 572.6 | 1.4485
|  | 572.7 | 1.445
|  | 572.8 | 1.4415
|  | 572.9 | 1.4423
|  | 573 | 1.4406
|  | 573.1 | 1.4388
|  | 573.2 | 1.4446
|  | 573.3 | 1.4464
|  | 573.4 | 1.4481
|  | 573.5 | 1.4531
|  | 573.6 | 1.4566
|  | 573.7 | 1.4533
|  | 573.8 | 1.4533
|  | 573.9 | 1.4533
|  | 574 | 1.4603
|  | 574.1 | 1.4621
|  | 574.2 | 1.4638
|  | 574.3 | 1.4567
|  | 574.4 | 1.4567
|  | 574.5 | 1.4567
|  | 574.6 | 1.4508
|  | 574.7 | 1.4455
|  | 574.8 | 1.4402
|  | 574.9 | 1.4511
|  | 575 | 1.4529
|  | 575.1 | 1.4546
|  | 575.2 | 1.4561
|  | 575.3 | 1.4596
|  | 575.4 | 1.4632
|  | 575.5 | 1.465
|  | 575.6 | 1.4668
|  | 575.7 | 1.4704
|  | 575.8 | 1.4739
|  | 575.9 | 1.4774
|  | 576 | 1.549
|  | 576.1 | 1.5842
|  | 576.2 | 1.6194
|  | 576.3 | 1.7991
|  | 576.4 | 1.9098
|  | 576.5 | 2.0205
|  | 576.6 | 2.1506
|  | 576.7 | 2.2862
|  | 576.8 | 2.4217
|  | 576.9 | 2.6403
|  | 577 | 2.8005
|  | 577.1 | 2.9607
|  | 577.2 | 2.6284
|  | 577.3 | 2.6267
|  | 577.4 | 2.3443
|  | 577.5 | 2.0979
|  | 577.6 | 1.8514
|  | 577.7 | 1.823
|  | 577.8 | 1.6628
|  | 577.9 | 1.5026
|  | 578 | 1.5667
|  | 578.1 | 1.5579
|  | 578.2 | 1.5491
|  | 578.3 | 1.6862
|  | 578.4 | 1.7479
|  | 578.5 | 1.8096
|  | 578.6 | 1.9848
|  | 578.7 | 2.1065
|  | 578.8 | 2.2282
|  | 578.9 | 2.3878
|  | 579 | 2.5286
|  | 579.1 | 2.5789
|  | 579.2 | 2.6742
|  | 579.3 | 2.7694
|  | 579.4 | 2.3873
|  | 579.5 | 2.2303
|  | 579.6 | 2.0732
|  | 579.7 | 1.9615
|  | 579.8 | 1.7392
|  | 579.9 | 1.5168
|  | 580 | 1.5769
|  | 580.1 | 1.5258
|  | 580.2 | 1.4747
|  | 580.3 | 1.5052
|  | 580.4 | 1.4981
|  | 580.5 | 1.4911
|  | 580.6 | 1.4885
|  | 580.7 | 1.4867
|  | 580.8 | 1.4832
|  | 580.9 | 1.4779
|  | 581 | 1.4726
|  | 581.1 | 1.4657
|  | 581.2 | 1.4604
|  | 581.3 | 1.4551
|  | 581.4 | 1.4633
|  | 581.5 | 1.4633
|  | 581.6 | 1.4633
|  | 581.7 | 1.4609
|  | 581.8 | 1.4627
|  | 581.9 | 1.4645
|  | 582 | 1.4621
|  | 582.1 | 1.4603
|  | 582.2 | 1.4586
|  | 582.3 | 1.4649
|  | 582.4 | 1.4666
|  | 582.5 | 1.4633
|  | 582.6 | 1.4633
|  | 582.7 | 1.4633
|  | 582.8 | 1.4667
|  | 582.9 | 1.4667
|  | 583 | 1.4667
|  | 583.1 | 1.4673
|  | 583.2 | 1.4691
|  | 583.3 | 1.4709
|  | 583.4 | 1.47
|  | 583.5 | 1.47
|  | 583.6 | 1.47
|  | 583.7 | 1.4654
|  | 583.8 | 1.4636
|  | 583.9 | 1.4618
|  | 584 | 1.4617
|  | 584.1 | 1.46
|  | 584.2 | 1.46
|  | 584.3 | 1.46
|  | 584.4 | 1.46
|  | 584.5 | 1.4524
|  | 584.6 | 1.4489
|  | 584.7 | 1.4453
|  | 584.8 | 1.4533
|  | 584.9 | 1.4533
|  | 585 | 1.4533
|  | 585.1 | 1.4467
|  | 585.2 | 1.4467
|  | 585.3 | 1.4467
|  | 585.4 | 1.4391
|  | 585.5 | 1.4338
|  | 585.6 | 1.4285
|  | 585.7 | 1.4367
|  | 585.8 | 1.4367
|  | 585.9 | 1.4333
|  | 586 | 1.4333
|  | 586.1 | 1.4333
|  | 586.2 | 1.4327
|  | 586.3 | 1.431
|  | 586.4 | 1.4292
|  | 586.5 | 1.4385
|  | 586.6 | 1.442
|  | 586.7 | 1.4456
|  | 586.8 | 1.4458
|  | 586.9 | 1.4493
|  | 587 | 1.4529
|  | 587.1 | 1.4598
|  | 587.2 | 1.4633
|  | 587.3 | 1.4635
|  | 587.4 | 1.4671
|  | 587.5 | 1.4706
|  | 587.6 | 1.4587
|  | 587.7 | 1.4534
|  | 587.8 | 1.4481
|  | 587.9 | 1.4396
|  | 588 | 1.4307
|  | 588.1 | 1.4218
|  | 588.2 | 1.4157
|  | 588.3 | 1.4086
|  | 588.4 | 1.4015
|  | 588.5 | 1.3896
|  | 588.6 | 1.3807
|  | 588.7 | 1.3719
|  | 588.8 | 1.3799
|  | 588.9 | 1.3763
|  | 589 | 1.3733
|  | 589.1 | 1.3733
|  | 589.2 | 1.3733
|  | 589.3 | 1.3767
|  | 589.4 | 1.3767
|  | 589.5 | 1.3767
|  | 589.6 | 1.3776
|  | 589.7 | 1.3794
|  | 589.8 | 1.3812
|  | 589.9 | 1.3754
|  | 590 | 1.3736
|  | 590.1 | 1.3718
|  | 590.2 | 1.3667
|  | 590.3 | 1.3632
|  | 590.4 | 1.3631
|  | 590.5 | 1.3614
|  | 590.6 | 1.3596
|  | 590.7 | 1.36
|  | 590.8 | 1.36
|  | 590.9 | 1.36
|  | 591 | 1.3558
|  | 591.1 | 1.354
|  | 591.2 | 1.3522
|  | 591.3 | 1.3612
|  | 591.4 | 1.363
|  | 591.5 | 1.3648
|  | 591.6 | 1.3567
|  | 591.7 | 1.3567
|  | 591.8 | 1.3599
|  | 591.9 | 1.3581
|  | 592 | 1.3563
|  | 592.1 | 1.3571
|  | 592.2 | 1.3589
|  | 592.3 | 1.3607
|  | 592.4 | 1.3559
|  | 592.5 | 1.3541
|  | 592.6 | 1.3523
|  | 592.7 | 1.3612
|  | 592.8 | 1.3629
|  | 592.9 | 1.3647
|  | 593 | 1.3615
|  | 593.1 | 1.3633
|  | 593.2 | 1.3599
|  | 593.3 | 1.3563
|  | 593.4 | 1.3527
|  | 593.5 | 1.3529
|  | 593.6 | 1.3511
|  | 593.7 | 1.3493
|  | 593.8 | 1.3377
|  | 593.9 | 1.3323
|  | 594 | 1.327
|  | 594.1 | 1.3266
|  | 594.2 | 1.3213
|  | 594.3 | 1.3159
|  | 594.4 | 1.3104
|  | 594.5 | 1.3068
|  | 594.6 | 1.3032
|  | 594.7 | 1.2978
|  | 594.8 | 1.2925
|  | 594.9 | 1.2892
|  | 595 | 1.2856
|  | 595.1 | 1.282
|  | 595.2 | 1.2826
|  | 595.3 | 1.2808
|  | 595.4 | 1.279
|  | 595.5 | 1.2755
|  | 595.6 | 1.2737
|  | 595.7 | 1.272
|  | 595.8 | 1.2718
|  | 595.9 | 1.27
|  | 596 | 1.2666
|  | 596.1 | 1.2648
|  | 596.2 | 1.263
|  | 596.3 | 1.2591
|  | 596.4 | 1.2556
|  | 596.5 | 1.252
|  | 596.6 | 1.2525
|  | 596.7 | 1.2507
|  | 596.8 | 1.249
|  | 596.9 | 1.2455
|  | 597 | 1.2437
|  | 597.1 | 1.2419
|  | 597.2 | 1.2418
|  | 597.3 | 1.24
|  | 597.4 | 1.2366
|  | 597.5 | 1.2348
|  | 597.6 | 1.233
|  | 597.7 | 1.2328
|  | 597.8 | 1.2311
|  | 597.9 | 1.2293
|  | 598 | 1.23
|  | 598.1 | 1.23
|  | 598.2 | 1.23
|  | 598.3 | 1.2346
|  | 598.4 | 1.2363
|  | 598.5 | 1.2381
|  | 598.6 | 1.2383
|  | 598.7 | 1.2401
|  | 598.8 | 1.2365
|  | 598.9 | 1.2347
|  | 599 | 1.2329
|  | 599.1 | 1.2328
|  | 599.2 | 1.231
|  | 599.3 | 1.2292
|  | 599.4 | 1.2343
|  | 599.5 | 1.2361
|  | 599.6 | 1.2378
|  | 599.7 | 1.224
|  | 599.8 | 1.2205
|  | 599.9 | 1.2169
|  | 600 | 1.2099
|  | 600.1 | 1.2027
|  | 600.2 | 1.203
|  | 600.3 | 1.2013
|  | 600.4 | 1.1995
|  | 600.5 | 1.2
|  | 600.6 | 1.2
|  | 600.7 | 1.2
|  | 600.8 | 1.1912
|  | 600.9 | 1.1877
|  | 601 | 1.1841
|  | 601.1 | 1.179
|  | 601.2 | 1.1737
|  | 601.3 | 1.17
|  | 601.4 | 1.1664
|  | 601.5 | 1.1628
|  | 601.6 | 1.1592
|  | 601.7 | 1.1556
|  | 601.8 | 1.152
|  | 601.9 | 1.1484
|  | 602 | 1.1448
|  | 602.1 | 1.1412
|  | 602.2 | 1.1421
|  | 602.3 | 1.1403
|  | 602.4 | 1.1385
|  | 602.5 | 1.14
|  | 602.6 | 1.14
|  | 602.7 | 1.14
|  | 602.8 | 1.14
|  | 602.9 | 1.14
|  | 603 | 1.14
|  | 603.1 | 1.14
|  | 603.2 | 1.14
|  | 603.3 | 1.1314
|  | 603.4 | 1.1278
|  | 603.5 | 1.1242
|  | 603.6 | 1.1239
|  | 603.7 | 1.1203
|  | 603.8 | 1.1167
|  | 603.9 | 1.1149
|  | 604 | 1.1131
|  | 604.1 | 1.113
|  | 604.2 | 1.1112
|  | 604.3 | 1.1094
|  | 604.4 | 1.11
|  | 604.5 | 1.11
|  | 604.6 | 1.11
|  | 604.7 | 1.1055
|  | 604.8 | 1.1037
|  | 604.9 | 1.1019
|  | 605 | 1.0969
|  | 605.1 | 1.0933
|  | 605.2 | 1.0932
|  | 605.3 | 1.0914
|  | 605.4 | 1.0896
|  | 605.5 | 1.09
|  | 605.6 | 1.09
|  | 605.7 | 1.09
|  | 605.8 | 1.0857
|  | 605.9 | 1.0839
|  | 606 | 1.0821
|  | 606.1 | 1.0772
|  | 606.2 | 1.0736
|  | 606.3 | 1.07
|  | 606.4 | 1.0715
|  | 606.5 | 1.0697
|  | 606.6 | 1.07
|  | 606.7 | 1.07
|  | 606.8 | 1.07
|  | 606.9 | 1.0617
|  | 607 | 1.0581
|  | 607.1 | 1.0545
|  | 607.2 | 1.0633
|  | 607.3 | 1.0633
|  | 607.4 | 1.0633
|  | 607.5 | 1.0616
|  | 607.6 | 1.0634
|  | 607.7 | 1.0631
|  | 607.8 | 1.0613
|  | 607.9 | 1.0595
|  | 608 | 1.06
|  | 608.1 | 1.06
|  | 608.2 | 1.06
|  | 608.3 | 1.06
|  | 608.4 | 1.06
|  | 608.5 | 1.06
|  | 608.6 | 1.06
|  | 608.7 | 1.06
|  | 608.8 | 1.0566
|  | 608.9 | 1.0548
|  | 609 | 1.0529
|  | 609.1 | 1.0644
|  | 609.2 | 1.068
|  | 609.3 | 1.0716
|  | 609.4 | 1.0729
|  | 609.5 | 1.0783
|  | 609.6 | 1.0837
|  | 609.7 | 1.0894
|  | 609.8 | 1.093
|  | 609.9 | 1.0967
|  | 610 | 1.1278
|  | 610.1 | 1.1422
|  | 610.2 | 1.1726
|  | 610.3 | 1.1979
|  | 610.4 | 1.2232
|  | 610.5 | 1.2829
|  | 610.6 | 1.3246
|  | 610.7 | 1.3662
|  | 610.8 | 1.456
|  | 610.9 | 1.5211
|  | 611 | 1.5862
|  | 611.1 | 1.6414
|  | 611.2 | 1.7065
|  | 611.3 | 1.7049
|  | 611.4 | 1.7321
|  | 611.5 | 1.7592
|  | 611.6 | 1.6681
|  | 611.7 | 1.6319
|  | 611.8 | 1.5957
|  | 611.9 | 1.4976
|  | 612 | 1.4234
|  | 612.1 | 1.3491
|  | 612.2 | 1.3212
|  | 612.3 | 1.2632
|  | 612.4 | 1.2459
|  | 612.5 | 1.217
|  | 612.6 | 1.188
|  | 612.7 | 1.1771
|  | 612.8 | 1.1608
|  | 612.9 | 1.1445
|  | 613 | 1.1443
|  | 613.1 | 1.1353
|  | 613.2 | 1.1262
|  | 613.3 | 1.1121
|  | 613.4 | 1.1031
|  | 613.5 | 1.099
|  | 613.6 | 1.0899
|  | 613.7 | 1.0809
|  | 613.8 | 1.0827
|  | 613.9 | 1.0809
|  | 614 | 1.0791
|  | 614.1 | 1.0756
|  | 614.2 | 1.0738
|  | 614.3 | 1.072
|  | 614.4 | 1.0669
|  | 614.5 | 1.0633
|  | 614.6 | 1.0562
|  | 614.7 | 1.0507
|  | 614.8 | 1.0453
|  | 614.9 | 1.0376
|  | 615 | 1.0303
|  | 615.1 | 1.0231
|  | 615.2 | 1.0114
|  | 615.3 | 1.0023
|  | 615.4 | 0.99321
|  | 615.5 | 0.98579
|  | 615.6 | 0.9767
|  | 615.7 | 0.96594
|  | 615.8 | 0.95685
|  | 615.9 | 0.94776
|  | 616 | 0.94427
|  | 616.1 | 0.93699
|  | 616.2 | 0.92972
|  | 616.3 | 0.9235
|  | 616.4 | 0.91805
|  | 616.5 | 0.91259
|  | 616.6 | 0.91365
|  | 616.7 | 0.91001
|  | 616.8 | 0.91
|  | 616.9 | 0.91
|  | 617 | 0.91
|  | 617.1 | 0.90213
|  | 617.2 | 0.8985
|  | 617.3 | 0.89486
|  | 617.4 | 0.89017
|  | 617.5 | 0.88472
|  | 617.6 | 0.87926
|  | 617.7 | 0.87696
|  | 617.8 | 0.87332
|  | 617.9 | 0.86619
|  | 618 | 0.86073
|  | 618.1 | 0.85526
|  | 618.2 | 0.85148
|  | 618.3 | 0.84602
|  | 618.4 | 0.84057
|  | 618.5 | 0.84225
|  | 618.6 | 0.84043
|  | 618.7 | 0.83861
|  | 618.8 | 0.84488
|  | 618.9 | 0.8467
|  | 619 | 0.84333
|  | 619.1 | 0.84334
|  | 619.2 | 0.84334
|  | 619.3 | 0.84268
|  | 619.4 | 0.84085
|  | 619.5 | 0.83903
|  | 619.6 | 0.83554
|  | 619.7 | 0.83371
|  | 619.8 | 0.83189
|  | 619.9 | 0.83667
|  | 620 | 0.83667
|  | 620.1 | 0.83691
|  | 620.2 | 0.83873
|  | 620.3 | 0.84055
|  | 620.4 | 0.83191
|  | 620.5 | 0.82826
|  | 620.6 | 0.82461
|  | 620.7 | 0.82429
|  | 620.8 | 0.82064
|  | 620.9 | 0.817
|  | 621 | 0.81001
|  | 621.1 | 0.80636
|  | 621.2 | 0.80603
|  | 621.3 | 0.80237
|  | 621.4 | 0.79872
|  | 621.5 | 0.8
|  | 621.6 | 0.8
|  | 621.7 | 0.8
|  | 621.8 | 0.8
|  | 621.9 | 0.8
|  | 622 | 0.8
|  | 622.1 | 0.78983
|  | 622.2 | 0.78618
|  | 622.3 | 0.78959
|  | 622.4 | 0.78776
|  | 622.5 | 0.78593
|  | 622.6 | 0.78333
|  | 622.7 | 0.78333
|  | 622.8 | 0.78333
|  | 622.9 | 0.78196
|  | 623 | 0.78013
|  | 623.1 | 0.78
|  | 623.2 | 0.78
|  | 623.3 | 0.78
|  | 623.4 | 0.77615
|  | 623.5 | 0.77432
|  | 623.6 | 0.77249
|  | 623.7 | 0.77233
|  | 623.8 | 0.77049
|  | 623.9 | 0.76866
|  | 624 | 0.76035
|  | 624.1 | 0.75669
|  | 624.2 | 0.75985
|  | 624.3 | 0.75802
|  | 624.4 | 0.75618
|  | 624.5 | 0.74936
|  | 624.6 | 0.74753
|  | 624.7 | 0.7457
|  | 624.8 | 0.73993
|  | 624.9 | 0.73443
|  | 625 | 0.72894
|  | 625.1 | 0.72674
|  | 625.2 | 0.72308
|  | 625.3 | 0.72667
|  | 625.4 | 0.72667
|  | 625.5 | 0.72667
|  | 625.6 | 0.72333
|  | 625.7 | 0.72333
|  | 625.8 | 0.72333
|  | 625.9 | 0.71742
|  | 626 | 0.71375
|  | 626.1 | 0.71008
|  | 626.2 | 0.73205
|  | 626.3 | 0.73756
|  | 626.4 | 0.71954
|  | 626.5 | 0.71771
|  | 626.6 | 0.71588
|  | 626.7 | 0.71616
|  | 626.8 | 0.70881
|  | 626.9 | 0.70146
|  | 627 | 0.70813
|  | 627.1 | 0.70996
|  | 627.2 | 0.70653
|  | 627.3 | 0.7047
|  | 627.4 | 0.70286
|  | 627.5 | 0.69476
|  | 627.6 | 0.68925
|  | 627.7 | 0.68375
|  | 627.8 | 0.68885
|  | 627.9 | 0.68702
|  | 628 | 0.68518
|  | 628.1 | 0.68333
|  | 628.2 | 0.68333
|  | 628.3 | 0.68301
|  | 628.4 | 0.68118
|  | 628.5 | 0.67934
|  | 628.6 | 0.67584
|  | 628.7 | 0.674
|  | 628.8 | 0.67217
|  | 628.9 | 0.66732
|  | 629 | 0.66364
|  | 629.1 | 0.66667
|  | 629.2 | 0.66667
|  | 629.3 | 0.66667
|  | 629.4 | 0.66333
|  | 629.5 | 0.66333
|  | 629.6 | 0.66333
|  | 629.7 | 0.66667
|  | 629.8 | 0.66667
|  | 629.9 | 0.66667
|  | 630 | 0.67312
|  | 630.1 | 0.6768
|  | 630.2 | 0.67691
|  | 630.3 | 0.67876
|  | 630.4 | 0.6806
|  | 630.5 | 0.6841
|  | 630.6 | 0.68594
|  | 630.7 | 0.68779
|  | 630.8 | 0.69719
|  | 630.9 | 0.70271
|  | 631 | 0.70822
|  | 631.1 | 0.69667
|  | 631.2 | 0.69668
|  | 631.3 | 0.68804
|  | 631.4 | 0.68067
|  | 631.5 | 0.6733
|  | 631.6 | 0.66157
|  | 631.7 | 0.65237
|  | 631.8 | 0.64316
|  | 631.9 | 0.63563
|  | 632 | 0.62642
|  | 632.1 | 0.62265
|  | 632.2 | 0.61713
|  | 632.3 | 0.6116
|  | 632.4 | 0.6044
|  | 632.5 | 0.59886
|  | 632.6 | 0.59333
|  | 632.7 | 0.58949
|  | 632.8 | 0.58396
|  | 632.9 | 0.57844
|  | 633 | 0.58666
|  | 633.1 | 0.58666
|  | 633.2 | 0.5795
|  | 633.3 | 0.57766
|  | 633.4 | 0.57581
|  | 633.5 | 0.57897
|  | 633.6 | 0.57712
|  | 633.7 | 0.57528
|  | 633.8 | 0.57333
|  | 633.9 | 0.57333
|  | 634 | 0.56948
|  | 634.1 | 0.56579
|  | 634.2 | 0.5621
|  | 634.3 | 0.56254
|  | 634.4 | 0.56069
|  | 634.5 | 0.55885
|  | 634.6 | 0.55533
|  | 634.7 | 0.55348
|  | 634.8 | 0.5533
|  | 634.9 | 0.55146
|  | 635 | 0.54961
|  | 635.1 | 0.54609
|  | 635.2 | 0.54425
|  | 635.3 | 0.5424
|  | 635.4 | 0.54222
|  | 635.5 | 0.54037
|  | 635.6 | 0.53852
|  | 635.7 | 0.54
|  | 635.8 | 0.54
|  | 635.9 | 0.53631
|  | 636 | 0.53446
|  | 636.1 | 0.53261
|  | 636.2 | 0.53243
|  | 636.3 | 0.53059
|  | 636.4 | 0.52874
|  | 636.5 | 0.52522
|  | 636.6 | 0.52337
|  | 636.7 | 0.52667
|  | 636.8 | 0.52667
|  | 636.9 | 0.52667
|  | 637 | 0.52333
|  | 637.1 | 0.52333
|  | 637.2 | 0.52333
|  | 637.3 | 0.52209
|  | 637.4 | 0.52024
|  | 637.5 | 0.51839
|  | 637.6 | 0.51487
|  | 637.7 | 0.51302
|  | 637.8 | 0.51283
|  | 637.9 | 0.51098
|  | 638 | 0.50913
|  | 638.1 | 0.50561
|  | 638.2 | 0.50376
|  | 638.3 | 0.50191
|  | 638.4 | 0.50172
|  | 638.5 | 0.49987
|  | 638.6 | 0.5
|  | 638.7 | 0.5
|  | 638.8 | 0.5
|  | 638.9 | 0.5
|  | 639 | 0.5
|  | 639.1 | 0.5
|  | 639.2 | 0.49522
|  | 639.3 | 0.49337
|  | 639.4 | 0.4897
|  | 639.5 | 0.486
|  | 639.6 | 0.4823
|  | 639.7 | 0.48262
|  | 639.8 | 0.48077
|  | 639.9 | 0.47891
|  | 640 | 0.48
|  | 640.1 | 0.48
|  | 640.2 | 0.48
|  | 640.3 | 0.47482
|  | 640.4 | 0.47297
|  | 640.5 | 0.4689
|  | 640.6 | 0.46519
|  | 640.7 | 0.46148
|  | 640.8 | 0.46221
|  | 640.9 | 0.46036
|  | 641 | 0.4585
|  | 641.1 | 0.46502
|  | 641.2 | 0.46687
|  | 641.3 | 0.46333
|  | 641.4 | 0.46334
|  | 641.5 | 0.46334
|  | 641.6 | 0.46237
|  | 641.7 | 0.46051
|  | 641.8 | 0.45865
|  | 641.9 | 0.46
|  | 642 | 0.46
|  | 642.1 | 0.45641
|  | 642.2 | 0.45455
|  | 642.3 | 0.45269
|  | 642.4 | 0.4525
|  | 642.5 | 0.45064
|  | 642.6 | 0.44878
|  | 642.7 | 0.45
|  | 642.8 | 0.45
|  | 642.9 | 0.45
|  | 643 | 0.45
|  | 643.1 | 0.45
|  | 643.2 | 0.44596
|  | 643.3 | 0.4441
|  | 643.4 | 0.44224
|  | 643.5 | 0.43742
|  | 643.6 | 0.43369
|  | 643.7 | 0.43667
|  | 643.8 | 0.43667
|  | 643.9 | 0.43667
|  | 644 | 0.43726
|  | 644.1 | 0.43912
|  | 644.2 | 0.44098
|  | 644.3 | 0.43098
|  | 644.4 | 0.42725
|  | 644.5 | 0.42353
|  | 644.6 | 0.42314
|  | 644.7 | 0.41942
|  | 644.8 | 0.42
|  | 644.9 | 0.42
|  | 645 | 0.42
|  | 645.1 | 0.42
|  | 645.2 | 0.42
|  | 645.3 | 0.42
|  | 645.4 | 0.42
|  | 645.5 | 0.42
|  | 645.6 | 0.41627
|  | 645.7 | 0.41441
|  | 645.8 | 0.41254
|  | 645.9 | 0.41234
|  | 646 | 0.41048
|  | 646.1 | 0.40861
|  | 646.2 | 0.41
|  | 646.3 | 0.41
|  | 646.4 | 0.41
|  | 646.5 | 0.41
|  | 646.6 | 0.41
|  | 646.7 | 0.40154
|  | 646.8 | 0.39781
|  | 646.9 | 0.39408
|  | 647 | 0.39368
|  | 647.1 | 0.38995
|  | 647.2 | 0.39356
|  | 647.3 | 0.39543
|  | 647.4 | 0.3973
|  | 647.5 | 0.38916
|  | 647.6 | 0.38729
|  | 647.7 | 0.38542
|  | 647.8 | 0.38856
|  | 647.9 | 0.3867
|  | 648 | 0.38683
|  | 648.1 | 0.3887
|  | 648.2 | 0.39057
|  | 648.3 | 0.39
|  | 648.4 | 0.39
|  | 648.5 | 0.39
|  | 648.6 | 0.39
|  | 648.7 | 0.39
|  | 648.8 | 0.39
|  | 648.9 | 0.39
|  | 649 | 0.39
|  | 649.1 | 0.38594
|  | 649.2 | 0.38407
|  | 649.3 | 0.3822
|  | 649.4 | 0.382
|  | 649.5 | 0.38013
|  | 649.6 | 0.38
|  | 649.7 | 0.38
|  | 649.8 | 0.38
|  | 649.9 | 0.38401
|  | 650 | 0.38588
|  | 650.1 | 0.38775
|  | 650.2 | 0.38796
|  | 650.3 | 0.38983
|  | 650.4 | 0.39
|  | 650.5 | 0.39
|  | 650.6 | 0.39
|  | 650.7 | 0.38601
|  | 650.8 | 0.38414
|  | 650.9 | 0.38227
|  | 651 | 0.37745
|  | 651.1 | 0.37371
|  | 651.2 | 0.36996
|  | 651.3 | 0.36622
|  | 651.4 | 0.36247
|  | 651.5 | 0.3627
|  | 651.6 | 0.36082
|  | 651.7 | 0.35895
|  | 651.8 | 0.36
|  | 651.9 | 0.36
|  | 652 | 0.36
|  | 652.1 | 0.36
|  | 652.2 | 0.36
|  | 652.3 | 0.36
|  | 652.4 | 0.36
|  | 652.5 | 0.36
|  | 652.6 | 0.36
|  | 652.7 | 0.36
|  | 652.8 | 0.36
|  | 652.9 | 0.34959
|  | 653 | 0.34584
|  | 653.1 | 0.34937
|  | 653.2 | 0.34749
|  | 653.3 | 0.34561
|  | 653.4 | 0.34333
|  | 653.5 | 0.34333
|  | 653.6 | 0.34333
|  | 653.7 | 0.34145
|  | 653.8 | 0.33957
|  | 653.9 | 0.33603
|  | 654 | 0.33416
|  | 654.1 | 0.33228
|  | 654.2 | 0.33207
|  | 654.3 | 0.33019
|  | 654.4 | 0.33
|  | 654.5 | 0.33
|  | 654.6 | 0.33
|  | 654.7 | 0.32601
|  | 654.8 | 0.32413
|  | 654.9 | 0.32225
|  | 655 | 0.31741
|  | 655.1 | 0.31365
|  | 655.2 | 0.31667
|  | 655.3 | 0.31667
|  | 655.4 | 0.31667
|  | 655.5 | 0.31333
|  | 655.6 | 0.31333
|  | 655.7 | 0.31333
|  | 655.8 | 0.312
|  | 655.9 | 0.31012
|  | 656 | 0.31
|  | 656.1 | 0.31
|  | 656.2 | 0.31
|  | 656.3 | 0.31
|  | 656.4 | 0.31
|  | 656.5 | 0.31
|  | 656.6 | 0.31
|  | 656.7 | 0.31
|  | 656.8 | 0.31
|  | 656.9 | 0.31
|  | 657 | 0.31
|  | 657.1 | 0.31
|  | 657.2 | 0.31
|  | 657.3 | 0.31
|  | 657.4 | 0.31
|  | 657.5 | 0.31
|  | 657.6 | 0.31
|  | 657.7 | 0.31
|  | 657.8 | 0.31
|  | 657.9 | 0.30165
|  | 658 | 0.29788
|  | 658.1 | 0.29411
|  | 658.2 | 0.2985
|  | 658.3 | 0.29662
|  | 658.4 | 0.29693
|  | 658.5 | 0.29882
|  | 658.6 | 0.3007
|  | 658.7 | 0.29575
|  | 658.8 | 0.29387
|  | 658.9 | 0.29198
|  | 659 | 0.29176
|  | 659.1 | 0.28988
|  | 659.2 | 0.29367
|  | 659.3 | 0.29556
|  | 659.4 | 0.29745
|  | 659.5 | 0.29333
|  | 659.6 | 0.29333
|  | 659.7 | 0.29333
|  | 659.8 | 0.28668
|  | 659.9 | 0.28291
|  | 660 | 0.2829
|  | 660.1 | 0.28101
|  | 660.2 | 0.27913
|  | 660.3 | 0.27557
|  | 660.4 | 0.27369
|  | 660.5 | 0.2718
|  | 660.6 | 0.27667
|  | 660.7 | 0.27667
|  | 660.8 | 0.27719
|  | 660.9 | 0.27908
|  | 661 | 0.28097
|  | 661.1 | 0.27094
|  | 661.2 | 0.26716
|  | 661.3 | 0.26338
|  | 661.4 | 0.27334
|  | 661.5 | 0.27334
|  | 661.6 | 0.27462
|  | 661.7 | 0.27839
|  | 661.8 | 0.28216
|  | 661.9 | 0.28
|  | 662 | 0.28
|  | 662.1 | 0.28
|  | 662.2 | 0.28
|  | 662.3 | 0.28
|  | 662.4 | 0.28
|  | 662.5 | 0.28
|  | 662.6 | 0.28
|  | 662.7 | 0.28
|  | 662.8 | 0.28
|  | 662.9 | 0.27289
|  | 663 | 0.26911
|  | 663.1 | 0.26533
|  | 663.2 | 0.2691
|  | 663.3 | 0.26721
|  | 663.4 | 0.26532
|  | 663.5 | 0.26333
|  | 663.6 | 0.26333
|  | 663.7 | 0.26297
|  | 663.8 | 0.26108
|  | 663.9 | 0.25919
|  | 664 | 0.25563
|  | 664.1 | 0.25373
|  | 664.2 | 0.25184
|  | 664.3 | 0.25667
|  | 664.4 | 0.25667
|  | 664.5 | 0.25718
|  | 664.6 | 0.25907
|  | 664.7 | 0.26097
|  | 664.8 | 0.25547
|  | 664.9 | 0.25358
|  | 665 | 0.25169
|  | 665.1 | 0.25146
|  | 665.2 | 0.24956
|  | 665.3 | 0.246
|  | 665.4 | 0.24411
|  | 665.5 | 0.24222
|  | 665.6 | 0.24198
|  | 665.7 | 0.24008
|  | 665.8 | 0.24
|  | 665.9 | 0.24
|  | 666 | 0.24
|  | 666.1 | 0.24
|  | 666.2 | 0.24
|  | 666.3 | 0.24
|  | 666.4 | 0.24
|  | 666.5 | 0.24
|  | 666.6 | 0.24
|  | 666.7 | 0.24
|  | 666.8 | 0.24
|  | 666.9 | 0.24
|  | 667 | 0.24
|  | 667.1 | 0.24
|  | 667.2 | 0.24505
|  | 667.3 | 0.24694
|  | 667.4 | 0.23949
|  | 667.5 | 0.23759
|  | 667.6 | 0.23568
|  | 667.7 | 0.23425
|  | 667.8 | 0.23046
|  | 667.9 | 0.22667
|  | 668 | 0.23
|  | 668.1 | 0.23
|  | 668.2 | 0.23
|  | 668.3 | 0.23
|  | 668.4 | 0.23
|  | 668.5 | 0.23
|  | 668.6 | 0.23
|  | 668.7 | 0.22646
|  | 668.8 | 0.22456
|  | 668.9 | 0.22266
|  | 669 | 0.22242
|  | 669.1 | 0.22052
|  | 669.2 | 0.21862
|  | 669.3 | 0.22
|  | 669.4 | 0.22
|  | 669.5 | 0.21624
|  | 669.6 | 0.21434
|  | 669.7 | 0.21243
|  | 669.8 | 0.21667
|  | 669.9 | 0.21667
|  | 670 | 0.21667
|  | 670.1 | 0.21333
|  | 670.2 | 0.21333
|  | 670.3 | 0.22065
|  | 670.4 | 0.22255
|  | 670.5 | 0.22445
|  | 670.6 | 0.22136
|  | 670.7 | 0.22326
|  | 670.8 | 0.21966
|  | 670.9 | 0.21586
|  | 671 | 0.21205
|  | 671.1 | 0.21245
|  | 671.2 | 0.21054
|  | 671.3 | 0.20864
|  | 671.4 | 0.21494
|  | 671.5 | 0.21684
|  | 671.6 | 0.21333
|  | 671.7 | 0.21333
|  | 671.8 | 0.21333
|  | 671.9 | 0.20774
|  | 672 | 0.20393
|  | 672.1 | 0.20012
|  | 672.2 | 0.20667
|  | 672.3 | 0.20667
|  | 672.4 | 0.20733
|  | 672.5 | 0.20924
|  | 672.6 | 0.21115
|  | 672.7 | 0.20527
|  | 672.8 | 0.20337
|  | 672.9 | 0.20312
|  | 673 | 0.20122
|  | 673.1 | 0.19931
|  | 673.2 | 0.2
|  | 673.3 | 0.2
|  | 673.4 | 0.2
|  | 673.5 | 0.2
|  | 673.6 | 0.2
|  | 673.7 | 0.19619
|  | 673.8 | 0.19428
|  | 673.9 | 0.19237
|  | 674 | 0.19667
|  | 674.1 | 0.19667
|  | 674.2 | 0.19669
|  | 674.3 | 0.19859
|  | 674.4 | 0.2005
|  | 674.5 | 0.2
|  | 674.6 | 0.2
|  | 674.7 | 0.2
|  | 674.8 | 0.19518
|  | 674.9 | 0.19327
|  | 675 | 0.19302
|  | 675.1 | 0.19111
|  | 675.2 | 0.18919
|  | 675.3 | 0.19438
|  | 675.4 | 0.19629
|  | 675.5 | 0.19821
|  | 675.6 | 0.19334
|  | 675.7 | 0.19334
|  | 675.8 | 0.19667
|  | 675.9 | 0.19667
|  | 676 | 0.19667
|  | 676.1 | 0.19801
|  | 676.2 | 0.19993
|  | 676.3 | 0.19299
|  | 676.4 | 0.18916
|  | 676.5 | 0.18533
|  | 676.6 | 0.18908
|  | 676.7 | 0.18716
|  | 676.8 | 0.18525
|  | 676.9 | 0.18833
|  | 677 | 0.19025
|  | 677.1 | 0.19
|  | 677.2 | 0.19
|  | 677.3 | 0.19
|  | 677.4 | 0.17631
|  | 677.5 | 0.17058
|  | 677.6 | 0.1834
|  | 677.7 | 0.18531
|  | 677.8 | 0.18722
|  | 677.9 | 0.17748
|  | 678 | 0.17939
|  | 678.1 | 0.1813
|  | 678.2 | 0.17529
|  | 678.3 | 0.16954
|  | 678.4 | 0.17
|  | 678.5 | 0.17
|  | 678.6 | 0.17
|  | 678.7 | 0.17448
|  | 678.8 | 0.1764
|  | 678.9 | 0.17831
|  | 679 | 0.17857
|  | 679.1 | 0.1805
|  | 679.2 | 0.18
|  | 679.3 | 0.18
|  | 679.4 | 0.18
|  | 679.5 | 0.17516
|  | 679.6 | 0.17325
|  | 679.7 | 0.17667
|  | 679.8 | 0.17667
|  | 679.9 | 0.17667
|  | 680 | 0.17333
|  | 680.1 | 0.17334
|  | 680.2 | 0.17334
|  | 680.3 | 0.17148
|  | 680.4 | 0.16956
|  | 680.5 | 0.17
|  | 680.6 | 0.17
|  | 680.7 | 0.17
|  | 680.8 | 0.17
|  | 680.9 | 0.17
|  | 681 | 0.17
|  | 681.1 | 0.17
|  | 681.2 | 0.17
|  | 681.3 | 0.1656
|  | 681.4 | 0.16368
|  | 681.5 | 0.16176
|  | 681.6 | 0.16666
|  | 681.7 | 0.16666
|  | 681.8 | 0.16333
|  | 681.9 | 0.16333
|  | 682 | 0.16333
|  | 682.1 | 0.16667
|  | 682.2 | 0.16667
|  | 682.3 | 0.16333
|  | 682.4 | 0.16334
|  | 682.5 | 0.16334
|  | 682.6 | 0.16226
|  | 682.7 | 0.16034
|  | 682.8 | 0.15841
|  | 682.9 | 0.16
|  | 683 | 0.16
|  | 683.1 | 0.16
|  | 683.2 | 0.16
|  | 683.3 | 0.16
|  | 683.4 | 0.1552
|  | 683.5 | 0.15327
|  | 683.6 | 0.14936
|  | 683.7 | 0.1455
|  | 683.8 | 0.14165
|  | 683.9 | 0.1511
|  | 684 | 0.15303
|  | 684.1 | 0.15495
|  | 684.2 | 0.15188
|  | 684.3 | 0.1538
|  | 684.4 | 0.15667
|  | 684.5 | 0.15666
|  | 684.6 | 0.15666
|  | 684.7 | 0.14848
|  | 684.8 | 0.14655
|  | 684.9 | 0.14963
|  | 685 | 0.1477
|  | 685.1 | 0.14577
|  | 685.2 | 0.14782
|  | 685.3 | 0.14976
|  | 685.4 | 0.14665
|  | 685.5 | 0.14472
|  | 685.6 | 0.14279
|  | 685.7 | 0.14252
|  | 685.8 | 0.14059
|  | 685.9 | 0.13866
|  | 686 | 0.14493
|  | 686.1 | 0.14686
|  | 686.2 | 0.14712
|  | 686.3 | 0.14904
|  | 686.4 | 0.15097
|  | 686.5 | 0.14542
|  | 686.6 | 0.14349
|  | 686.7 | 0.15011
|  | 686.8 | 0.15205
|  | 686.9 | 0.15399
|  | 687 | 0.15515
|  | 687.1 | 0.15901
|  | 687.2 | 0.16287
|  | 687.3 | 0.16503
|  | 687.4 | 0.16696
|  | 687.5 | 0.15944
|  | 687.6 | 0.15751
|  | 687.7 | 0.15558
|  | 687.8 | 0.15394
|  | 687.9 | 0.15007
|  | 688 | 0.15356
|  | 688.1 | 0.15549
|  | 688.2 | 0.15743
|  | 688.3 | 0.15333
|  | 688.4 | 0.15334
|  | 688.5 | 0.15334
|  | 688.6 | 0.15667
|  | 688.7 | 0.15667
|  | 688.8 | 0.1493
|  | 688.9 | 0.14737
|  | 689 | 0.14544
|  | 689.1 | 0.14366
|  | 689.2 | 0.1398
|  | 689.3 | 0.14
|  | 689.4 | 0.14
|  | 689.5 | 0.14
|  | 689.6 | 0.14
|  | 689.7 | 0.14
|  | 689.8 | 0.14
|  | 689.9 | 0.14
|  | 690 | 0.14
|  | 690.1 | 0.14
|  | 690.2 | 0.14
|  | 690.3 | 0.14
|  | 690.4 | 0.17001
|  | 690.5 | 0.18161
|  | 690.6 | 0.18713
|  | 690.7 | 0.2007
|  | 690.8 | 0.21426
|  | 690.9 | 0.17988
|  | 691 | 0.17019
|  | 691.1 | 0.16142
|  | 691.2 | 0.14588
|  | 691.3 | 0.13034
|  | 691.4 | 0.13455
|  | 691.5 | 0.13066
|  | 691.6 | 0.12678
|  | 691.7 | 0.13
|  | 691.8 | 0.13
|  | 691.9 | 0.13
|  | 692 | 0.13
|  | 692.1 | 0.13
|  | 692.2 | 0.12509
|  | 692.3 | 0.12315
|  | 692.4 | 0.12667
|  | 692.5 | 0.12667
|  | 692.6 | 0.12667
|  | 692.7 | 0.12334
|  | 692.8 | 0.12334
|  | 692.9 | 0.13367
|  | 693 | 0.13756
|  | 693.1 | 0.14145
|  | 693.2 | 0.13433
|  | 693.3 | 0.13628
|  | 693.4 | 0.13823
|  | 693.5 | 0.13818
|  | 693.6 | 0.13625
|  | 693.7 | 0.13333
|  | 693.8 | 0.13334
|  | 693.9 | 0.13334
|  | 694 | 0.1318
|  | 694.1 | 0.12985
|  | 694.2 | 0.13
|  | 694.3 | 0.13
|  | 694.4 | 0.13
|  | 694.5 | 0.12541
|  | 694.6 | 0.12347
|  | 694.7 | 0.11972
|  | 694.8 | 0.11583
|  | 694.9 | 0.11195
|  | 695 | 0.11667
|  | 695.1 | 0.11667
|  | 695.2 | 0.11667
|  | 695.3 | 0.12362
|  | 695.4 | 0.12751
|  | 695.5 | 0.12333
|  | 695.6 | 0.12333
|  | 695.7 | 0.12333
|  | 695.8 | 0.12179
|  | 695.9 | 0.11984
|  | 696 | 0.12378
|  | 696.1 | 0.12573
|  | 696.2 | 0.12768
|  | 696.3 | 0.13719
|  | 696.4 | 0.14302
|  | 696.5 | 0.14018
|  | 696.6 | 0.14212
|  | 696.7 | 0.14407
|  | 696.8 | 0.13796
|  | 696.9 | 0.13406
|  | 697 | 0.13016
|  | 697.1 | 0.12107
|  | 697.2 | 0.11522
|  | 697.3 | 0.11923
|  | 697.4 | 0.11728
|  | 697.5 | 0.11533
|  | 697.6 | 0.11828
|  | 697.7 | 0.12023
|  | 697.8 | 0.12
|  | 697.9 | 0.12
|  | 698 | 0.12
|  | 698.1 | 0.1153
|  | 698.2 | 0.11335
|  | 698.3 | 0.11307
|  | 698.4 | 0.11112
|  | 698.5 | 0.10917
|  | 698.6 | 0.11
|  | 698.7 | 0.11
|  | 698.8 | 0.11
|  | 698.9 | 0.11
|  | 699 | 0.11
|  | 699.1 | 0.11
|  | 699.2 | 0.11
|  | 699.3 | 0.11
|  | 699.4 | 0.11
|  | 699.5 | 0.11
|  | 699.6 | 0.11
|  | 699.7 | 0.11
|  | 699.8 | 0.11
|  | 699.9 | 0.11
|  | 700 | 0.11
|  | 700.1 | 0.11
|  | 700.2 | 0.11
|  | 700.3 | 0.11
|  | 700.4 | 0.1054
|  | 700.5 | 0.10344
|  | 700.6 | 0.10667
|  | 700.7 | 0.10666
|  | 700.8 | 0.10666
|  | 700.9 | 0.10772
|  | 701 | 0.10967
|  | 701.1 | 0.11163
|  | 701.2 | 0.099505
|  | 701.3 | 0.095599
|  | 701.4 | 0.10749
|  | 701.5 | 0.10944
|  | 701.6 | 0.11139
|  | 701.7 | 0.10502
|  | 701.8 | 0.10697
|  | 701.9 | 0.10939
|  | 702 | 0.10742
|  | 702.1 | 0.10546
|  | 702.2 | 0.10815
|  | 702.3 | 0.1101
|  | 702.4 | 0.11
|  | 702.5 | 0.11
|  | 702.6 | 0.11
|  | 702.7 | 0.10539
|  | 702.8 | 0.10343
|  | 702.9 | 0.10314
|  | 703 | 0.10118
|  | 703.1 | 0.099216
|  | 703.2 | 0.1044
|  | 703.3 | 0.10636
|  | 703.4 | 0.10832
|  | 703.5 | 0.10334
|  | 703.6 | 0.10334
|  | 703.7 | 0.10247
|  | 703.8 | 0.10051
|  | 703.9 | 0.098549
|  | 704 | 0.1
|  | 704.1 | 0.1
|  | 704.2 | 0.104
|  | 704.3 | 0.10596
|  | 704.4 | 0.10792
|  | 704.5 | 0.10823
|  | 704.6 | 0.11019
|  | 704.7 | 0.11
|  | 704.8 | 0.11
|  | 704.9 | 0.11
|  | 705 | 0.10529
|  | 705.1 | 0.10333
|  | 705.2 | 0.10667
|  | 705.3 | 0.10667
|  | 705.4 | 0.10667
|  | 705.5 | 0.11237
|  | 705.6 | 0.1163
|  | 705.7 | 0.12022
|  | 705.8 | 0.12415
|  | 705.9 | 0.12808
|  | 706 | 0.12768
|  | 706.1 | 0.12964
|  | 706.2 | 0.13161
|  | 706.3 | 0.1457
|  | 706.4 | 0.1516
|  | 706.5 | 0.15666
|  | 706.6 | 0.16452
|  | 706.7 | 0.17238
|  | 706.8 | 0.15827
|  | 706.9 | 0.15631
|  | 707 | 0.15137
|  | 707.1 | 0.14547
|  | 707.2 | 0.13958
|  | 707.3 | 0.14178
|  | 707.4 | 0.13981
|  | 707.5 | 0.13235
|  | 707.6 | 0.12841
|  | 707.7 | 0.12448
|  | 707.8 | 0.12861
|  | 707.9 | 0.12665
|  | 708 | 0.12699
|  | 708.1 | 0.12895
|  | 708.2 | 0.13092
|  | 708.3 | 0.12545
|  | 708.4 | 0.12348
|  | 708.5 | 0.13016
|  | 708.6 | 0.13213
|  | 708.7 | 0.13409
|  | 708.8 | 0.14425
|  | 708.9 | 0.15213
|  | 709 | 0.16
|  | 709.1 | 0.15198
|  | 709.2 | 0.15396
|  | 709.3 | 0.1439
|  | 709.4 | 0.13799
|  | 709.5 | 0.13207
|  | 709.6 | 0.11751
|  | 709.7 | 0.10767
|  | 709.8 | 0.1077
|  | 709.9 | 0.10179
|  | 710 | 0.095886
|  | 710.1 | 0.1
|  | 710.2 | 0.1
|  | 710.3 | 0.10792
|  | 710.4 | 0.11186
|  | 710.5 | 0.11581
|  | 710.6 | 0.10667
|  | 710.7 | 0.10668
|  | 710.8 | 0.10569
|  | 710.9 | 0.10175
|  | 711 | 0.09781
|  | 711.1 | 0.1
|  | 711.2 | 0.1
|  | 711.3 | 0.17009
|  | 711.4 | 0.20764
|  | 711.5 | 0.24519
|  | 711.6 | 0.17713
|  | 711.7 | 0.183
|  | 711.8 | 0.18331
|  | 711.9 | 0.1537
|  | 712 | 0.1241
|  | 712.1 | 0.13781
|  | 712.2 | 0.13979
|  | 712.3 | 0.13657
|  | 712.4 | 0.13459
|  | 712.5 | 0.13262
|  | 712.6 | 0.12358
|  | 712.7 | 0.11765
|  | 712.8 | 0.11173
|  | 712.9 | 0.10217
|  | 713 | 0.094269
|  | 713.1 | 0.099091
|  | 713.2 | 0.097115
|  | 713.3 | 0.095138
|  | 713.4 | 0.098505
|  | 713.5 | 0.10048
|  | 713.6 | 0.1
|  | 713.7 | 0.1
|  | 713.8 | 0.1
|  | 713.9 | 0.094947
|  | 714 | 0.092971
|  | 714.1 | 0.08864
|  | 714.2 | 0.08468
|  | 714.3 | 0.080719
|  | 714.4 | 0.086669
|  | 714.5 | 0.086671
|  | 714.6 | 0.087248
|  | 714.7 | 0.089229
|  | 714.8 | 0.091211
|  | 714.9 | 0.085148
|  | 715 | 0.083166
|  | 715.1 | 0.082851
|  | 715.2 | 0.08087
|  | 715.3 | 0.078888
|  | 715.4 | 0.075247
|  | 715.5 | 0.073265
|  | 715.6 | 0.076666
|  | 715.7 | 0.076664
|  | 715.8 | 0.076661
|  | 715.9 | 0.077986
|  | 716 | 0.079965
|  | 716.1 | 0.08
|  | 716.2 | 0.08
|  | 716.3 | 0.08
|  | 716.4 | 0.084563
|  | 716.5 | 0.086548
|  | 716.6 | 0.086865
|  | 716.7 | 0.088849
|  | 716.8 | 0.090833
|  | 716.9 | 0.085516
|  | 717 | 0.083532
|  | 717.1 | 0.083214
|  | 717.2 | 0.08123
|  | 717.3 | 0.079246
|  | 717.4 | 0.075595
|  | 717.5 | 0.073611
|  | 717.6 | 0.076667
|  | 717.7 | 0.076667
|  | 717.8 | 0.076667
|  | 717.9 | 0.073333
|  | 718 | 0.073333
|  | 718.1 | 0.073333
|  | 718.2 | 0.071389
|  | 718.3 | 0.069405
|  | 718.4 | 0.07
|  | 718.5 | 0.07
|  | 718.6 | 0.07
|  | 718.7 | 0.07
|  | 718.8 | 0.07
|  | 718.9 | 0.074187
|  | 719 | 0.076171
|  | 719.1 | 0.078155
|  | 719.2 | 0.078484
|  | 719.3 | 0.080473
|  | 719.4 | 0.08
|  | 719.5 | 0.08
|  | 719.6 | 0.08
|  | 719.7 | 0.08
|  | 719.8 | 0.08
|  | 719.9 | 0.08
|  | 720 | 0.08
|  | 720.1 | 0.08
|  | 720.2 | 0.08
|  | 720.3 | 0.08
|  | 720.4 | 0.08
|  | 720.5 | 0.08
|  | 720.6 | 0.08
|  | 720.7 | 0.08
|  | 720.8 | 0.08
|  | 720.9 | 0.076044
|  | 721 | 0.074057
|  | 721.1 | 0.072071
|  | 721.2 | 0.07174
|  | 721.3 | 0.069748
|  | 721.4 | 0.073911
|  | 721.5 | 0.075903
|  | 721.6 | 0.077895
|  | 721.7 | 0.07822
|  | 721.8 | 0.080212
|  | 721.9 | 0.076129
|  | 722 | 0.074137
|  | 722.1 | 0.072145
|  | 722.2 | 0.076667
|  | 722.3 | 0.076667
|  | 722.4 | 0.069502
|  | 722.5 | 0.06751
|  | 722.6 | 0.065518
|  | 722.7 | 0.063697
|  | 722.8 | 0.059703
|  | 722.9 | 0.06
|  | 723 | 0.06
|  | 723.1 | 0.06
|  | 723.2 | 0.06
|  | 723.3 | 0.06
|  | 723.4 | 0.06
|  | 723.5 | 0.06
|  | 723.6 | 0.06
|  | 723.7 | 0.064765
|  | 723.8 | 0.066762
|  | 723.9 | 0.067093
|  | 724 | 0.06909
|  | 724.1 | 0.071087
|  | 724.2 | 0.07
|  | 724.3 | 0.07
|  | 724.4 | 0.07
|  | 724.5 | 0.07
|  | 724.6 | 0.07
|  | 724.7 | 0.07
|  | 724.8 | 0.07
|  | 724.9 | 0.07
|  | 725 | 0.07
|  | 725.1 | 0.07
|  | 725.2 | 0.065287
|  | 725.3 | 0.063287
|  | 725.4 | 0.062953
|  | 725.5 | 0.060953
|  | 725.6 | 0.058953
|  | 725.7 | 0.069427
|  | 725.8 | 0.073427
|  | 725.9 | 0.062953
|  | 726 | 0.060953
|  | 726.1 | 0.058953
|  | 726.2 | 0.07
|  | 726.3 | 0.07
|  | 726.4 | 0.063713
|  | 726.5 | 0.065713
|  | 726.6 | 0.067713
|  | 726.7 | 0.073333
|  | 726.8 | 0.073333
|  | 726.9 | 0.062953
|  | 727 | 0.060953
|  | 727.1 | 0.058953
|  | 727.2 | 0.065287
|  | 727.3 | 0.063287
|  | 727.4 | 0.06038
|  | 727.5 | 0.06238
|  | 727.6 | 0.06438
|  | 727.7 | 0.061943
|  | 727.8 | 0.059938
|  | 727.9 | 0.067455
|  | 728 | 0.071466
|  | 728.1 | 0.075477
|  | 728.2 | 0.061929
|  | 728.3 | 0.059931
|  | 728.4 | 0.070002
|  | 728.5 | 0.07001
|  | 728.6 | 0.070018
|  | 728.7 | 0.06953
|  | 728.8 | 0.073543
|  | 728.9 | 0.072899
|  | 729 | 0.070897
|  | 729.1 | 0.068894
|  | 729.2 | 0.065218
|  | 729.3 | 0.063215
|  | 729.4 | 0.062871
|  | 729.5 | 0.060863
|  | 729.6 | 0.058855
|  | 729.7 | 0.06
|  | 729.8 | 0.06
|  | 729.9 | 0.06
|  | 730 | 0.06
|  | 730.1 | 0.06
|  | 730.2 | 0.055158
|  | 730.3 | 0.053155
|  | 730.4 | 0.052811
|  | 730.5 | 0.050803
|  | 730.6 | 0.048795
|  | 730.7 | 0.05488
|  | 730.8 | 0.056888
|  | 730.9 | 0.061124
|  | 731 | 0.065141
|  | 731.1 | 0.069157
|  | 731.2 | 0.063333
|  | 731.3 | 0.063333
|  | 731.4 | 0.062723
|  | 731.5 | 0.060709
|  | 731.6 | 0.058696
|  | 731.7 | 0.06
|  | 731.8 | 0.06
|  | 731.9 | 0.063996
|  | 732 | 0.066004
|  | 732.1 | 0.068012
|  | 732.2 | 0.063333
|  | 732.3 | 0.063333
|  | 732.4 | 0.058578
|  | 732.5 | 0.054554
|  | 732.6 | 0.05053
|  | 732.7 | 0.061747
|  | 732.8 | 0.063763
|  | 732.9 | 0.060773
|  | 733 | 0.062789
|  | 733.1 | 0.064805
|  | 733.2 | 0.06153
|  | 733.3 | 0.05952
|  | 733.4 | 0.06
|  | 733.5 | 0.06
|  | 733.6 | 0.06
|  | 733.7 | 0.06
|  | 733.8 | 0.06
|  | 733.9 | 0.055767
|  | 734 | 0.053754
|  | 734.1 | 0.051741
|  | 734.2 | 0.051394
|  | 734.3 | 0.04938
|  | 734.4 | 0.05
|  | 734.5 | 0.05
|  | 734.6 | 0.05
|  | 734.7 | 0.05
|  | 734.8 | 0.05
|  | 734.9 | 0.054382
|  | 735 | 0.056398
|  | 735.1 | 0.056747
|  | 735.2 | 0.058763
|  | 735.3 | 0.06078
|  | 735.4 | 0.055538
|  | 735.5 | 0.053522
|  | 735.6 | 0.063656
|  | 735.7 | 0.067688
|  | 735.8 | 0.07172
|  | 735.9 | 0.069086
|  | 736 | 0.073118
|  | 736.1 | 0.065941
|  | 736.2 | 0.059892
|  | 736.3 | 0.053844
|  | 736.4 | 0.054066
|  | 736.5 | 0.050022
|  | 736.6 | 0.05367
|  | 736.7 | 0.055692
|  | 736.8 | 0.057713
|  | 736.9 | 0.058058
|  | 737 | 0.060074
|  | 737.1 | 0.06
|  | 737.2 | 0.06
|  | 737.3 | 0.06
|  | 737.4 | 0.06
|  | 737.5 | 0.06
|  | 737.6 | 0.056135
|  | 737.7 | 0.054116
|  | 737.8 | 0.052097
|  | 737.9 | 0.061599
|  | 738 | 0.063623
|  | 738.1 | 0.080506
|  | 738.2 | 0.09263
|  | 738.3 | 0.10475
|  | 738.4 | 0.09845
|  | 738.5 | 0.10452
|  | 738.6 | 0.090337
|  | 738.7 | 0.08224
|  | 738.8 | 0.074143
|  | 738.9 | 0.074534
|  | 739 | 0.068462
|  | 739.1 | 0.065796
|  | 739.2 | 0.063772
|  | 739.3 | 0.061748
|  | 739.4 | 0.066669
|  | 739.5 | 0.066672
|  | 739.6 | 0.059026
|  | 739.7 | 0.057009
|  | 739.8 | 0.053252
|  | 739.9 | 0.047166
|  | 740 | 0.041079
|  | 740.1 | 0.046668
|  | 740.2 | 0.046671
|  | 740.3 | 0.046829
|  | 740.4 | 0.048853
|  | 740.5 | 0.050877
|  | 740.6 | 0.05
|  | 740.7 | 0.05
|  | 740.8 | 0.053617
|  | 740.9 | 0.055641
|  | 741 | 0.057665
|  | 741.1 | 0.058033
|  | 741.2 | 0.060063
|  | 741.3 | 0.06
|  | 741.4 | 0.06
|  | 741.5 | 0.06
|  | 741.6 | 0.055169
|  | 741.7 | 0.053144
|  | 741.8 | 0.052779
|  | 741.9 | 0.050749
|  | 742 | 0.048719
|  | 742.1 | 0.05498
|  | 742.2 | 0.057012
|  | 742.3 | 0.053332
|  | 742.4 | 0.05333
|  | 742.5 | 0.053327
|  | 742.6 | 0.05154
|  | 742.7 | 0.049511
|  | 742.8 | 0.05
|  | 742.9 | 0.05
|  | 743 | 0.05
|  | 743.1 | 0.05
|  | 743.2 | 0.05
|  | 743.3 | 0.054333
|  | 743.4 | 0.05636
|  | 743.5 | 0.056728
|  | 743.6 | 0.05876
|  | 743.7 | 0.060793
|  | 743.8 | 0.051016
|  | 743.9 | 0.046951
|  | 744 | 0.049776
|  | 744.1 | 0.047744
|  | 744.2 | 0.045711
|  | 744.3 | 0.047988
|  | 744.4 | 0.05002
|  | 744.5 | 0.04628
|  | 744.6 | 0.044248
|  | 744.7 | 0.042215
|  | 744.8 | 0.046667
|  | 744.9 | 0.046667
|  | 745 | 0.047215
|  | 745.1 | 0.049248
|  | 745.2 | 0.05128
|  | 745.3 | 0.05
|  | 745.4 | 0.05
|  | 745.5 | 0.05406
|  | 745.6 | 0.056098
|  | 745.7 | 0.058136
|  | 745.8 | 0.053333
|  | 745.9 | 0.053333
|  | 746 | 0.05243
|  | 746.1 | 0.050392
|  | 746.2 | 0.048354
|  | 746.3 | 0.05
|  | 746.4 | 0.05
|  | 746.5 | 0.058838
|  | 746.6 | 0.062908
|  | 746.7 | 0.056666
|  | 746.8 | 0.056661
|  | 746.9 | 0.056655
|  | 747 | 0.049524
|  | 747.1 | 0.043413
|  | 747.2 | 0.042986
|  | 747.3 | 0.040946
|  | 747.4 | 0.038905
|  | 747.5 | 0.049605
|  | 747.6 | 0.053687
|  | 747.7 | 0.046665
|  | 747.8 | 0.04666
|  | 747.9 | 0.046654
|  | 748 | 0.04834
|  | 748.1 | 0.046299
|  | 748.2 | 0.043333
|  | 748.3 | 0.043333
|  | 748.4 | 0.043333
|  | 748.5 | 0.046667
|  | 748.6 | 0.046667
|  | 748.7 | 0.047605
|  | 748.8 | 0.049646
|  | 748.9 | 0.046646
|  | 749 | 0.044605
|  | 749.1 | 0.042565
|  | 749.2 | 0.046667
|  | 749.3 | 0.046667
|  | 749.4 | 0.043334
|  | 749.5 | 0.043336
|  | 749.6 | 0.043339
|  | 749.7 | 0.041971
|  | 749.8 | 0.039927
|  | 749.9 | 0.043782
|  | 750 | 0.045823
|  | 750.1 | 0.047864
|  | 750.2 | 0.072762
|  | 750.3 | 0.085007
|  | 750.4 | 0.078653
|  | 750.5 | 0.084798
|  | 750.6 | 0.090943
|  | 750.7 | 0.072829
|  | 750.8 | 0.064649
|  | 750.9 | 0.069112
|  | 751 | 0.067063
|  | 751.1 | 0.065014
|  | 751.2 | 0.08474
|  | 751.3 | 0.092923
|  | 751.4 | 0.097779
|  | 751.5 | 0.10596
|  | 751.6 | 0.092719
|  | 751.7 | 0.086571
|  | 751.8 | 0.080423
|  | 751.9 | 0.073288
|  | 752 | 0.063068
|  | 752.1 | 0.059127
|  | 752.2 | 0.055037
|  | 752.3 | 0.050948
|  | 752.4 | 0.046844
|  | 752.5 | 0.042746
|  | 752.6 | 0.046667
|  | 752.7 | 0.046667
|  | 752.8 | 0.046667
|  | 752.9 | 0.04849
|  | 753 | 0.05054
|  | 753.1 | 0.05
|  | 753.2 | 0.05
|  | 753.3 | 0.05
|  | 753.4 | 0.05
|  | 753.5 | 0.05
|  | 753.6 | 0.054501
|  | 753.7 | 0.056551
|  | 753.8 | 0.049734
|  | 753.9 | 0.047684
|  | 754 | 0.045635
|  | 754.1 | 0.048586
|  | 754.2 | 0.046537
|  | 754.3 | 0.043334
|  | 754.4 | 0.043337
|  | 754.5 | 0.04334
|  | 754.6 | 0.046669
|  | 754.7 | 0.046672
|  | 754.8 | 0.043333
|  | 754.9 | 0.043333
|  | 755 | 0.043333
|  | 755.1 | 0.051933
|  | 755.2 | 0.053982
|  | 755.3 | 0.051038
|  | 755.4 | 0.053087
|  | 755.5 | 0.053183
|  | 755.6 | 0.051128
|  | 755.7 | 0.049073
|  | 755.8 | 0.045358
|  | 755.9 | 0.043304
|  | 756 | 0.039165
|  | 756.1 | 0.035058
|  | 756.2 | 0.030951
|  | 756.3 | 0.036667
|  | 756.4 | 0.036667
|  | 756.5 | 0.033333
|  | 756.6 | 0.033333
|  | 756.7 | 0.033333
|  | 756.8 | 0.036669
|  | 756.9 | 0.036672
|  | 757 | 0.041944
|  | 757.1 | 0.046051
|  | 757.2 | 0.043333
|  | 757.3 | 0.043333
|  | 757.4 | 0.043333
|  | 757.5 | 0.046667
|  | 757.6 | 0.046667
|  | 757.7 | 0.03963
|  | 757.8 | 0.037572
|  | 757.9 | 0.035514
|  | 758 | 0.038457
|  | 758.1 | 0.036399
|  | 758.2 | 0.033333
|  | 758.3 | 0.033333
|  | 758.4 | 0.033333
|  | 758.5 | 0.03149
|  | 758.6 | 0.029427
|  | 758.7 | 0.038592
|  | 758.8 | 0.042719
|  | 758.9 | 0.046914
|  | 759 | 0.053086
|  | 759.1 | 0.059259
|  | 759.2 | 0.048745
|  | 759.3 | 0.046687
|  | 759.4 | 0.042199
|  | 759.5 | 0.036012
|  | 759.6 | 0.029825
|  | 759.7 | 0.036669
|  | 759.8 | 0.036672
|  | 759.9 | 0.037354
|  | 760 | 0.039418
|  | 760.1 | 0.041481
|  | 760.2 | 0.04
|  | 760.3 | 0.04
|  | 760.4 | 0.035671
|  | 760.5 | 0.03361
|  | 760.6 | 0.033209
|  | 760.7 | 0.031143
|  | 760.8 | 0.029077
|  | 760.9 | 0.034647
|  | 761 | 0.03671
|  | 761.1 | 0.033334
|  | 761.2 | 0.033337
|  | 761.3 | 0.03334
|  | 761.4 | 0.036667
|  | 761.5 | 0.036667
|  | 761.6 | 0.041529
|  | 761.7 | 0.045661
|  | 761.8 | 0.049793
|  | 761.9 | 0.053926
|  | 762 | 0.058058
|  | 762.1 | 0.048905
|  | 762.2 | 0.046839
|  | 762.3 | 0.042652
|  | 762.4 | 0.036453
|  | 762.5 | 0.030255
|  | 762.6 | 0.031908
|  | 762.7 | 0.029842
|  | 762.8 | 0.037782
|  | 762.9 | 0.041915
|  | 763 | 0.046047
|  | 763.1 | 0.062114
|  | 763.2 | 0.072445
|  | 763.3 | 0.13614
|  | 763.4 | 0.17339
|  | 763.5 | 0.16705
|  | 763.6 | 0.18979
|  | 763.7 | 0.21254
|  | 763.8 | 0.15722
|  | 763.9 | 0.13036
|  | 764 | 0.10802
|  | 764.1 | 0.079027
|  | 764.2 | 0.050036
|  | 764.3 | 0.053814
|  | 764.4 | 0.041397
|  | 764.5 | 0.05073
|  | 764.6 | 0.052796
|  | 764.7 | 0.054862
|  | 764.8 | 0.041387
|  | 764.9 | 0.03931
|  | 765 | 0.045579
|  | 765.1 | 0.043502
|  | 765.2 | 0.036667
|  | 765.3 | 0.036667
|  | 765.4 | 0.036667
|  | 765.5 | 0.038561
|  | 765.6 | 0.036495
|  | 765.7 | 0.041173
|  | 765.8 | 0.045314
|  | 765.9 | 0.049455
|  | 766 | 0.038195
|  | 766.1 | 0.03612
|  | 766.2 | 0.039046
|  | 766.3 | 0.036971
|  | 766.4 | 0.029896
|  | 766.5 | 0.027822
|  | 766.6 | 0.025747
|  | 766.7 | 0.028672
|  | 766.8 | 0.026598
|  | 766.9 | 0.027144
|  | 767 | 0.029219
|  | 767.1 | 0.031293
|  | 767.2 | 0.040069
|  | 767.3 | 0.044219
|  | 767.4 | 0.036667
|  | 767.5 | 0.036667
|  | 767.6 | 0.036667
|  | 767.7 | 0.037925
|  | 767.8 | 0.035851
|  | 767.9 | 0.037901
|  | 768 | 0.039982
|  | 768.1 | 0.036278
|  | 768.2 | 0.034198
|  | 768.3 | 0.032117
|  | 768.4 | 0.031715
|  | 768.5 | 0.02964
|  | 768.6 | 0.025891
|  | 768.7 | 0.023813
|  | 768.8 | 0.021736
|  | 768.9 | 0.032014
|  | 769 | 0.034097
|  | 769.1 | 0.031174
|  | 769.2 | 0.033257
|  | 769.3 | 0.036667
|  | 769.4 | 0.03667
|  | 769.5 | 0.036673
|  | 769.6 | 0.033336
|  | 769.7 | 0.033338
|  | 769.8 | 0.028542
|  | 769.9 | 0.024375
|  | 770 | 0.020208
|  | 770.1 | 0.031972
|  | 770.2 | 0.034056
|  | 770.3 | 0.040056
|  | 770.4 | 0.046295
|  | 770.5 | 0.036083
|  | 770.6 | 0.031917
|  | 770.7 | 0.02775
|  | 770.8 | 0.025375
|  | 770.9 | 0.019125
|  | 771 | 0.024042
|  | 771.1 | 0.026125
|  | 771.2 | 0.028208
|  | 771.3 | 0.033917
|  | 771.4 | 0.038083
|  | 771.5 | 0.037792
|  | 771.6 | 0.039875
|  | 771.7 | 0.04
|  | 771.8 | 0.04
|  | 771.9 | 0.04
|  | 772 | 0.04
|  | 772.1 | 0.04
|  | 772.2 | 0.06425
|  | 772.3 | 0.07675
|  | 772.4 | 0.08925
|  | 772.5 | 0.097126
|  | 772.6 | 0.11175
|  | 772.7 | 0.094381
|  | 772.8 | 0.0902
|  | 772.9 | 0.081771
|  | 773 | 0.071354
|  | 773.1 | 0.060938
|  | 773.2 | 0.053709
|  | 773.3 | 0.045358
|  | 773.4 | 0.0411
|  | 773.5 | 0.034838
|  | 773.6 | 0.028577
|  | 773.7 | 0.031321
|  | 773.8 | 0.029232
|  | 773.9 | 0.025483
|  | 774 | 0.023394
|  | 774.1 | 0.034056
|  | 774.2 | 0.038226
|  | 774.3 | 0.042395
|  | 774.4 | 0.03
|  | 774.5 | 0.03
|  | 774.6 | 0.031721
|  | 774.7 | 0.027552
|  | 774.8 | 0.023382
|  | 774.9 | 0.023336
|  | 775 | 0.023339
|  | 775.1 | 0.026667
|  | 775.2 | 0.026667
|  | 775.3 | 0.023333
|  | 775.4 | 0.023333
|  | 775.5 | 0.023333
|  | 775.6 | 0.021618
|  | 775.7 | 0.019526
|  | 775.8 | 0.024233
|  | 775.9 | 0.026325
|  | 776 | 0.030167
|  | 776.1 | 0.034351
|  | 776.2 | 0.038536
|  | 776.3 | 0.038026
|  | 776.4 | 0.040119
|  | 776.5 | 0.04
|  | 776.6 | 0.04
|  | 776.7 | 0.04
|  | 776.8 | 0.029674
|  | 776.9 | 0.025484
|  | 777 | 0.028982
|  | 777.1 | 0.02689
|  | 777.2 | 0.023333
|  | 777.3 | 0.023333
|  | 777.4 | 0.023333
|  | 777.5 | 0.021837
|  | 777.6 | 0.019739
|  | 777.7 | 0.02
|  | 777.8 | 0.02
|  | 777.9 | 0.02
|  | 778 | 0.02
|  | 778.1 | 0.02
|  | 778.2 | 0.02
|  | 778.3 | 0.02
|  | 778.4 | 0.027381
|  | 778.5 | 0.031583
|  | 778.6 | 0.035784
|  | 778.7 | 0.031646
|  | 778.8 | 0.033738
|  | 778.9 | 0.032487
|  | 779 | 0.030389
|  | 779.1 | 0.026625
|  | 779.2 | 0.024524
|  | 779.3 | 0.022423
|  | 779.4 | 0.026669
|  | 779.5 | 0.026671
|  | 779.6 | 0.023334
|  | 779.7 | 0.023337
|  | 779.8 | 0.02334
|  | 779.9 | 0.026667
|  | 780 | 0.026667
|  | 780.1 | 0.027696
|  | 780.2 | 0.029797
|  | 780.3 | 0.03
|  | 780.4 | 0.03
|  | 780.5 | 0.03
|  | 780.6 | 0.020266
|  | 780.7 | 0.016064
|  | 780.8 | 0.023333
|  | 780.9 | 0.023333
|  | 781 | 0.023333
|  | 781.1 | 0.027409
|  | 781.2 | 0.031611
|  | 781.3 | 0.03
|  | 781.4 | 0.03
|  | 781.5 | 0.02621
|  | 781.6 | 0.024103
|  | 781.7 | 0.021997
|  | 781.8 | 0.026667
|  | 781.9 | 0.026667
|  | 782 | 0.027643
|  | 782.1 | 0.02975
|  | 782.2 | 0.03
|  | 782.3 | 0.03
|  | 782.4 | 0.03
|  | 782.5 | 0.025175
|  | 782.6 | 0.023074
|  | 782.7 | 0.022631
|  | 782.8 | 0.020524
|  | 782.9 | 0.018418
|  | 783 | 0.02
|  | 783.1 | 0.02
|  | 783.2 | 0.02
|  | 783.3 | 0.02
|  | 783.4 | 0.023776
|  | 783.5 | 0.025886
|  | 783.6 | 0.027996
|  | 783.7 | 0.033544
|  | 783.8 | 0.037764
|  | 783.9 | 0.029015
|  | 784 | 0.026906
|  | 784.1 | 0.029803
|  | 784.2 | 0.027693
|  | 784.3 | 0.025584
|  | 784.4 | 0.028186
|  | 784.5 | 0.030295
|  | 784.6 | 0.03
|  | 784.7 | 0.03
|  | 784.8 | 0.03
|  | 784.9 | 0.019198
|  | 785 | 0.014979
|  | 785.1 | 0.023333
|  | 785.2 | 0.023333
|  | 785.3 | 0.020506
|  | 785.4 | 0.022616
|  | 785.5 | 0.024726
|  | 785.6 | 0.021498
|  | 785.7 | 0.019388
|  | 785.8 | 0.02
|  | 785.9 | 0.02
|  | 786 | 0.023615
|  | 786.1 | 0.025728
|  | 786.2 | 0.02784
|  | 786.3 | 0.038197
|  | 786.4 | 0.044541
|  | 786.5 | 0.028312
|  | 786.6 | 0.024093
|  | 786.7 | 0.026455
|  | 786.8 | 0.020105
|  | 786.9 | 0.013755
|  | 787 | 0.018075
|  | 787.1 | 0.020188
|  | 787.2 | 0.023974
|  | 787.3 | 0.026087
|  | 787.4 | 0.0282
|  | 787.5 | 0.033961
|  | 787.6 | 0.038189
|  | 787.7 | 0.03788
|  | 787.8 | 0.039995
|  | 787.9 | 0.04
|  | 788 | 0.04
|  | 788.1 | 0.04
|  | 788.2 | 0.01954
|  | 788.3 | 0.011089
|  | 788.4 | 0.013616
|  | 788.5 | 0.0072599
|  | 788.6 | 0.0069209
|  | 788.7 | 0.0090395
|  | 788.8 | 0.011158
|  | 788.9 | 0.01
|  | 789 | 0.01
|  | 789.1 | 0.01
|  | 789.2 | 0.01
|  | 789.3 | 0.01
|  | 789.4 | 0.01
|  | 789.5 | 0.01
|  | 789.6 | 0.01
|  | 789.7 | 0.01
|  | 789.8 | 0.017996
|  | 789.9 | 0.022239
|  | 790 | 0.026483
|  | 790.1 | 0.027393
|  | 790.2 | 0.031636
|  | 790.3 | 0.020791
|  | 790.4 | 0.016554
|  | 790.5 | 0.019492
|  | 790.6 | 0.017373
|  | 790.7 | 0.015254
|  | 790.8 | 0.013333
|  | 790.9 | 0.013333
|  | 791 | 0.012221
|  | 791.1 | 0.010097
|  | 791.2 | 0.017375
|  | 791.3 | 0.021624
|  | 791.4 | 0.025873
|  | 791.5 | 0.026766
|  | 791.6 | 0.031003
|  | 791.7 | 0.025705
|  | 791.8 | 0.023583
|  | 791.9 | 0.01959
|  | 792 | 0.015343
|  | 792.1 | 0.011097
|  | 792.2 | 0.021575
|  | 792.3 | 0.023693
|  | 792.4 | 0.024993
|  | 792.5 | 0.029248
|  | 792.6 | 0.023178
|  | 792.7 | 0.018928
|  | 792.8 | 0.014679
|  | 792.9 | 0.018546
|  | 793 | 0.016418
|  | 793.1 | 0.013333
|  | 793.2 | 0.013333
|  | 793.3 | 0.013333
|  | 793.4 | 0.011248
|  | 793.5 | 0.0091206
|  | 793.6 | 0.014674
|  | 793.7 | 0.016801
|  | 793.8 | 0.021191
|  | 793.9 | 0.025447
|  | 794 | 0.029702
|  | 794.1 | 0.023333
|  | 794.2 | 0.023333
|  | 794.3 | 0.017532
|  | 794.4 | 0.013277
|  | 794.5 | 0.016667
|  | 794.6 | 0.016667
|  | 794.7 | 0.016667
|  | 794.8 | 0.049773
|  | 794.9 | 0.064667
|  | 795 | 0.071227
|  | 795.1 | 0.086121
|  | 795.2 | 0.064752
|  | 795.3 | 0.054113
|  | 795.4 | 0.043475
|  | 795.5 | 0.04444
|  | 795.6 | 0.033769
|  | 795.7 | 0.024588
|  | 795.8 | 0.020318
|  | 795.9 | 0.022738
|  | 796 | 0.018482
|  | 796.1 | 0.014227
|  | 796.2 | 0.03343
|  | 796.3 | 0.041956
|  | 796.4 | 0.029996
|  | 796.5 | 0.029987
|  | 796.6 | 0.032892
|  | 796.7 | 0.028622
|  | 796.8 | 0.024351
|  | 796.9 | 0.018383
|  | 797 | 0.016246
|  | 797.1 | 0.019117
|  | 797.2 | 0.01698
|  | 797.3 | 0.020299
|  | 797.4 | 0.024573
|  | 797.5 | 0.028846
|  | 797.6 | 0.028226
|  | 797.7 | 0.030363
|  | 797.8 | 0.03
|  | 797.9 | 0.03
|  | 798 | 0.023149
|  | 798.1 | 0.018887
|  | 798.2 | 0.014626
|  | 798.3 | 0.023333
|  | 798.4 | 0.023333
|  | 798.5 | 0.016667
|  | 798.6 | 0.016667
|  | 798.7 | 0.023333
|  | 798.8 | 0.023333
|  | 798.9 | 0.023333
|  | 799 | 0.016667
|  | 799.1 | 0.016667
|  | 799.2 | 0.023333
|  | 799.3 | 0.023333
|  | 799.4 | 0.023333
|  | 799.5 | 0.033077
|  | 799.6 | 0.039487
|  | 799.7 | 0.038088
|  | 799.8 | 0.040231
|  | 799.9 | 0.04
|  | 800 | 0.04
|  | 800.1 | 0.04
|  | 800.2 | 0.023654
|  | 800.3 | 0.017244
|  | 800.4 | 0.020531
|  | 800.5 | 0.016243
|  | 800.6 | 0.021356
|  | 800.7 | 0.025639
|  | 800.8 | 0.029921
|  | 800.9 | 0.055897
|  | 801 | 0.068718
|  | 801.1 | 0.051386
|  | 801.2 | 0.053542
|  | 801.3 | 0.050709
|  | 801.4 | 0.042149
|  | 801.5 | 0.03359
|  | 801.6 | 0.072496
|  | 801.7 | 0.085335
|  | 801.8 | 0.097654
|  | 801.9 | 0.11482
|  | 802 | 0.086712
|  | 802.1 | 0.075992
|  | 802.2 | 0.065272
|  | 802.3 | 0.037069
|  | 802.4 | 0.01351
|  | 802.5 | 0.022546
|  | 802.6 | 0.016109
|  | 802.7 | 0.02133
|  | 802.8 | 0.025622
|  | 802.9 | 0.029914
|  | 803 | 0.02877
|  | 803.1 | 0.030916
|  | 803.2 | 0.020544
|  | 803.3 | 0.016252
|  | 803.4 | 0.023333
|  | 803.5 | 0.023333
|  | 803.6 | 0.023333
|  | 803.7 | 0.027582
|  | 803.8 | 0.031874
|  | 803.9 | 0.020501
|  | 804 | 0.016209
|  | 804.1 | 0.019292
|  | 804.2 | 0.017146
|  | 804.3 | 0.015
|  | 804.4 | 0.018813
|  | 804.5 | 0.020959
|  | 804.6 | 0.02
|  | 804.7 | 0.02
|  | 804.8 | 0.024079
|  | 804.9 | 0.026231
|  | 805 | 0.02671
|  | 805.1 | 0.028856
|  | 805.2 | 0.031001
|  | 805.3 | 0.020353
|  | 805.4 | 0.016055
|  | 805.5 | 0.019213
|  | 805.6 | 0.017067
|  | 805.7 | 0.020186
|  | 805.8 | 0.024487
|  | 805.9 | 0.028789
|  | 806 | 0.018455
|  | 806.1 | 0.0163
|  | 806.2 | 0.02751
|  | 806.3 | 0.029652
|  | 806.4 | 0.016516
|  | 806.5 | 0.014361
|  | 806.6 | 0.012205
|  | 806.7 | 0.021736
|  | 806.8 | 0.019596
|  | 806.9 | 0.014247
|  | 807 | 0.016405
|  | 807.1 | 0.012687
|  | 807.2 | 0.0062213
|  | 807.3 | -0.00024425
|  | 807.4 | 0.0016523
|  | 807.5 | -0.00050287
|  | 807.6 | 0.0043247
|  | 807.7 | 0.0064799

## radiometer factor calibration
wavelength | raw blanked output (averaged) | new wavelength | forecast over 1, 1/2 nm increments | new wavelength | forecast over 1, 0.2 nm increments | new wavelength | forecast over 1, 0.5 nm increments | new wavelength | forecast over 1, 1 nm increments | 1 | RLE is shown in column k
195.4234 | 0 | 199.8 | 0 | 199.8 | 0 | 199.5 | 0 | 200 | 0.00036314 | 0.00036314
195.7963 | 0 | 199.9 | 0.00028704 | 200 | 0.00039945 | 200 | 0.00032429 | 201 | 0.00075267 | 0.00075267
196.1692 | 0 | 200 | 0.00039945 | 200.2 | 0.00057342 | 200.5 | 0.00076513 | 202 | 0.00068699 | 0.00068699
196.5419 | 0 | 200.1 | 0.00051187 | 200.4 | 0.00069187 | 201 | 0.00076874 | 203 | 0.00062421 | 0.00062421
196.9145 | 0 | 200.2 | 0.00062429 | 200.6 | 0.00083439 | 201.5 | 0.00072413 | 204 | 0.00057621 | 0.00057621
197.2871 | 0 | 200.3 | 0.00058411 | 200.8 | 0.00078701 | 202 | 0.00068469 | 205 | 0.00053669 | 0.00053669
197.6596 | 0 | 200.4 | 0.00069187 | 201 | 0.00076873 | 202.5 | 0.00065217 | 206 | 0.00050618 | 0.00050618
198.032 | 0 | 200.5 | 0.00079963 | 201.2 | 0.00075049 | 203 | 0.00062357 | 207 | 0.00048188 | 0.00048188
198.4043 | 0 | 200.6 | 0.00090738 | 201.4 | 0.00073291 | 203.5 | 0.0005969 | 208 | 0.00045999 | 0.00045999
198.7765 | 0 | 200.7 | 0.00079615 | 201.6 | 0.00071534 | 204 | 0.00057563 | 209 | 0.00043995 | 0.00043995
199.1486 | 0 | 200.8 | 0.00078701 | 201.8 | 0.00069948 | 204.5 | 0.00055609 | 210 | 0.0004115 | 0.0004115
199.5207 | 0 | 200.9 | 0.00077788 | 202 | 0.00068411 | 205 | 0.00053567 | 211 | 0.00037684 | 0.00037684
199.8927 | 0 | 201 | 0.00076875 | 202.2 | 0.00067047 | 205.5 | 0.00051831 | 212 | 0.00034194 | 0.00034194
200.2645 | 0.00083622 | 201.1 | 0.00075955 | 202.4 | 0.00065791 | 206 | 0.00050657 | 213 | 0.00030763 | 0.00030763
200.6363 | 0.00080132 | 201.2 | 0.00075049 | 202.6 | 0.00064627 | 206.5 | 0.00049366 | 214 | 0.00027735 | 0.00027735
201.008 | 0.00076834 | 201.3 | 0.00074143 | 202.8 | 0.0006346 | 207 | 0.00048116 | 215 | 0.00025107 | 0.00025107
201.3796 | 0.00073398 | 201.4 | 0.00073309 | 203 | 0.00062352 | 207.5 | 0.00047083 | 216 | 0.0002287 | 0.0002287
201.7512 | 0.00070239 | 201.5 | 0.00072422 | 203.2 | 0.00061258 | 208 | 0.00045981 | 217 | 0.0002113 | 0.0002113
202.1226 | 0.00067511 | 201.6 | 0.00071534 | 203.4 | 0.00060188 | 208.5 | 0.00044932 | 218 | 0.00019875 | 0.00019875
202.494 | 0.00065152 | 201.7 | 0.00070647 | 203.6 | 0.00059182 | 209 | 0.00044107 | 219 | 0.00018969 | 0.00018969
202.8652 | 0.00063105 | 201.8 | 0.00069995 | 203.8 | 0.00058329 | 209.5 | 0.00042946 | 220 | 0.00018126 | 0.00018126
203.2364 | 0.0006104 | 201.9 | 0.00069203 | 204 | 0.00057541 | 210 | 0.00041124 | 221 | 0.0001719 | 0.0001719
203.6075 | 0.0005908 | 202 | 0.00068411 | 204.2 | 0.00056817 | 210.5 | 0.00039378 | 222 | 0.00016117 | 0.00016117
203.9785 | 0.00057569 | 202.1 | 0.00067618 | 204.4 | 0.00056008 | 211 | 0.00037717 | 223 | 0.00014812 | 0.00014812
204.3495 | 0.00056296 | 202.2 | 0.00067104 | 204.6 | 0.00055208 | 211.5 | 0.00035957 | 224 | 0.00013406 | 0.00013406
204.7203 | 0.00054683 | 202.3 | 0.00066419 | 204.8 | 0.00054389 | 212 | 0.00034194 | 225 | 0.00011941 | 0.00011941
205.0911 | 0.0005318 | 202.4 | 0.00065734 | 205 | 0.00053555 | 212.5 | 0.00032431 | 226 | 0.00010461 | 0.00010461
205.4617 | 0.00051715 | 202.5 | 0.0006522 | 205.2 | 0.00052756 | 213 | 0.00030723 | 227 | 9.1111e-05 | 9.1111e-05
205.8323 | 0.00051023 | 202.6 | 0.00064627 | 205.4 | 0.00052058 | 213.5 | 0.00029136 | 228 | 7.9268e-05 | 7.9268e-05
206.2028 | 0.00050301 | 202.7 | 0.00064033 | 205.6 | 0.0005157 | 214 | 0.00027699 | 229 | 6.9196e-05 | 6.9196e-05
206.5732 | 0.00049058 | 202.8 | 0.0006344 | 205.8 | 0.0005105 | 214.5 | 0.00026369 | 230 | 6.0962e-05 | 6.0962e-05
206.9436 | 0.00048248 | 202.9 | 0.00062906 | 206 | 0.00050693 | 215 | 0.00025071 | 231 | 5.4336e-05 | 5.4336e-05
207.3138 | 0.00047417 | 203 | 0.00062352 | 206.2 | 0.00050228 | 215.5 | 0.00023881 | 232 | 4.8904e-05 | 4.8904e-05
207.684 | 0.0004677 | 203.1 | 0.00061799 | 206.4 | 0.00049604 | 216 | 0.00022837 | 233 | 4.4466e-05 | 4.4466e-05
208.054 | 0.00045809 | 203.2 | 0.00061245 | 206.6 | 0.00049106 | 216.5 | 0.00021891 | 234 | 4.0772e-05 | 4.0772e-05
208.424 | 0.00045067 | 203.3 | 0.0006073 | 206.8 | 0.00048574 | 217 | 0.0002109 | 235 | 3.7666e-05 | 3.7666e-05
208.7939 | 0.00044373 | 203.4 | 0.00060188 | 207 | 0.00048102 | 217.5 | 0.00020408 | 236 | 3.4983e-05 | 3.4983e-05
209.1637 | 0.00043947 | 203.5 | 0.00059646 | 207.2 | 0.00047673 | 218 | 0.00019846 | 237 | 3.2681e-05 | 3.2681e-05
209.5335 | 0.0004298 | 203.6 | 0.00059103 | 207.4 | 0.00047288 | 218.5 | 0.0001937 | 238 | 3.072e-05 | 3.072e-05
209.9031 | 0.00041375 | 203.7 | 0.00058797 | 207.6 | 0.00046882 | 219 | 0.00018966 | 239 | 2.9055e-05 | 2.9055e-05
210.2727 | 0.00040176 | 203.8 | 0.00058329 | 207.8 | 0.00046413 | 219.5 | 0.00018571 | 240 | 2.7621e-05 | 2.7621e-05
210.6421 | 0.0003888 | 203.9 | 0.00057862 | 208 | 0.00045984 | 220 | 0.00018137 | 241 | 2.6408e-05 | 2.6408e-05
211.0115 | 0.00037653 | 204 | 0.00057568 | 208.2 | 0.00045546 | 220.5 | 0.0001767 | 242 | 2.5382e-05 | 2.5382e-05
211.3808 | 0.00036525 | 204.1 | 0.00057193 | 208.4 | 0.00045113 | 221 | 0.00017195 | 243 | 2.4488e-05 | 2.4488e-05
211.75 | 0.00034957 | 204.2 | 0.00056817 | 208.6 | 0.00044742 | 221.5 | 0.00016704 | 244 | 2.3705e-05 | 2.3705e-05
212.1192 | 0.00033853 | 204.3 | 0.00056442 | 208.8 | 0.00044434 | 222 | 0.00016144 | 245 | 2.3022e-05 | 2.3022e-05
212.4882 | 0.00032394 | 204.4 | 0.00055986 | 209 | 0.00044151 | 222.5 | 0.00015504 | 246 | 2.2434e-05 | 2.2434e-05
212.8572 | 0.00031218 | 204.5 | 0.00055597 | 209.2 | 0.00043736 | 223 | 0.00014814 | 247 | 2.1911e-05 | 2.1911e-05
213.226 | 0.00029958 | 204.6 | 0.00055208 | 209.4 | 0.00043322 | 223.5 | 0.00014118 | 248 | 2.1447e-05 | 2.1447e-05
213.5948 | 0.00028804 | 204.7 | 0.00054819 | 209.6 | 0.00042619 | 224 | 0.00013405 | 249 | 2.103e-05 | 2.103e-05
213.9635 | 0.0002784 | 204.8 | 0.00054384 | 209.8 | 0.0004184 | 224.5 | 0.00012694 | 250 | 2.0679e-05 | 2.0679e-05
214.3322 | 0.00026727 | 204.9 | 0.00053964 | 210 | 0.00041133 | 225 | 0.00011944 | 251 | 2.0386e-05 | 2.0386e-05
214.7007 | 0.00025899 | 205 | 0.00053544 | 210.2 | 0.000404 | 225.5 | 0.00011185 | 252 | 2.0134e-05 | 2.0134e-05
215.0691 | 0.00024861 | 205.1 | 0.00053157 | 210.4 | 0.00039714 | 226 | 0.00010445 | 253 | 1.9937e-05 | 1.9937e-05
215.4375 | 0.0002395 | 205.2 | 0.00052756 | 210.6 | 0.0003904 | 226.5 | 9.7519e-05 | 254 | 1.9764e-05 | 1.9764e-05
215.8058 | 0.0002327 | 205.3 | 0.00052356 | 210.8 | 0.00038364 | 227 | 9.099e-05 | 255 | 1.9642e-05 | 1.9642e-05
216.1739 | 0.00022443 | 205.4 | 0.00051956 | 211 | 0.00037702 | 227.5 | 8.4823e-05 | 256 | 1.9598e-05 | 1.9598e-05
216.542 | 0.0002181 | 205.5 | 0.00051861 | 211.2 | 0.00037085 | 228 | 7.9136e-05 | 257 | 1.9597e-05 | 1.9597e-05
216.9101 | 0.00021216 | 205.6 | 0.0005157 | 211.4 | 0.00036339 | 228.5 | 7.3845e-05 | 258 | 1.9625e-05 | 1.9625e-05
217.278 | 0.00020663 | 205.7 | 0.00051279 | 211.6 | 0.00035578 | 229 | 6.9066e-05 | 259 | 1.9668e-05 | 1.9668e-05
217.6459 | 0.00020225 | 205.8 | 0.00050988 | 211.8 | 0.00034905 | 229.5 | 6.4676e-05 | 260 | 1.9759e-05 | 1.9759e-05
218.0136 | 0.0001981 | 205.9 | 0.00050884 | 212 | 0.00034207 | 230 | 6.083e-05 | 261 | 1.9902e-05 | 1.9902e-05
218.3813 | 0.00019478 | 206 | 0.00050693 | 212.2 | 0.00033469 | 230.5 | 5.7379e-05 | 262 | 2.0071e-05 | 2.0071e-05
218.7489 | 0.00019144 | 206.1 | 0.00050502 | 212.4 | 0.00032771 | 231 | 5.4234e-05 | 263 | 2.0263e-05 | 2.0263e-05
219.1164 | 0.00018881 | 206.2 | 0.00050311 | 212.6 | 0.00032089 | 231.5 | 5.1396e-05 | 264 | 2.043e-05 | 2.043e-05
219.4838 | 0.00018603 | 206.3 | 0.0004987 | 212.8 | 0.00031385 | 232 | 4.8819e-05 | 265 | 2.0594e-05 | 2.0594e-05
219.8512 | 0.00018265 | 206.4 | 0.00049604 | 213 | 0.00030718 | 232.5 | 4.6496e-05 | 266 | 2.0768e-05 | 2.0768e-05
220.2184 | 0.00017949 | 206.5 | 0.00049339 | 213.2 | 0.00030066 | 233 | 4.4408e-05 | 267 | 2.0924e-05 | 2.0924e-05
220.5856 | 0.00017587 | 206.6 | 0.00049128 | 213.4 | 0.00029424 | 233.5 | 4.2493e-05 | 268 | 2.1062e-05 | 2.1062e-05
220.9527 | 0.00017226 | 206.7 | 0.00048851 | 213.6 | 0.00028838 | 234 | 4.0714e-05 | 269 | 2.1154e-05 | 2.1154e-05
221.3197 | 0.00016912 | 206.8 | 0.00048574 | 213.8 | 0.00028278 | 234.5 | 3.9109e-05 | 270 | 2.1238e-05 | 2.1238e-05
221.6866 | 0.00016503 | 206.9 | 0.00048296 | 214 | 0.00027695 | 235 | 3.7642e-05 | 271 | 2.1319e-05 | 2.1319e-05
222.0534 | 0.00016106 | 207 | 0.00048116 | 214.2 | 0.00027124 | 235.5 | 3.6247e-05 | 272 | 2.1362e-05 | 2.1362e-05
222.4201 | 0.00015618 | 207.1 | 0.00047894 | 214.4 | 0.00026622 | 236 | 3.4943e-05 | 273 | 2.1372e-05 | 2.1372e-05
222.7868 | 0.00015107 | 207.2 | 0.00047673 | 214.6 | 0.00026117 | 236.5 | 3.3758e-05 | 274 | 2.1308e-05 | 2.1308e-05
223.1534 | 0.00014603 | 207.3 | 0.00047451 | 214.8 | 0.00025585 | 237 | 3.2656e-05 | 275 | 2.1213e-05 | 2.1213e-05
223.5198 | 0.00014097 | 207.4 | 0.00047306 | 215 | 0.00025072 | 237.5 | 3.1628e-05 | 276 | 2.1104e-05 | 2.1104e-05
223.8862 | 0.0001356 | 207.5 | 0.00047107 | 215.2 | 0.00024557 | 238 | 3.0693e-05 | 277 | 2.0946e-05 | 2.0946e-05
224.2526 | 0.00013049 | 207.6 | 0.00046907 | 215.4 | 0.00024071 | 238.5 | 2.9837e-05 | 278 | 2.0736e-05 | 2.0736e-05
224.6188 | 0.00012529 | 207.7 | 0.00046631 | 215.6 | 0.00023676 | 239 | 2.9038e-05 | 279 | 2.045e-05 | 2.045e-05
224.9849 | 0.0001197 | 207.8 | 0.00046413 | 215.8 | 0.00023244 | 239.5 | 2.8288e-05 | 280 | 2.0112e-05 | 2.0112e-05
225.351 | 0.00011405 | 207.9 | 0.00046196 | 216 | 0.00022824 | 240 | 2.7599e-05 | 281 | 1.9709e-05 | 1.9709e-05
225.717 | 0.00010861 | 208 | 0.00045979 | 216.2 | 0.00022444 | 240.5 | 2.6974e-05 | 282 | 1.9226e-05 | 1.9226e-05
226.0829 | 0.00010309 | 208.1 | 0.00045776 | 216.4 | 0.00022059 | 241 | 2.639e-05 | 283 | 1.8706e-05 | 1.8706e-05
226.4487 | 9.8163e-05 | 208.2 | 0.00045546 | 216.6 | 0.00021716 | 241.5 | 2.5861e-05 | 284 | 1.8125e-05 | 1.8125e-05
226.8144 | 9.3355e-05 | 208.3 | 0.00045316 | 216.8 | 0.00021393 | 242 | 2.5373e-05 | 285 | 1.7508e-05 | 1.7508e-05
227.18 | 8.8631e-05 | 208.4 | 0.00045086 | 217 | 0.00021083 | 242.5 | 2.4911e-05 | 286 | 1.6878e-05 | 1.6878e-05
227.5456 | 8.4209e-05 | 208.5 | 0.00044936 | 217.2 | 0.00020794 | 243 | 2.4479e-05 | 287 | 1.622e-05 | 1.622e-05
227.9111 | 8.0112e-05 | 208.6 | 0.00044742 | 217.4 | 0.00020537 | 243.5 | 2.4075e-05 | 288 | 1.5552e-05 | 1.5552e-05
228.2765 | 7.6047e-05 | 208.7 | 0.00044548 | 217.6 | 0.0002028 | 244 | 2.3698e-05 | 289 | 1.4894e-05 | 1.4894e-05
228.6418 | 7.2371e-05 | 208.8 | 0.00044453 | 217.8 | 0.00020053 | 244.5 | 2.3343e-05 | 290 | 1.4259e-05 | 1.4259e-05
229.007 | 6.8962e-05 | 208.9 | 0.00044302 | 218 | 0.00019836 | 245 | 2.3013e-05 | 291 | 1.3652e-05 | 1.3652e-05
229.3721 | 6.5691e-05 | 209 | 0.00044151 | 218.2 | 0.00019648 | 245.5 | 2.2709e-05 | 292 | 1.3078e-05 | 1.3078e-05
229.7371 | 6.2698e-05 | 209.1 | 0.00043999 | 218.4 | 0.00019459 | 246 | 2.2431e-05 | 293 | 1.2534e-05 | 1.2534e-05
230.1021 | 6.008e-05 | 209.2 | 0.00043699 | 218.6 | 0.0001928 | 246.5 | 2.2163e-05 | 294 | 1.2021e-05 | 1.2021e-05
230.467 | 5.7564e-05 | 209.3 | 0.0004351 | 218.8 | 0.0001912 | 247 | 2.1906e-05 | 295 | 1.1547e-05 | 1.1547e-05
230.8318 | 5.5177e-05 | 209.4 | 0.00043322 | 219 | 0.00018964 | 247.5 | 2.1666e-05 | 296 | 1.1116e-05 | 1.1116e-05
231.1965 | 5.3056e-05 | 209.5 | 0.00043133 | 219.2 | 0.00018813 | 248 | 2.1448e-05 | 297 | 1.0731e-05 | 1.0731e-05
231.5611 | 5.1029e-05 | 209.6 | 0.00042536 | 219.4 | 0.00018659 | 248.5 | 2.1229e-05 | 298 | 1.0373e-05 | 1.0373e-05
231.9256 | 4.9147e-05 | 209.7 | 0.00042188 | 219.6 | 0.00018485 | 249 | 2.1023e-05 | 299 | 1.0042e-05 | 1.0042e-05
232.2901 | 4.7413e-05 | 209.8 | 0.0004184 | 219.8 | 0.00018316 | 249.5 | 2.0839e-05 | 300 | 9.7416e-06 | 9.7416e-06
232.6544 | 4.5785e-05 | 209.9 | 0.00041492 | 220 | 0.0001814 | 250 | 2.0676e-05 | 301 | 9.471e-06 | 9.471e-06
233.0187 | 4.4313e-05 | 210 | 0.00041143 | 220.2 | 0.00017957 | 250.5 | 2.0524e-05 | 302 | 9.2378e-06 | 9.2378e-06
233.3829 | 4.29e-05 | 210.1 | 0.00040764 | 220.4 | 0.00017766 | 251 | 2.0384e-05 | 303 | 9.0158e-06 | 9.0158e-06
233.7471 | 4.159e-05 | 210.2 | 0.00040384 | 220.6 | 0.00017574 | 251.5 | 2.0249e-05 | 304 | 8.801e-06 | 8.801e-06
234.1111 | 4.0307e-05 | 210.3 | 0.00040051 | 220.8 | 0.00017376 | 252 | 2.0127e-05 | 305 | 8.6111e-06 | 8.6111e-06
234.475 | 3.916e-05 | 210.4 | 0.00039714 | 221 | 0.00017194 | 252.5 | 2.0024e-05 | 306 | 8.4457e-06 | 8.4457e-06
234.8389 | 3.8094e-05 | 210.5 | 0.00039376 | 221.2 | 0.00017014 | 253 | 1.9936e-05 | 307 | 8.2961e-06 | 8.2961e-06
235.2027 | 3.7066e-05 | 210.6 | 0.00039038 | 221.4 | 0.00016809 | 253.5 | 1.985e-05 | 308 | 8.1553e-06 | 8.1553e-06
235.5664 | 3.6049e-05 | 210.7 | 0.00038705 | 221.6 | 0.000166 | 254 | 1.976e-05 | 309 | 8.0215e-06 | 8.0215e-06
235.93 | 3.5102e-05 | 210.8 | 0.00038364 | 221.8 | 0.00016383 | 254.5 | 1.9683e-05 | 310 | 7.8939e-06 | 7.8939e-06
236.2935 | 3.4214e-05 | 210.9 | 0.00038022 | 222 | 0.00016152 | 255 | 1.9634e-05 | 311 | 7.7766e-06 | 7.7766e-06
236.6569 | 3.3399e-05 | 211 | 0.0003768 | 222.2 | 0.00015899 | 255.5 | 1.961e-05 | 312 | 7.6703e-06 | 7.6703e-06
237.0203 | 3.26e-05 | 211.1 | 0.00037404 | 222.4 | 0.00015646 | 256 | 1.9596e-05 | 313 | 7.5662e-06 | 7.5662e-06
237.3836 | 3.1848e-05 | 211.2 | 0.00037085 | 222.6 | 0.00015366 | 256.5 | 1.9589e-05 | 314 | 7.4689e-06 | 7.4689e-06
237.7467 | 3.1142e-05 | 211.3 | 0.00036766 | 222.8 | 0.00015091 | 257 | 1.9594e-05 | 315 | 7.3775e-06 | 7.3775e-06
238.1098 | 3.0492e-05 | 211.4 | 0.00036308 | 223 | 0.00014814 | 257.5 | 1.9609e-05 | 316 | 7.3057e-06 | 7.3057e-06
238.4729 | 2.987e-05 | 211.5 | 0.00035943 | 223.2 | 0.00014538 | 258 | 1.9625e-05 | 317 | 7.2485e-06 | 7.2485e-06
238.8358 | 2.9289e-05 | 211.6 | 0.00035578 | 223.4 | 0.00014262 | 258.5 | 1.9642e-05 | 318 | 7.1932e-06 | 7.1932e-06
239.1986 | 2.8724e-05 | 211.7 | 0.00035213 | 223.6 | 0.00013976 | 259 | 1.9664e-05 | 319 | 7.1367e-06 | 7.1367e-06
239.5614 | 2.8195e-05 | 211.8 | 0.00034931 | 223.8 | 0.00013689 | 259.5 | 1.9697e-05 | 320 | 7.081e-06 | 7.081e-06
239.9241 | 2.7684e-05 | 211.9 | 0.00034569 | 224 | 0.00013406 | 260 | 1.9754e-05 | 321 | 7.0134e-06 | 7.0134e-06
240.2867 | 2.7232e-05 | 212 | 0.00034207 | 224.2 | 0.0001312 | 260.5 | 1.9827e-05 | 322 | 6.9527e-06 | 6.9527e-06
240.6492 | 2.679e-05 | 212.1 | 0.00033846 | 224.4 | 0.00012838 | 261 | 1.9902e-05 | 323 | 6.9004e-06 | 6.9004e-06
241.0116 | 2.6366e-05 | 212.2 | 0.00033454 | 224.6 | 0.0001255 | 261.5 | 1.9977e-05 | 324 | 6.854e-06 | 6.854e-06
241.374 | 2.5984e-05 | 212.3 | 0.00033107 | 224.8 | 0.00012249 | 262 | 2.0066e-05 | 325 | 6.822e-06 | 6.822e-06
241.7362 | 2.5622e-05 | 212.4 | 0.00032759 | 225 | 0.00011946 | 262.5 | 2.0169e-05 | 326 | 6.7948e-06 | 6.7948e-06
242.0984 | 2.5277e-05 | 212.5 | 0.00032446 | 225.2 | 0.00011638 | 263 | 2.0265e-05 | 327 | 6.7753e-06 | 6.7753e-06
242.4605 | 2.4939e-05 | 212.6 | 0.00032089 | 225.4 | 0.00011336 | 263.5 | 2.0354e-05 | 328 | 6.7567e-06 | 6.7567e-06
242.8225 | 2.4627e-05 | 212.7 | 0.00031732 | 225.6 | 0.00011035 | 264 | 2.0428e-05 | 329 | 6.7209e-06 | 6.7209e-06
243.1844 | 2.4319e-05 | 212.8 | 0.00031375 | 225.8 | 0.00010734 | 264.5 | 2.0507e-05 | 330 | 6.6951e-06 | 6.6951e-06
243.5463 | 2.4039e-05 | 212.9 | 0.00031049 | 226 | 0.00010441 | 265 | 2.0593e-05 | 331 | 6.6708e-06 | 6.6708e-06
243.908 | 2.376e-05 | 213 | 0.00030718 | 226.2 | 0.00010162 | 265.5 | 2.0681e-05 | 332 | 6.6465e-06 | 6.6465e-06
244.2697 | 2.3503e-05 | 213.1 | 0.00030388 | 226.4 | 9.8823e-05 | 266 | 2.0771e-05 | 333 | 6.638e-06 | 6.638e-06
244.6313 | 2.3249e-05 | 213.2 | 0.00030058 | 226.6 | 9.6189e-05 | 266.5 | 2.0853e-05 | 334 | 6.6221e-06 | 6.6221e-06
244.9928 | 2.3013e-05 | 213.3 | 0.00029751 | 226.8 | 9.355e-05 | 267 | 2.0924e-05 | 335 | 6.6068e-06 | 6.6068e-06
245.3542 | 2.2791e-05 | 213.4 | 0.00029424 | 227 | 9.0964e-05 | 267.5 | 2.0995e-05 | 336 | 6.5968e-06 | 6.5968e-06
245.7156 | 2.2585e-05 | 213.5 | 0.00029097 | 227.2 | 8.8458e-05 | 268 | 2.1073e-05 | 337 | 6.5881e-06 | 6.5881e-06
246.0768 | 2.2387e-05 | 213.6 | 0.00028853 | 227.4 | 8.598e-05 | 268.5 | 2.1118e-05 | 338 | 6.5785e-06 | 6.5785e-06
246.438 | 2.2198e-05 | 213.7 | 0.00028565 | 227.6 | 8.3643e-05 | 269 | 2.115e-05 | 339 | 6.5716e-06 | 6.5716e-06
246.799 | 2.2003e-05 | 213.8 | 0.00028278 | 227.8 | 8.1352e-05 | 269.5 | 2.1195e-05 | 340 | 6.5694e-06 | 6.5694e-06
247.1601 | 2.1825e-05 | 213.9 | 0.00027991 | 228 | 7.911e-05 | 270 | 2.1239e-05 | 341 | 6.5746e-06 | 6.5746e-06
247.521 | 2.1649e-05 | 214 | 0.00027688 | 228.2 | 7.6945e-05 | 270.5 | 2.1281e-05 | 342 | 6.5888e-06 | 6.5888e-06
247.8818 | 2.1505e-05 | 214.1 | 0.00027406 | 228.4 | 7.4868e-05 | 271 | 2.1324e-05 | 343 | 6.5979e-06 | 6.5979e-06
248.2425 | 2.134e-05 | 214.2 | 0.00027124 | 228.6 | 7.2817e-05 | 271.5 | 2.1352e-05 | 344 | 6.5999e-06 | 6.5999e-06
248.6032 | 2.1182e-05 | 214.3 | 0.00026843 | 228.8 | 7.0925e-05 | 272 | 2.1354e-05 | 345 | 6.6e-06 | 6.6e-06
248.9638 | 2.1037e-05 | 214.4 | 0.00026643 | 229 | 6.9032e-05 | 272.5 | 2.138e-05 | 346 | 6.5955e-06 | 6.5955e-06
249.3243 | 2.0893e-05 | 214.5 | 0.0002638 | 229.2 | 6.7242e-05 | 273 | 2.1388e-05 | 347 | 6.5982e-06 | 6.5982e-06
249.6847 | 2.0776e-05 | 214.6 | 0.00026117 | 229.4 | 6.5519e-05 | 273.5 | 2.135e-05 | 348 | 6.6098e-06 | 6.6098e-06
250.045 | 2.0661e-05 | 214.7 | 0.00025854 | 229.6 | 6.3828e-05 | 274 | 2.1309e-05 | 349 | 6.6215e-06 | 6.6215e-06
250.4053 | 2.0546e-05 | 214.8 | 0.00025577 | 229.8 | 6.2293e-05 | 274.5 | 2.1264e-05 | 350 | 6.6365e-06 | 6.6365e-06
250.7654 | 2.045e-05 | 214.9 | 0.00025324 | 230 | 6.0802e-05 | 275 | 2.1211e-05 | 351 | 6.6646e-06 | 6.6646e-06
251.1255 | 2.0347e-05 | 215 | 0.00025071 | 230.2 | 5.9393e-05 | 275.5 | 2.1166e-05 | 352 | 6.6963e-06 | 6.6963e-06
251.4855 | 2.0254e-05 | 215.1 | 0.00024822 | 230.4 | 5.8041e-05 | 276 | 2.1111e-05 | 353 | 6.733e-06 | 6.733e-06
251.8454 | 2.0158e-05 | 215.2 | 0.00024557 | 230.6 | 5.6713e-05 | 276.5 | 2.1036e-05 | 354 | 6.7653e-06 | 6.7653e-06
252.2053 | 2.0081e-05 | 215.3 | 0.00024293 | 230.8 | 5.5418e-05 | 277 | 2.0949e-05 | 355 | 6.7864e-06 | 6.7864e-06
252.565 | 2.0011e-05 | 215.4 | 0.00024028 | 231 | 5.4226e-05 | 277.5 | 2.0855e-05 | 356 | 6.8075e-06 | 6.8075e-06
252.9247 | 1.9943e-05 | 215.5 | 0.00023892 | 231.2 | 5.3058e-05 | 278 | 2.0739e-05 | 357 | 6.8202e-06 | 6.8202e-06
253.2842 | 1.9896e-05 | 215.6 | 0.00023676 | 231.4 | 5.1929e-05 | 278.5 | 2.0614e-05 | 358 | 6.8324e-06 | 6.8324e-06
253.6438 | 1.9821e-05 | 215.7 | 0.0002346 | 231.6 | 5.0854e-05 | 279 | 2.045e-05 | 359 | 6.8379e-06 | 6.8379e-06
254.0032 | 1.9756e-05 | 215.8 | 0.00023244 | 231.8 | 4.9796e-05 | 279.5 | 2.0284e-05 | 360 | 6.8249e-06 | 6.8249e-06
254.3625 | 1.9703e-05 | 215.9 | 0.00023028 | 232 | 4.8806e-05 | 280 | 2.0121e-05 | 361 | 6.8229e-06 | 6.8229e-06
254.7217 | 1.965e-05 | 216 | 0.00022824 | 232.2 | 4.7856e-05 | 280.5 | 1.993e-05 | 362 | 6.8217e-06 | 6.8217e-06
255.0809 | 1.9627e-05 | 216.1 | 0.00022619 | 232.4 | 4.6941e-05 | 281 | 1.9716e-05 | 363 | 6.8158e-06 | 6.8158e-06
255.44 | 1.9606e-05 | 216.2 | 0.00022456 | 232.6 | 4.6047e-05 | 281.5 | 1.9481e-05 | 364 | 6.8146e-06 | 6.8146e-06
255.799 | 1.9604e-05 | 216.3 | 0.00022258 | 232.8 | 4.5218e-05 | 282 | 1.9224e-05 | 365 | 6.8053e-06 | 6.8053e-06
256.1579 | 1.959e-05 | 216.4 | 0.00022059 | 233 | 4.439e-05 | 282.5 | 1.8974e-05 | 366 | 6.7819e-06 | 6.7819e-06
256.5167 | 1.9589e-05 | 216.5 | 0.00021861 | 233.2 | 4.3615e-05 | 283 | 1.8715e-05 | 367 | 6.7476e-06 | 6.7476e-06
256.8755 | 1.9588e-05 | 216.6 | 0.00021726 | 233.4 | 4.2862e-05 | 283.5 | 1.843e-05 | 368 | 6.7221e-06 | 6.7221e-06
257.2341 | 1.9599e-05 | 216.7 | 0.00021559 | 233.6 | 4.2123e-05 | 284 | 1.8128e-05 | 369 | 6.7126e-06 | 6.7126e-06
257.5927 | 1.9614e-05 | 216.8 | 0.00021393 | 233.8 | 4.1404e-05 | 284.5 | 1.7816e-05 | 370 | 6.7137e-06 | 6.7137e-06
257.9512 | 1.9623e-05 | 216.9 | 0.00021226 | 234 | 4.0698e-05 | 285 | 1.7508e-05 | 371 | 6.7218e-06 | 6.7218e-06
258.3096 | 1.9635e-05 | 217 | 0.00021089 | 234.2 | 4.0039e-05 | 285.5 | 1.7198e-05 | 372 | 6.7122e-06 | 6.7122e-06
258.6679 | 1.9647e-05 | 217.1 | 0.00020933 | 234.4 | 3.9407e-05 | 286 | 1.6885e-05 | 373 | 6.6689e-06 | 6.6689e-06
259.0262 | 1.9662e-05 | 217.2 | 0.00020777 | 234.6 | 3.8807e-05 | 286.5 | 1.6551e-05 | 374 | 6.5933e-06 | 6.5933e-06
259.3843 | 1.969e-05 | 217.3 | 0.00020671 | 234.8 | 3.8211e-05 | 287 | 1.6222e-05 | 375 | 6.5251e-06 | 6.5251e-06
259.7424 | 1.9714e-05 | 217.4 | 0.00020537 | 235 | 3.7643e-05 | 287.5 | 1.5887e-05 | 376 | 6.4776e-06 | 6.4776e-06
260.1004 | 1.9771e-05 | 217.5 | 0.00020402 | 235.2 | 3.7073e-05 | 288 | 1.555e-05 | 377 | 6.4352e-06 | 6.4352e-06
260.4583 | 1.9816e-05 | 217.6 | 0.00020267 | 235.4 | 3.6515e-05 | 288.5 | 1.5219e-05 | 378 | 6.3929e-06 | 6.3929e-06
260.8161 | 1.9877e-05 | 217.7 | 0.00020169 | 235.6 | 3.5976e-05 | 289 | 1.4892e-05 | 379 | 6.3479e-06 | 6.3479e-06
261.1739 | 1.9927e-05 | 217.8 | 0.00020053 | 235.8 | 3.5441e-05 | 289.5 | 1.457e-05 | 380 | 6.3001e-06 | 6.3001e-06
261.5315 | 1.998e-05 | 217.9 | 0.00019937 | 236 | 3.4936e-05 | 290 | 1.4256e-05 | 381 | 6.2482e-06 | 6.2482e-06
261.8891 | 2.0044e-05 | 218 | 0.00019821 | 236.2 | 3.4452e-05 | 290.5 | 1.3951e-05 | 382 | 6.1982e-06 | 6.1982e-06
262.2466 | 2.0111e-05 | 218.1 | 0.0001975 | 236.4 | 3.3989e-05 | 291 | 1.3648e-05 | 383 | 6.1493e-06 | 6.1493e-06
262.604 | 2.0194e-05 | 218.2 | 0.00019648 | 236.6 | 3.3527e-05 | 291.5 | 1.3356e-05 | 384 | 6.1006e-06 | 6.1006e-06
262.9613 | 2.0259e-05 | 218.3 | 0.00019547 | 236.8 | 3.3086e-05 | 292 | 1.3077e-05 | 385 | 6.0549e-06 | 6.0549e-06
263.3186 | 2.0322e-05 | 218.4 | 0.0001946 | 237 | 3.2651e-05 | 292.5 | 1.2801e-05 | 386 | 6.0072e-06 | 6.0072e-06
263.6758 | 2.0389e-05 | 218.5 | 0.0001937 | 237.2 | 3.2232e-05 | 293 | 1.2533e-05 | 387 | 5.9639e-06 | 5.9639e-06
264.0328 | 2.0428e-05 | 218.6 | 0.0001928 | 237.4 | 3.1826e-05 | 293.5 | 1.2269e-05 | 388 | 5.9217e-06 | 5.9217e-06
264.3898 | 2.0488e-05 | 218.7 | 0.00019189 | 237.6 | 3.1429e-05 | 294 | 1.2019e-05 | 389 | 5.8816e-06 | 5.8816e-06
264.7468 | 2.055e-05 | 218.8 | 0.00019126 | 237.8 | 3.1055e-05 | 294.5 | 1.1777e-05 | 390 | 5.8402e-06 | 5.8402e-06
265.1036 | 2.0612e-05 | 218.9 | 0.00019045 | 238 | 3.0688e-05 | 295 | 1.1542e-05 | 391 | 5.8007e-06 | 5.8007e-06
265.4603 | 2.0671e-05 | 219 | 0.00018964 | 238.2 | 3.0337e-05 | 295.5 | 1.1323e-05 | 392 | 5.7597e-06 | 5.7597e-06
265.817 | 2.0741e-05 | 219.1 | 0.00018883 | 238.4 | 3e-05 | 296 | 1.1112e-05 | 393 | 5.7194e-06 | 5.7194e-06
266.1736 | 2.0802e-05 | 219.2 | 0.00018814 | 238.6 | 2.9673e-05 | 296.5 | 1.0914e-05 | 394 | 5.681e-06 | 5.681e-06
266.5301 | 2.086e-05 | 219.3 | 0.00018741 | 238.8 | 2.9347e-05 | 297 | 1.0731e-05 | 395 | 5.6419e-06 | 5.6419e-06
266.8865 | 2.0905e-05 | 219.4 | 0.00018667 | 239 | 2.9035e-05 | 297.5 | 1.0549e-05 | 396 | 5.6037e-06 | 5.6037e-06
267.2428 | 2.0963e-05 | 219.5 | 0.00018569 | 239.2 | 2.8731e-05 | 298 | 1.0369e-05 | 397 | 5.567e-06 | 5.567e-06
267.5991 | 2.1002e-05 | 219.6 | 0.00018485 | 239.4 | 2.8432e-05 | 298.5 | 1.0201e-05 | 398 | 5.5335e-06 | 5.5335e-06
267.9552 | 2.1076e-05 | 219.7 | 0.00018402 | 239.6 | 2.8143e-05 | 299 | 1.0041e-05 | 399 | 5.5024e-06 | 5.5024e-06
268.3113 | 2.111e-05 | 219.8 | 0.00018318 | 239.8 | 2.7859e-05 | 299.5 | 9.8849e-06 | 400 | 5.4705e-06 | 5.4705e-06
268.6673 | 2.1124e-05 | 219.9 | 0.00018229 | 240 | 2.7596e-05 | 300 | 9.741e-06 | 401 | 5.4361e-06 | 5.4361e-06
269.0233 | 2.1147e-05 | 220 | 0.0001814 | 240.2 | 2.7342e-05 | 300.5 | 9.5988e-06 | 402 | 5.4017e-06 | 5.4017e-06
269.3791 | 2.1186e-05 | 220.1 | 0.0001805 | 240.4 | 2.7096e-05 | 301 | 9.4661e-06 | 403 | 5.3687e-06 | 5.3687e-06
269.7348 | 2.1215e-05 | 220.2 | 0.00017961 | 240.6 | 2.6852e-05 | 301.5 | 9.3482e-06 | 404 | 5.3362e-06 | 5.3362e-06
270.0905 | 2.1249e-05 | 220.3 | 0.00017858 | 240.8 | 2.6616e-05 | 302 | 9.2373e-06 | 405 | 5.3014e-06 | 5.3014e-06
270.4461 | 2.1273e-05 | 220.4 | 0.00017766 | 241 | 2.6385e-05 | 302.5 | 9.1278e-06 | 406 | 5.2656e-06 | 5.2656e-06
270.8016 | 2.1309e-05 | 220.5 | 0.00017674 | 241.2 | 2.6171e-05 | 303 | 9.0159e-06 | 407 | 5.228e-06 | 5.228e-06
271.157 | 2.1338e-05 | 220.6 | 0.00017573 | 241.4 | 2.5961e-05 | 303.5 | 8.9036e-06 | 408 | 5.1934e-06 | 5.1934e-06
271.5124 | 2.1363e-05 | 220.7 | 0.00017474 | 241.6 | 2.5759e-05 | 304 | 8.799e-06 | 409 | 5.1585e-06 | 5.1585e-06
271.8676 | 2.1339e-05 | 220.8 | 0.00017376 | 241.8 | 2.5563e-05 | 304.5 | 8.7003e-06 | 410 | 5.1231e-06 | 5.1231e-06
272.2228 | 2.1362e-05 | 220.9 | 0.00017278 | 242 | 2.5372e-05 | 305 | 8.6085e-06 | 411 | 5.0904e-06 | 5.0904e-06
272.5779 | 2.1388e-05 | 221 | 0.00017198 | 242.2 | 2.5184e-05 | 305.5 | 8.5245e-06 | 412 | 5.0602e-06 | 5.0602e-06
272.9329 | 2.14e-05 | 221.1 | 0.00017106 | 242.4 | 2.4999e-05 | 306 | 8.446e-06 | 413 | 5.031e-06 | 5.031e-06
273.2878 | 2.1375e-05 | 221.2 | 0.00017014 | 242.6 | 2.4823e-05 | 306.5 | 8.3665e-06 | 414 | 5.0016e-06 | 5.0016e-06
273.6427 | 2.1333e-05 | 221.3 | 0.00016922 | 242.8 | 2.4646e-05 | 307 | 8.2973e-06 | 415 | 4.971e-06 | 4.971e-06
273.9975 | 2.1308e-05 | 221.4 | 0.00016801 | 243 | 2.4476e-05 | 307.5 | 8.2244e-06 | 416 | 4.9406e-06 | 4.9406e-06
274.3521 | 2.1287e-05 | 221.5 | 0.00016703 | 243.2 | 2.4314e-05 | 308 | 8.1545e-06 | 417 | 4.9111e-06 | 4.9111e-06
274.7067 | 2.1238e-05 | 221.6 | 0.00016604 | 243.4 | 2.4153e-05 | 308.5 | 8.0871e-06 | 418 | 4.8839e-06 | 4.8839e-06
275.0612 | 2.1203e-05 | 221.7 | 0.00016492 | 243.6 | 2.3997e-05 | 309 | 8.0206e-06 | 419 | 4.8589e-06 | 4.8589e-06
275.4157 | 2.1171e-05 | 221.8 | 0.00016383 | 243.8 | 2.3843e-05 | 309.5 | 7.9566e-06 | 420 | 4.8365e-06 | 4.8365e-06
275.77 | 2.1145e-05 | 221.9 | 0.00016273 | 244 | 2.3697e-05 | 310 | 7.893e-06 | 421 | 4.8165e-06 | 4.8165e-06
276.1243 | 2.1096e-05 | 222 | 0.00016163 | 244.2 | 2.3553e-05 | 310.5 | 7.8321e-06 | 422 | 4.7958e-06 | 4.7958e-06
276.4785 | 2.1039e-05 | 222.1 | 0.0001602 | 244.4 | 2.3412e-05 | 311 | 7.7745e-06 | 423 | 4.7844e-06 | 4.7844e-06
276.8326 | 2.0982e-05 | 222.2 | 0.00015899 | 244.6 | 2.3273e-05 | 311.5 | 7.7231e-06 | 424 | 4.8096e-06 | 4.8096e-06
277.1866 | 2.0911e-05 | 222.3 | 0.00015778 | 244.8 | 2.3141e-05 | 312 | 7.6698e-06 | 425 | 4.8951e-06 | 4.8951e-06
277.5405 | 2.0856e-05 | 222.4 | 0.00015658 | 245 | 2.3012e-05 | 312.5 | 7.6178e-06 | 426 | 5.0166e-06 | 5.0166e-06
277.8944 | 2.0761e-05 | 222.5 | 0.00015502 | 245.2 | 2.2887e-05 | 313 | 7.5652e-06 | 427 | 5.1004e-06 | 5.1004e-06
278.2482 | 2.068e-05 | 222.6 | 0.00015366 | 245.4 | 2.2768e-05 | 313.5 | 7.5156e-06 | 428 | 5.0907e-06 | 5.0907e-06
278.6019 | 2.0595e-05 | 222.7 | 0.00015229 | 245.6 | 2.2651e-05 | 314 | 7.4701e-06 | 429 | 4.9762e-06 | 4.9762e-06
278.9555 | 2.0461e-05 | 222.8 | 0.00015091 | 245.8 | 2.2539e-05 | 314.5 | 7.421e-06 | 430 | 4.8484e-06 | 4.8484e-06
279.309 | 2.0347e-05 | 222.9 | 0.00014953 | 246 | 2.243e-05 | 315 | 7.3742e-06 | 431 | 4.7886e-06 | 4.7886e-06
279.6624 | 2.023e-05 | 223 | 0.00014814 | 246.2 | 2.2324e-05 | 315.5 | 7.3373e-06 | 432 | 4.7913e-06 | 4.7913e-06
280.0158 | 2.0125e-05 | 223.1 | 0.00014676 | 246.4 | 2.2217e-05 | 316 | 7.3051e-06 | 433 | 4.846e-06 | 4.846e-06
280.3691 | 1.9982e-05 | 223.2 | 0.00014538 | 246.6 | 2.211e-05 | 316.5 | 7.2746e-06 | 434 | 4.9043e-06 | 4.9043e-06
280.7223 | 1.9841e-05 | 223.3 | 0.000144 | 246.8 | 2.2007e-05 | 317 | 7.2488e-06 | 435 | 4.9423e-06 | 4.9423e-06
281.0754 | 1.9689e-05 | 223.4 | 0.00014262 | 247 | 2.1905e-05 | 317.5 | 7.222e-06 | 436 | 4.9698e-06 | 4.9698e-06
281.4284 | 1.9515e-05 | 223.5 | 0.00014124 | 247.2 | 2.1805e-05 | 318 | 7.1951e-06 | 437 | 4.9761e-06 | 4.9761e-06
281.7814 | 1.9345e-05 | 223.6 | 0.00013972 | 247.4 | 2.1708e-05 | 318.5 | 7.1625e-06 | 438 | 4.964e-06 | 4.964e-06
282.1343 | 1.9149e-05 | 223.7 | 0.0001383 | 247.6 | 2.1621e-05 | 319 | 7.137e-06 | 439 | 4.9442e-06 | 4.9442e-06
282.4871 | 1.8977e-05 | 223.8 | 0.00013688 | 247.8 | 2.1535e-05 | 319.5 | 7.1106e-06 | 440 | 4.9226e-06 | 4.9226e-06
282.8397 | 1.8813e-05 | 223.9 | 0.00013549 | 248 | 2.1447e-05 | 320 | 7.0825e-06 | 441 | 4.9003e-06 | 4.9003e-06
283.1924 | 1.8606e-05 | 224 | 0.00013406 | 248.2 | 2.1361e-05 | 320.5 | 7.0498e-06 | 442 | 4.8799e-06 | 4.8799e-06
283.5449 | 1.8409e-05 | 224.1 | 0.00013263 | 248.4 | 2.1272e-05 | 321 | 7.0116e-06 | 443 | 4.8635e-06 | 4.8635e-06
283.8974 | 1.8187e-05 | 224.2 | 0.0001312 | 248.6 | 2.1185e-05 | 321.5 | 6.9788e-06 | 444 | 4.8511e-06 | 4.8511e-06
284.2498 | 1.7979e-05 | 224.3 | 0.00012979 | 248.8 | 2.1104e-05 | 322 | 6.9528e-06 | 445 | 4.8401e-06 | 4.8401e-06
284.6021 | 1.7748e-05 | 224.4 | 0.00012838 | 249 | 2.1023e-05 | 322.5 | 6.9264e-06 | 446 | 4.8293e-06 | 4.8293e-06
284.9543 | 1.7539e-05 | 224.5 | 0.00012697 | 249.2 | 2.0943e-05 | 323 | 6.8989e-06 | 447 | 4.8162e-06 | 4.8162e-06
285.3064 | 1.7318e-05 | 224.6 | 0.00012557 | 249.4 | 2.0872e-05 | 323.5 | 6.8759e-06 | 448 | 4.8044e-06 | 4.8044e-06
285.6584 | 1.7097e-05 | 224.7 | 0.00012396 | 249.6 | 2.0805e-05 | 324 | 6.8512e-06 | 449 | 4.7937e-06 | 4.7937e-06
286.0104 | 1.6888e-05 | 224.8 | 0.00012249 | 249.8 | 2.074e-05 | 324.5 | 6.8349e-06 | 450 | 4.7853e-06 | 4.7853e-06
286.3623 | 1.6642e-05 | 224.9 | 0.00012102 | 250 | 2.0676e-05 | 325 | 6.8237e-06 | 451 | 4.7792e-06 | 4.7792e-06
286.7141 | 1.6406e-05 | 225 | 0.00011945 | 250.2 | 2.0612e-05 | 325.5 | 6.8073e-06 | 452 | 4.7757e-06 | 4.7757e-06
287.0658 | 1.6183e-05 | 225.1 | 0.00011791 | 250.4 | 2.0551e-05 | 326 | 6.7923e-06 | 453 | 4.7721e-06 | 4.7721e-06
287.4175 | 1.5945e-05 | 225.2 | 0.00011638 | 250.6 | 2.0496e-05 | 326.5 | 6.7849e-06 | 454 | 4.7687e-06 | 4.7687e-06
287.769 | 1.5701e-05 | 225.3 | 0.00011484 | 250.8 | 2.0438e-05 | 327 | 6.7734e-06 | 455 | 4.7654e-06 | 4.7654e-06
288.1205 | 1.547e-05 | 225.4 | 0.00011338 | 251 | 2.0383e-05 | 327.5 | 6.7675e-06 | 456 | 4.763e-06 | 4.763e-06
288.4719 | 1.5238e-05 | 225.5 | 0.00011186 | 251.2 | 2.0329e-05 | 328 | 6.7614e-06 | 457 | 4.7621e-06 | 4.7621e-06
288.8232 | 1.5006e-05 | 225.6 | 0.00011035 | 251.4 | 2.0276e-05 | 328.5 | 6.7412e-06 | 458 | 4.7621e-06 | 4.7621e-06
289.1744 | 1.4777e-05 | 225.7 | 0.00010883 | 251.6 | 2.0223e-05 | 329 | 6.7169e-06 | 459 | 4.7647e-06 | 4.7647e-06
289.5256 | 1.4555e-05 | 225.8 | 0.00010734 | 251.8 | 2.0173e-05 | 329.5 | 6.7046e-06 | 460 | 4.7692e-06 | 4.7692e-06
289.8766 | 1.4326e-05 | 225.9 | 0.00010584 | 252 | 2.0127e-05 | 330 | 6.6955e-06 | 461 | 4.774e-06 | 4.774e-06
290.2276 | 1.4118e-05 | 226 | 0.00010435 | 252.2 | 2.0082e-05 | 330.5 | 6.6852e-06 | 462 | 4.782e-06 | 4.782e-06
290.5785 | 1.3902e-05 | 226.1 | 0.00010304 | 252.4 | 2.0044e-05 | 331 | 6.6726e-06 | 463 | 4.7908e-06 | 4.7908e-06
290.9294 | 1.3692e-05 | 226.2 | 0.00010162 | 252.6 | 2.0005e-05 | 331.5 | 6.6548e-06 | 464 | 4.8009e-06 | 4.8009e-06
291.2801 | 1.3477e-05 | 226.3 | 0.00010019 | 252.8 | 1.9967e-05 | 332 | 6.6431e-06 | 465 | 4.8107e-06 | 4.8107e-06
291.6308 | 1.3282e-05 | 226.4 | 9.8761e-05 | 253 | 1.9936e-05 | 332.5 | 6.6417e-06 | 466 | 4.8175e-06 | 4.8175e-06
291.9813 | 1.3089e-05 | 226.5 | 9.752e-05 | 253.2 | 1.9904e-05 | 333 | 6.6406e-06 | 467 | 4.8227e-06 | 4.8227e-06
292.3318 | 1.2893e-05 | 226.6 | 9.6189e-05 | 253.4 | 1.9867e-05 | 333.5 | 6.6318e-06 | 468 | 4.8287e-06 | 4.8287e-06
292.6823 | 1.2698e-05 | 226.7 | 9.4858e-05 | 253.6 | 1.9832e-05 | 334 | 6.6212e-06 | 469 | 4.8371e-06 | 4.8371e-06
293.0326 | 1.2518e-05 | 226.8 | 9.3527e-05 | 253.8 | 1.9794e-05 | 334.5 | 6.6134e-06 | 470 | 4.8483e-06 | 4.8483e-06
293.3828 | 1.2328e-05 | 226.9 | 9.2267e-05 | 254 | 1.9758e-05 | 335 | 6.6056e-06 | 471 | 4.8612e-06 | 4.8612e-06
293.733 | 1.2147e-05 | 227 | 9.0964e-05 | 254.2 | 1.9728e-05 | 335.5 | 6.6014e-06 | 472 | 4.8735e-06 | 4.8735e-06
294.0831 | 1.1979e-05 | 227.1 | 8.966e-05 | 254.4 | 1.9697e-05 | 336 | 6.5976e-06 | 473 | 4.8858e-06 | 4.8858e-06
294.4331 | 1.1808e-05 | 227.2 | 8.8482e-05 | 254.6 | 1.9668e-05 | 336.5 | 6.5915e-06 | 474 | 4.8985e-06 | 4.8985e-06
294.783 | 1.164e-05 | 227.3 | 8.7231e-05 | 254.8 | 1.9648e-05 | 337 | 6.5889e-06 | 475 | 4.9121e-06 | 4.9121e-06
295.1329 | 1.1478e-05 | 227.4 | 8.598e-05 | 255 | 1.9632e-05 | 337.5 | 6.5838e-06 | 476 | 4.9264e-06 | 4.9264e-06
295.4826 | 1.1331e-05 | 227.5 | 8.4729e-05 | 255.2 | 1.962e-05 | 338 | 6.5785e-06 | 477 | 4.9413e-06 | 4.9413e-06
295.8323 | 1.1182e-05 | 227.6 | 8.3683e-05 | 255.4 | 1.9611e-05 | 338.5 | 6.5733e-06 | 478 | 4.9561e-06 | 4.9561e-06
296.1819 | 1.1033e-05 | 227.7 | 8.2518e-05 | 255.6 | 1.9607e-05 | 339 | 6.5723e-06 | 479 | 4.9712e-06 | 4.9712e-06
296.5314 | 1.0901e-05 | 227.8 | 8.1352e-05 | 255.8 | 1.9601e-05 | 339.5 | 6.5692e-06 | 480 | 4.9873e-06 | 4.9873e-06
296.8809 | 1.0774e-05 | 227.9 | 8.0187e-05 | 256 | 1.9596e-05 | 340 | 6.569e-06 | 481 | 5.0025e-06 | 5.0025e-06
297.2302 | 1.0648e-05 | 228 | 7.9129e-05 | 256.2 | 1.9593e-05 | 340.5 | 6.5701e-06 | 482 | 5.018e-06 | 5.018e-06
297.5795 | 1.0521e-05 | 228.1 | 7.8013e-05 | 256.4 | 1.9589e-05 | 341 | 6.5738e-06 | 483 | 5.0341e-06 | 5.0341e-06
297.9287 | 1.0391e-05 | 228.2 | 7.6896e-05 | 256.6 | 1.9588e-05 | 341.5 | 6.5798e-06 | 484 | 5.0512e-06 | 5.0512e-06
298.2778 | 1.0272e-05 | 228.3 | 7.5927e-05 | 256.8 | 1.959e-05 | 342 | 6.5894e-06 | 485 | 5.068e-06 | 5.068e-06
298.6268 | 1.0159e-05 | 228.4 | 7.4868e-05 | 257 | 1.9594e-05 | 342.5 | 6.5972e-06 | 486 | 5.0847e-06 | 5.0847e-06
298.9758 | 1.0049e-05 | 228.5 | 7.3808e-05 | 257.2 | 1.9598e-05 | 343 | 6.5973e-06 | 487 | 5.0999e-06 | 5.0999e-06
299.3246 | 9.938e-06 | 228.6 | 7.2749e-05 | 257.4 | 1.9606e-05 | 343.5 | 6.5992e-06 | 488 | 5.1144e-06 | 5.1144e-06
299.6734 | 9.8299e-06 | 228.7 | 7.1895e-05 | 257.6 | 1.9612e-05 | 344 | 6.6e-06 | 489 | 5.1301e-06 | 5.1301e-06
300.0221 | 9.7366e-06 | 228.8 | 7.0925e-05 | 257.8 | 1.9619e-05 | 344.5 | 6.6005e-06 | 490 | 5.1456e-06 | 5.1456e-06
300.3707 | 9.635e-06 | 228.9 | 6.9955e-05 | 258 | 1.9625e-05 | 345 | 6.6001e-06 | 491 | 5.1597e-06 | 5.1597e-06
300.7193 | 9.5364e-06 | 229 | 6.8985e-05 | 258.2 | 1.9631e-05 | 345.5 | 6.5993e-06 | 492 | 5.1736e-06 | 5.1736e-06
301.0677 | 9.4452e-06 | 229.1 | 6.8157e-05 | 258.4 | 1.9638e-05 | 346 | 6.5926e-06 | 493 | 5.1881e-06 | 5.1881e-06
301.4161 | 9.3674e-06 | 229.2 | 6.7242e-05 | 258.6 | 1.9645e-05 | 346.5 | 6.5947e-06 | 494 | 5.2033e-06 | 5.2033e-06
301.7644 | 9.2887e-06 | 229.3 | 6.6328e-05 | 258.8 | 1.9653e-05 | 347 | 6.5975e-06 | 495 | 5.2182e-06 | 5.2182e-06
302.1126 | 9.2128e-06 | 229.4 | 6.5544e-05 | 259 | 1.9663e-05 | 347.5 | 6.6024e-06 | 496 | 5.232e-06 | 5.232e-06
302.4608 | 9.1349e-06 | 229.5 | 6.4686e-05 | 259.2 | 1.9677e-05 | 348 | 6.6091e-06 | 497 | 5.2452e-06 | 5.2452e-06
302.8088 | 9.0618e-06 | 229.6 | 6.3828e-05 | 259.4 | 1.969e-05 | 348.5 | 6.6178e-06 | 498 | 5.2589e-06 | 5.2589e-06
303.1568 | 8.9782e-06 | 229.7 | 6.297e-05 | 259.6 | 1.9704e-05 | 349 | 6.6211e-06 | 499 | 5.2715e-06 | 5.2715e-06
303.5047 | 8.9022e-06 | 229.8 | 6.2339e-05 | 259.8 | 1.9728e-05 | 349.5 | 6.6257e-06 | 500 | 5.2843e-06 | 5.2843e-06
303.8525 | 8.8285e-06 | 229.9 | 6.1571e-05 | 260 | 1.9754e-05 | 350 | 6.635e-06 | 501 | 5.2965e-06 | 5.2965e-06
304.2002 | 8.758e-06 | 230 | 6.0802e-05 | 260.2 | 1.9781e-05 | 350.5 | 6.6487e-06 | 502 | 5.3088e-06 | 5.3088e-06
304.5479 | 8.6892e-06 | 230.1 | 6.0033e-05 | 260.4 | 1.9811e-05 | 351 | 6.6638e-06 | 503 | 5.3204e-06 | 5.3204e-06
304.8955 | 8.6291e-06 | 230.2 | 5.9425e-05 | 260.6 | 1.9843e-05 | 351.5 | 6.6814e-06 | 504 | 5.3319e-06 | 5.3319e-06
305.243 | 8.5622e-06 | 230.3 | 5.8721e-05 | 260.8 | 1.9872e-05 | 352 | 6.6959e-06 | 505 | 5.3433e-06 | 5.3433e-06
305.5904 | 8.5107e-06 | 230.4 | 5.8018e-05 | 261 | 1.9902e-05 | 352.5 | 6.7116e-06 | 506 | 5.3553e-06 | 5.3553e-06
305.9377 | 8.4569e-06 | 230.5 | 5.7385e-05 | 261.2 | 1.9932e-05 | 353 | 6.7336e-06 | 507 | 5.3678e-06 | 5.3678e-06
306.2849 | 8.3996e-06 | 230.6 | 5.6713e-05 | 261.4 | 1.996e-05 | 353.5 | 6.7537e-06 | 508 | 5.3797e-06 | 5.3797e-06
306.6321 | 8.3443e-06 | 230.7 | 5.6041e-05 | 261.6 | 1.9994e-05 | 354 | 6.7659e-06 | 509 | 5.3911e-06 | 5.3911e-06
306.9792 | 8.3015e-06 | 230.8 | 5.5369e-05 | 261.8 | 2.0029e-05 | 354.5 | 6.7763e-06 | 510 | 5.4018e-06 | 5.4018e-06
307.3262 | 8.2524e-06 | 230.9 | 5.4844e-05 | 262 | 2.0065e-05 | 355 | 6.7863e-06 | 511 | 5.4123e-06 | 5.4123e-06
307.6731 | 8.1956e-06 | 231 | 5.4226e-05 | 262.2 | 2.0104e-05 | 355.5 | 6.7965e-06 | 512 | 5.4229e-06 | 5.4229e-06
308.02 | 8.1515e-06 | 231.1 | 5.3608e-05 | 262.4 | 2.0149e-05 | 356 | 6.8091e-06 | 513 | 5.4333e-06 | 5.4333e-06
308.3667 | 8.1075e-06 | 231.2 | 5.3067e-05 | 262.6 | 2.0189e-05 | 356.5 | 6.8169e-06 | 514 | 5.4439e-06 | 5.4439e-06
308.7134 | 8.0565e-06 | 231.3 | 5.2498e-05 | 262.8 | 2.0229e-05 | 357 | 6.8206e-06 | 515 | 5.4542e-06 | 5.4542e-06
309.06 | 8.0124e-06 | 231.4 | 5.1929e-05 | 263 | 2.0266e-05 | 357.5 | 6.8231e-06 | 516 | 5.4643e-06 | 5.4643e-06
309.4065 | 7.9694e-06 | 231.5 | 5.1361e-05 | 263.2 | 2.0301e-05 | 358 | 6.8323e-06 | 517 | 5.4752e-06 | 5.4752e-06
309.7529 | 7.9236e-06 | 231.6 | 5.0868e-05 | 263.4 | 2.0338e-05 | 358.5 | 6.8417e-06 | 518 | 5.486e-06 | 5.486e-06
310.0993 | 7.8789e-06 | 231.7 | 5.0332e-05 | 263.6 | 2.0371e-05 | 359 | 6.841e-06 | 519 | 5.4961e-06 | 5.4961e-06
310.4456 | 7.8405e-06 | 231.8 | 4.9796e-05 | 263.8 | 2.0398e-05 | 359.5 | 6.8309e-06 | 520 | 5.5054e-06 | 5.5054e-06
310.7918 | 7.7956e-06 | 231.9 | 4.926e-05 | 264 | 2.0428e-05 | 360 | 6.8232e-06 | 521 | 5.5142e-06 | 5.5142e-06
311.1379 | 7.7586e-06 | 232 | 4.8827e-05 | 264.2 | 2.0458e-05 | 360.5 | 6.8205e-06 | 522 | 5.5249e-06 | 5.5249e-06
311.4839 | 7.7265e-06 | 232.1 | 4.8331e-05 | 264.4 | 2.049e-05 | 361 | 6.8221e-06 | 523 | 5.5371e-06 | 5.5371e-06
311.8299 | 7.6885e-06 | 232.2 | 4.7835e-05 | 264.6 | 2.0524e-05 | 361.5 | 6.8259e-06 | 524 | 5.5501e-06 | 5.5501e-06
312.1757 | 7.6496e-06 | 232.3 | 4.7402e-05 | 264.8 | 2.0559e-05 | 362 | 6.8235e-06 | 525 | 5.5628e-06 | 5.5628e-06
312.5215 | 7.6157e-06 | 232.4 | 4.6941e-05 | 265 | 2.0594e-05 | 362.5 | 6.8158e-06 | 526 | 5.5747e-06 | 5.5747e-06
312.8673 | 7.582e-06 | 232.5 | 4.648e-05 | 265.2 | 2.0628e-05 | 363 | 6.8157e-06 | 527 | 5.5869e-06 | 5.5869e-06
313.2129 | 7.5397e-06 | 232.6 | 4.6019e-05 | 265.4 | 2.0662e-05 | 363.5 | 6.8157e-06 | 528 | 5.5989e-06 | 5.5989e-06
313.5584 | 7.5102e-06 | 232.7 | 4.5643e-05 | 265.6 | 2.07e-05 | 364 | 6.815e-06 | 529 | 5.6097e-06 | 5.6097e-06
313.9039 | 7.4814e-06 | 232.8 | 4.5218e-05 | 265.8 | 2.0736e-05 | 364.5 | 6.813e-06 | 530 | 5.6206e-06 | 5.6206e-06
314.2493 | 7.4436e-06 | 232.9 | 4.4793e-05 | 266 | 2.0772e-05 | 365 | 6.8066e-06 | 531 | 5.6334e-06 | 5.6334e-06
314.5946 | 7.413e-06 | 233 | 4.4367e-05 | 266.2 | 2.0806e-05 | 365.5 | 6.7963e-06 | 532 | 5.6469e-06 | 5.6469e-06
314.9398 | 7.3776e-06 | 233.1 | 4.4011e-05 | 266.4 | 2.0839e-05 | 366 | 6.7829e-06 | 533 | 5.661e-06 | 5.661e-06
315.285 | 7.3487e-06 | 233.2 | 4.3615e-05 | 266.6 | 2.0867e-05 | 366.5 | 6.7664e-06 | 534 | 5.6751e-06 | 5.6751e-06
315.63 | 7.3301e-06 | 233.3 | 4.3219e-05 | 266.8 | 2.0895e-05 | 367 | 6.7459e-06 | 535 | 5.6887e-06 | 5.6887e-06
315.975 | 7.3057e-06 | 233.4 | 4.2871e-05 | 267 | 2.0925e-05 | 367.5 | 6.7305e-06 | 536 | 5.702e-06 | 5.702e-06
316.32 | 7.2855e-06 | 233.5 | 4.2497e-05 | 267.2 | 2.0953e-05 | 368 | 6.7201e-06 | 537 | 5.7164e-06 | 5.7164e-06
316.6648 | 7.2638e-06 | 233.6 | 4.2123e-05 | 267.4 | 2.0978e-05 | 368.5 | 6.7159e-06 | 538 | 5.7315e-06 | 5.7315e-06
317.0095 | 7.2495e-06 | 233.7 | 4.1749e-05 | 267.6 | 2.1012e-05 | 369 | 6.7116e-06 | 539 | 5.7475e-06 | 5.7475e-06
317.3542 | 7.23e-06 | 233.8 | 4.141e-05 | 267.8 | 2.1046e-05 | 369.5 | 6.7103e-06 | 540 | 5.7632e-06 | 5.7632e-06
317.6988 | 7.2104e-06 | 233.9 | 4.1054e-05 | 268 | 2.1072e-05 | 370 | 6.7116e-06 | 541 | 5.7787e-06 | 5.7787e-06
318.0433 | 7.1955e-06 | 234 | 4.0698e-05 | 268.2 | 2.11e-05 | 370.5 | 6.7191e-06 | 542 | 5.7952e-06 | 5.7952e-06
318.3877 | 7.1696e-06 | 234.1 | 4.0342e-05 | 268.4 | 2.1113e-05 | 371 | 6.7233e-06 | 543 | 5.8109e-06 | 5.8109e-06
318.7321 | 7.1466e-06 | 234.2 | 4.0055e-05 | 268.6 | 2.1122e-05 | 371.5 | 6.7231e-06 | 544 | 5.8266e-06 | 5.8266e-06
319.0763 | 7.1377e-06 | 234.3 | 3.9721e-05 | 268.8 | 2.1134e-05 | 372 | 6.7165e-06 | 545 | 5.842e-06 | 5.842e-06
319.4205 | 7.1131e-06 | 234.4 | 3.9388e-05 | 269 | 2.1147e-05 | 372.5 | 6.6971e-06 | 546 | 5.8589e-06 | 5.8589e-06
319.7646 | 7.0965e-06 | 234.5 | 3.9111e-05 | 269.2 | 2.1168e-05 | 373 | 6.6743e-06 | 547 | 5.8761e-06 | 5.8761e-06
320.1086 | 7.0766e-06 | 234.6 | 3.8807e-05 | 269.4 | 2.1185e-05 | 373.5 | 6.6352e-06 | 548 | 5.8941e-06 | 5.8941e-06
320.4526 | 7.0551e-06 | 234.7 | 3.8503e-05 | 269.6 | 2.1204e-05 | 374 | 6.5914e-06 | 549 | 5.9118e-06 | 5.9118e-06
320.7964 | 7.0272e-06 | 234.8 | 3.8199e-05 | 269.8 | 2.1222e-05 | 374.5 | 6.5533e-06 | 550 | 5.9293e-06 | 5.9293e-06
321.1402 | 6.9993e-06 | 234.9 | 3.7931e-05 | 270 | 2.1239e-05 | 375 | 6.5225e-06 | 551 | 5.948e-06 | 5.948e-06
321.4839 | 6.9796e-06 | 235 | 3.7643e-05 | 270.2 | 2.1254e-05 | 375.5 | 6.4994e-06 | 552 | 5.9667e-06 | 5.9667e-06
321.8275 | 6.9598e-06 | 235.1 | 3.7355e-05 | 270.4 | 2.1272e-05 | 376 | 6.4775e-06 | 553 | 5.9856e-06 | 5.9856e-06
322.1711 | 6.9457e-06 | 235.2 | 3.7067e-05 | 270.6 | 2.129e-05 | 376.5 | 6.456e-06 | 554 | 6.0054e-06 | 6.0054e-06
322.5145 | 6.9269e-06 | 235.3 | 3.6796e-05 | 270.8 | 2.1307e-05 | 377 | 6.4352e-06 | 555 | 6.0256e-06 | 6.0256e-06
322.8579 | 6.9018e-06 | 235.4 | 3.6515e-05 | 271 | 2.1325e-05 | 377.5 | 6.4143e-06 | 556 | 6.0455e-06 | 6.0455e-06
323.2012 | 6.8914e-06 | 235.5 | 3.6234e-05 | 271.2 | 2.134e-05 | 378 | 6.3932e-06 | 557 | 6.0645e-06 | 6.0645e-06
323.5444 | 6.8738e-06 | 235.6 | 3.5981e-05 | 271.4 | 2.1355e-05 | 378.5 | 6.3713e-06 | 558 | 6.0849e-06 | 6.0849e-06
323.8876 | 6.856e-06 | 235.7 | 3.5711e-05 | 271.6 | 2.1352e-05 | 379 | 6.3477e-06 | 559 | 6.1059e-06 | 6.1059e-06
324.2307 | 6.8392e-06 | 235.8 | 3.5441e-05 | 271.8 | 2.135e-05 | 379.5 | 6.3248e-06 | 560 | 6.1269e-06 | 6.1269e-06
324.5736 | 6.8333e-06 | 235.9 | 3.5171e-05 | 272 | 2.1355e-05 | 380 | 6.3011e-06 | 561 | 6.148e-06 | 6.148e-06
324.9165 | 6.8272e-06 | 236 | 3.4945e-05 | 272.2 | 2.1359e-05 | 380.5 | 6.2744e-06 | 562 | 6.1694e-06 | 6.1694e-06
325.2594 | 6.8162e-06 | 236.1 | 3.4692e-05 | 272.4 | 2.1375e-05 | 381 | 6.2479e-06 | 563 | 6.1906e-06 | 6.1906e-06
325.6021 | 6.8037e-06 | 236.2 | 3.444e-05 | 272.6 | 2.1385e-05 | 381.5 | 6.2223e-06 | 564 | 6.2111e-06 | 6.2111e-06
325.9448 | 6.7903e-06 | 236.3 | 3.4223e-05 | 272.8 | 2.1395e-05 | 382 | 6.1976e-06 | 565 | 6.2316e-06 | 6.2316e-06
326.2874 | 6.7886e-06 | 236.4 | 3.3989e-05 | 273 | 2.139e-05 | 382.5 | 6.1746e-06 | 566 | 6.2511e-06 | 6.2511e-06
326.6298 | 6.7839e-06 | 236.5 | 3.3754e-05 | 273.2 | 2.1378e-05 | 383 | 6.1502e-06 | 567 | 6.2704e-06 | 6.2704e-06
326.9723 | 6.7724e-06 | 236.6 | 3.352e-05 | 273.4 | 2.1359e-05 | 383.5 | 6.123e-06 | 568 | 6.2907e-06 | 6.2907e-06
327.3146 | 6.7674e-06 | 236.7 | 3.3308e-05 | 273.6 | 2.1341e-05 | 384 | 6.0999e-06 | 569 | 6.313e-06 | 6.313e-06
327.6569 | 6.7675e-06 | 236.8 | 3.3086e-05 | 273.8 | 2.1324e-05 | 384.5 | 6.0788e-06 | 570 | 6.3364e-06 | 6.3364e-06
327.9991 | 6.7656e-06 | 236.9 | 3.2864e-05 | 274 | 2.1309e-05 | 385 | 6.0552e-06 | 571 | 6.3599e-06 | 6.3599e-06
328.3412 | 6.75e-06 | 237 | 3.2642e-05 | 274.2 | 2.1296e-05 | 385.5 | 6.0307e-06 | 572 | 6.3828e-06 | 6.3828e-06
328.6832 | 6.7313e-06 | 237.1 | 3.2445e-05 | 274.4 | 2.1275e-05 | 386 | 6.0066e-06 | 573 | 6.4051e-06 | 6.4051e-06
329.0251 | 6.7136e-06 | 237.2 | 3.2232e-05 | 274.6 | 2.1253e-05 | 386.5 | 5.9843e-06 | 574 | 6.4271e-06 | 6.4271e-06
329.367 | 6.7037e-06 | 237.3 | 3.2018e-05 | 274.8 | 2.1231e-05 | 387 | 5.9642e-06 | 575 | 6.4504e-06 | 6.4504e-06
329.7088 | 6.7034e-06 | 237.4 | 3.183e-05 | 275 | 2.1209e-05 | 387.5 | 5.9431e-06 | 576 | 6.4739e-06 | 6.4739e-06
330.0505 | 6.694e-06 | 237.5 | 3.163e-05 | 275.2 | 2.1191e-05 | 388 | 5.9206e-06 | 577 | 6.4977e-06 | 6.4977e-06
330.3921 | 6.6869e-06 | 237.6 | 3.1429e-05 | 275.4 | 2.1173e-05 | 388.5 | 5.9015e-06 | 578 | 6.5219e-06 | 6.5219e-06
330.7337 | 6.6808e-06 | 237.7 | 3.1228e-05 | 275.6 | 2.1158e-05 | 389 | 5.8819e-06 | 579 | 6.5467e-06 | 6.5467e-06
331.0751 | 6.6716e-06 | 237.8 | 3.1061e-05 | 275.8 | 2.1136e-05 | 389.5 | 5.8614e-06 | 580 | 6.5712e-06 | 6.5712e-06
331.4165 | 6.6584e-06 | 237.9 | 3.0874e-05 | 276 | 2.1113e-05 | 390 | 5.8399e-06 | 581 | 6.5964e-06 | 6.5964e-06
331.7578 | 6.6441e-06 | 238 | 3.0688e-05 | 276.2 | 2.1084e-05 | 390.5 | 5.8192e-06 | 582 | 6.6229e-06 | 6.6229e-06
332.099 | 6.6394e-06 | 238.1 | 3.0501e-05 | 276.4 | 2.1052e-05 | 391 | 5.8017e-06 | 583 | 6.6515e-06 | 6.6515e-06
332.4402 | 6.6426e-06 | 238.2 | 3.0343e-05 | 276.6 | 2.102e-05 | 391.5 | 5.7814e-06 | 584 | 6.6802e-06 | 6.6802e-06
332.7813 | 6.6417e-06 | 238.3 | 3.0168e-05 | 276.8 | 2.0985e-05 | 392 | 5.7591e-06 | 585 | 6.7102e-06 | 6.7102e-06
333.1223 | 6.6414e-06 | 238.4 | 2.9993e-05 | 277 | 2.0947e-05 | 392.5 | 5.7386e-06 | 586 | 6.7431e-06 | 6.7431e-06
333.4632 | 6.6323e-06 | 238.5 | 2.9839e-05 | 277.2 | 2.0913e-05 | 393 | 5.7192e-06 | 587 | 6.7795e-06 | 6.7795e-06
333.804 | 6.6249e-06 | 238.6 | 2.9673e-05 | 277.4 | 2.0878e-05 | 393.5 | 5.7003e-06 | 588 | 6.8178e-06 | 6.8178e-06
334.1447 | 6.6181e-06 | 238.7 | 2.9507e-05 | 277.6 | 2.0833e-05 | 394 | 5.6816e-06 | 589 | 6.8578e-06 | 6.8578e-06
334.4854 | 6.613e-06 | 238.8 | 2.9342e-05 | 277.8 | 2.0787e-05 | 394.5 | 5.6611e-06 | 590 | 6.8982e-06 | 6.8982e-06
334.826 | 6.6097e-06 | 238.9 | 2.9193e-05 | 278 | 2.0739e-05 | 395 | 5.6417e-06 | 591 | 6.9378e-06 | 6.9378e-06
335.1665 | 6.6008e-06 | 239 | 2.9035e-05 | 278.2 | 2.069e-05 | 395.5 | 5.623e-06 | 592 | 6.9731e-06 | 6.9731e-06
335.507 | 6.6013e-06 | 239.1 | 2.8877e-05 | 278.4 | 2.0643e-05 | 396 | 5.6033e-06 | 593 | 7.0045e-06 | 7.0045e-06
335.8473 | 6.6029e-06 | 239.2 | 2.8734e-05 | 278.6 | 2.0588e-05 | 396.5 | 5.5848e-06 | 594 | 7.0314e-06 | 7.0314e-06
336.1876 | 6.5925e-06 | 239.3 | 2.8583e-05 | 278.8 | 2.0517e-05 | 397 | 5.5668e-06 | 595 | 7.0544e-06 | 7.0544e-06
336.5278 | 6.5911e-06 | 239.4 | 2.8432e-05 | 279 | 2.0452e-05 | 397.5 | 5.5495e-06 | 596 | 7.0726e-06 | 7.0726e-06
336.8679 | 6.591e-06 | 239.5 | 2.8281e-05 | 279.2 | 2.0382e-05 | 398 | 5.5334e-06 | 597 | 7.0834e-06 | 7.0834e-06
337.2079 | 6.5871e-06 | 239.6 | 2.8146e-05 | 279.4 | 2.0316e-05 | 398.5 | 5.5177e-06 | 598 | 7.0871e-06 | 7.0871e-06
337.5479 | 6.5829e-06 | 239.7 | 2.8002e-05 | 279.6 | 2.0252e-05 | 399 | 5.5023e-06 | 599 | 7.0823e-06 | 7.0823e-06
337.8878 | 6.5803e-06 | 239.8 | 2.7859e-05 | 279.8 | 2.0191e-05 | 399.5 | 5.4872e-06 | 600 | 7.0694e-06 | 7.0694e-06
338.2276 | 6.5757e-06 | 239.9 | 2.7716e-05 | 280 | 2.0123e-05 | 400 | 5.4708e-06 | 601 | 7.0475e-06 | 7.0475e-06
338.5673 | 6.5714e-06 | 240 | 2.7603e-05 | 280.2 | 2.0048e-05 | 400.5 | 5.4535e-06 | 602 | 7.0163e-06 | 7.0163e-06
338.9069 | 6.573e-06 | 240.1 | 2.747e-05 | 280.4 | 1.9971e-05 | 401 | 5.4362e-06 | 603 | 6.9761e-06 | 6.9761e-06
339.2465 | 6.5729e-06 | 240.2 | 2.7337e-05 | 280.6 | 1.989e-05 | 401.5 | 5.4187e-06 | 604 | 6.9278e-06 | 6.9278e-06
339.586 | 6.5666e-06 | 240.3 | 2.7219e-05 | 280.8 | 1.9806e-05 | 402 | 5.4016e-06 | 605 | 6.8765e-06 | 6.8765e-06
339.9254 | 6.5684e-06 | 240.4 | 2.7096e-05 | 281 | 1.9719e-05 | 402.5 | 5.3847e-06 | 606 | 6.8254e-06 | 6.8254e-06
340.2647 | 6.5716e-06 | 240.5 | 2.6972e-05 | 281.2 | 1.9624e-05 | 403 | 5.3685e-06 | 607 | 6.7748e-06 | 6.7748e-06
340.6039 | 6.5688e-06 | 240.6 | 2.6849e-05 | 281.4 | 1.953e-05 | 403.5 | 5.353e-06 | 608 | 6.7267e-06 | 6.7267e-06
340.9431 | 6.5727e-06 | 240.7 | 2.6735e-05 | 281.6 | 1.9433e-05 | 404 | 5.3365e-06 | 609 | 6.6828e-06 | 6.6828e-06
341.2822 | 6.5779e-06 | 240.8 | 2.6616e-05 | 281.8 | 1.9329e-05 | 404.5 | 5.319e-06 | 610 | 6.6433e-06 | 6.6433e-06
341.6212 | 6.5804e-06 | 240.9 | 2.6496e-05 | 282 | 1.9223e-05 | 405 | 5.3014e-06 | 611 | 6.609e-06 | 6.609e-06
341.9601 | 6.5884e-06 | 241 | 2.6377e-05 | 282.2 | 1.9122e-05 | 405.5 | 5.2839e-06 | 612 | 6.5802e-06 | 6.5802e-06
342.2989 | 6.5971e-06 | 241.1 | 2.6282e-05 | 282.4 | 1.9021e-05 | 406 | 5.266e-06 | 613 | 6.5584e-06 | 6.5584e-06
342.6377 | 6.5987e-06 | 241.2 | 2.6171e-05 | 282.6 | 1.8926e-05 | 406.5 | 5.2467e-06 | 614 | 6.5415e-06 | 6.5415e-06
342.9764 | 6.5964e-06 | 241.3 | 2.6059e-05 | 282.8 | 1.8825e-05 | 407 | 5.2274e-06 | 615 | 6.5297e-06 | 6.5297e-06
343.315 | 6.5977e-06 | 241.4 | 2.5964e-05 | 283 | 1.8714e-05 | 407.5 | 5.2099e-06 | 616 | 6.5228e-06 | 6.5228e-06
343.6535 | 6.6011e-06 | 241.5 | 2.5861e-05 | 283.2 | 1.8605e-05 | 408 | 5.1936e-06 | 617 | 6.5224e-06 | 6.5224e-06
343.992 | 6.5997e-06 | 241.6 | 2.5759e-05 | 283.4 | 1.8491e-05 | 408.5 | 5.1767e-06 | 618 | 6.5291e-06 | 6.5291e-06
344.3304 | 6.6e-06 | 241.7 | 2.5656e-05 | 283.6 | 1.837e-05 | 409 | 5.1586e-06 | 619 | 6.5462e-06 | 6.5462e-06
344.6687 | 6.6017e-06 | 241.8 | 2.5565e-05 | 283.8 | 1.8249e-05 | 409.5 | 5.1401e-06 | 620 | 6.5754e-06 | 6.5754e-06
345.0069 | 6.598e-06 | 241.9 | 2.5468e-05 | 284 | 1.8129e-05 | 410 | 5.1228e-06 | 621 | 6.6154e-06 | 6.6154e-06
345.345 | 6.6026e-06 | 242 | 2.537e-05 | 284.2 | 1.8005e-05 | 410.5 | 5.1063e-06 | 622 | 6.6664e-06 | 6.6664e-06
345.6831 | 6.5982e-06 | 242.1 | 2.5278e-05 | 284.4 | 1.7878e-05 | 411 | 5.0901e-06 | 623 | 6.7252e-06 | 6.7252e-06
346.0211 | 6.5889e-06 | 242.2 | 2.5184e-05 | 284.6 | 1.7754e-05 | 411.5 | 5.0749e-06 | 624 | 6.7889e-06 | 6.7889e-06
346.359 | 6.593e-06 | 242.3 | 2.5089e-05 | 284.8 | 1.7632e-05 | 412 | 5.0602e-06 | 625 | 6.8545e-06 | 6.8545e-06
346.6968 | 6.5977e-06 | 242.4 | 2.4995e-05 | 285 | 1.7508e-05 | 412.5 | 5.0455e-06 | 626 | 6.9234e-06 | 6.9234e-06
347.0345 | 6.5976e-06 | 242.5 | 2.4912e-05 | 285.2 | 1.7385e-05 | 413 | 5.0309e-06 | 627 | 6.9945e-06 | 6.9945e-06
347.3722 | 6.5981e-06 | 242.6 | 2.4823e-05 | 285.4 | 1.726e-05 | 413.5 | 5.0164e-06 | 628 | 7.0637e-06 | 7.0637e-06
347.7098 | 6.6068e-06 | 242.7 | 2.4733e-05 | 285.6 | 1.7135e-05 | 414 | 5.0016e-06 | 629 | 7.1263e-06 | 7.1263e-06
348.0473 | 6.6081e-06 | 242.8 | 2.4643e-05 | 285.8 | 1.7014e-05 | 414.5 | 4.9869e-06 | 630 | 7.1809e-06 | 7.1809e-06
348.3847 | 6.6161e-06 | 242.9 | 2.4562e-05 | 286 | 1.6887e-05 | 415 | 4.9709e-06 | 631 | 7.2286e-06 | 7.2286e-06
348.7221 | 6.6219e-06 | 243 | 2.4476e-05 | 286.2 | 1.6753e-05 | 415.5 | 4.9553e-06 | 632 | 7.2675e-06 | 7.2675e-06
349.0593 | 6.6208e-06 | 243.1 | 2.4391e-05 | 286.4 | 1.6619e-05 | 416 | 4.9407e-06 | 633 | 7.2951e-06 | 7.2951e-06
349.3965 | 6.6214e-06 | 243.2 | 2.4316e-05 | 286.6 | 1.6482e-05 | 416.5 | 4.9258e-06 | 634 | 7.3088e-06 | 7.3088e-06
349.7336 | 6.6308e-06 | 243.3 | 2.4234e-05 | 286.8 | 1.6352e-05 | 417 | 4.9106e-06 | 635 | 7.3155e-06 | 7.3155e-06
350.0707 | 6.635e-06 | 243.4 | 2.4153e-05 | 287 | 1.6223e-05 | 417.5 | 4.8968e-06 | 636 | 7.3188e-06 | 7.3188e-06
350.4076 | 6.6451e-06 | 243.5 | 2.4072e-05 | 287.2 | 1.609e-05 | 418 | 4.8836e-06 | 637 | 7.317e-06 | 7.317e-06
350.7445 | 6.6573e-06 | 243.6 | 2.3998e-05 | 287.4 | 1.5957e-05 | 418.5 | 4.8712e-06 | 638 | 7.3035e-06 | 7.3035e-06
351.0813 | 6.6654e-06 | 243.7 | 2.3921e-05 | 287.6 | 1.5818e-05 | 419 | 4.8588e-06 | 639 | 7.278e-06 | 7.278e-06
351.418 | 6.6766e-06 | 243.8 | 2.3843e-05 | 287.8 | 1.5683e-05 | 419.5 | 4.8467e-06 | 640 | 7.2413e-06 | 7.2413e-06
351.7547 | 6.6929e-06 | 243.9 | 2.3766e-05 | 288 | 1.5549e-05 | 420 | 4.836e-06 | 641 | 7.1922e-06 | 7.1922e-06
352.0912 | 6.6955e-06 | 244 | 2.3699e-05 | 288.2 | 1.5417e-05 | 420.5 | 4.8268e-06 | 642 | 7.1347e-06 | 7.1347e-06
352.4277 | 6.7084e-06 | 244.1 | 2.3625e-05 | 288.4 | 1.5285e-05 | 421 | 4.8167e-06 | 643 | 7.081e-06 | 7.081e-06
352.7641 | 6.7223e-06 | 244.2 | 2.3551e-05 | 288.6 | 1.5153e-05 | 421.5 | 4.8059e-06 | 644 | 7.0354e-06 | 7.0354e-06
353.1004 | 6.7392e-06 | 244.3 | 2.3483e-05 | 288.8 | 1.5022e-05 | 422 | 4.7953e-06 | 645 | 6.9919e-06 | 6.9919e-06
353.4367 | 6.7509e-06 | 244.4 | 2.3412e-05 | 289 | 1.4891e-05 | 422.5 | 4.7863e-06 | 646 | 6.9484e-06 | 6.9484e-06
353.7729 | 6.7642e-06 | 244.5 | 2.3341e-05 | 289.2 | 1.4762e-05 | 423 | 4.7809e-06 | 647 | 6.9073e-06 | 6.9073e-06
354.1089 | 6.7667e-06 | 244.6 | 2.327e-05 | 289.4 | 1.4635e-05 | 423.5 | 4.7862e-06 | 648 | 6.8732e-06 | 6.8732e-06
354.445 | 6.7745e-06 | 244.7 | 2.3208e-05 | 289.6 | 1.4506e-05 | 424 | 4.8049e-06 | 649 | 6.8424e-06 | 6.8424e-06
354.7809 | 6.7836e-06 | 244.8 | 2.3141e-05 | 289.8 | 1.4379e-05 | 424.5 | 4.8375e-06 | 650 | 6.812e-06 | 6.812e-06
355.1168 | 6.7875e-06 | 244.9 | 2.3073e-05 | 290 | 1.4256e-05 | 425 | 4.8903e-06 | 651 | 6.7815e-06 | 6.7815e-06
355.4525 | 6.7954e-06 | 245 | 2.3013e-05 | 290.2 | 1.4133e-05 | 425.5 | 4.9576e-06 | 652 | 6.7544e-06 | 6.7544e-06
355.7882 | 6.803e-06 | 245.1 | 2.295e-05 | 290.4 | 1.4011e-05 | 426 | 5.0199e-06 | 653 | 6.7332e-06 | 6.7332e-06
356.1239 | 6.8143e-06 | 245.2 | 2.2887e-05 | 290.6 | 1.3891e-05 | 426.5 | 5.0722e-06 | 654 | 6.7159e-06 | 6.7159e-06
356.4594 | 6.814e-06 | 245.3 | 2.2823e-05 | 290.8 | 1.3769e-05 | 427 | 5.1085e-06 | 655 | 6.704e-06 | 6.704e-06
356.7948 | 6.8221e-06 | 245.4 | 2.2769e-05 | 291 | 1.3648e-05 | 427.5 | 5.1205e-06 | 656 | 6.6963e-06 | 6.6963e-06
357.1302 | 6.8214e-06 | 245.5 | 2.271e-05 | 291.2 | 1.3528e-05 | 428 | 5.1005e-06 | 657 | 6.6903e-06 | 6.6903e-06
357.4656 | 6.8178e-06 | 245.6 | 2.2651e-05 | 291.4 | 1.3413e-05 | 428.5 | 5.0512e-06 | 658 | 6.6832e-06 | 6.6832e-06
357.8008 | 6.8302e-06 | 245.7 | 2.2591e-05 | 291.6 | 1.3298e-05 | 429 | 4.9765e-06 | 659 | 6.6805e-06 | 6.6805e-06
358.1359 | 6.8342e-06 | 245.8 | 2.2541e-05 | 291.8 | 1.3189e-05 | 429.5 | 4.9009e-06 | 660 | 6.6827e-06 | 6.6827e-06
358.471 | 6.8425e-06 | 245.9 | 2.2485e-05 | 292 | 1.3077e-05 | 430 | 4.841e-06 | 661 | 6.6906e-06 | 6.6906e-06
358.806 | 6.8463e-06 | 246 | 2.2429e-05 | 292.2 | 1.2966e-05 | 430.5 | 4.8032e-06 | 662 | 6.6966e-06 | 6.6966e-06
359.1409 | 6.8377e-06 | 246.1 | 2.2377e-05 | 292.4 | 1.2855e-05 | 431 | 4.7842e-06 | 663 | 6.7055e-06 | 6.7055e-06
359.4758 | 6.8325e-06 | 246.2 | 2.2324e-05 | 292.6 | 1.2746e-05 | 431.5 | 4.7783e-06 | 664 | 6.7195e-06 | 6.7195e-06
359.8105 | 6.8235e-06 | 246.3 | 2.227e-05 | 292.8 | 1.264e-05 | 432 | 4.7861e-06 | 665 | 6.7377e-06 | 6.7377e-06
360.1452 | 6.8223e-06 | 246.4 | 2.2217e-05 | 293 | 1.2533e-05 | 432.5 | 4.8093e-06 | 666 | 6.7568e-06 | 6.7568e-06
360.4798 | 6.8188e-06 | 246.5 | 2.2163e-05 | 293.2 | 1.2426e-05 | 433 | 4.8463e-06 | 667 | 6.7774e-06 | 6.7774e-06
360.8143 | 6.8212e-06 | 246.6 | 2.211e-05 | 293.4 | 1.2322e-05 | 433.5 | 4.8824e-06 | 668 | 6.8016e-06 | 6.8016e-06
361.1488 | 6.8226e-06 | 246.7 | 2.2056e-05 | 293.6 | 1.2216e-05 | 434 | 4.9066e-06 | 669 | 6.8256e-06 | 6.8256e-06
361.4831 | 6.8271e-06 | 246.8 | 2.2008e-05 | 293.8 | 1.2116e-05 | 434.5 | 4.9239e-06 | 670 | 6.8505e-06 | 6.8505e-06
361.8174 | 6.8272e-06 | 246.9 | 2.1956e-05 | 294 | 1.2019e-05 | 435 | 4.9428e-06 | 671 | 6.8788e-06 | 6.8788e-06
362.1517 | 6.8216e-06 | 247 | 2.1905e-05 | 294.2 | 1.1921e-05 | 435.5 | 4.9603e-06 | 672 | 6.913e-06 | 6.913e-06
362.4858 | 6.8132e-06 | 247.1 | 2.1853e-05 | 294.4 | 1.1825e-05 | 436 | 4.9717e-06 | 673 | 6.9484e-06 | 6.9484e-06
362.8198 | 6.8139e-06 | 247.2 | 2.1806e-05 | 294.6 | 1.1728e-05 | 436.5 | 4.9775e-06 | 674 | 6.9803e-06 | 6.9803e-06
363.1538 | 6.8188e-06 | 247.3 | 2.1757e-05 | 294.8 | 1.1634e-05 | 437 | 4.9779e-06 | 675 | 7.0112e-06 | 7.0112e-06
363.4877 | 6.814e-06 | 247.4 | 2.1708e-05 | 295 | 1.154e-05 | 437.5 | 4.9731e-06 | 676 | 7.044e-06 | 7.044e-06
363.8216 | 6.8156e-06 | 247.5 | 2.1659e-05 | 295.2 | 1.1452e-05 | 438 | 4.9645e-06 | 677 | 7.079e-06 | 7.079e-06
364.1553 | 6.8153e-06 | 247.6 | 2.1625e-05 | 295.4 | 1.1366e-05 | 438.5 | 4.9545e-06 | 678 | 7.1106e-06 | 7.1106e-06
364.489 | 6.8133e-06 | 247.7 | 2.158e-05 | 295.6 | 1.1281e-05 | 439 | 4.9441e-06 | 679 | 7.1435e-06 | 7.1435e-06
364.8225 | 6.8107e-06 | 247.8 | 2.1536e-05 | 295.8 | 1.1196e-05 | 439.5 | 4.934e-06 | 680 | 7.184e-06 | 7.184e-06
365.1561 | 6.8031e-06 | 247.9 | 2.149e-05 | 296 | 1.1111e-05 | 440 | 4.9225e-06 | 681 | 7.2258e-06 | 7.2258e-06
365.4895 | 6.7991e-06 | 248 | 2.1447e-05 | 296.2 | 1.103e-05 | 440.5 | 4.9113e-06 | 682 | 7.2662e-06 | 7.2662e-06
365.8229 | 6.7868e-06 | 248.1 | 2.1405e-05 | 296.4 | 1.0951e-05 | 441 | 4.9001e-06 | 683 | 7.3062e-06 | 7.3062e-06
366.1561 | 6.7783e-06 | 248.2 | 2.1362e-05 | 296.6 | 1.0876e-05 | 441.5 | 4.8894e-06 | 684 | 7.3504e-06 | 7.3504e-06
366.4893 | 6.7722e-06 | 248.3 | 2.1316e-05 | 296.8 | 1.0804e-05 | 442 | 4.8798e-06 | 685 | 7.3889e-06 | 7.3889e-06
366.8225 | 6.7491e-06 | 248.4 | 2.1272e-05 | 297 | 1.0731e-05 | 442.5 | 4.8705e-06 | 686 | 7.4224e-06 | 7.4224e-06
367.1555 | 6.7401e-06 | 248.5 | 2.1227e-05 | 297.2 | 1.0659e-05 | 443 | 4.863e-06 | 687 | 7.4608e-06 | 7.4608e-06
367.4885 | 6.7326e-06 | 248.6 | 2.1182e-05 | 297.4 | 1.0586e-05 | 443.5 | 4.8571e-06 | 688 | 7.5049e-06 | 7.5049e-06
367.8214 | 6.7195e-06 | 248.7 | 2.1146e-05 | 297.6 | 1.0512e-05 | 444 | 4.8511e-06 | 689 | 7.5457e-06 | 7.5457e-06
368.1542 | 6.7188e-06 | 248.8 | 2.1104e-05 | 297.8 | 1.0439e-05 | 444.5 | 4.8452e-06 | 690 | 7.5798e-06 | 7.5798e-06
368.4869 | 6.7167e-06 | 248.9 | 2.1062e-05 | 298 | 1.0368e-05 | 445 | 4.84e-06 | 691 | 7.6166e-06 | 7.6166e-06
368.8196 | 6.7126e-06 | 249 | 2.1023e-05 | 298.2 | 1.0299e-05 | 445.5 | 4.835e-06 | 692 | 7.6615e-06 | 7.6615e-06
369.1521 | 6.7106e-06 | 249.1 | 2.0983e-05 | 298.4 | 1.0233e-05 | 446 | 4.8299e-06 | 693 | 7.7069e-06 | 7.7069e-06
369.4846 | 6.7076e-06 | 249.2 | 2.0943e-05 | 298.6 | 1.0168e-05 | 446.5 | 4.8229e-06 | 694 | 7.7486e-06 | 7.7486e-06
369.8171 | 6.7131e-06 | 249.3 | 2.0903e-05 | 298.8 | 1.0104e-05 | 447 | 4.8159e-06 | 695 | 7.7876e-06 | 7.7876e-06
370.1494 | 6.7089e-06 | 249.4 | 2.0875e-05 | 299 | 1.0041e-05 | 447.5 | 4.81e-06 | 696 | 7.8337e-06 | 7.8337e-06
370.4817 | 6.7202e-06 | 249.5 | 2.0839e-05 | 299.2 | 9.9775e-06 | 448 | 4.8045e-06 | 697 | 7.884e-06 | 7.884e-06
370.8139 | 6.7254e-06 | 249.6 | 2.0803e-05 | 299.4 | 9.915e-06 | 448.5 | 4.7988e-06 | 698 | 7.9342e-06 | 7.9342e-06
371.146 | 6.7214e-06 | 249.7 | 2.0772e-05 | 299.6 | 9.8545e-06 | 449 | 4.7935e-06 | 699 | 7.9806e-06 | 7.9806e-06
371.478 | 6.7243e-06 | 249.8 | 2.074e-05 | 299.8 | 9.7983e-06 | 449.5 | 4.789e-06 | 700 | 8.0263e-06 | 8.0263e-06
371.81 | 6.7226e-06 | 249.9 | 2.0708e-05 | 300 | 9.7406e-06 | 450 | 4.7851e-06 | 701 | 8.072e-06 | 8.072e-06
372.1419 | 6.7139e-06 | 250 | 2.0675e-05 | 300.2 | 9.6841e-06 | 450.5 | 4.7818e-06 | 702 | 8.1138e-06 | 8.1138e-06
372.4737 | 6.6983e-06 | 250.1 | 2.0644e-05 | 300.4 | 9.6275e-06 | 451 | 4.7787e-06 | 703 | 8.1508e-06 | 8.1508e-06
372.8054 | 6.6832e-06 | 250.2 | 2.0612e-05 | 300.6 | 9.5702e-06 | 451.5 | 4.7769e-06 | 704 | 8.1908e-06 | 8.1908e-06
373.137 | 6.6724e-06 | 250.3 | 2.058e-05 | 300.8 | 9.516e-06 | 452 | 4.7759e-06 | 705 | 8.2402e-06 | 8.2402e-06
373.4686 | 6.6365e-06 | 250.4 | 2.0548e-05 | 301 | 9.4646e-06 | 452.5 | 4.7743e-06 | 706 | 8.2933e-06 | 8.2933e-06
373.8001 | 6.6075e-06 | 250.5 | 2.0525e-05 | 301.2 | 9.4176e-06 | 453 | 4.7718e-06 | 707 | 8.3413e-06 | 8.3413e-06
374.1315 | 6.5801e-06 | 250.6 | 2.0496e-05 | 301.4 | 9.3702e-06 | 453.5 | 4.7701e-06 | 708 | 8.3829e-06 | 8.3829e-06
374.4629 | 6.5569e-06 | 250.7 | 2.0466e-05 | 301.6 | 9.3258e-06 | 454 | 4.7686e-06 | 709 | 8.4277e-06 | 8.4277e-06
374.7941 | 6.5309e-06 | 250.8 | 2.0438e-05 | 301.8 | 9.2815e-06 | 454.5 | 4.7673e-06 | 710 | 8.4834e-06 | 8.4834e-06
375.1253 | 6.516e-06 | 250.9 | 2.0411e-05 | 302 | 9.2373e-06 | 455 | 4.7652e-06 | 711 | 8.5476e-06 | 8.5476e-06
375.4564 | 6.5011e-06 | 251 | 2.0383e-05 | 302.2 | 9.1929e-06 | 455.5 | 4.7638e-06 | 712 | 8.6076e-06 | 8.6076e-06
375.7874 | 6.487e-06 | 251.1 | 2.0355e-05 | 302.4 | 9.1491e-06 | 456 | 4.7628e-06 | 713 | 8.661e-06 | 8.661e-06
376.1184 | 6.4721e-06 | 251.2 | 2.033e-05 | 302.6 | 9.1063e-06 | 456.5 | 4.7623e-06 | 714 | 8.7215e-06 | 8.7215e-06
376.4493 | 6.4579e-06 | 251.3 | 2.0303e-05 | 302.8 | 9.0618e-06 | 457 | 4.7622e-06 | 715 | 8.7946e-06 | 8.7946e-06
376.7801 | 6.4443e-06 | 251.4 | 2.0275e-05 | 303 | 9.0153e-06 | 457.5 | 4.7618e-06 | 716 | 8.873e-06 | 8.873e-06
377.1108 | 6.4305e-06 | 251.5 | 2.0249e-05 | 303.2 | 8.9705e-06 | 458 | 4.7617e-06 | 717 | 8.9436e-06 | 8.9436e-06
377.4414 | 6.4175e-06 | 251.6 | 2.0223e-05 | 303.4 | 8.925e-06 | 458.5 | 4.7627e-06 | 718 | 9.0088e-06 | 9.0088e-06
377.772 | 6.4023e-06 | 251.7 | 2.0196e-05 | 303.6 | 8.8818e-06 | 459 | 4.7645e-06 | 719 | 9.0798e-06 | 9.0798e-06
378.1025 | 6.3894e-06 | 251.8 | 2.017e-05 | 303.8 | 8.84e-06 | 459.5 | 4.7669e-06 | 720 | 9.1649e-06 | 9.1649e-06
378.4329 | 6.3738e-06 | 251.9 | 2.0151e-05 | 304 | 8.799e-06 | 460 | 4.7698e-06 | 721 | 9.2626e-06 | 9.2626e-06
378.7632 | 6.3599e-06 | 252 | 2.0127e-05 | 304.2 | 8.7581e-06 | 460.5 | 4.7709e-06 | 722 | 9.3558e-06 | 9.3558e-06
379.0935 | 6.3424e-06 | 252.1 | 2.0103e-05 | 304.4 | 8.7185e-06 | 461 | 4.7735e-06 | 723 | 9.4349e-06 | 9.4349e-06
379.4237 | 6.3279e-06 | 252.2 | 2.0079e-05 | 304.6 | 8.6816e-06 | 461.5 | 4.7776e-06 | 724 | 9.4953e-06 | 9.4953e-06
379.7538 | 6.3139e-06 | 252.3 | 2.0064e-05 | 304.8 | 8.6451e-06 | 462 | 4.7821e-06 | 725 | 9.5542e-06 | 9.5542e-06
380.0838 | 6.298e-06 | 252.4 | 2.0044e-05 | 305 | 8.6077e-06 | 462.5 | 4.7863e-06 | 726 | 9.6149e-06 | 9.6149e-06
380.4137 | 6.2785e-06 | 252.5 | 2.0023e-05 | 305.2 | 8.5728e-06 | 463 | 4.7908e-06 | 727 | 9.6747e-06 | 9.6747e-06
380.7436 | 6.2615e-06 | 252.6 | 2.0005e-05 | 305.4 | 8.5406e-06 | 463.5 | 4.7952e-06 | 728 | 9.7311e-06 | 9.7311e-06
381.0734 | 6.2435e-06 | 252.7 | 1.9986e-05 | 305.6 | 8.5084e-06 | 464 | 4.8011e-06 | 729 | 9.7833e-06 | 9.7833e-06
381.4031 | 6.2278e-06 | 252.8 | 1.9967e-05 | 305.8 | 8.4781e-06 | 464.5 | 4.8065e-06 | 730 | 9.8298e-06 | 9.8298e-06
381.7328 | 6.2099e-06 | 252.9 | 1.9948e-05 | 306 | 8.4462e-06 | 465 | 4.8109e-06 | 731 | 9.8681e-06 | 9.8681e-06
382.0623 | 6.1947e-06 | 253 | 1.9938e-05 | 306.2 | 8.4138e-06 | 465.5 | 4.8146e-06 | 732 | 9.909e-06 | 9.909e-06
382.3918 | 6.1789e-06 | 253.1 | 1.9922e-05 | 306.4 | 8.3816e-06 | 466 | 4.8178e-06 | 733 | 9.9593e-06 | 9.9593e-06
382.7212 | 6.165e-06 | 253.2 | 1.9906e-05 | 306.6 | 8.3511e-06 | 466.5 | 4.8201e-06 | 734 | 1.0018e-05 | 1.0018e-05
383.0506 | 6.1482e-06 | 253.3 | 1.9884e-05 | 306.8 | 8.3247e-06 | 467 | 4.8226e-06 | 735 | 1.0076e-05 | 1.0076e-05
383.3798 | 6.1296e-06 | 253.4 | 1.9867e-05 | 307 | 8.2969e-06 | 467.5 | 4.8254e-06 | 736 | 1.0131e-05 | 1.0131e-05
383.709 | 6.1115e-06 | 253.5 | 1.985e-05 | 307.2 | 8.2702e-06 | 468 | 4.8284e-06 | 737 | 1.0194e-05 | 1.0194e-05
384.0381 | 6.098e-06 | 253.6 | 1.9833e-05 | 307.4 | 8.2396e-06 | 468.5 | 4.8323e-06 | 738 | 1.0266e-05 | 1.0266e-05
384.3671 | 6.0853e-06 | 253.7 | 1.9814e-05 | 307.6 | 8.2091e-06 | 469 | 4.8369e-06 | 739 | 1.0332e-05 | 1.0332e-05
384.6961 | 6.0702e-06 | 253.8 | 1.9794e-05 | 307.8 | 8.1814e-06 | 469.5 | 4.8421e-06 | 740 | 1.0385e-05 | 1.0385e-05
385.0249 | 6.054e-06 | 253.9 | 1.9775e-05 | 308 | 8.1535e-06 | 470 | 4.8482e-06 | 741 | 1.0437e-05 | 1.0437e-05
385.3537 | 6.0378e-06 | 254 | 1.9755e-05 | 308.2 | 8.1286e-06 | 470.5 | 4.8548e-06 | 742 | 1.0505e-05 | 1.0505e-05
385.6825 | 6.0219e-06 | 254.1 | 1.9744e-05 | 308.4 | 8.1011e-06 | 471 | 4.8613e-06 | 743 | 1.0588e-05 | 1.0588e-05
386.0111 | 6.0059e-06 | 254.2 | 1.9728e-05 | 308.6 | 8.0732e-06 | 471.5 | 4.8674e-06 | 744 | 1.0665e-05 | 1.0665e-05
386.3397 | 5.9906e-06 | 254.3 | 1.9711e-05 | 308.8 | 8.0466e-06 | 472 | 4.8734e-06 | 745 | 1.0725e-05 | 1.0725e-05
386.6682 | 5.9772e-06 | 254.4 | 1.9697e-05 | 309 | 8.0202e-06 | 472.5 | 4.8796e-06 | 746 | 1.0789e-05 | 1.0789e-05
386.9966 | 5.9642e-06 | 254.5 | 1.9683e-05 | 309.2 | 7.9952e-06 | 473 | 4.8859e-06 | 747 | 1.0872e-05 | 1.0872e-05
387.3249 | 5.9515e-06 | 254.6 | 1.9668e-05 | 309.4 | 7.9697e-06 | 473.5 | 4.8919e-06 | 748 | 1.0965e-05 | 1.0965e-05
387.6532 | 5.936e-06 | 254.7 | 1.9653e-05 | 309.6 | 7.9436e-06 | 474 | 4.8984e-06 | 749 | 1.1047e-05 | 1.1047e-05
387.9813 | 5.9207e-06 | 254.8 | 1.9651e-05 | 309.8 | 7.9178e-06 | 474.5 | 4.9052e-06 | 750 | 1.1117e-05 | 1.1117e-05
388.3094 | 5.9079e-06 | 254.9 | 1.9641e-05 | 310 | 7.8924e-06 | 475 | 4.9119e-06 | 751 | 1.1191e-05 | 1.1191e-05
388.6375 | 5.8973e-06 | 255 | 1.963e-05 | 310.2 | 7.8689e-06 | 475.5 | 4.9191e-06 | 752 | 1.1274e-05 | 1.1274e-05
388.9654 | 5.8835e-06 | 255.1 | 1.9626e-05 | 310.4 | 7.8445e-06 | 476 | 4.9262e-06 | 753 | 1.1354e-05 | 1.1354e-05
389.2933 | 5.8693e-06 | 255.2 | 1.962e-05 | 310.6 | 7.8197e-06 | 476.5 | 4.9338e-06 | 754 | 1.1425e-05 | 1.1425e-05
389.6211 | 5.857e-06 | 255.3 | 1.9614e-05 | 310.8 | 7.7968e-06 | 477 | 4.9415e-06 | 755 | 1.1506e-05 | 1.1506e-05
389.9488 | 5.8421e-06 | 255.4 | 1.9608e-05 | 311 | 7.7736e-06 | 477.5 | 4.9487e-06 | 756 | 1.1606e-05 | 1.1606e-05
390.2765 | 5.8272e-06 | 255.5 | 1.961e-05 | 311.2 | 7.7533e-06 | 478 | 4.9561e-06 | 757 | 1.1699e-05 | 1.1699e-05
390.604 | 5.8151e-06 | 255.6 | 1.9607e-05 | 311.4 | 7.7336e-06 | 478.5 | 4.9635e-06 | 758 | 1.1784e-05 | 1.1784e-05
390.9315 | 5.8046e-06 | 255.7 | 1.9604e-05 | 311.6 | 7.7128e-06 | 479 | 4.9709e-06 | 759 | 1.1871e-05 | 1.1871e-05
391.2589 | 5.7925e-06 | 255.8 | 1.96e-05 | 311.8 | 7.6919e-06 | 479.5 | 4.9792e-06 | 760 | 1.1972e-05 | 1.1972e-05
391.5863 | 5.7776e-06 | 255.9 | 1.9598e-05 | 312 | 7.6693e-06 | 480 | 4.9874e-06 | 761 | 1.2073e-05 | 1.2073e-05
391.9135 | 5.7628e-06 | 256 | 1.9596e-05 | 312.2 | 7.6484e-06 | 480.5 | 4.9953e-06 | 762 | 1.2165e-05 | 1.2165e-05
392.2407 | 5.7483e-06 | 256.1 | 1.9593e-05 | 312.4 | 7.6277e-06 | 481 | 5.0023e-06 | 763 | 1.2247e-05 | 1.2247e-05
392.5678 | 5.7359e-06 | 256.2 | 1.9593e-05 | 312.6 | 7.6079e-06 | 481.5 | 5.0101e-06 | 764 | 1.2336e-05 | 1.2336e-05
392.8949 | 5.7241e-06 | 256.3 | 1.9591e-05 | 312.8 | 7.5875e-06 | 482 | 5.0181e-06 | 765 | 1.2438e-05 | 1.2438e-05
393.2218 | 5.7099e-06 | 256.4 | 1.9589e-05 | 313 | 7.5645e-06 | 482.5 | 5.0259e-06 | 766 | 1.2537e-05 | 1.2537e-05
393.5487 | 5.6987e-06 | 256.5 | 1.9587e-05 | 313.2 | 7.5437e-06 | 483 | 5.034e-06 | 767 | 1.2626e-05 | 1.2626e-05
393.8755 | 5.6866e-06 | 256.6 | 1.9589e-05 | 313.4 | 7.5245e-06 | 483.5 | 5.0424e-06 | 768 | 1.2724e-05 | 1.2724e-05
394.2022 | 5.6741e-06 | 256.7 | 1.9589e-05 | 313.6 | 7.5065e-06 | 484 | 5.0512e-06 | 769 | 1.2844e-05 | 1.2844e-05
394.5289 | 5.6593e-06 | 256.8 | 1.9588e-05 | 313.8 | 7.4901e-06 | 484.5 | 5.0599e-06 | 770 | 1.2958e-05 | 1.2958e-05
394.8555 | 5.6472e-06 | 256.9 | 1.9592e-05 | 314 | 7.4701e-06 | 485 | 5.0678e-06 | 771 | 1.3049e-05 | 1.3049e-05
395.182 | 5.6347e-06 | 257 | 1.9594e-05 | 314.2 | 7.4501e-06 | 485.5 | 5.0762e-06 | 772 | 1.3141e-05 | 1.3141e-05
395.5084 | 5.6232e-06 | 257.1 | 1.9595e-05 | 314.4 | 7.4311e-06 | 486 | 5.0848e-06 | 773 | 1.3258e-05 | 1.3258e-05
395.8347 | 5.6095e-06 | 257.2 | 1.9596e-05 | 314.6 | 7.4111e-06 | 486.5 | 5.093e-06 | 774 | 1.3371e-05 | 1.3371e-05
396.161 | 5.5969e-06 | 257.3 | 1.9603e-05 | 314.8 | 7.3918e-06 | 487 | 5.0999e-06 | 775 | 1.3469e-05 | 1.3469e-05
396.4872 | 5.5851e-06 | 257.4 | 1.9606e-05 | 315 | 7.3737e-06 | 487.5 | 5.1068e-06 | 776 | 1.3582e-05 | 1.3582e-05
396.8133 | 5.5738e-06 | 257.5 | 1.961e-05 | 315.2 | 7.3572e-06 | 488 | 5.1142e-06 | 777 | 1.3717e-05 | 1.3717e-05
397.1394 | 5.5613e-06 | 257.6 | 1.9612e-05 | 315.4 | 7.3442e-06 | 488.5 | 5.1221e-06 | 778 | 1.3832e-05 | 1.3832e-05
397.4653 | 5.5505e-06 | 257.7 | 1.9615e-05 | 315.6 | 7.3305e-06 | 489 | 5.13e-06 | 779 | 1.3932e-05 | 1.3932e-05
397.7912 | 5.5399e-06 | 257.8 | 1.9619e-05 | 315.8 | 7.3176e-06 | 489.5 | 5.1382e-06 | 780 | 1.4052e-05 | 1.4052e-05
398.117 | 5.5299e-06 | 257.9 | 1.9622e-05 | 316 | 7.3053e-06 | 490 | 5.1457e-06 | 781 | 1.4179e-05 | 1.4179e-05
398.4428 | 5.519e-06 | 258 | 1.9625e-05 | 316.2 | 7.2926e-06 | 490.5 | 5.153e-06 | 782 | 1.4303e-05 | 1.4303e-05
398.7684 | 5.5098e-06 | 258.1 | 1.9628e-05 | 316.4 | 7.2801e-06 | 491 | 5.1598e-06 | 783 | 1.4436e-05 | 1.4436e-05
399.094 | 5.4989e-06 | 258.2 | 1.9631e-05 | 316.6 | 7.2688e-06 | 491.5 | 5.1665e-06 | 784 | 1.4552e-05 | 1.4552e-05
399.4195 | 5.4898e-06 | 258.3 | 1.9634e-05 | 316.8 | 7.2592e-06 | 492 | 5.1736e-06 | 785 | 1.4671e-05 | 1.4671e-05
399.745 | 5.4799e-06 | 258.4 | 1.9638e-05 | 317 | 7.2487e-06 | 492.5 | 5.1808e-06 | 786 | 1.482e-05 | 1.482e-05
400.0703 | 5.4683e-06 | 258.5 | 1.9641e-05 | 317.2 | 7.2384e-06 | 493 | 5.1879e-06 | 787 | 1.4965e-05 | 1.4965e-05
400.3956 | 5.4571e-06 | 258.6 | 1.9645e-05 | 317.4 | 7.2275e-06 | 493.5 | 5.1957e-06 | 788 | 1.5099e-05 | 1.5099e-05
400.7208 | 5.4459e-06 | 258.7 | 1.9649e-05 | 317.6 | 7.2165e-06 | 494 | 5.2033e-06 | 789 | 1.524e-05 | 1.524e-05
401.0459 | 5.4347e-06 | 258.8 | 1.9653e-05 | 317.8 | 7.2069e-06 | 494.5 | 5.2109e-06 | 790 | 1.5372e-05 | 1.5372e-05
401.371 | 5.4232e-06 | 258.9 | 1.9657e-05 | 318 | 7.1957e-06 | 495 | 5.2181e-06 | 791 | 1.5496e-05 | 1.5496e-05
401.6959 | 5.4117e-06 | 259 | 1.9661e-05 | 318.2 | 7.1825e-06 | 495.5 | 5.2255e-06 | 792 | 1.5643e-05 | 1.5643e-05
402.0208 | 5.401e-06 | 259.1 | 1.9671e-05 | 318.4 | 7.1696e-06 | 496 | 5.2321e-06 | 793 | 1.5795e-05 | 1.5795e-05
402.3457 | 5.3899e-06 | 259.2 | 1.9677e-05 | 318.6 | 7.1555e-06 | 496.5 | 5.2382e-06 | 794 | 1.5944e-05 | 1.5944e-05
402.6704 | 5.3789e-06 | 259.3 | 1.9683e-05 | 318.8 | 7.1467e-06 | 497 | 5.2451e-06 | 795 | 1.6088e-05 | 1.6088e-05
402.9951 | 5.3682e-06 | 259.4 | 1.969e-05 | 319 | 7.1379e-06 | 497.5 | 5.2522e-06 | 796 | 1.6217e-05 | 1.6217e-05
403.3197 | 5.3589e-06 | 259.5 | 1.9697e-05 | 319.2 | 7.1265e-06 | 498 | 5.2591e-06 | 797 | 1.6357e-05 | 1.6357e-05
403.6442 | 5.3484e-06 | 259.6 | 1.9704e-05 | 319.4 | 7.1164e-06 | 498.5 | 5.2653e-06 | 798 | 1.6497e-05 | 1.6497e-05
403.9687 | 5.3377e-06 | 259.7 | 1.9711e-05 | 319.6 | 7.1051e-06 | 499 | 5.2712e-06 | 799 | 1.6658e-05 | 1.6658e-05
404.293 | 5.3265e-06 | 259.8 | 1.9731e-05 | 319.8 | 7.0936e-06 | 499.5 | 5.2779e-06 | 800 | 1.6822e-05 | 1.6822e-05
404.6173 | 5.3148e-06 | 259.9 | 1.9743e-05 | 320 | 7.083e-06 | 500 | 5.2845e-06 | 801 | 1.6998e-05 | 1.6998e-05
404.9416 | 5.3032e-06 | 260 | 1.9754e-05 | 320.2 | 7.0709e-06 | 500.5 | 5.2906e-06 | 802 | 1.7178e-05 | 1.7178e-05
405.2657 | 5.2925e-06 | 260.1 | 1.9765e-05 | 320.4 | 7.0576e-06 | 501 | 5.2962e-06 | 803 | 1.7384e-05 | 1.7384e-05
405.5898 | 5.2805e-06 | 260.2 | 1.9781e-05 | 320.6 | 7.0424e-06 | 501.5 | 5.3026e-06 | 804 | 1.7561e-05 | 1.7561e-05
405.9137 | 5.2697e-06 | 260.3 | 1.9796e-05 | 320.8 | 7.027e-06 | 502 | 5.3089e-06 | 805 | 1.7884e-05 | 1.7884e-05
406.2377 | 5.2569e-06 | 260.4 | 1.981e-05 | 321 | 7.0107e-06 | 502.5 | 5.3149e-06 | 806 | 2.284e-05 | 2.284e-05
406.5615 | 5.2446e-06 | 260.5 | 1.9828e-05 | 321.2 | 6.9971e-06 | 503 | 5.3205e-06 | 807 | 3.9438e-05 | 3.9438e-05
406.8853 | 5.2312e-06 | 260.6 | 1.9843e-05 | 321.4 | 6.9845e-06 | 503.5 | 5.3258e-06
407.209 | 5.2196e-06 | 260.7 | 1.9858e-05 | 321.6 | 6.9729e-06 | 504 | 5.3321e-06
407.5326 | 5.2088e-06 | 260.8 | 1.9872e-05 | 321.8 | 6.9622e-06 | 504.5 | 5.3378e-06
407.8561 | 5.1984e-06 | 260.9 | 1.9886e-05 | 322 | 6.9532e-06 | 505 | 5.3431e-06
408.1796 | 5.1879e-06 | 261 | 1.9902e-05 | 322.2 | 6.943e-06 | 505.5 | 5.349e-06
408.503 | 5.1767e-06 | 261.1 | 1.9917e-05 | 322.4 | 6.9332e-06 | 506 | 5.3552e-06
408.8263 | 5.165e-06 | 261.2 | 1.9932e-05 | 322.6 | 6.9202e-06 | 506.5 | 5.3616e-06
409.1495 | 5.1532e-06 | 261.3 | 1.9946e-05 | 322.8 | 6.908e-06 | 507 | 5.368e-06
409.4727 | 5.1407e-06 | 261.4 | 1.996e-05 | 323 | 6.8994e-06 | 507.5 | 5.3737e-06
409.7957 | 5.1295e-06 | 261.5 | 1.9974e-05 | 323.2 | 6.8894e-06 | 508 | 5.3796e-06
410.1188 | 5.1189e-06 | 261.6 | 1.9995e-05 | 323.4 | 6.8809e-06 | 508.5 | 5.3857e-06
410.4417 | 5.1082e-06 | 261.7 | 2.0011e-05 | 323.6 | 6.8711e-06 | 509 | 5.3912e-06
410.7646 | 5.0975e-06 | 261.8 | 2.0028e-05 | 323.8 | 6.8607e-06 | 509.5 | 5.3965e-06
411.0873 | 5.0872e-06 | 261.9 | 2.0047e-05 | 324 | 6.8506e-06 | 510 | 5.402e-06
411.4101 | 5.0773e-06 | 262 | 2.0065e-05 | 324.2 | 6.8422e-06 | 510.5 | 5.4069e-06
411.7327 | 5.0683e-06 | 262.1 | 2.0084e-05 | 324.4 | 6.8372e-06 | 511 | 5.4123e-06
412.0552 | 5.0587e-06 | 262.2 | 2.0102e-05 | 324.6 | 6.8325e-06 | 511.5 | 5.4177e-06
412.3777 | 5.0486e-06 | 262.3 | 2.0128e-05 | 324.8 | 6.8292e-06 | 512 | 5.423e-06
412.7001 | 5.0401e-06 | 262.4 | 2.0149e-05 | 325 | 6.824e-06 | 512.5 | 5.4282e-06
413.0224 | 5.0298e-06 | 262.5 | 2.017e-05 | 325.2 | 6.8179e-06 | 513 | 5.4334e-06
413.3447 | 5.0213e-06 | 262.6 | 2.0191e-05 | 325.4 | 6.8109e-06 | 513.5 | 5.4384e-06
413.6669 | 5.0115e-06 | 262.7 | 2.0208e-05 | 325.6 | 6.8037e-06 | 514 | 5.444e-06
413.989 | 5.0013e-06 | 262.8 | 2.0229e-05 | 325.8 | 6.7959e-06 | 514.5 | 5.4494e-06
414.311 | 4.993e-06 | 262.9 | 2.0249e-05 | 326 | 6.792e-06 | 515 | 5.454e-06
414.633 | 4.9829e-06 | 263 | 2.0265e-05 | 326.2 | 6.7889e-06 | 515.5 | 5.4591e-06
414.9549 | 4.9726e-06 | 263.1 | 2.0283e-05 | 326.4 | 6.7865e-06 | 516 | 5.4643e-06
415.2767 | 4.9614e-06 | 263.2 | 2.0301e-05 | 326.6 | 6.7834e-06 | 516.5 | 5.4695e-06
415.5984 | 4.9524e-06 | 263.3 | 2.0319e-05 | 326.8 | 6.7776e-06 | 517 | 5.4753e-06
415.92 | 4.9433e-06 | 263.4 | 2.0338e-05 | 327 | 6.7735e-06 | 517.5 | 5.4807e-06
416.2416 | 4.9334e-06 | 263.5 | 2.0356e-05 | 327.2 | 6.7691e-06 | 518 | 5.4861e-06
416.5631 | 4.9242e-06 | 263.6 | 2.0375e-05 | 327.4 | 6.7676e-06 | 518.5 | 5.4913e-06
416.8845 | 4.9138e-06 | 263.7 | 2.0383e-05 | 327.6 | 6.7672e-06 | 519 | 5.4959e-06
417.2059 | 4.9043e-06 | 263.8 | 2.0398e-05 | 327.8 | 6.7664e-06 | 519.5 | 5.501e-06
417.5272 | 4.8963e-06 | 263.9 | 2.0413e-05 | 328 | 6.7619e-06 | 520 | 5.5055e-06
417.8484 | 4.8876e-06 | 264 | 2.0428e-05 | 328.2 | 6.7559e-06 | 520.5 | 5.5096e-06
418.1695 | 4.8789e-06 | 264.1 | 2.0444e-05 | 328.4 | 6.7468e-06 | 521 | 5.514e-06
418.4906 | 4.8713e-06 | 264.2 | 2.0458e-05 | 328.6 | 6.7359e-06 | 521.5 | 5.519e-06
418.8116 | 4.8639e-06 | 264.3 | 2.0472e-05 | 328.8 | 6.7254e-06 | 522 | 5.5248e-06
419.1325 | 4.8551e-06 | 264.4 | 2.049e-05 | 329 | 6.716e-06 | 522.5 | 5.531e-06
419.4533 | 4.8479e-06 | 264.5 | 2.0507e-05 | 329.2 | 6.7092e-06 | 523 | 5.5369e-06
419.774 | 4.8402e-06 | 264.6 | 2.0524e-05 | 329.4 | 6.7055e-06 | 523.5 | 5.5433e-06
420.0947 | 4.8339e-06 | 264.7 | 2.0541e-05 | 329.6 | 6.7034e-06 | 524 | 5.5501e-06
420.4153 | 4.8283e-06 | 264.8 | 2.0559e-05 | 329.8 | 6.6996e-06 | 524.5 | 5.5569e-06
420.7358 | 4.8226e-06 | 264.9 | 2.0576e-05 | 330 | 6.6958e-06 | 525 | 5.563e-06
421.0563 | 4.8156e-06 | 265 | 2.0594e-05 | 330.2 | 6.6912e-06 | 525.5 | 5.5686e-06
421.3767 | 4.8086e-06 | 265.1 | 2.0611e-05 | 330.4 | 6.687e-06 | 526 | 5.5746e-06
421.697 | 4.8016e-06 | 265.2 | 2.0627e-05 | 330.6 | 6.6832e-06 | 526.5 | 5.5808e-06
422.0172 | 4.7946e-06 | 265.3 | 2.0644e-05 | 330.8 | 6.6786e-06 | 527 | 5.5871e-06
422.3374 | 4.7887e-06 | 265.4 | 2.0661e-05 | 331 | 6.6731e-06 | 527.5 | 5.5929e-06
422.6574 | 4.7838e-06 | 265.5 | 2.0682e-05 | 331.2 | 6.6662e-06 | 528 | 5.5991e-06
422.9774 | 4.7786e-06 | 265.6 | 2.07e-05 | 331.4 | 6.6591e-06 | 528.5 | 5.6048e-06
423.2974 | 4.7803e-06 | 265.7 | 2.0718e-05 | 331.6 | 6.6506e-06 | 529 | 5.6099e-06
423.6172 | 4.7884e-06 | 265.8 | 2.0736e-05 | 331.8 | 6.6454e-06 | 529.5 | 5.6145e-06
423.937 | 4.8002e-06 | 265.9 | 2.0753e-05 | 332 | 6.642e-06 | 530 | 5.6203e-06
424.2567 | 4.8174e-06 | 266 | 2.0772e-05 | 332.2 | 6.6418e-06 | 530.5 | 5.6268e-06
424.5763 | 4.8415e-06 | 266.1 | 2.079e-05 | 332.4 | 6.6415e-06 | 531 | 5.6334e-06
424.8959 | 4.8737e-06 | 266.2 | 2.0805e-05 | 332.6 | 6.6418e-06 | 531.5 | 5.64e-06
425.2154 | 4.9173e-06 | 266.3 | 2.0822e-05 | 332.8 | 6.6419e-06 | 532 | 5.6467e-06
425.5348 | 4.9639e-06 | 266.4 | 2.0839e-05 | 333 | 6.6415e-06 | 532.5 | 5.6539e-06
425.8541 | 5.0044e-06 | 266.5 | 2.0856e-05 | 333.2 | 6.6382e-06 | 533 | 5.6611e-06
426.1734 | 5.0401e-06 | 266.6 | 2.0866e-05 | 333.4 | 6.6342e-06 | 533.5 | 5.6681e-06
426.4926 | 5.074e-06 | 266.7 | 2.088e-05 | 333.6 | 6.6296e-06 | 534 | 5.6751e-06
426.8117 | 5.1002e-06 | 266.8 | 2.0894e-05 | 333.8 | 6.625e-06 | 534.5 | 5.6822e-06
427.1307 | 5.1169e-06 | 266.9 | 2.0911e-05 | 334 | 6.621e-06 | 535 | 5.6888e-06
427.4497 | 5.1244e-06 | 267 | 2.0925e-05 | 334.2 | 6.6175e-06 | 535.5 | 5.6951e-06
427.7686 | 5.1184e-06 | 267.1 | 2.094e-05 | 334.4 | 6.6145e-06 | 536 | 5.702e-06
428.0874 | 5.0988e-06 | 267.2 | 2.0954e-05 | 334.6 | 6.6122e-06 | 536.5 | 5.7088e-06
428.4061 | 5.0662e-06 | 267.3 | 2.0964e-05 | 334.8 | 6.609e-06 | 537 | 5.7164e-06
428.7248 | 5.0209e-06 | 267.4 | 2.0978e-05 | 335 | 6.6047e-06 | 537.5 | 5.7241e-06
429.0433 | 4.9689e-06 | 267.5 | 2.0991e-05 | 335.2 | 6.6029e-06 | 538 | 5.7314e-06
429.3619 | 4.9194e-06 | 267.6 | 2.1014e-05 | 335.4 | 6.601e-06 | 538.5 | 5.7392e-06
429.6803 | 4.8744e-06 | 267.7 | 2.103e-05 | 335.6 | 6.6014e-06 | 539 | 5.7473e-06
429.9987 | 4.8367e-06 | 267.8 | 2.1046e-05 | 335.8 | 6.601e-06 | 539.5 | 5.7559e-06
430.317 | 4.8123e-06 | 267.9 | 2.1062e-05 | 336 | 6.5969e-06 | 540 | 5.7633e-06
430.6352 | 4.7953e-06 | 268 | 2.107e-05 | 336.2 | 6.5948e-06 | 540.5 | 5.7704e-06
430.9533 | 4.7839e-06 | 268.1 | 2.1085e-05 | 336.4 | 6.5918e-06 | 541 | 5.7786e-06
431.2714 | 4.7769e-06 | 268.2 | 2.11e-05 | 336.6 | 6.5909e-06 | 541.5 | 5.7873e-06
431.5894 | 4.7771e-06 | 268.3 | 2.1115e-05 | 336.8 | 6.5905e-06 | 542 | 5.7954e-06
431.9073 | 4.782e-06 | 268.4 | 2.1109e-05 | 337 | 6.5889e-06 | 542.5 | 5.803e-06
432.2251 | 4.7927e-06 | 268.5 | 2.1116e-05 | 337.2 | 6.5873e-06 | 543 | 5.8106e-06
432.5429 | 4.8103e-06 | 268.6 | 2.1123e-05 | 337.4 | 6.5847e-06 | 543.5 | 5.8189e-06
432.8606 | 4.8345e-06 | 268.7 | 2.1128e-05 | 337.6 | 6.5828e-06 | 544 | 5.8269e-06
433.1782 | 4.8611e-06 | 268.8 | 2.1134e-05 | 337.8 | 6.5808e-06 | 544.5 | 5.8339e-06
433.4958 | 4.8851e-06 | 268.9 | 2.1139e-05 | 338 | 6.5785e-06 | 545 | 5.8418e-06
433.8132 | 4.9002e-06 | 269 | 2.1144e-05 | 338.2 | 6.5762e-06 | 545.5 | 5.8503e-06
434.1306 | 4.9115e-06 | 269.1 | 2.1159e-05 | 338.4 | 6.5736e-06 | 546 | 5.8588e-06
434.448 | 4.9214e-06 | 269.2 | 2.1168e-05 | 338.6 | 6.5728e-06 | 546.5 | 5.8677e-06
434.7652 | 4.9336e-06 | 269.3 | 2.1177e-05 | 338.8 | 6.5724e-06 | 547 | 5.8758e-06
435.0824 | 4.9462e-06 | 269.4 | 2.1185e-05 | 339 | 6.5725e-06 | 547.5 | 5.8849e-06
435.3995 | 4.9581e-06 | 269.5 | 2.1194e-05 | 339.2 | 6.572e-06 | 548 | 5.8941e-06
435.7165 | 4.9666e-06 | 269.6 | 2.1204e-05 | 339.4 | 6.5694e-06 | 548.5 | 5.9033e-06
436.0335 | 4.9728e-06 | 269.7 | 2.1213e-05 | 339.6 | 6.5687e-06 | 549 | 5.9118e-06
436.3503 | 4.9771e-06 | 269.8 | 2.1222e-05 | 339.8 | 6.5679e-06 | 549.5 | 5.9203e-06
436.6672 | 4.9787e-06 | 269.9 | 2.1231e-05 | 340 | 6.569e-06 | 550 | 5.929e-06
436.9839 | 4.979e-06 | 270 | 2.124e-05 | 340.2 | 6.5702e-06 | 550.5 | 5.9386e-06
437.3005 | 4.9759e-06 | 270.1 | 2.1246e-05 | 340.4 | 6.5697e-06 | 551 | 5.948e-06
437.6171 | 4.9715e-06 | 270.2 | 2.1254e-05 | 340.6 | 6.5703e-06 | 551.5 | 5.9573e-06
437.9336 | 4.9659e-06 | 270.3 | 2.1263e-05 | 340.8 | 6.5714e-06 | 552 | 5.9666e-06
438.2501 | 4.9598e-06 | 270.4 | 2.1271e-05 | 341 | 6.5736e-06 | 552.5 | 5.9763e-06
438.5664 | 4.9533e-06 | 270.5 | 2.1281e-05 | 341.2 | 6.5764e-06 | 553 | 5.9854e-06
438.8827 | 4.9461e-06 | 270.6 | 2.129e-05 | 341.4 | 6.5783e-06 | 553.5 | 5.995e-06
439.1989 | 4.9403e-06 | 270.7 | 2.1298e-05 | 341.6 | 6.5812e-06 | 554 | 6.0056e-06
439.5151 | 4.934e-06 | 270.8 | 2.1307e-05 | 341.8 | 6.585e-06 | 554.5 | 6.0157e-06
439.8311 | 4.9264e-06 | 270.9 | 2.1315e-05 | 342 | 6.5894e-06 | 555 | 6.0258e-06
440.1471 | 4.919e-06 | 271 | 2.1325e-05 | 342.2 | 6.5938e-06 | 555.5 | 6.0354e-06
440.4631 | 4.9121e-06 | 271.1 | 2.1334e-05 | 342.4 | 6.5963e-06 | 556 | 6.0456e-06
440.7789 | 4.9052e-06 | 271.2 | 2.134e-05 | 342.6 | 6.5982e-06 | 556.5 | 6.0556e-06
441.0947 | 4.898e-06 | 271.3 | 2.1348e-05 | 342.8 | 6.5973e-06 | 557 | 6.0642e-06
441.4104 | 4.8909e-06 | 271.4 | 2.1355e-05 | 343 | 6.5974e-06 | 557.5 | 6.0737e-06
441.726 | 4.8851e-06 | 271.5 | 2.1363e-05 | 343.2 | 6.5973e-06 | 558 | 6.085e-06
442.0415 | 4.8788e-06 | 271.6 | 2.1347e-05 | 343.4 | 6.5986e-06 | 558.5 | 6.096e-06
442.357 | 4.8732e-06 | 271.7 | 2.1347e-05 | 343.6 | 6.5999e-06 | 559 | 6.1057e-06
442.6724 | 4.867e-06 | 271.8 | 2.1347e-05 | 343.8 | 6.5999e-06 | 559.5 | 6.1159e-06
442.9878 | 4.8631e-06 | 271.9 | 2.1355e-05 | 344 | 6.6002e-06 | 560 | 6.1268e-06
443.303 | 4.8593e-06 | 272 | 2.1355e-05 | 344.2 | 6.5999e-06 | 560.5 | 6.1379e-06
443.6182 | 4.8559e-06 | 272.1 | 2.1354e-05 | 344.4 | 6.6005e-06 | 561 | 6.1481e-06
443.9333 | 4.8519e-06 | 272.2 | 2.1354e-05 | 344.6 | 6.6007e-06 | 561.5 | 6.1582e-06
444.2483 | 4.8481e-06 | 272.3 | 2.1368e-05 | 344.8 | 6.5995e-06 | 562 | 6.1693e-06
444.5633 | 4.8442e-06 | 272.4 | 2.1375e-05 | 345 | 6.5997e-06 | 562.5 | 6.1805e-06
444.8782 | 4.8412e-06 | 272.5 | 2.1382e-05 | 345.2 | 6.6011e-06 | 563 | 6.1906e-06
445.193 | 4.838e-06 | 272.6 | 2.1384e-05 | 345.4 | 6.6002e-06 | 563.5 | 6.2008e-06
445.5077 | 4.8349e-06 | 272.7 | 2.139e-05 | 345.6 | 6.5985e-06 | 564 | 6.2112e-06
445.8224 | 4.8319e-06 | 272.8 | 2.1395e-05 | 345.8 | 6.5942e-06 | 564.5 | 6.2215e-06
446.137 | 4.8288e-06 | 272.9 | 2.14e-05 | 346 | 6.5917e-06 | 565 | 6.2317e-06
446.4515 | 4.8234e-06 | 273 | 2.1386e-05 | 346.2 | 6.592e-06 | 565.5 | 6.2415e-06
446.766 | 4.8188e-06 | 273.1 | 2.1384e-05 | 346.4 | 6.5933e-06 | 566 | 6.2509e-06
447.0804 | 4.8148e-06 | 273.2 | 2.1383e-05 | 346.6 | 6.5958e-06 | 566.5 | 6.2608e-06
447.3947 | 4.8108e-06 | 273.3 | 2.1368e-05 | 346.8 | 6.5968e-06 | 567 | 6.2705e-06
447.7089 | 4.808e-06 | 273.4 | 2.1359e-05 | 347 | 6.5978e-06 | 567.5 | 6.28e-06
448.023 | 4.8042e-06 | 273.5 | 2.1349e-05 | 347.2 | 6.5979e-06 | 568 | 6.2907e-06
448.3371 | 4.8006e-06 | 273.6 | 2.134e-05 | 347.4 | 6.6006e-06 | 568.5 | 6.3015e-06
448.6511 | 4.797e-06 | 273.7 | 2.1333e-05 | 347.6 | 6.604e-06 | 569 | 6.3128e-06
448.9651 | 4.7938e-06 | 273.8 | 2.1324e-05 | 347.8 | 6.6061e-06 | 569.5 | 6.3247e-06
449.2789 | 4.7908e-06 | 273.9 | 2.1314e-05 | 348 | 6.609e-06 | 570 | 6.3363e-06
449.5927 | 4.7881e-06 | 274 | 2.1309e-05 | 348.2 | 6.6124e-06 | 570.5 | 6.3482e-06
449.9064 | 4.7857e-06 | 274.1 | 2.1302e-05 | 348.4 | 6.6157e-06 | 571 | 6.3601e-06
450.2201 | 4.7836e-06 | 274.2 | 2.1296e-05 | 348.6 | 6.6198e-06 | 571.5 | 6.3715e-06
450.5336 | 4.7815e-06 | 274.3 | 2.1289e-05 | 348.8 | 6.6209e-06 | 572 | 6.3828e-06
450.8471 | 4.7798e-06 | 274.4 | 2.1273e-05 | 349 | 6.6212e-06 | 572.5 | 6.3942e-06
451.1605 | 4.7774e-06 | 274.5 | 2.1263e-05 | 349.2 | 6.6212e-06 | 573 | 6.405e-06
451.4739 | 4.777e-06 | 274.6 | 2.1253e-05 | 349.4 | 6.6238e-06 | 573.5 | 6.4161e-06
451.7871 | 4.7762e-06 | 274.7 | 2.1243e-05 | 349.6 | 6.6274e-06 | 574 | 6.4269e-06
452.1003 | 4.7758e-06 | 274.8 | 2.1232e-05 | 349.8 | 6.6306e-06 | 574.5 | 6.4383e-06
452.4135 | 4.7747e-06 | 274.9 | 2.122e-05 | 350 | 6.6348e-06 | 575 | 6.4505e-06
452.7265 | 4.7733e-06 | 275 | 2.1208e-05 | 350.2 | 6.6397e-06 | 575.5 | 6.4624e-06
453.0395 | 4.7713e-06 | 275.1 | 2.12e-05 | 350.4 | 6.6449e-06 | 576 | 6.4738e-06
453.3524 | 4.7705e-06 | 275.2 | 2.1191e-05 | 350.6 | 6.6521e-06 | 576.5 | 6.4855e-06
453.6653 | 4.7698e-06 | 275.3 | 2.1182e-05 | 350.8 | 6.6579e-06 | 577 | 6.4977e-06
453.978 | 4.7684e-06 | 275.4 | 2.1172e-05 | 351 | 6.6637e-06 | 577.5 | 6.5097e-06
454.2907 | 4.7679e-06 | 275.5 | 2.1166e-05 | 351.2 | 6.6698e-06 | 578 | 6.5219e-06
454.6033 | 4.7672e-06 | 275.6 | 2.1158e-05 | 351.4 | 6.6766e-06 | 578.5 | 6.5341e-06
454.9159 | 4.7653e-06 | 275.7 | 2.115e-05 | 351.6 | 6.6857e-06 | 579 | 6.5468e-06
455.2283 | 4.7643e-06 | 275.8 | 2.1134e-05 | 351.8 | 6.6906e-06 | 579.5 | 6.5591e-06
455.5407 | 4.7636e-06 | 275.9 | 2.1124e-05 | 352 | 6.6956e-06 | 580 | 6.5712e-06
455.8531 | 4.7633e-06 | 276 | 2.1113e-05 | 352.2 | 6.7014e-06 | 580.5 | 6.5835e-06
456.1653 | 4.7623e-06 | 276.1 | 2.1103e-05 | 352.4 | 6.7071e-06 | 581 | 6.5963e-06
456.4775 | 4.7623e-06 | 276.2 | 2.1082e-05 | 352.6 | 6.7156e-06 | 581.5 | 6.6093e-06
456.7896 | 4.7623e-06 | 276.3 | 2.1067e-05 | 352.8 | 6.7247e-06 | 582 | 6.6223e-06
457.1016 | 4.7622e-06 | 276.4 | 2.1052e-05 | 353 | 6.7341e-06 | 582.5 | 6.6371e-06
457.4136 | 4.7621e-06 | 276.5 | 2.1036e-05 | 353.2 | 6.7421e-06 | 583 | 6.6518e-06
457.7255 | 4.7613e-06 | 276.6 | 2.102e-05 | 353.4 | 6.75e-06 | 583.5 | 6.6657e-06
458.0373 | 4.7616e-06 | 276.7 | 2.1003e-05 | 353.6 | 6.7575e-06 | 584 | 6.6802e-06
458.349 | 4.7623e-06 | 276.8 | 2.0987e-05 | 353.8 | 6.762e-06 | 584.5 | 6.6946e-06
458.6607 | 4.7631e-06 | 276.9 | 2.0965e-05 | 354 | 6.766e-06 | 585 | 6.71e-06
458.9723 | 4.7642e-06 | 277 | 2.0947e-05 | 354.2 | 6.7699e-06 | 585.5 | 6.7259e-06
459.2838 | 4.7658e-06 | 277.1 | 2.0929e-05 | 354.4 | 6.7735e-06 | 586 | 6.7424e-06
459.5953 | 4.7673e-06 | 277.2 | 2.0914e-05 | 354.6 | 6.7788e-06 | 586.5 | 6.761e-06
459.9066 | 4.7693e-06 | 277.3 | 2.0896e-05 | 354.8 | 6.7826e-06 | 587 | 6.7794e-06
460.2179 | 4.7714e-06 | 277.4 | 2.0878e-05 | 355 | 6.7861e-06 | 587.5 | 6.7981e-06
460.5292 | 4.77e-06 | 277.5 | 2.0861e-05 | 355.2 | 6.7902e-06 | 588 | 6.8176e-06
460.8403 | 4.7723e-06 | 277.6 | 2.083e-05 | 355.4 | 6.7941e-06 | 588.5 | 6.8376e-06
461.1514 | 4.7746e-06 | 277.7 | 2.0809e-05 | 355.6 | 6.7987e-06 | 589 | 6.8578e-06
461.4624 | 4.7772e-06 | 277.8 | 2.0788e-05 | 355.8 | 6.8043e-06 | 589.5 | 6.8779e-06
461.7734 | 4.78e-06 | 277.9 | 2.0764e-05 | 356 | 6.8102e-06 | 590 | 6.8981e-06
462.0842 | 4.7829e-06 | 278 | 2.0739e-05 | 356.2 | 6.8127e-06 | 590.5 | 6.9185e-06
462.3951 | 4.7857e-06 | 278.1 | 2.0714e-05 | 356.4 | 6.8152e-06 | 591 | 6.9383e-06
462.7058 | 4.7878e-06 | 278.2 | 2.069e-05 | 356.6 | 6.8184e-06 | 591.5 | 6.9566e-06
463.0164 | 4.7911e-06 | 278.3 | 2.0666e-05 | 356.8 | 6.8197e-06 | 592 | 6.9734e-06
463.327 | 4.7939e-06 | 278.4 | 2.0643e-05 | 357 | 6.8214e-06 | 592.5 | 6.9893e-06
463.6375 | 4.7958e-06 | 278.5 | 2.0619e-05 | 357.2 | 6.8206e-06 | 593 | 7.005e-06
463.9479 | 4.8011e-06 | 278.6 | 2.0596e-05 | 357.4 | 6.8205e-06 | 593.5 | 7.0191e-06
464.2583 | 4.8039e-06 | 278.7 | 2.0548e-05 | 357.6 | 6.8249e-06 | 594 | 7.0316e-06
464.5686 | 4.8074e-06 | 278.8 | 2.0517e-05 | 357.8 | 6.8278e-06 | 594.5 | 7.0434e-06
464.8788 | 4.81e-06 | 278.9 | 2.0486e-05 | 358 | 6.8323e-06 | 595 | 7.0545e-06
465.1889 | 4.8125e-06 | 279 | 2.0452e-05 | 358.2 | 6.8367e-06 | 595.5 | 7.0653e-06
465.499 | 4.8147e-06 | 279.1 | 2.0417e-05 | 358.4 | 6.8402e-06 | 596 | 7.0731e-06
465.809 | 4.8167e-06 | 279.2 | 2.0382e-05 | 358.6 | 6.8433e-06 | 596.5 | 7.0793e-06
466.1189 | 4.8187e-06 | 279.3 | 2.0347e-05 | 358.8 | 6.8445e-06 | 597 | 7.084e-06
466.4288 | 4.8199e-06 | 279.4 | 2.0316e-05 | 359 | 6.8408e-06 | 597.5 | 7.0869e-06
466.7386 | 4.821e-06 | 279.5 | 2.0284e-05 | 359.2 | 6.8378e-06 | 598 | 7.088e-06
467.0483 | 4.8228e-06 | 279.6 | 2.0251e-05 | 359.4 | 6.8332e-06 | 598.5 | 7.0864e-06
467.3579 | 4.8247e-06 | 279.7 | 2.0222e-05 | 359.6 | 6.8286e-06 | 599 | 7.0829e-06
467.6675 | 4.8264e-06 | 279.8 | 2.0191e-05 | 359.8 | 6.8252e-06 | 599.5 | 7.0774e-06
467.977 | 4.8281e-06 | 279.9 | 2.0159e-05 | 360 | 6.8232e-06 | 600 | 7.0702e-06
468.2864 | 4.8303e-06 | 280 | 2.0128e-05 | 360.2 | 6.8211e-06 | 600.5 | 7.0606e-06
468.5958 | 4.8332e-06 | 280.1 | 2.0083e-05 | 360.4 | 6.8203e-06 | 601 | 7.0482e-06
468.905 | 4.8359e-06 | 280.2 | 2.0048e-05 | 360.6 | 6.8206e-06 | 601.5 | 7.0337e-06
469.2142 | 4.8391e-06 | 280.3 | 2.0012e-05 | 360.8 | 6.8207e-06 | 602 | 7.0169e-06
469.5234 | 4.8423e-06 | 280.4 | 1.997e-05 | 361 | 6.8219e-06 | 602.5 | 6.9982e-06
469.8324 | 4.8458e-06 | 280.5 | 1.993e-05 | 361.2 | 6.8238e-06 | 603 | 6.977e-06
470.1414 | 4.8501e-06 | 280.6 | 1.989e-05 | 361.4 | 6.8255e-06 | 603.5 | 6.9532e-06
470.4503 | 4.854e-06 | 280.7 | 1.9849e-05 | 361.6 | 6.8264e-06 | 604 | 6.9281e-06
470.7592 | 4.8582e-06 | 280.8 | 1.9805e-05 | 361.8 | 6.8265e-06 | 604.5 | 6.9022e-06
471.0679 | 4.8623e-06 | 280.9 | 1.9763e-05 | 362 | 6.8238e-06 | 605 | 6.8764e-06
471.3766 | 4.8659e-06 | 281 | 1.9722e-05 | 362.2 | 6.8201e-06 | 605.5 | 6.8509e-06
471.6853 | 4.8698e-06 | 281.1 | 1.967e-05 | 362.4 | 6.8164e-06 | 606 | 6.8252e-06
471.9938 | 4.8732e-06 | 281.2 | 1.9624e-05 | 362.6 | 6.8149e-06 | 606.5 | 6.8e-06
472.3023 | 4.8772e-06 | 281.3 | 1.9578e-05 | 362.8 | 6.8141e-06 | 607 | 6.7743e-06
472.6107 | 4.8808e-06 | 281.4 | 1.9531e-05 | 363 | 6.8168e-06 | 607.5 | 6.75e-06
472.9191 | 4.885e-06 | 281.5 | 1.9481e-05 | 363.2 | 6.8163e-06 | 608 | 6.7266e-06
473.2273 | 4.8888e-06 | 281.6 | 1.9433e-05 | 363.4 | 6.8158e-06 | 608.5 | 6.7036e-06
473.5355 | 4.8922e-06 | 281.7 | 1.9384e-05 | 363.6 | 6.8156e-06 | 609 | 6.6825e-06
473.8437 | 4.8962e-06 | 281.8 | 1.9327e-05 | 363.8 | 6.815e-06 | 609.5 | 6.6621e-06
474.1517 | 4.9004e-06 | 281.9 | 1.9275e-05 | 364 | 6.8153e-06 | 610 | 6.643e-06
474.4597 | 4.9048e-06 | 282 | 1.9223e-05 | 364.2 | 6.8148e-06 | 610.5 | 6.6248e-06
474.7676 | 4.9088e-06 | 282.1 | 1.9171e-05 | 364.4 | 6.8137e-06 | 611 | 6.6085e-06
475.0754 | 4.9129e-06 | 282.2 | 1.9123e-05 | 364.6 | 6.8123e-06 | 611.5 | 6.5936e-06
475.3832 | 4.9172e-06 | 282.3 | 1.9071e-05 | 364.8 | 6.8102e-06 | 612 | 6.5795e-06
475.6909 | 4.9221e-06 | 282.4 | 1.9018e-05 | 365 | 6.8063e-06 | 612.5 | 6.5676e-06
475.9985 | 4.9261e-06 | 282.5 | 1.8973e-05 | 365.2 | 6.8034e-06 | 613 | 6.5581e-06
476.3061 | 4.9305e-06 | 282.6 | 1.8926e-05 | 365.4 | 6.7993e-06 | 613.5 | 6.5494e-06
476.6135 | 4.9358e-06 | 282.7 | 1.8878e-05 | 365.6 | 6.7936e-06 | 614 | 6.5409e-06
476.9209 | 4.9402e-06 | 282.8 | 1.883e-05 | 365.8 | 6.7885e-06 | 614.5 | 6.5342e-06
477.2283 | 4.9449e-06 | 282.9 | 1.8767e-05 | 366 | 6.7825e-06 | 615 | 6.5293e-06
477.5355 | 4.9492e-06 | 283 | 1.8714e-05 | 366.2 | 6.7778e-06 | 615.5 | 6.5255e-06
477.8427 | 4.9536e-06 | 283.1 | 1.8662e-05 | 366.4 | 6.7719e-06 | 616 | 6.5222e-06
478.1498 | 4.9584e-06 | 283.2 | 1.8605e-05 | 366.6 | 6.7617e-06 | 616.5 | 6.5207e-06
478.4569 | 4.9631e-06 | 283.3 | 1.8548e-05 | 366.8 | 6.7534e-06 | 617 | 6.522e-06
478.7639 | 4.9673e-06 | 283.4 | 1.8491e-05 | 367 | 6.7452e-06 | 617.5 | 6.5244e-06
479.0708 | 4.9716e-06 | 283.5 | 1.8433e-05 | 367.2 | 6.739e-06 | 618 | 6.5282e-06
479.3776 | 4.9773e-06 | 283.6 | 1.8368e-05 | 367.4 | 6.734e-06 | 618.5 | 6.5346e-06
479.6844 | 4.9823e-06 | 283.7 | 1.8309e-05 | 367.6 | 6.7273e-06 | 619 | 6.5451e-06
479.991 | 4.9872e-06 | 283.8 | 1.8249e-05 | 367.8 | 6.7224e-06 | 619.5 | 6.5589e-06
480.2977 | 4.9923e-06 | 283.9 | 1.819e-05 | 368 | 6.7199e-06 | 620 | 6.5747e-06
480.6042 | 4.9969e-06 | 284 | 1.8129e-05 | 368.2 | 6.7179e-06 | 620.5 | 6.5925e-06
480.9107 | 5.0008e-06 | 284.1 | 1.8068e-05 | 368.4 | 6.717e-06 | 621 | 6.6145e-06
481.2171 | 5.0056e-06 | 284.2 | 1.8007e-05 | 368.6 | 6.715e-06 | 621.5 | 6.6391e-06
481.5234 | 5.0104e-06 | 284.3 | 1.794e-05 | 368.8 | 6.7133e-06 | 622 | 6.6659e-06
481.8297 | 5.0153e-06 | 284.4 | 1.7878e-05 | 369 | 6.7116e-06 | 622.5 | 6.6943e-06
482.1359 | 5.0205e-06 | 284.5 | 1.7816e-05 | 369.2 | 6.7099e-06 | 623 | 6.7247e-06
482.442 | 5.0249e-06 | 284.6 | 1.7753e-05 | 369.4 | 6.7093e-06 | 623.5 | 6.7567e-06
482.748 | 5.0297e-06 | 284.7 | 1.7694e-05 | 369.6 | 6.7108e-06 | 624 | 6.7888e-06
483.054 | 5.0349e-06 | 284.8 | 1.7632e-05 | 369.8 | 6.7109e-06 | 624.5 | 6.8213e-06
483.3599 | 5.0399e-06 | 284.9 | 1.7569e-05 | 370 | 6.7102e-06 | 625 | 6.8543e-06
483.6657 | 5.0454e-06 | 285 | 1.7507e-05 | 370.2 | 6.7136e-06 | 625.5 | 6.8879e-06
483.9715 | 5.0505e-06 | 285.1 | 1.7446e-05 | 370.4 | 6.717e-06 | 626 | 6.9231e-06
484.2772 | 5.0564e-06 | 285.2 | 1.7385e-05 | 370.6 | 6.7211e-06 | 626.5 | 6.9591e-06
484.5828 | 5.0614e-06 | 285.3 | 1.7324e-05 | 370.8 | 6.7241e-06 | 627 | 6.9945e-06
484.8883 | 5.0661e-06 | 285.4 | 1.7259e-05 | 371 | 6.7227e-06 | 627.5 | 7.0298e-06
485.1938 | 5.0707e-06 | 285.5 | 1.7197e-05 | 371.2 | 6.7233e-06 | 628 | 7.0643e-06
485.4992 | 5.0762e-06 | 285.6 | 1.7134e-05 | 371.4 | 6.7232e-06 | 628.5 | 7.0972e-06
485.8046 | 5.0814e-06 | 285.7 | 1.7076e-05 | 371.6 | 6.723e-06 | 629 | 7.1271e-06
486.1098 | 5.0867e-06 | 285.8 | 1.7014e-05 | 371.8 | 6.7218e-06 | 629.5 | 7.1546e-06
486.415 | 5.0918e-06 | 285.9 | 1.6953e-05 | 372 | 6.7173e-06 | 630 | 7.1815e-06
486.7201 | 5.0964e-06 | 286 | 1.6892e-05 | 372.2 | 6.7103e-06 | 630.5 | 7.2066e-06
487.0252 | 5.1001e-06 | 286.1 | 1.6817e-05 | 372.4 | 6.7017e-06 | 631 | 7.2294e-06
487.3301 | 5.1044e-06 | 286.2 | 1.6753e-05 | 372.6 | 6.6926e-06 | 631.5 | 7.2499e-06
487.635 | 5.1087e-06 | 286.3 | 1.6688e-05 | 372.8 | 6.6841e-06 | 632 | 7.2683e-06
487.9399 | 5.1132e-06 | 286.4 | 1.6619e-05 | 373 | 6.6771e-06 | 632.5 | 7.2842e-06
488.2446 | 5.1179e-06 | 286.5 | 1.6551e-05 | 373.2 | 6.6618e-06 | 633 | 7.2964e-06
488.5493 | 5.1232e-06 | 286.6 | 1.6482e-05 | 373.4 | 6.6446e-06 | 633.5 | 7.3047e-06
488.8539 | 5.1274e-06 | 286.7 | 1.6414e-05 | 373.6 | 6.6259e-06 | 634 | 7.3094e-06
489.1585 | 5.1326e-06 | 286.8 | 1.6354e-05 | 373.8 | 6.6073e-06 | 634.5 | 7.3123e-06
489.463 | 5.1377e-06 | 286.9 | 1.6289e-05 | 374 | 6.591e-06 | 635 | 7.3158e-06
489.7674 | 5.1426e-06 | 287 | 1.6223e-05 | 374.2 | 6.5758e-06 | 635.5 | 7.3183e-06
490.0717 | 5.1466e-06 | 287.1 | 1.6156e-05 | 374.4 | 6.5609e-06 | 636 | 7.3187e-06
490.376 | 5.1512e-06 | 287.2 | 1.609e-05 | 374.6 | 6.5457e-06 | 636.5 | 7.3193e-06
490.6802 | 5.1557e-06 | 287.3 | 1.6025e-05 | 374.8 | 6.5335e-06 | 637 | 7.318e-06
490.9843 | 5.1596e-06 | 287.4 | 1.5959e-05 | 375 | 6.5219e-06 | 637.5 | 7.3137e-06
491.2884 | 5.1634e-06 | 287.5 | 1.5886e-05 | 375.2 | 6.5122e-06 | 638 | 7.3043e-06
491.5923 | 5.1678e-06 | 287.6 | 1.5818e-05 | 375.4 | 6.5038e-06 | 638.5 | 7.2927e-06
491.8963 | 5.172e-06 | 287.7 | 1.5749e-05 | 375.6 | 6.4951e-06 | 639 | 7.279e-06
492.2001 | 5.1766e-06 | 287.8 | 1.5684e-05 | 375.8 | 6.4863e-06 | 639.5 | 7.2622e-06
492.5039 | 5.1808e-06 | 287.9 | 1.5617e-05 | 376 | 6.4774e-06 | 640 | 7.2424e-06
492.8076 | 5.1849e-06 | 288 | 1.5549e-05 | 376.2 | 6.4687e-06 | 640.5 | 7.2192e-06
493.1112 | 5.1894e-06 | 288.1 | 1.5481e-05 | 376.4 | 6.4601e-06 | 641 | 7.1933e-06
493.4147 | 5.1945e-06 | 288.2 | 1.5417e-05 | 376.6 | 6.4518e-06 | 641.5 | 7.164e-06
493.7182 | 5.199e-06 | 288.3 | 1.5351e-05 | 376.8 | 6.4434e-06 | 642 | 7.1342e-06
494.0216 | 5.2036e-06 | 288.4 | 1.5286e-05 | 377 | 6.4351e-06 | 642.5 | 7.1057e-06
494.325 | 5.2084e-06 | 288.5 | 1.5219e-05 | 377.2 | 6.4271e-06 | 643 | 7.0801e-06
494.6283 | 5.2127e-06 | 288.6 | 1.5153e-05 | 377.4 | 6.4188e-06 | 643.5 | 7.0573e-06
494.9315 | 5.2173e-06 | 288.7 | 1.5087e-05 | 377.6 | 6.41e-06 | 644 | 7.0353e-06
495.2346 | 5.2214e-06 | 288.8 | 1.5021e-05 | 377.8 | 6.4017e-06 | 644.5 | 7.0135e-06
495.5376 | 5.2261e-06 | 288.9 | 1.4956e-05 | 378 | 6.3934e-06 | 645 | 6.9919e-06
495.8406 | 5.2305e-06 | 289 | 1.4891e-05 | 378.2 | 6.3844e-06 | 645.5 | 6.9704e-06
496.1436 | 5.2337e-06 | 289.1 | 1.4825e-05 | 378.4 | 6.3756e-06 | 646 | 6.9486e-06
496.4464 | 5.2374e-06 | 289.2 | 1.4763e-05 | 378.6 | 6.3669e-06 | 646.5 | 6.9264e-06
496.7492 | 5.2415e-06 | 289.3 | 1.4699e-05 | 378.8 | 6.3572e-06 | 647 | 6.9063e-06
497.0519 | 5.2458e-06 | 289.4 | 1.4635e-05 | 379 | 6.3476e-06 | 647.5 | 6.8891e-06
497.3545 | 5.25e-06 | 289.5 | 1.457e-05 | 379.2 | 6.3382e-06 | 648 | 6.8732e-06
497.6571 | 5.2545e-06 | 289.6 | 1.4505e-05 | 379.4 | 6.3289e-06 | 648.5 | 6.8574e-06
497.9596 | 5.2589e-06 | 289.7 | 1.4441e-05 | 379.6 | 6.3205e-06 | 649 | 6.8424e-06
498.262 | 5.2623e-06 | 289.8 | 1.4377e-05 | 379.8 | 6.3113e-06 | 649.5 | 6.8274e-06
498.5644 | 5.2662e-06 | 289.9 | 1.4319e-05 | 380 | 6.3015e-06 | 650 | 6.8119e-06
498.8667 | 5.2695e-06 | 290 | 1.4256e-05 | 380.2 | 6.2905e-06 | 650.5 | 6.7968e-06
499.1689 | 5.273e-06 | 290.1 | 1.4194e-05 | 380.4 | 6.2798e-06 | 651 | 6.7817e-06
499.471 | 5.2779e-06 | 290.2 | 1.4132e-05 | 380.6 | 6.269e-06 | 651.5 | 6.766e-06
499.7731 | 5.2813e-06 | 290.3 | 1.4072e-05 | 380.8 | 6.2582e-06 | 652 | 6.7534e-06
500.0751 | 5.2855e-06 | 290.4 | 1.4011e-05 | 381 | 6.2478e-06 | 652.5 | 6.7437e-06
500.377 | 5.2897e-06 | 290.5 | 1.3951e-05 | 381.2 | 6.2378e-06 | 653 | 6.7332e-06
500.6789 | 5.2925e-06 | 290.6 | 1.3891e-05 | 381.4 | 6.2275e-06 | 653.5 | 6.7227e-06
500.9807 | 5.2956e-06 | 290.7 | 1.383e-05 | 381.6 | 6.2171e-06 | 654 | 6.7149e-06
501.2824 | 5.2997e-06 | 290.8 | 1.3769e-05 | 381.8 | 6.2073e-06 | 654.5 | 6.71e-06
501.5841 | 5.3039e-06 | 290.9 | 1.3709e-05 | 382 | 6.1975e-06 | 655 | 6.7037e-06
501.8856 | 5.3074e-06 | 291 | 1.3647e-05 | 382.2 | 6.188e-06 | 655.5 | 6.6982e-06
502.1872 | 5.3113e-06 | 291.1 | 1.3587e-05 | 382.4 | 6.1791e-06 | 656 | 6.6959e-06
502.4886 | 5.315e-06 | 291.2 | 1.3526e-05 | 382.6 | 6.1702e-06 | 656.5 | 6.6949e-06
502.79 | 5.3178e-06 | 291.3 | 1.3472e-05 | 382.8 | 6.1606e-06 | 657 | 6.6909e-06
503.0913 | 5.322e-06 | 291.4 | 1.3413e-05 | 383 | 6.1506e-06 | 657.5 | 6.6851e-06
503.3925 | 5.3244e-06 | 291.5 | 1.3355e-05 | 383.2 | 6.1396e-06 | 658 | 6.6826e-06
503.6936 | 5.3276e-06 | 291.6 | 1.3296e-05 | 383.4 | 6.1286e-06 | 658.5 | 6.6818e-06
503.9947 | 5.3325e-06 | 291.7 | 1.3244e-05 | 383.6 | 6.1175e-06 | 659 | 6.6805e-06
504.2958 | 5.3358e-06 | 291.8 | 1.3189e-05 | 383.8 | 6.1082e-06 | 659.5 | 6.6792e-06
504.5967 | 5.3389e-06 | 291.9 | 1.3133e-05 | 384 | 6.0996e-06 | 660 | 6.6816e-06
504.8976 | 5.3414e-06 | 292 | 1.3077e-05 | 384.2 | 6.0918e-06 | 660.5 | 6.6874e-06
505.1984 | 5.3456e-06 | 292.1 | 1.3022e-05 | 384.4 | 6.0833e-06 | 661 | 6.6911e-06
505.4991 | 5.3491e-06 | 292.2 | 1.2966e-05 | 384.6 | 6.0744e-06 | 661.5 | 6.6933e-06
505.7998 | 5.3524e-06 | 292.3 | 1.2911e-05 | 384.8 | 6.0649e-06 | 662 | 6.6964e-06
506.1004 | 5.3568e-06 | 292.4 | 1.2855e-05 | 385 | 6.0553e-06 | 662.5 | 6.7002e-06
506.4009 | 5.3601e-06 | 292.5 | 1.28e-05 | 385.2 | 6.0454e-06 | 663 | 6.705e-06
506.7014 | 5.3643e-06 | 292.6 | 1.2744e-05 | 385.4 | 6.0356e-06 | 663.5 | 6.7112e-06
507.0018 | 5.3683e-06 | 292.7 | 1.2694e-05 | 385.6 | 6.0259e-06 | 664 | 6.7192e-06
507.3021 | 5.3713e-06 | 292.8 | 1.264e-05 | 385.8 | 6.0161e-06 | 664.5 | 6.7281e-06
507.6024 | 5.3749e-06 | 292.9 | 1.2587e-05 | 386 | 6.0066e-06 | 665 | 6.7375e-06
507.9025 | 5.3781e-06 | 293 | 1.2533e-05 | 386.2 | 5.9972e-06 | 665.5 | 6.7475e-06
508.2026 | 5.3826e-06 | 293.1 | 1.2479e-05 | 386.4 | 5.9884e-06 | 666 | 6.7569e-06
508.5027 | 5.3858e-06 | 293.2 | 1.2426e-05 | 386.6 | 5.9801e-06 | 666.5 | 6.7662e-06
508.8026 | 5.3887e-06 | 293.3 | 1.2373e-05 | 386.8 | 5.9721e-06 | 667 | 6.777e-06
509.1025 | 5.3926e-06 | 293.4 | 1.2322e-05 | 387 | 5.9642e-06 | 667.5 | 6.7891e-06
509.4024 | 5.3955e-06 | 293.5 | 1.2269e-05 | 387.2 | 5.9563e-06 | 668 | 6.8017e-06
509.7021 | 5.3986e-06 | 293.6 | 1.2216e-05 | 387.4 | 5.9476e-06 | 668.5 | 6.8139e-06
510.0018 | 5.4022e-06 | 293.7 | 1.2163e-05 | 387.6 | 5.9386e-06 | 669 | 6.8255e-06
510.3014 | 5.4051e-06 | 293.8 | 1.2118e-05 | 387.8 | 5.9292e-06 | 669.5 | 6.8376e-06
510.601 | 5.4078e-06 | 293.9 | 1.2068e-05 | 388 | 5.9206e-06 | 670 | 6.8502e-06
510.9005 | 5.4112e-06 | 294 | 1.2018e-05 | 388.2 | 5.9122e-06 | 670.5 | 6.8636e-06
511.1999 | 5.4147e-06 | 294.1 | 1.197e-05 | 388.4 | 5.905e-06 | 671 | 6.8782e-06
511.4992 | 5.4176e-06 | 294.2 | 1.1921e-05 | 388.6 | 5.898e-06 | 671.5 | 6.8945e-06
511.7985 | 5.4208e-06 | 294.3 | 1.1873e-05 | 388.8 | 5.8902e-06 | 672 | 6.9129e-06
512.0977 | 5.4241e-06 | 294.4 | 1.1824e-05 | 389 | 5.882e-06 | 672.5 | 6.9315e-06
512.3968 | 5.427e-06 | 294.5 | 1.1777e-05 | 389.2 | 5.8735e-06 | 673 | 6.949e-06
512.6959 | 5.4301e-06 | 294.6 | 1.1728e-05 | 389.4 | 5.8656e-06 | 673.5 | 6.9648e-06
512.9949 | 5.4336e-06 | 294.7 | 1.168e-05 | 389.6 | 5.8573e-06 | 674 | 6.98e-06
513.2938 | 5.4364e-06 | 294.8 | 1.1634e-05 | 389.8 | 5.8487e-06 | 674.5 | 6.9961e-06
513.5926 | 5.439e-06 | 294.9 | 1.1587e-05 | 390 | 5.8399e-06 | 675 | 7.0109e-06
513.8914 | 5.443e-06 | 295 | 1.154e-05 | 390.2 | 5.831e-06 | 675.5 | 7.0266e-06
514.1901 | 5.446e-06 | 295.1 | 1.1493e-05 | 390.4 | 5.823e-06 | 676 | 7.0435e-06
514.4888 | 5.4495e-06 | 295.2 | 1.1454e-05 | 390.6 | 5.8154e-06 | 676.5 | 7.0619e-06
514.7874 | 5.4521e-06 | 295.3 | 1.1409e-05 | 390.8 | 5.8089e-06 | 677 | 7.0793e-06
515.0859 | 5.4545e-06 | 295.4 | 1.1365e-05 | 391 | 5.8018e-06 | 677.5 | 7.0958e-06
515.3843 | 5.4579e-06 | 295.5 | 1.1323e-05 | 391.2 | 5.7943e-06 | 678 | 7.1106e-06
515.6827 | 5.4611e-06 | 295.6 | 1.1281e-05 | 391.4 | 5.7857e-06 | 678.5 | 7.1256e-06
515.981 | 5.4642e-06 | 295.7 | 1.1238e-05 | 391.6 | 5.777e-06 | 679 | 7.1426e-06
516.2792 | 5.4669e-06 | 295.8 | 1.1196e-05 | 391.8 | 5.7679e-06 | 679.5 | 7.1624e-06
516.5773 | 5.4702e-06 | 295.9 | 1.1153e-05 | 392 | 5.759e-06 | 680 | 7.1844e-06
516.8754 | 5.4742e-06 | 296 | 1.1111e-05 | 392.2 | 5.7504e-06 | 680.5 | 7.2052e-06
517.1735 | 5.477e-06 | 296.1 | 1.1068e-05 | 392.4 | 5.7424e-06 | 681 | 7.2258e-06
517.4714 | 5.4806e-06 | 296.2 | 1.1031e-05 | 392.6 | 5.7348e-06 | 681.5 | 7.2465e-06
517.7693 | 5.4833e-06 | 296.3 | 1.0991e-05 | 392.8 | 5.7273e-06 | 682 | 7.2666e-06
518.0671 | 5.4869e-06 | 296.4 | 1.0951e-05 | 393 | 5.7191e-06 | 682.5 | 7.2853e-06
518.3648 | 5.4901e-06 | 296.5 | 1.0911e-05 | 393.2 | 5.7114e-06 | 683 | 7.305e-06
518.6625 | 5.4928e-06 | 296.6 | 1.0877e-05 | 393.4 | 5.704e-06 | 683.5 | 7.3284e-06
518.9601 | 5.4953e-06 | 296.7 | 1.084e-05 | 393.6 | 5.6965e-06 | 684 | 7.3517e-06
519.2576 | 5.4984e-06 | 296.8 | 1.0803e-05 | 393.8 | 5.6893e-06 | 684.5 | 7.3711e-06
519.5551 | 5.5018e-06 | 296.9 | 1.0767e-05 | 394 | 5.6818e-06 | 685 | 7.3895e-06
519.8525 | 5.5042e-06 | 297 | 1.0731e-05 | 394.2 | 5.6738e-06 | 685.5 | 7.406e-06
520.1498 | 5.5067e-06 | 297.1 | 1.0695e-05 | 394.4 | 5.6651e-06 | 686 | 7.422e-06
520.4471 | 5.5094e-06 | 297.2 | 1.0659e-05 | 394.6 | 5.6571e-06 | 686.5 | 7.4392e-06
520.7443 | 5.5112e-06 | 297.3 | 1.0622e-05 | 394.8 | 5.6492e-06 | 687 | 7.4599e-06
521.0414 | 5.5144e-06 | 297.4 | 1.0586e-05 | 395 | 5.6416e-06 | 687.5 | 7.4833e-06
521.3384 | 5.5172e-06 | 297.5 | 1.055e-05 | 395.2 | 5.6343e-06 | 688 | 7.5048e-06
521.6354 | 5.5205e-06 | 297.6 | 1.0512e-05 | 395.4 | 5.627e-06 | 688.5 | 7.5266e-06
521.9323 | 5.5239e-06 | 297.7 | 1.0475e-05 | 395.6 | 5.6191e-06 | 689 | 7.5465e-06
522.2292 | 5.5275e-06 | 297.8 | 1.0439e-05 | 395.8 | 5.6112e-06 | 689.5 | 7.564e-06
522.5259 | 5.5315e-06 | 297.9 | 1.0402e-05 | 396 | 5.6032e-06 | 690 | 7.5795e-06
522.8226 | 5.5346e-06 | 298 | 1.0369e-05 | 396.2 | 5.5956e-06 | 690.5 | 7.5958e-06
523.1193 | 5.5384e-06 | 298.1 | 1.0334e-05 | 396.4 | 5.5883e-06 | 691 | 7.6158e-06
523.4158 | 5.5425e-06 | 298.2 | 1.0298e-05 | 396.6 | 5.5813e-06 | 691.5 | 7.6382e-06
523.7123 | 5.5458e-06 | 298.3 | 1.0267e-05 | 396.8 | 5.5741e-06 | 692 | 7.6616e-06
524.0087 | 5.5503e-06 | 298.4 | 1.0233e-05 | 397 | 5.5666e-06 | 692.5 | 7.6846e-06
524.3051 | 5.5543e-06 | 298.5 | 1.02e-05 | 397.2 | 5.5596e-06 | 693 | 7.707e-06
524.6014 | 5.5582e-06 | 298.6 | 1.0167e-05 | 397.4 | 5.5527e-06 | 693.5 | 7.7292e-06
524.8976 | 5.5619e-06 | 298.7 | 1.0136e-05 | 397.6 | 5.5462e-06 | 694 | 7.7495e-06
525.1938 | 5.5652e-06 | 298.8 | 1.0104e-05 | 397.8 | 5.5398e-06 | 694.5 | 7.7671e-06
525.4898 | 5.5685e-06 | 298.9 | 1.0072e-05 | 398 | 5.5335e-06 | 695 | 7.7866e-06
525.7858 | 5.5717e-06 | 299 | 1.0041e-05 | 398.2 | 5.527e-06 | 695.5 | 7.8092e-06
526.0818 | 5.5758e-06 | 299.1 | 1.0009e-05 | 398.4 | 5.5207e-06 | 696 | 7.8336e-06
526.3777 | 5.5794e-06 | 299.2 | 9.9775e-06 | 398.6 | 5.5147e-06 | 696.5 | 7.8584e-06
526.6735 | 5.5828e-06 | 299.3 | 9.9459e-06 | 398.8 | 5.5083e-06 | 697 | 7.8834e-06
526.9692 | 5.5872e-06 | 299.4 | 9.9152e-06 | 399 | 5.5022e-06 | 697.5 | 7.9102e-06
527.2649 | 5.5902e-06 | 299.5 | 9.8838e-06 | 399.2 | 5.4963e-06 | 698 | 7.9348e-06
527.5605 | 5.5933e-06 | 299.6 | 9.8525e-06 | 399.4 | 5.4902e-06 | 698.5 | 7.9576e-06
527.856 | 5.5973e-06 | 299.7 | 9.8271e-06 | 399.6 | 5.4843e-06 | 699 | 7.9802e-06
528.1515 | 5.6011e-06 | 299.8 | 9.7983e-06 | 399.8 | 5.4777e-06 | 699.5 | 8.0039e-06
528.4469 | 5.6041e-06 | 299.9 | 9.7694e-06 | 400 | 5.4708e-06 | 700 | 8.0263e-06
528.7422 | 5.6076e-06 | 300 | 9.7405e-06 | 400.2 | 5.4639e-06 | 700.5 | 8.0486e-06
529.0375 | 5.6105e-06 | 300.1 | 9.712e-06 | 400.4 | 5.4569e-06 | 701 | 8.0727e-06
529.3326 | 5.6125e-06 | 300.2 | 9.6841e-06 | 400.6 | 5.45e-06 | 701.5 | 8.0946e-06
529.6277 | 5.6159e-06 | 300.3 | 9.6561e-06 | 400.8 | 5.4432e-06 | 702 | 8.1142e-06
529.9228 | 5.6191e-06 | 300.4 | 9.6276e-06 | 401 | 5.4362e-06 | 702.5 | 8.1327e-06
530.2178 | 5.6232e-06 | 300.5 | 9.5989e-06 | 401.2 | 5.4292e-06 | 703 | 8.1512e-06
530.5127 | 5.627e-06 | 300.6 | 9.5702e-06 | 401.4 | 5.4222e-06 | 703.5 | 8.1686e-06
530.8075 | 5.6308e-06 | 300.7 | 9.5414e-06 | 401.6 | 5.4152e-06 | 704 | 8.1897e-06
531.1023 | 5.6348e-06 | 300.8 | 9.5169e-06 | 401.8 | 5.4084e-06 | 704.5 | 8.2139e-06
531.397 | 5.6384e-06 | 300.9 | 9.4896e-06 | 402 | 5.4016e-06 | 705 | 8.24e-06
531.6917 | 5.6426e-06 | 301 | 9.4624e-06 | 402.2 | 5.3948e-06 | 705.5 | 8.2668e-06
531.9862 | 5.6466e-06 | 301.1 | 9.4418e-06 | 402.4 | 5.3881e-06 | 706 | 8.2935e-06
532.2807 | 5.6503e-06 | 301.2 | 9.4176e-06 | 402.6 | 5.3813e-06 | 706.5 | 8.3194e-06
532.5752 | 5.6554e-06 | 301.3 | 9.3933e-06 | 402.8 | 5.3746e-06 | 707 | 8.3422e-06
532.8695 | 5.659e-06 | 301.4 | 9.369e-06 | 403 | 5.3684e-06 | 707.5 | 8.3623e-06
533.1638 | 5.6637e-06 | 301.5 | 9.3482e-06 | 403.2 | 5.3624e-06 | 708 | 8.3821e-06
533.458 | 5.6673e-06 | 301.6 | 9.3258e-06 | 403.4 | 5.3561e-06 | 708.5 | 8.4041e-06
533.7522 | 5.6716e-06 | 301.7 | 9.3033e-06 | 403.6 | 5.3498e-06 | 709 | 8.427e-06
534.0463 | 5.6759e-06 | 301.8 | 9.2817e-06 | 403.8 | 5.3433e-06 | 709.5 | 8.4518e-06
534.3403 | 5.6798e-06 | 301.9 | 9.2595e-06 | 404 | 5.3365e-06 | 710 | 8.4829e-06
534.6343 | 5.6843e-06 | 302 | 9.2373e-06 | 404.2 | 5.3297e-06 | 710.5 | 8.5154e-06
534.9282 | 5.688e-06 | 302.1 | 9.2152e-06 | 404.4 | 5.3226e-06 | 711 | 8.5478e-06
535.222 | 5.6913e-06 | 302.2 | 9.1928e-06 | 404.6 | 5.3155e-06 | 711.5 | 8.5795e-06
535.5157 | 5.6954e-06 | 302.3 | 9.1708e-06 | 404.8 | 5.3083e-06 | 712 | 8.6089e-06
535.8094 | 5.6994e-06 | 302.4 | 9.1487e-06 | 405 | 5.3014e-06 | 712.5 | 8.6343e-06
536.103 | 5.7036e-06 | 302.5 | 9.128e-06 | 405.2 | 5.2945e-06 | 713 | 8.6596e-06
536.3966 | 5.7069e-06 | 302.6 | 9.1063e-06 | 405.4 | 5.2874e-06 | 713.5 | 8.689e-06
536.69 | 5.7117e-06 | 302.7 | 9.0846e-06 | 405.6 | 5.2805e-06 | 714 | 8.7199e-06
536.9834 | 5.716e-06 | 302.8 | 9.063e-06 | 405.8 | 5.2735e-06 | 714.5 | 8.7557e-06
537.2768 | 5.721e-06 | 302.9 | 9.0378e-06 | 406 | 5.266e-06 | 715 | 8.7938e-06
537.5701 | 5.7252e-06 | 303 | 9.0153e-06 | 406.2 | 5.2585e-06 | 715.5 | 8.8344e-06
537.8632 | 5.729e-06 | 303.1 | 8.9928e-06 | 406.4 | 5.2507e-06 | 716 | 8.8737e-06
538.1564 | 5.7338e-06 | 303.2 | 8.9708e-06 | 406.6 | 5.2428e-06 | 716.5 | 8.9109e-06
538.4494 | 5.7383e-06 | 303.3 | 8.9479e-06 | 406.8 | 5.2349e-06 | 717 | 8.9443e-06
538.7424 | 5.743e-06 | 303.4 | 8.925e-06 | 407 | 5.2274e-06 | 717.5 | 8.9755e-06
539.0354 | 5.7479e-06 | 303.5 | 8.902e-06 | 407.2 | 5.2199e-06 | 718 | 9.0083e-06
539.3282 | 5.7529e-06 | 303.6 | 8.8824e-06 | 407.4 | 5.2132e-06 | 718.5 | 9.0427e-06
539.621 | 5.7583e-06 | 303.7 | 8.8609e-06 | 407.6 | 5.2066e-06 | 719 | 9.0786e-06
539.9138 | 5.7617e-06 | 303.8 | 8.8394e-06 | 407.8 | 5.2002e-06 | 719.5 | 9.118e-06
540.2064 | 5.7665e-06 | 303.9 | 8.8197e-06 | 408 | 5.1937e-06 | 720 | 9.1635e-06
540.499 | 5.7702e-06 | 304 | 8.799e-06 | 408.2 | 5.187e-06 | 720.5 | 9.2132e-06
540.7915 | 5.7746e-06 | 304.1 | 8.7782e-06 | 408.4 | 5.1803e-06 | 721 | 9.2633e-06
541.084 | 5.7803e-06 | 304.2 | 8.7575e-06 | 408.6 | 5.1732e-06 | 721.5 | 9.3112e-06
541.3764 | 5.7852e-06 | 304.3 | 8.7385e-06 | 408.8 | 5.166e-06 | 722 | 9.3558e-06
541.6687 | 5.7901e-06 | 304.4 | 8.7185e-06 | 409 | 5.1586e-06 | 722.5 | 9.4004e-06
541.9609 | 5.795e-06 | 304.5 | 8.6985e-06 | 409.2 | 5.1511e-06 | 723 | 9.438e-06
542.2531 | 5.7992e-06 | 304.6 | 8.6824e-06 | 409.4 | 5.1436e-06 | 723.5 | 9.4662e-06
542.5452 | 5.8038e-06 | 304.7 | 8.6639e-06 | 409.6 | 5.1365e-06 | 724 | 9.4938e-06
542.8373 | 5.808e-06 | 304.8 | 8.6453e-06 | 409.8 | 5.1295e-06 | 724.5 | 9.5258e-06
543.1292 | 5.8127e-06 | 304.9 | 8.626e-06 | 410 | 5.1228e-06 | 725 | 9.5548e-06
543.4212 | 5.8173e-06 | 305 | 8.6077e-06 | 410.2 | 5.1162e-06 | 725.5 | 9.5821e-06
543.713 | 5.8229e-06 | 305.1 | 8.5895e-06 | 410.4 | 5.1096e-06 | 726 | 9.6143e-06
544.0048 | 5.8273e-06 | 305.2 | 8.5712e-06 | 410.6 | 5.1029e-06 | 726.5 | 9.6482e-06
544.2965 | 5.8307e-06 | 305.3 | 8.5576e-06 | 410.8 | 5.0964e-06 | 727 | 9.6748e-06
544.5881 | 5.8352e-06 | 305.4 | 8.5406e-06 | 411 | 5.09e-06 | 727.5 | 9.7011e-06
544.8797 | 5.8397e-06 | 305.5 | 8.5235e-06 | 411.2 | 5.0838e-06 | 728 | 9.7318e-06
545.1712 | 5.8444e-06 | 305.6 | 8.5084e-06 | 411.4 | 5.0777e-06 | 728.5 | 9.7603e-06
545.4626 | 5.8498e-06 | 305.7 | 8.4933e-06 | 411.6 | 5.072e-06 | 729 | 9.7831e-06
545.754 | 5.8545e-06 | 305.8 | 8.4781e-06 | 411.8 | 5.0662e-06 | 729.5 | 9.8064e-06
546.0453 | 5.8593e-06 | 305.9 | 8.4629e-06 | 412 | 5.0603e-06 | 730 | 9.8313e-06
546.3365 | 5.8649e-06 | 306 | 8.4458e-06 | 412.2 | 5.0541e-06 | 730.5 | 9.8517e-06
546.6277 | 5.8701e-06 | 306.1 | 8.4298e-06 | 412.4 | 5.0484e-06 | 731 | 9.8667e-06
546.9188 | 5.8742e-06 | 306.2 | 8.4138e-06 | 412.6 | 5.0427e-06 | 731.5 | 9.886e-06
547.2098 | 5.8793e-06 | 306.3 | 8.3978e-06 | 412.8 | 5.0367e-06 | 732 | 9.9087e-06
547.5008 | 5.8852e-06 | 306.4 | 8.3816e-06 | 413 | 5.0308e-06 | 732.5 | 9.9324e-06
547.7917 | 5.8899e-06 | 306.5 | 8.3654e-06 | 413.2 | 5.0252e-06 | 733 | 9.9578e-06
548.0825 | 5.8958e-06 | 306.6 | 8.3492e-06 | 413.4 | 5.0194e-06 | 733.5 | 9.9877e-06
548.3733 | 5.9012e-06 | 306.7 | 8.3389e-06 | 413.6 | 5.0135e-06 | 734 | 1.0019e-05
548.664 | 5.9061e-06 | 306.8 | 8.3247e-06 | 413.8 | 5.0073e-06 | 734.5 | 1.0048e-05
548.9546 | 5.9109e-06 | 306.9 | 8.3106e-06 | 414 | 5.0015e-06 | 735 | 1.0076e-05
549.2451 | 5.916e-06 | 307 | 8.2967e-06 | 414.2 | 4.9959e-06 | 735.5 | 1.0103e-05
549.5356 | 5.9211e-06 | 307.1 | 8.2834e-06 | 414.4 | 4.9899e-06 | 736 | 1.0129e-05
549.8261 | 5.9255e-06 | 307.2 | 8.2702e-06 | 414.6 | 4.9839e-06 | 736.5 | 1.016e-05
550.1164 | 5.9312e-06 | 307.3 | 8.257e-06 | 414.8 | 4.9775e-06 | 737 | 1.0193e-05
550.4067 | 5.937e-06 | 307.4 | 8.2386e-06 | 415 | 4.9709e-06 | 737.5 | 1.0228e-05
550.6969 | 5.9424e-06 | 307.5 | 8.2233e-06 | 415.2 | 4.9643e-06 | 738 | 1.0267e-05
550.9871 | 5.9476e-06 | 307.6 | 8.208e-06 | 415.4 | 4.9582e-06 | 738.5 | 1.0303e-05
551.2771 | 5.9536e-06 | 307.7 | 8.1959e-06 | 415.6 | 4.9523e-06 | 739 | 1.0334e-05
551.5672 | 5.9584e-06 | 307.8 | 8.1814e-06 | 415.8 | 4.9467e-06 | 739.5 | 1.036e-05
551.8571 | 5.9636e-06 | 307.9 | 8.1668e-06 | 416 | 4.9408e-06 | 740 | 1.0385e-05
552.147 | 5.9696e-06 | 308 | 8.1523e-06 | 416.2 | 4.9348e-06 | 740.5 | 1.041e-05
552.4368 | 5.9754e-06 | 308.1 | 8.1413e-06 | 416.4 | 4.9289e-06 | 741 | 1.0434e-05
552.7266 | 5.9804e-06 | 308.2 | 8.1286e-06 | 416.6 | 4.9227e-06 | 741.5 | 1.0466e-05
553.0162 | 5.9856e-06 | 308.3 | 8.1159e-06 | 416.8 | 4.9166e-06 | 742 | 1.0504e-05
553.3058 | 5.991e-06 | 308.4 | 8.1006e-06 | 417 | 4.9105e-06 | 742.5 | 1.0546e-05
553.5954 | 5.9967e-06 | 308.5 | 8.0869e-06 | 417.2 | 4.9047e-06 | 743 | 1.0588e-05
553.8849 | 6.0031e-06 | 308.6 | 8.0732e-06 | 417.4 | 4.8995e-06 | 743.5 | 1.0631e-05
554.1743 | 6.0098e-06 | 308.7 | 8.0595e-06 | 417.6 | 4.8942e-06 | 744 | 1.0668e-05
554.4636 | 6.0147e-06 | 308.8 | 8.0469e-06 | 417.8 | 4.8889e-06 | 744.5 | 1.0697e-05
554.7529 | 6.0209e-06 | 308.9 | 8.0332e-06 | 418 | 4.8835e-06 | 745 | 1.0725e-05
555.0421 | 6.0269e-06 | 309 | 8.0195e-06 | 418.2 | 4.8784e-06 | 745.5 | 1.0754e-05
555.3313 | 6.0321e-06 | 309.1 | 8.0078e-06 | 418.4 | 4.8735e-06 | 746 | 1.0788e-05
555.6203 | 6.0376e-06 | 309.2 | 7.9952e-06 | 418.6 | 4.8688e-06 | 746.5 | 1.0826e-05
555.9093 | 6.0437e-06 | 309.3 | 7.9826e-06 | 418.8 | 4.8639e-06 | 747 | 1.0871e-05
556.1983 | 6.0499e-06 | 309.4 | 7.9701e-06 | 419 | 4.8587e-06 | 747.5 | 1.0919e-05
556.4872 | 6.0554e-06 | 309.5 | 7.9565e-06 | 419.2 | 4.8538e-06 | 748 | 1.0966e-05
556.776 | 6.0607e-06 | 309.6 | 7.9436e-06 | 419.4 | 4.849e-06 | 748.5 | 1.1008e-05
557.0647 | 6.0649e-06 | 309.7 | 7.9308e-06 | 419.6 | 4.8443e-06 | 749 | 1.1047e-05
557.3534 | 6.0703e-06 | 309.8 | 7.9178e-06 | 419.8 | 4.84e-06 | 749.5 | 1.1085e-05
557.642 | 6.0768e-06 | 309.9 | 7.9047e-06 | 420 | 4.8359e-06 | 750 | 1.1117e-05
557.9305 | 6.0835e-06 | 310 | 7.8917e-06 | 420.2 | 4.8322e-06 | 750.5 | 1.1151e-05
558.219 | 6.09e-06 | 310.1 | 7.8809e-06 | 420.4 | 4.8285e-06 | 751 | 1.1189e-05
558.5074 | 6.0963e-06 | 310.2 | 7.8689e-06 | 420.6 | 4.825e-06 | 751.5 | 1.1233e-05
558.7957 | 6.1018e-06 | 310.3 | 7.8569e-06 | 420.8 | 4.821e-06 | 752 | 1.1275e-05
559.084 | 6.1072e-06 | 310.4 | 7.8449e-06 | 421 | 4.8168e-06 | 752.5 | 1.1315e-05
559.3722 | 6.1128e-06 | 310.5 | 7.8318e-06 | 421.2 | 4.8124e-06 | 753 | 1.1354e-05
559.6604 | 6.1196e-06 | 310.6 | 7.8197e-06 | 421.4 | 4.8081e-06 | 753.5 | 1.1392e-05
559.9484 | 6.1256e-06 | 310.7 | 7.8077e-06 | 421.6 | 4.8037e-06 | 754 | 1.1424e-05
560.2364 | 6.132e-06 | 310.8 | 7.7972e-06 | 421.8 | 4.7993e-06 | 754.5 | 1.1459e-05
560.5244 | 6.1386e-06 | 310.9 | 7.7854e-06 | 422 | 4.7951e-06 | 755 | 1.1504e-05
560.8122 | 6.1444e-06 | 311 | 7.7736e-06 | 422.2 | 4.7913e-06 | 755.5 | 1.1556e-05
561.1 | 6.15e-06 | 311.1 | 7.7617e-06 | 422.4 | 4.7879e-06 | 756 | 1.1607e-05
561.3878 | 6.1559e-06 | 311.2 | 7.754e-06 | 422.6 | 4.7847e-06 | 756.5 | 1.1654e-05
561.6754 | 6.1617e-06 | 311.3 | 7.744e-06 | 422.8 | 4.7815e-06 | 757 | 1.17e-05
561.9631 | 6.1684e-06 | 311.4 | 7.734e-06 | 423 | 4.7803e-06 | 757.5 | 1.1744e-05
562.2506 | 6.1753e-06 | 311.5 | 7.7229e-06 | 423.2 | 4.7808e-06 | 758 | 1.1785e-05
562.5381 | 6.1816e-06 | 311.6 | 7.7128e-06 | 423.4 | 4.784e-06 | 758.5 | 1.1823e-05
562.8255 | 6.187e-06 | 311.7 | 7.7026e-06 | 423.6 | 4.7883e-06 | 759 | 1.1869e-05
563.1128 | 6.1929e-06 | 311.8 | 7.6925e-06 | 423.8 | 4.7953e-06 | 759.5 | 1.1922e-05
563.4001 | 6.1988e-06 | 311.9 | 7.6804e-06 | 424 | 4.8042e-06 | 760 | 1.1972e-05
563.6873 | 6.2044e-06 | 312 | 7.6693e-06 | 424.2 | 4.8153e-06 | 760.5 | 1.2022e-05
563.9744 | 6.2105e-06 | 312.1 | 7.6582e-06 | 424.4 | 4.829e-06 | 761 | 1.2074e-05
564.2615 | 6.217e-06 | 312.2 | 7.6487e-06 | 424.6 | 4.8456e-06 | 761.5 | 1.2124e-05
564.5485 | 6.2224e-06 | 312.3 | 7.6382e-06 | 424.8 | 4.8657e-06 | 762 | 1.2165e-05
564.8354 | 6.2283e-06 | 312.4 | 7.6277e-06 | 425 | 4.8899e-06 | 762.5 | 1.2205e-05
565.1223 | 6.2342e-06 | 312.5 | 7.6172e-06 | 425.2 | 4.9152e-06 | 763 | 1.2248e-05
565.4091 | 6.2397e-06 | 312.6 | 7.6081e-06 | 425.4 | 4.9444e-06 | 763.5 | 1.2288e-05
565.6958 | 6.2454e-06 | 312.7 | 7.5983e-06 | 425.6 | 4.9712e-06 | 764 | 1.2334e-05
565.9825 | 6.2503e-06 | 312.8 | 7.5885e-06 | 425.8 | 4.9969e-06 | 764.5 | 1.2386e-05
566.2691 | 6.2561e-06 | 312.9 | 7.5755e-06 | 426 | 5.0202e-06 | 765 | 1.2439e-05
566.5557 | 6.2619e-06 | 313 | 7.5645e-06 | 426.2 | 5.0426e-06 | 765.5 | 1.2489e-05
566.8421 | 6.2678e-06 | 313.1 | 7.5535e-06 | 426.4 | 5.0632e-06 | 766 | 1.2539e-05
567.1285 | 6.2729e-06 | 313.2 | 7.5426e-06 | 426.6 | 5.0816e-06 | 766.5 | 1.2584e-05
567.4149 | 6.2778e-06 | 313.3 | 7.5349e-06 | 426.8 | 5.0981e-06 | 767 | 1.2625e-05
567.7011 | 6.2843e-06 | 313.4 | 7.5245e-06 | 427 | 5.1097e-06 | 767.5 | 1.267e-05
567.9873 | 6.2903e-06 | 313.5 | 7.5142e-06 | 427.2 | 5.1177e-06 | 768 | 1.2722e-05
568.2735 | 6.297e-06 | 313.6 | 7.507e-06 | 427.4 | 5.1215e-06 | 768.5 | 1.278e-05
568.5596 | 6.3025e-06 | 313.7 | 7.4985e-06 | 427.6 | 5.1203e-06 | 769 | 1.2844e-05
568.8456 | 6.3091e-06 | 313.8 | 7.4901e-06 | 427.8 | 5.1139e-06 | 769.5 | 1.2907e-05
569.1315 | 6.3159e-06 | 313.9 | 7.4817e-06 | 428 | 5.1022e-06 | 770 | 1.2961e-05
569.4174 | 6.3226e-06 | 314 | 7.4692e-06 | 428.2 | 5.0852e-06 | 770.5 | 1.3006e-05
569.7032 | 6.3296e-06 | 314.1 | 7.4595e-06 | 428.4 | 5.0654e-06 | 771 | 1.3049e-05
569.9889 | 6.3361e-06 | 314.2 | 7.4499e-06 | 428.6 | 5.0383e-06 | 771.5 | 1.3092e-05
570.2746 | 6.3425e-06 | 314.3 | 7.441e-06 | 428.8 | 5.0084e-06 | 772 | 1.3138e-05
570.5602 | 6.3498e-06 | 314.4 | 7.4311e-06 | 429 | 4.9764e-06 | 772.5 | 1.3195e-05
570.8458 | 6.3563e-06 | 314.5 | 7.4212e-06 | 429.2 | 4.9448e-06 | 773 | 1.3258e-05
571.1312 | 6.3635e-06 | 314.6 | 7.4109e-06 | 429.4 | 4.9148e-06 | 773.5 | 1.332e-05
571.4166 | 6.3694e-06 | 314.7 | 7.4013e-06 | 429.6 | 4.8867e-06 | 774 | 1.3373e-05
571.702 | 6.3762e-06 | 314.8 | 7.3918e-06 | 429.8 | 4.8613e-06 | 774.5 | 1.342e-05
571.9873 | 6.3823e-06 | 314.9 | 7.3822e-06 | 430 | 4.8402e-06 | 775 | 1.3467e-05
572.2725 | 6.3891e-06 | 315 | 7.3742e-06 | 430.2 | 4.8215e-06 | 775.5 | 1.3521e-05
572.5576 | 6.3955e-06 | 315.1 | 7.3648e-06 | 430.4 | 4.808e-06 | 776 | 1.3579e-05
572.8427 | 6.4017e-06 | 315.2 | 7.3555e-06 | 430.6 | 4.7978e-06 | 776.5 | 1.3645e-05
573.1277 | 6.4075e-06 | 315.3 | 7.3511e-06 | 430.8 | 4.7898e-06 | 777 | 1.3719e-05
573.4127 | 6.4144e-06 | 315.4 | 7.3442e-06 | 431 | 4.7835e-06 | 777.5 | 1.3786e-05
573.6975 | 6.4204e-06 | 315.5 | 7.3373e-06 | 431.2 | 4.7795e-06 | 778 | 1.3834e-05
573.9824 | 6.4265e-06 | 315.6 | 7.3304e-06 | 431.4 | 4.7779e-06 | 778.5 | 1.3877e-05
574.2671 | 6.4326e-06 | 315.7 | 7.3238e-06 | 431.6 | 4.7784e-06 | 779 | 1.393e-05
574.5518 | 6.4395e-06 | 315.8 | 7.3176e-06 | 431.8 | 4.7804e-06 | 779.5 | 1.3989e-05
574.8364 | 6.4464e-06 | 315.9 | 7.3114e-06 | 432 | 4.7853e-06 | 780 | 1.405e-05
575.121 | 6.4536e-06 | 316 | 7.3055e-06 | 432.2 | 4.7927e-06 | 780.5 | 1.4118e-05
575.4055 | 6.4602e-06 | 316.1 | 7.299e-06 | 432.4 | 4.8028e-06 | 781 | 1.4181e-05
575.6899 | 6.467e-06 | 316.2 | 7.2926e-06 | 432.6 | 4.8154e-06 | 781.5 | 1.4237e-05
575.9742 | 6.4731e-06 | 316.3 | 7.2861e-06 | 432.8 | 4.8302e-06 | 782 | 1.4301e-05
576.2585 | 6.4798e-06 | 316.4 | 7.2801e-06 | 433 | 4.8464e-06 | 782.5 | 1.4372e-05
576.5427 | 6.4865e-06 | 316.5 | 7.274e-06 | 433.2 | 4.8621e-06 | 783 | 1.4437e-05
576.8269 | 6.4934e-06 | 316.6 | 7.268e-06 | 433.4 | 4.8767e-06 | 783.5 | 1.4498e-05
577.111 | 6.5007e-06 | 316.7 | 7.2644e-06 | 433.6 | 4.8886e-06 | 784 | 1.4552e-05
577.395 | 6.5069e-06 | 316.8 | 7.2592e-06 | 433.8 | 4.8994e-06 | 784.5 | 1.4605e-05
577.679 | 6.5141e-06 | 316.9 | 7.254e-06 | 434 | 4.9067e-06 | 785 | 1.4668e-05
577.9629 | 6.5211e-06 | 317 | 7.2488e-06 | 434.2 | 4.9136e-06 | 785.5 | 1.474e-05
578.2467 | 6.5279e-06 | 317.1 | 7.2433e-06 | 434.4 | 4.9202e-06 | 786 | 1.4821e-05
578.5305 | 6.5347e-06 | 317.2 | 7.2384e-06 | 434.6 | 4.9275e-06 | 786.5 | 1.4898e-05
578.8142 | 6.542e-06 | 317.3 | 7.2335e-06 | 434.8 | 4.935e-06 | 787 | 1.4968e-05
579.0978 | 6.5496e-06 | 317.4 | 7.2274e-06 | 435 | 4.9428e-06 | 787.5 | 1.5031e-05
579.3814 | 6.5563e-06 | 317.5 | 7.2217e-06 | 435.2 | 4.9505e-06 | 788 | 1.5098e-05
579.6649 | 6.5628e-06 | 317.6 | 7.216e-06 | 435.4 | 4.9572e-06 | 788.5 | 1.517e-05
579.9483 | 6.5702e-06 | 317.7 | 7.2119e-06 | 435.6 | 4.9634e-06 | 789 | 1.5242e-05
580.2317 | 6.5765e-06 | 317.8 | 7.2069e-06 | 435.8 | 4.9682e-06 | 789.5 | 1.5308e-05
580.515 | 6.5841e-06 | 317.9 | 7.2019e-06 | 436 | 4.9719e-06 | 790 | 1.5371e-05
580.7982 | 6.5908e-06 | 318 | 7.1969e-06 | 436.2 | 4.9749e-06 | 790.5 | 1.5435e-05
581.0814 | 6.5987e-06 | 318.1 | 7.1885e-06 | 436.4 | 4.9769e-06 | 791 | 1.5494e-05
581.3645 | 6.6054e-06 | 318.2 | 7.1825e-06 | 436.6 | 4.9781e-06 | 791.5 | 1.556e-05
581.6475 | 6.6134e-06 | 318.3 | 7.1766e-06 | 436.8 | 4.9787e-06 | 792 | 1.5644e-05
581.9305 | 6.6202e-06 | 318.4 | 7.1697e-06 | 437 | 4.978e-06 | 792.5 | 1.5724e-05
582.2134 | 6.6278e-06 | 318.5 | 7.1626e-06 | 437.2 | 4.9769e-06 | 793 | 1.5795e-05
582.4963 | 6.6371e-06 | 318.6 | 7.1555e-06 | 437.4 | 4.9746e-06 | 793.5 | 1.5867e-05
582.779 | 6.6458e-06 | 318.7 | 7.1484e-06 | 437.6 | 4.9716e-06 | 794 | 1.5942e-05
583.0617 | 6.6536e-06 | 318.8 | 7.1482e-06 | 437.8 | 4.9682e-06 | 794.5 | 1.6021e-05
583.3444 | 6.6612e-06 | 318.9 | 7.1435e-06 | 438 | 4.9646e-06 | 795 | 1.6091e-05
583.627 | 6.6692e-06 | 319 | 7.1389e-06 | 438.2 | 4.9607e-06 | 795.5 | 1.615e-05
583.9095 | 6.6778e-06 | 319.1 | 7.1313e-06 | 438.4 | 4.9567e-06 | 796 | 1.6211e-05
584.192 | 6.6855e-06 | 319.2 | 7.1265e-06 | 438.6 | 4.9524e-06 | 796.5 | 1.6288e-05
584.4743 | 6.6939e-06 | 319.3 | 7.1216e-06 | 438.8 | 4.9481e-06 | 797 | 1.636e-05
584.7567 | 6.7021e-06 | 319.4 | 7.1167e-06 | 439 | 4.9442e-06 | 797.5 | 1.6421e-05
585.0389 | 6.7112e-06 | 319.5 | 7.111e-06 | 439.2 | 4.9401e-06 | 798 | 1.6492e-05
585.3211 | 6.7203e-06 | 319.6 | 7.1051e-06 | 439.4 | 4.9363e-06 | 798.5 | 1.6577e-05
585.6032 | 6.7292e-06 | 319.7 | 7.0991e-06 | 439.6 | 4.9319e-06 | 799 | 1.6662e-05
585.8853 | 6.738e-06 | 319.8 | 7.0936e-06 | 439.8 | 4.9272e-06 | 799.5 | 1.6735e-05
586.1673 | 6.7481e-06 | 319.9 | 7.0883e-06 | 440 | 4.9225e-06 | 800 | 1.6818e-05
586.4492 | 6.7591e-06 | 320 | 7.083e-06 | 440.2 | 4.9179e-06 | 800.5 | 1.6913e-05
586.7311 | 6.7698e-06 | 320.1 | 7.0777e-06 | 440.4 | 4.9135e-06 | 801 | 1.7003e-05
587.0129 | 6.7797e-06 | 320.2 | 7.0706e-06 | 440.6 | 4.9091e-06 | 801.5 | 1.7079e-05
587.2946 | 6.7903e-06 | 320.3 | 7.0646e-06 | 440.8 | 4.9046e-06 | 802 | 1.7175e-05
587.5763 | 6.801e-06 | 320.4 | 7.0585e-06 | 441 | 4.9001e-06 | 802.5 | 1.7282e-05
587.8579 | 6.8117e-06 | 320.5 | 7.0496e-06 | 441.2 | 4.8956e-06 | 803 | 1.7385e-05
588.1394 | 6.8233e-06 | 320.6 | 7.0424e-06 | 441.4 | 4.8913e-06 | 803.5 | 1.7484e-05
588.4209 | 6.8345e-06 | 320.7 | 7.0352e-06 | 441.6 | 4.8875e-06 | 804 | 1.7572e-05
588.7023 | 6.8455e-06 | 320.8 | 7.0269e-06 | 441.8 | 4.8835e-06 | 804.5 | 1.7626e-05
588.9837 | 6.8573e-06 | 320.9 | 7.0188e-06 | 442 | 4.8798e-06 | 805 | 1.7664e-05
589.2649 | 6.8685e-06 | 321 | 7.0107e-06 | 442.2 | 4.876e-06 | 805.5 | 1.836e-05
589.5461 | 6.8798e-06 | 321.1 | 7.0026e-06 | 442.4 | 4.8722e-06 | 806 | 2.1786e-05
589.8273 | 6.8908e-06 | 321.2 | 6.9979e-06 | 442.6 | 4.8687e-06 | 806.5 | 2.8374e-05
590.1084 | 6.9028e-06 | 321.3 | 6.991e-06 | 442.8 | 4.8657e-06 | 807 | 3.5886e-05
590.3894 | 6.9137e-06 | 321.4 | 6.984e-06 | 443 | 4.8629e-06 | 807.5 | 5.4055e-05
590.6703 | 6.9257e-06 | 321.5 | 6.9786e-06 | 443.2 | 4.8605e-06
590.9512 | 6.9367e-06 | 321.6 | 6.9729e-06 | 443.4 | 4.8583e-06
591.2321 | 6.947e-06 | 321.7 | 6.9671e-06 | 443.6 | 4.856e-06
591.5128 | 6.957e-06 | 321.8 | 6.9614e-06 | 443.8 | 4.8536e-06
591.7935 | 6.9671e-06 | 321.9 | 6.9581e-06 | 444 | 4.8512e-06
592.0741 | 6.9756e-06 | 322 | 6.9532e-06 | 444.2 | 4.8487e-06
592.3547 | 6.9848e-06 | 322.1 | 6.9483e-06 | 444.4 | 4.8462e-06
592.6352 | 6.9935e-06 | 322.2 | 6.9427e-06 | 444.6 | 4.8441e-06
592.9156 | 7.0024e-06 | 322.3 | 6.938e-06 | 444.8 | 4.842e-06
593.196 | 7.0112e-06 | 322.4 | 6.9332e-06 | 445 | 4.84e-06
593.4763 | 7.0191e-06 | 322.5 | 6.9284e-06 | 445.2 | 4.838e-06
593.7565 | 7.0254e-06 | 322.6 | 6.9193e-06 | 445.4 | 4.836e-06
594.0367 | 7.0328e-06 | 322.7 | 6.913e-06 | 445.6 | 4.834e-06
594.3168 | 7.0392e-06 | 322.8 | 6.9066e-06 | 445.8 | 4.8321e-06
594.5969 | 7.0458e-06 | 322.9 | 6.9045e-06 | 446 | 4.8302e-06
594.8769 | 7.0514e-06 | 323 | 6.8994e-06 | 446.2 | 4.8274e-06
595.1568 | 7.0582e-06 | 323.1 | 6.8942e-06 | 446.4 | 4.8244e-06
595.4366 | 7.0643e-06 | 323.2 | 6.889e-06 | 446.6 | 4.8213e-06
595.7164 | 7.0696e-06 | 323.3 | 6.885e-06 | 446.8 | 4.8185e-06
595.9961 | 7.0733e-06 | 323.4 | 6.8809e-06 | 447 | 4.8159e-06
596.2758 | 7.0764e-06 | 323.5 | 6.8768e-06 | 447.2 | 4.8133e-06
596.5554 | 7.0803e-06 | 323.6 | 6.8709e-06 | 447.4 | 4.8111e-06
596.8349 | 7.0828e-06 | 323.7 | 6.8657e-06 | 447.6 | 4.8089e-06
597.1144 | 7.0851e-06 | 323.8 | 6.8606e-06 | 447.8 | 4.8067e-06
597.3938 | 7.0863e-06 | 323.9 | 6.8557e-06 | 448 | 4.8045e-06
597.6731 | 7.0878e-06 | 324 | 6.8506e-06 | 448.2 | 4.8022e-06
597.9524 | 7.0886e-06 | 324.1 | 6.8456e-06 | 448.4 | 4.7999e-06
598.2316 | 7.0878e-06 | 324.2 | 6.8406e-06 | 448.6 | 4.7976e-06
598.5107 | 7.0865e-06 | 324.3 | 6.8405e-06 | 448.8 | 4.7955e-06
598.7898 | 7.0847e-06 | 324.4 | 6.8372e-06 | 449 | 4.7935e-06
599.0688 | 7.0826e-06 | 324.5 | 6.8339e-06 | 449.2 | 4.7916e-06
599.3478 | 7.0798e-06 | 324.6 | 6.8327e-06 | 449.4 | 4.7898e-06
599.6267 | 7.0754e-06 | 324.7 | 6.831e-06 | 449.6 | 4.7881e-06
599.9055 | 7.0721e-06 | 324.8 | 6.8292e-06 | 449.8 | 4.7865e-06
600.1842 | 7.0674e-06 | 324.9 | 6.8275e-06 | 450 | 4.7851e-06
600.4629 | 7.0621e-06 | 325 | 6.8234e-06 | 450.2 | 4.7837e-06
600.7415 | 7.0549e-06 | 325.1 | 6.821e-06 | 450.4 | 4.7824e-06
601.0201 | 7.0478e-06 | 325.2 | 6.8185e-06 | 450.6 | 4.7812e-06
601.2986 | 7.0402e-06 | 325.3 | 6.8143e-06 | 450.8 | 4.78e-06
601.577 | 7.0316e-06 | 325.4 | 6.8109e-06 | 451 | 4.7785e-06
601.8554 | 7.022e-06 | 325.5 | 6.8074e-06 | 451.2 | 4.7777e-06
602.1337 | 7.0127e-06 | 325.6 | 6.804e-06 | 451.4 | 4.7771e-06
602.412 | 7.0015e-06 | 325.7 | 6.7997e-06 | 451.6 | 4.7767e-06
602.6901 | 6.991e-06 | 325.8 | 6.7959e-06 | 451.8 | 4.7763e-06
602.9682 | 6.9791e-06 | 325.9 | 6.7922e-06 | 452 | 4.7759e-06
603.2463 | 6.9655e-06 | 326 | 6.793e-06 | 452.2 | 4.7754e-06
603.5243 | 6.952e-06 | 326.1 | 6.7908e-06 | 452.4 | 4.7747e-06
603.8022 | 6.9384e-06 | 326.2 | 6.7886e-06 | 452.6 | 4.7739e-06
604.0801 | 6.924e-06 | 326.3 | 6.7875e-06 | 452.8 | 4.7728e-06
604.3578 | 6.9097e-06 | 326.4 | 6.7865e-06 | 453 | 4.7717e-06
604.6356 | 6.8949e-06 | 326.5 | 6.7856e-06 | 453.2 | 4.771e-06
604.9132 | 6.8809e-06 | 326.6 | 6.7847e-06 | 453.4 | 4.7704e-06
605.1908 | 6.8663e-06 | 326.7 | 6.7799e-06 | 453.6 | 4.7699e-06
605.4684 | 6.853e-06 | 326.8 | 6.7776e-06 | 453.8 | 4.7691e-06
605.7458 | 6.8381e-06 | 326.9 | 6.7752e-06 | 454 | 4.7686e-06
606.0232 | 6.8238e-06 | 327 | 6.7739e-06 | 454.2 | 4.7681e-06
606.3006 | 6.8102e-06 | 327.1 | 6.7715e-06 | 454.4 | 4.7676e-06
606.5779 | 6.7963e-06 | 327.2 | 6.7691e-06 | 454.6 | 4.767e-06
606.8551 | 6.7816e-06 | 327.3 | 6.7667e-06 | 454.8 | 4.7659e-06
607.1322 | 6.7672e-06 | 327.4 | 6.7685e-06 | 455 | 4.7652e-06
607.4093 | 6.7544e-06 | 327.5 | 6.7677e-06 | 455.2 | 4.7644e-06
607.6863 | 6.741e-06 | 327.6 | 6.767e-06 | 455.4 | 4.7639e-06
607.9633 | 6.7282e-06 | 327.7 | 6.7667e-06 | 455.6 | 4.7636e-06
608.2402 | 6.7157e-06 | 327.8 | 6.7664e-06 | 455.8 | 4.7633e-06
608.517 | 6.7025e-06 | 327.9 | 6.7662e-06 | 456 | 4.7628e-06
608.7938 | 6.6907e-06 | 328 | 6.761e-06 | 456.2 | 4.7625e-06
609.0705 | 6.6799e-06 | 328.1 | 6.7584e-06 | 456.4 | 4.7623e-06
609.3472 | 6.6681e-06 | 328.2 | 6.7559e-06 | 456.6 | 4.7623e-06
609.6237 | 6.6571e-06 | 328.3 | 6.7533e-06 | 456.8 | 4.7622e-06
609.9002 | 6.6465e-06 | 328.4 | 6.746e-06 | 457 | 4.7622e-06
610.1767 | 6.6368e-06 | 328.5 | 6.741e-06 | 457.2 | 4.7622e-06
610.4531 | 6.626e-06 | 328.6 | 6.736e-06 | 457.4 | 4.762e-06
610.7294 | 6.617e-06 | 328.7 | 6.7308e-06 | 457.6 | 4.7616e-06
611.0057 | 6.6082e-06 | 328.8 | 6.7254e-06 | 457.8 | 4.7615e-06
611.2818 | 6.5999e-06 | 328.9 | 6.7201e-06 | 458 | 4.7616e-06
611.558 | 6.5916e-06 | 329 | 6.7148e-06 | 458.2 | 4.762e-06
611.8341 | 6.5843e-06 | 329.1 | 6.7132e-06 | 458.4 | 4.7624e-06
612.1101 | 6.576e-06 | 329.2 | 6.7092e-06 | 458.6 | 4.763e-06
612.386 | 6.5695e-06 | 329.3 | 6.7051e-06 | 458.8 | 4.7636e-06
612.6619 | 6.5643e-06 | 329.4 | 6.7064e-06 | 459 | 4.7645e-06
612.9377 | 6.5593e-06 | 329.5 | 6.7049e-06 | 459.2 | 4.7654e-06
613.2134 | 6.554e-06 | 329.6 | 6.7034e-06 | 459.4 | 4.7664e-06
613.4891 | 6.5498e-06 | 329.7 | 6.7019e-06 | 459.6 | 4.7675e-06
613.7647 | 6.5448e-06 | 329.8 | 6.6991e-06 | 459.8 | 4.7686e-06
614.0403 | 6.5398e-06 | 329.9 | 6.6977e-06 | 460 | 4.7699e-06
614.3158 | 6.5362e-06 | 330 | 6.6963e-06 | 460.2 | 4.7708e-06
614.5912 | 6.5332e-06 | 330.1 | 6.6936e-06 | 460.4 | 4.7705e-06
614.8666 | 6.5304e-06 | 330.2 | 6.6912e-06 | 460.6 | 4.7712e-06
615.1419 | 6.5278e-06 | 330.3 | 6.6888e-06 | 460.8 | 4.772e-06
615.4171 | 6.5264e-06 | 330.4 | 6.6871e-06 | 461 | 4.7735e-06
615.6923 | 6.5239e-06 | 330.5 | 6.6852e-06 | 461.2 | 4.775e-06
615.9674 | 6.5221e-06 | 330.6 | 6.6832e-06 | 461.4 | 4.7767e-06
616.2425 | 6.521e-06 | 330.7 | 6.6813e-06 | 461.6 | 4.7785e-06
616.5175 | 6.5204e-06 | 330.8 | 6.6783e-06 | 461.8 | 4.7803e-06
616.7924 | 6.5209e-06 | 330.9 | 6.6761e-06 | 462 | 4.7821e-06
617.0673 | 6.5224e-06 | 331 | 6.6738e-06 | 462.2 | 4.7839e-06
617.3421 | 6.5235e-06 | 331.1 | 6.6695e-06 | 462.4 | 4.7856e-06
617.6168 | 6.525e-06 | 331.2 | 6.6662e-06 | 462.6 | 4.7871e-06
617.8915 | 6.5268e-06 | 331.3 | 6.6629e-06 | 462.8 | 4.7889e-06
618.1661 | 6.5297e-06 | 331.4 | 6.6596e-06 | 463 | 4.7908e-06
618.4406 | 6.5331e-06 | 331.5 | 6.6547e-06 | 463.2 | 4.7927e-06
618.7151 | 6.5383e-06 | 331.6 | 6.6506e-06 | 463.4 | 4.7942e-06
618.9895 | 6.5444e-06 | 331.7 | 6.6466e-06 | 463.6 | 4.7961e-06
619.2639 | 6.5516e-06 | 331.8 | 6.6461e-06 | 463.8 | 4.7988e-06
619.5382 | 6.5599e-06 | 331.9 | 6.6434e-06 | 464 | 4.801e-06
619.8124 | 6.5687e-06 | 332 | 6.6406e-06 | 464.2 | 4.8034e-06
620.0866 | 6.5775e-06 | 332.1 | 6.642e-06 | 464.4 | 4.8055e-06
620.3607 | 6.5865e-06 | 332.2 | 6.6418e-06 | 464.6 | 4.8074e-06
620.6347 | 6.5976e-06 | 332.3 | 6.6416e-06 | 464.8 | 4.8093e-06
620.9087 | 6.6099e-06 | 332.4 | 6.6414e-06 | 465 | 4.8109e-06
621.1826 | 6.6232e-06 | 332.5 | 6.6414e-06 | 465.2 | 4.8125e-06
621.4565 | 6.6366e-06 | 332.6 | 6.6418e-06 | 465.4 | 4.8139e-06
621.7302 | 6.651e-06 | 332.7 | 6.6421e-06 | 465.6 | 4.8153e-06
622.004 | 6.666e-06 | 332.8 | 6.6419e-06 | 465.8 | 4.8166e-06
622.2776 | 6.6812e-06 | 332.9 | 6.6417e-06 | 466 | 4.8179e-06
622.5512 | 6.697e-06 | 333 | 6.6415e-06 | 466.2 | 4.8189e-06
622.8248 | 6.7137e-06 | 333.1 | 6.6414e-06 | 466.4 | 4.8198e-06
623.0982 | 6.7304e-06 | 333.2 | 6.6374e-06 | 466.6 | 4.8205e-06
623.3717 | 6.7487e-06 | 333.3 | 6.636e-06 | 466.8 | 4.8214e-06
623.645 | 6.766e-06 | 333.4 | 6.6346e-06 | 467 | 4.8225e-06
623.9183 | 6.7835e-06 | 333.5 | 6.632e-06 | 467.2 | 4.8237e-06
624.1915 | 6.801e-06 | 333.6 | 6.6296e-06 | 467.4 | 4.8249e-06
624.4647 | 6.819e-06 | 333.7 | 6.6272e-06 | 467.6 | 4.826e-06
624.7378 | 6.8369e-06 | 333.8 | 6.6247e-06 | 467.8 | 4.8271e-06
625.0108 | 6.8549e-06 | 333.9 | 6.6231e-06 | 468 | 4.8284e-06
625.2838 | 6.8732e-06 | 334 | 6.621e-06 | 468.2 | 4.8298e-06
625.5567 | 6.8915e-06 | 334.1 | 6.6189e-06 | 468.4 | 4.8315e-06
625.8295 | 6.9107e-06 | 334.2 | 6.6177e-06 | 468.6 | 4.8332e-06
626.1023 | 6.9304e-06 | 334.3 | 6.6159e-06 | 468.8 | 4.8349e-06
626.375 | 6.9503e-06 | 334.4 | 6.6142e-06 | 469 | 4.8369e-06
626.6477 | 6.9698e-06 | 334.5 | 6.6134e-06 | 469.2 | 4.8389e-06
626.9203 | 6.9886e-06 | 334.6 | 6.6122e-06 | 469.4 | 4.841e-06
627.1928 | 7.0084e-06 | 334.7 | 6.6109e-06 | 469.6 | 4.8432e-06
627.4653 | 7.0274e-06 | 334.8 | 6.6097e-06 | 469.8 | 4.8455e-06
627.7377 | 7.0463e-06 | 334.9 | 6.6065e-06 | 470 | 4.8482e-06
628.01 | 7.0654e-06 | 335 | 6.6047e-06 | 470.2 | 4.8508e-06
628.2823 | 7.0832e-06 | 335.1 | 6.6029e-06 | 470.4 | 4.8534e-06
628.5545 | 7.1009e-06 | 335.2 | 6.6035e-06 | 470.6 | 4.8561e-06
628.8267 | 7.1178e-06 | 335.3 | 6.6023e-06 | 470.8 | 4.8587e-06
629.0988 | 7.1328e-06 | 335.4 | 6.601e-06 | 471 | 4.8613e-06
629.3708 | 7.1474e-06 | 335.5 | 6.5998e-06 | 471.2 | 4.8638e-06
629.6428 | 7.1624e-06 | 335.6 | 6.602e-06 | 471.4 | 4.8662e-06
629.9147 | 7.1773e-06 | 335.7 | 6.6023e-06 | 471.6 | 4.8687e-06
630.1865 | 7.1914e-06 | 335.8 | 6.6026e-06 | 471.8 | 4.871e-06
630.4583 | 7.2048e-06 | 335.9 | 6.5982e-06 | 472 | 4.8735e-06
630.73 | 7.2176e-06 | 336 | 6.5969e-06 | 472.2 | 4.8759e-06
631.0017 | 7.2295e-06 | 336.1 | 6.5956e-06 | 472.4 | 4.8783e-06
631.2733 | 7.2415e-06 | 336.2 | 6.5953e-06 | 472.6 | 4.8808e-06
631.5448 | 7.2518e-06 | 336.3 | 6.5936e-06 | 472.8 | 4.8834e-06
631.8163 | 7.2616e-06 | 336.4 | 6.5918e-06 | 473 | 4.886e-06
632.0877 | 7.2719e-06 | 336.5 | 6.5901e-06 | 473.2 | 4.8884e-06
632.359 | 7.2806e-06 | 336.6 | 6.5914e-06 | 473.4 | 4.8907e-06
632.6303 | 7.288e-06 | 336.7 | 6.5911e-06 | 473.6 | 4.8931e-06
632.9015 | 7.2947e-06 | 336.8 | 6.5909e-06 | 473.8 | 4.8957e-06
633.1727 | 7.3005e-06 | 336.9 | 6.5895e-06 | 474 | 4.8984e-06
633.4438 | 7.304e-06 | 337 | 6.5889e-06 | 474.2 | 4.9011e-06
633.7148 | 7.3075e-06 | 337.1 | 6.5883e-06 | 474.4 | 4.9039e-06
633.9858 | 7.3098e-06 | 337.2 | 6.5877e-06 | 474.6 | 4.9066e-06
634.2567 | 7.3107e-06 | 337.3 | 6.5859e-06 | 474.8 | 4.9093e-06
634.5276 | 7.3124e-06 | 337.4 | 6.5847e-06 | 475 | 4.9119e-06
634.7983 | 7.3141e-06 | 337.5 | 6.5835e-06 | 475.2 | 4.9147e-06
635.0691 | 7.3165e-06 | 337.6 | 6.5829e-06 | 475.4 | 4.9176e-06
635.3397 | 7.3183e-06 | 337.7 | 6.5819e-06 | 475.6 | 4.9206e-06
635.6103 | 7.3186e-06 | 337.8 | 6.5809e-06 | 475.8 | 4.9234e-06
635.8809 | 7.3185e-06 | 337.9 | 6.5795e-06 | 476 | 4.9262e-06
636.1513 | 7.3186e-06 | 338 | 6.5785e-06 | 476.2 | 4.929e-06
636.4217 | 7.3197e-06 | 338.1 | 6.5774e-06 | 476.4 | 4.9321e-06
636.6921 | 7.3194e-06 | 338.2 | 6.5764e-06 | 476.6 | 4.9354e-06
636.9624 | 7.3186e-06 | 338.3 | 6.5749e-06 | 476.8 | 4.9385e-06
637.2326 | 7.3169e-06 | 338.4 | 6.5736e-06 | 477 | 4.9415e-06
637.5028 | 7.314e-06 | 338.5 | 6.5723e-06 | 477.2 | 4.9444e-06
637.7729 | 7.3096e-06 | 338.6 | 6.5732e-06 | 477.4 | 4.9473e-06
638.0429 | 7.3034e-06 | 338.7 | 6.5728e-06 | 477.6 | 4.9502e-06
638.3129 | 7.2972e-06 | 338.8 | 6.5724e-06 | 477.8 | 4.953e-06
638.5828 | 7.2909e-06 | 338.9 | 6.572e-06 | 478 | 4.9561e-06
638.8527 | 7.2838e-06 | 339 | 6.5726e-06 | 478.2 | 4.9591e-06
639.1225 | 7.2754e-06 | 339.1 | 6.5728e-06 | 478.4 | 4.9621e-06
639.3922 | 7.2666e-06 | 339.2 | 6.573e-06 | 478.6 | 4.965e-06
639.6619 | 7.2563e-06 | 339.3 | 6.5703e-06 | 478.8 | 4.9678e-06
639.9315 | 7.2458e-06 | 339.4 | 6.5694e-06 | 479 | 4.9708e-06
640.201 | 7.2338e-06 | 339.5 | 6.5684e-06 | 479.2 | 4.9741e-06
640.4705 | 7.221e-06 | 339.6 | 6.5692e-06 | 479.4 | 4.9775e-06
640.7399 | 7.2073e-06 | 339.7 | 6.5686e-06 | 479.6 | 4.9809e-06
641.0093 | 7.1936e-06 | 339.8 | 6.5679e-06 | 479.8 | 4.9841e-06
641.2786 | 7.1773e-06 | 339.9 | 6.5673e-06 | 480 | 4.9874e-06
641.5478 | 7.1612e-06 | 340 | 6.5694e-06 | 480.2 | 4.9906e-06
641.817 | 7.145e-06 | 340.1 | 6.5702e-06 | 480.4 | 4.9938e-06
642.0861 | 7.1289e-06 | 340.2 | 6.5709e-06 | 480.6 | 4.9968e-06
642.3551 | 7.1134e-06 | 340.3 | 6.5696e-06 | 480.8 | 4.9994e-06
642.6241 | 7.0989e-06 | 340.4 | 6.5697e-06 | 481 | 5.0023e-06
642.8931 | 7.0847e-06 | 340.5 | 6.5698e-06 | 481.2 | 5.0053e-06
643.1619 | 7.0721e-06 | 340.6 | 6.5698e-06 | 481.4 | 5.0085e-06
643.4307 | 7.0603e-06 | 340.7 | 6.5712e-06 | 481.6 | 5.0117e-06
643.6995 | 7.0486e-06 | 340.8 | 6.5714e-06 | 481.8 | 5.0149e-06
643.9682 | 7.0366e-06 | 340.9 | 6.5715e-06 | 482 | 5.0182e-06
644.2368 | 7.025e-06 | 341 | 6.5739e-06 | 482.2 | 5.0213e-06
644.5053 | 7.0131e-06 | 341.1 | 6.5753e-06 | 482.4 | 5.0243e-06
644.7738 | 7.0018e-06 | 341.2 | 6.5766e-06 | 482.6 | 5.0274e-06
645.0423 | 6.9902e-06 | 341.3 | 6.5772e-06 | 482.8 | 5.0306e-06
645.3107 | 6.978e-06 | 341.4 | 6.5783e-06 | 483 | 5.034e-06
645.579 | 6.9673e-06 | 341.5 | 6.5795e-06 | 483.2 | 5.0373e-06
645.8472 | 6.9558e-06 | 341.6 | 6.5806e-06 | 483.4 | 5.0407e-06
646.1154 | 6.9433e-06 | 341.7 | 6.5835e-06 | 483.6 | 5.0442e-06
646.3835 | 6.9311e-06 | 341.8 | 6.585e-06 | 483.8 | 5.0476e-06
646.6516 | 6.9197e-06 | 341.9 | 6.5865e-06 | 484 | 5.0512e-06
646.9196 | 6.9089e-06 | 342 | 6.5896e-06 | 484.2 | 5.0548e-06
647.1875 | 6.8992e-06 | 342.1 | 6.5921e-06 | 484.4 | 5.0583e-06
647.4554 | 6.8905e-06 | 342.2 | 6.5946e-06 | 484.6 | 5.0616e-06
647.7233 | 6.8819e-06 | 342.3 | 6.5948e-06 | 484.8 | 5.0647e-06
647.991 | 6.8734e-06 | 342.4 | 6.5963e-06 | 485 | 5.0677e-06
648.2587 | 6.8651e-06 | 342.5 | 6.5978e-06 | 485.2 | 5.0711e-06
648.5264 | 6.8565e-06 | 342.6 | 6.5994e-06 | 485.4 | 5.0744e-06
648.7939 | 6.8483e-06 | 342.7 | 6.5974e-06 | 485.6 | 5.0779e-06
649.0615 | 6.8404e-06 | 342.8 | 6.5973e-06 | 485.8 | 5.0814e-06
649.3289 | 6.8332e-06 | 342.9 | 6.5971e-06 | 486 | 5.0848e-06
649.5963 | 6.8241e-06 | 343 | 6.5976e-06 | 486.2 | 5.0882e-06
649.8636 | 6.8159e-06 | 343.1 | 6.5974e-06 | 486.4 | 5.0915e-06
650.1309 | 6.808e-06 | 343.2 | 6.5973e-06 | 486.6 | 5.0946e-06
650.3981 | 6.8e-06 | 343.3 | 6.5971e-06 | 486.8 | 5.0973e-06
650.6653 | 6.7917e-06 | 343.4 | 6.599e-06 | 487 | 5.0999e-06
650.9323 | 6.784e-06 | 343.5 | 6.5997e-06 | 487.2 | 5.1026e-06
651.1994 | 6.7757e-06 | 343.6 | 6.6004e-06 | 487.4 | 5.1054e-06
651.4663 | 6.767e-06 | 343.7 | 6.5996e-06 | 487.6 | 5.1082e-06
651.7332 | 6.7587e-06 | 343.8 | 6.5999e-06 | 487.8 | 5.1112e-06
652.0001 | 6.7534e-06 | 343.9 | 6.6002e-06 | 488 | 5.1142e-06
652.2669 | 6.7478e-06 | 344 | 6.6002e-06 | 488.2 | 5.1173e-06
652.5336 | 6.7436e-06 | 344.1 | 6.6001e-06 | 488.4 | 5.1206e-06
652.8003 | 6.7375e-06 | 344.2 | 6.5999e-06 | 488.6 | 5.1237e-06
653.0669 | 6.7317e-06 | 344.3 | 6.5997e-06 | 488.8 | 5.1268e-06
653.3334 | 6.7258e-06 | 344.4 | 6.6007e-06 | 489 | 5.13e-06
653.5999 | 6.7205e-06 | 344.5 | 6.601e-06 | 489.2 | 5.1333e-06
653.8663 | 6.7162e-06 | 344.6 | 6.6013e-06 | 489.4 | 5.1366e-06
654.1326 | 6.7129e-06 | 344.7 | 6.5998e-06 | 489.6 | 5.1399e-06
654.3989 | 6.711e-06 | 344.8 | 6.5995e-06 | 489.8 | 5.1429e-06
654.6652 | 6.7089e-06 | 344.9 | 6.5992e-06 | 490 | 5.1457e-06
654.9313 | 6.7046e-06 | 345 | 6.599e-06 | 490.2 | 5.1486e-06
655.1975 | 6.7007e-06 | 345.1 | 6.6009e-06 | 490.4 | 5.1515e-06
655.4635 | 6.698e-06 | 345.2 | 6.6011e-06 | 490.6 | 5.1544e-06
655.7295 | 6.6968e-06 | 345.3 | 6.6012e-06 | 490.8 | 5.1571e-06
655.9954 | 6.6959e-06 | 345.4 | 6.5996e-06 | 491 | 5.1598e-06
656.2613 | 6.695e-06 | 345.5 | 6.5997e-06 | 491.2 | 5.1624e-06
656.5271 | 6.6954e-06 | 345.6 | 6.5997e-06 | 491.4 | 5.1651e-06
656.7929 | 6.6932e-06 | 345.7 | 6.5962e-06 | 491.6 | 5.1679e-06
657.0585 | 6.6907e-06 | 345.8 | 6.5942e-06 | 491.8 | 5.1707e-06
657.3242 | 6.6869e-06 | 345.9 | 6.5922e-06 | 492 | 5.1736e-06
657.5897 | 6.6835e-06 | 346 | 6.5901e-06 | 492.2 | 5.1765e-06
657.8552 | 6.6826e-06 | 346.1 | 6.5927e-06 | 492.4 | 5.1794e-06
658.1207 | 6.6825e-06 | 346.2 | 6.592e-06 | 492.6 | 5.1821e-06
658.3861 | 6.6821e-06 | 346.3 | 6.5912e-06 | 492.8 | 5.1849e-06
658.6514 | 6.6816e-06 | 346.4 | 6.5937e-06 | 493 | 5.1878e-06
658.9166 | 6.681e-06 | 346.5 | 6.595e-06 | 493.2 | 5.1909e-06
659.1818 | 6.6798e-06 | 346.6 | 6.5963e-06 | 493.4 | 5.1941e-06
659.447 | 6.679e-06 | 346.7 | 6.5961e-06 | 493.6 | 5.1972e-06
659.7121 | 6.6792e-06 | 346.8 | 6.5968e-06 | 493.8 | 5.2003e-06
659.9771 | 6.6809e-06 | 346.9 | 6.5975e-06 | 494 | 5.2033e-06
660.242 | 6.6838e-06 | 347 | 6.5982e-06 | 494.2 | 5.2064e-06
660.5069 | 6.6877e-06 | 347.1 | 6.5979e-06 | 494.4 | 5.2094e-06
660.7718 | 6.6905e-06 | 347.2 | 6.5979e-06 | 494.6 | 5.2123e-06
661.0366 | 6.6913e-06 | 347.3 | 6.598e-06 | 494.8 | 5.2153e-06
661.3013 | 6.6924e-06 | 347.4 | 6.6012e-06 | 495 | 5.2181e-06
661.5659 | 6.6934e-06 | 347.5 | 6.6026e-06 | 495.2 | 5.221e-06
661.8305 | 6.6949e-06 | 347.6 | 6.604e-06 | 495.4 | 5.224e-06
662.0951 | 6.6971e-06 | 347.7 | 6.6054e-06 | 495.6 | 5.227e-06
662.3596 | 6.6991e-06 | 347.8 | 6.6057e-06 | 495.8 | 5.2298e-06
662.624 | 6.7012e-06 | 347.9 | 6.6072e-06 | 496 | 5.2321e-06
662.8883 | 6.7034e-06 | 348 | 6.6086e-06 | 496.2 | 5.2345e-06
663.1526 | 6.7066e-06 | 348.1 | 6.6111e-06 | 496.4 | 5.2369e-06
663.4168 | 6.7099e-06 | 348.2 | 6.6124e-06 | 496.6 | 5.2395e-06
663.681 | 6.7134e-06 | 348.3 | 6.6138e-06 | 496.8 | 5.2423e-06
663.9451 | 6.7183e-06 | 348.4 | 6.6157e-06 | 497 | 5.2451e-06
664.2092 | 6.7228e-06 | 348.5 | 6.6177e-06 | 497.2 | 5.2478e-06
664.4732 | 6.7276e-06 | 348.6 | 6.6198e-06 | 497.4 | 5.2507e-06
664.7371 | 6.7324e-06 | 348.7 | 6.6218e-06 | 497.6 | 5.2536e-06
665.001 | 6.7375e-06 | 348.8 | 6.6201e-06 | 497.8 | 5.2566e-06
665.2648 | 6.7428e-06 | 348.9 | 6.6208e-06 | 498 | 5.2592e-06
665.5286 | 6.7483e-06 | 349 | 6.6215e-06 | 498.2 | 5.2617e-06
665.7923 | 6.7529e-06 | 349.1 | 6.6213e-06 | 498.4 | 5.2641e-06
666.0559 | 6.7581e-06 | 349.2 | 6.6212e-06 | 498.6 | 5.2665e-06
666.3195 | 6.7625e-06 | 349.3 | 6.6212e-06 | 498.8 | 5.2688e-06
666.583 | 6.7676e-06 | 349.4 | 6.6244e-06 | 499 | 5.2711e-06
666.8464 | 6.7734e-06 | 349.5 | 6.6259e-06 | 499.2 | 5.2738e-06
667.1098 | 6.7794e-06 | 349.6 | 6.6274e-06 | 499.4 | 5.2766e-06
667.3732 | 6.7856e-06 | 349.7 | 6.6289e-06 | 499.6 | 5.2792e-06
667.6364 | 6.7926e-06 | 349.8 | 6.6304e-06 | 499.8 | 5.2819e-06
667.8996 | 6.7994e-06 | 349.9 | 6.6324e-06 | 500 | 5.2845e-06
668.1628 | 6.8057e-06 | 350 | 6.6345e-06 | 500.2 | 5.2872e-06
668.4259 | 6.8123e-06 | 350.1 | 6.6376e-06 | 500.4 | 5.2896e-06
668.6889 | 6.8183e-06 | 350.2 | 6.6397e-06 | 500.6 | 5.2917e-06
668.9519 | 6.8239e-06 | 350.3 | 6.6418e-06 | 500.8 | 5.2938e-06
669.2148 | 6.8308e-06 | 350.4 | 6.6439e-06 | 501 | 5.2961e-06
669.4777 | 6.8372e-06 | 350.5 | 6.6488e-06 | 501.2 | 5.2986e-06
669.7405 | 6.8433e-06 | 350.6 | 6.6521e-06 | 501.4 | 5.3014e-06
670.0032 | 6.8503e-06 | 350.7 | 6.6554e-06 | 501.6 | 5.3039e-06
670.2659 | 6.8571e-06 | 350.8 | 6.6576e-06 | 501.8 | 5.3065e-06
670.5285 | 6.8642e-06 | 350.9 | 6.6606e-06 | 502 | 5.309e-06
670.791 | 6.872e-06 | 351 | 6.6636e-06 | 502.2 | 5.3114e-06
671.0535 | 6.8797e-06 | 351.1 | 6.667e-06 | 502.4 | 5.3138e-06
671.316 | 6.8875e-06 | 351.2 | 6.6698e-06 | 502.6 | 5.3159e-06
671.5783 | 6.8973e-06 | 351.3 | 6.6727e-06 | 502.8 | 5.3183e-06
671.8407 | 6.9071e-06 | 351.4 | 6.6756e-06 | 503 | 5.3206e-06
672.1029 | 6.9166e-06 | 351.5 | 6.6816e-06 | 503.2 | 5.3226e-06
672.3651 | 6.9267e-06 | 351.6 | 6.6857e-06 | 503.4 | 5.3247e-06
672.6273 | 6.9363e-06 | 351.7 | 6.6898e-06 | 503.6 | 5.3268e-06
672.8893 | 6.945e-06 | 351.8 | 6.6896e-06 | 503.8 | 5.3296e-06
673.1513 | 6.9547e-06 | 351.9 | 6.6924e-06 | 504 | 5.3321e-06
673.4133 | 6.9625e-06 | 352 | 6.6952e-06 | 504.2 | 5.3347e-06
673.6752 | 6.9695e-06 | 352.1 | 6.6991e-06 | 504.4 | 5.3368e-06
673.937 | 6.9776e-06 | 352.2 | 6.7014e-06 | 504.6 | 5.3388e-06
674.1988 | 6.9867e-06 | 352.3 | 6.7037e-06 | 504.8 | 5.3407e-06
674.4605 | 6.9951e-06 | 352.4 | 6.706e-06 | 505 | 5.3431e-06
674.7222 | 7.0028e-06 | 352.5 | 6.7116e-06 | 505.2 | 5.3454e-06
674.9838 | 7.0104e-06 | 352.6 | 6.7156e-06 | 505.4 | 5.3479e-06
675.2453 | 7.018e-06 | 352.7 | 6.7196e-06 | 505.6 | 5.3502e-06
675.5068 | 7.0266e-06 | 352.8 | 6.7249e-06 | 505.8 | 5.3527e-06
675.7682 | 7.036e-06 | 352.9 | 6.7295e-06 | 506 | 5.3553e-06
676.0296 | 7.0438e-06 | 353 | 6.7341e-06 | 506.2 | 5.3577e-06
676.2909 | 7.0539e-06 | 353.1 | 6.7387e-06 | 506.4 | 5.3603e-06
676.5521 | 7.0644e-06 | 353.2 | 6.7417e-06 | 506.6 | 5.3629e-06
676.8133 | 7.0727e-06 | 353.3 | 6.7459e-06 | 506.8 | 5.3656e-06
677.0744 | 7.0818e-06 | 353.4 | 6.7502e-06 | 507 | 5.3681e-06
677.3355 | 7.0906e-06 | 353.5 | 6.7538e-06 | 507.2 | 5.3703e-06
677.5965 | 7.0992e-06 | 353.6 | 6.7575e-06 | 507.4 | 5.3726e-06
677.8575 | 7.1064e-06 | 353.7 | 6.7612e-06 | 507.6 | 5.3748e-06
678.1184 | 7.1138e-06 | 353.8 | 6.7613e-06 | 507.8 | 5.377e-06
678.3792 | 7.122e-06 | 353.9 | 6.7636e-06 | 508 | 5.3797e-06
678.64 | 7.1295e-06 | 354 | 6.766e-06 | 508.2 | 5.3823e-06
678.9007 | 7.1392e-06 | 354.1 | 6.7683e-06 | 508.4 | 5.3847e-06
679.1613 | 7.1481e-06 | 354.2 | 6.7699e-06 | 508.6 | 5.3868e-06
679.4219 | 7.1586e-06 | 354.3 | 6.7714e-06 | 508.8 | 5.3889e-06
679.6824 | 7.1705e-06 | 354.4 | 6.7729e-06 | 509 | 5.3912e-06
679.9429 | 7.1819e-06 | 354.5 | 6.7763e-06 | 509.2 | 5.3934e-06
680.2033 | 7.1934e-06 | 354.6 | 6.7788e-06 | 509.4 | 5.3955e-06
680.4637 | 7.2034e-06 | 354.7 | 6.7814e-06 | 509.6 | 5.3975e-06
680.724 | 7.2144e-06 | 354.8 | 6.7823e-06 | 509.8 | 5.3998e-06
680.9842 | 7.2255e-06 | 354.9 | 6.7842e-06 | 510 | 5.4021e-06
681.2444 | 7.2355e-06 | 355 | 6.7861e-06 | 510.2 | 5.4041e-06
681.5045 | 7.2471e-06 | 355.1 | 6.7881e-06 | 510.4 | 5.406e-06
681.7646 | 7.2571e-06 | 355.2 | 6.7903e-06 | 510.6 | 5.4079e-06
682.0246 | 7.2682e-06 | 355.3 | 6.7921e-06 | 510.8 | 5.41e-06
682.2845 | 7.2773e-06 | 355.4 | 6.7938e-06 | 511 | 5.4123e-06
682.5444 | 7.2871e-06 | 355.5 | 6.7964e-06 | 511.2 | 5.4145e-06
682.8042 | 7.2962e-06 | 355.6 | 6.7987e-06 | 511.4 | 5.4166e-06
683.064 | 7.3071e-06 | 355.7 | 6.801e-06 | 511.6 | 5.4188e-06
683.3237 | 7.3197e-06 | 355.8 | 6.8046e-06 | 511.8 | 5.4209e-06
683.5833 | 7.3326e-06 | 355.9 | 6.8074e-06 | 512 | 5.423e-06
683.8429 | 7.3453e-06 | 356 | 6.8102e-06 | 512.2 | 5.4251e-06
684.1024 | 7.3567e-06 | 356.1 | 6.813e-06 | 512.4 | 5.4271e-06
684.3619 | 7.3658e-06 | 356.2 | 6.8117e-06 | 512.6 | 5.4292e-06
684.6213 | 7.3757e-06 | 356.3 | 6.8133e-06 | 512.8 | 5.4314e-06
684.8807 | 7.3853e-06 | 356.4 | 6.815e-06 | 513 | 5.4335e-06
685.14 | 7.3948e-06 | 356.5 | 6.8173e-06 | 513.2 | 5.4355e-06
685.3992 | 7.4028e-06 | 356.6 | 6.8184e-06 | 513.4 | 5.4373e-06
685.6584 | 7.4111e-06 | 356.7 | 6.8196e-06 | 513.6 | 5.4395e-06
685.9175 | 7.419e-06 | 356.8 | 6.8192e-06 | 513.8 | 5.4417e-06
686.1766 | 7.4277e-06 | 356.9 | 6.8203e-06 | 514 | 5.4439e-06
686.4356 | 7.4366e-06 | 357 | 6.8214e-06 | 514.2 | 5.4463e-06
686.6945 | 7.4459e-06 | 357.1 | 6.8225e-06 | 514.4 | 5.4484e-06
686.9534 | 7.4571e-06 | 357.2 | 6.82e-06 | 514.6 | 5.4503e-06
687.2122 | 7.4698e-06 | 357.3 | 6.8193e-06 | 514.8 | 5.4522e-06
687.471 | 7.4828e-06 | 357.4 | 6.8187e-06 | 515 | 5.4539e-06
687.7297 | 7.4932e-06 | 357.5 | 6.8236e-06 | 515.2 | 5.456e-06
687.9883 | 7.5037e-06 | 357.6 | 6.8249e-06 | 515.4 | 5.458e-06
688.2469 | 7.5159e-06 | 357.7 | 6.8262e-06 | 515.6 | 5.4602e-06
688.5054 | 7.527e-06 | 357.8 | 6.8275e-06 | 515.8 | 5.4623e-06
688.7639 | 7.5373e-06 | 357.9 | 6.8298e-06 | 516 | 5.4643e-06
689.0223 | 7.5476e-06 | 358 | 6.8323e-06 | 516.2 | 5.4662e-06
689.2807 | 7.5573e-06 | 358.1 | 6.8347e-06 | 516.4 | 5.4683e-06
689.539 | 7.5653e-06 | 358.2 | 6.8368e-06 | 516.6 | 5.4707e-06
689.7972 | 7.5735e-06 | 358.3 | 6.8386e-06 | 516.8 | 5.4731e-06
690.0554 | 7.5818e-06 | 358.4 | 6.8404e-06 | 517 | 5.4752e-06
690.3135 | 7.5883e-06 | 358.5 | 6.8415e-06 | 517.2 | 5.4775e-06
690.5716 | 7.5981e-06 | 358.6 | 6.8433e-06 | 517.4 | 5.4797e-06
690.8296 | 7.6092e-06 | 358.7 | 6.8451e-06 | 517.6 | 5.4817e-06
691.0875 | 7.6186e-06 | 358.8 | 6.847e-06 | 517.8 | 5.4839e-06
691.3454 | 7.6305e-06 | 358.9 | 6.8415e-06 | 518 | 5.486e-06
691.6032 | 7.6434e-06 | 359 | 6.8408e-06 | 518.2 | 5.4883e-06
691.861 | 7.6556e-06 | 359.1 | 6.8401e-06 | 518.4 | 5.4903e-06
692.1187 | 7.6666e-06 | 359.2 | 6.8376e-06 | 518.6 | 5.4922e-06
692.3764 | 7.6792e-06 | 359.3 | 6.8355e-06 | 518.8 | 5.494e-06
692.634 | 7.6909e-06 | 359.4 | 6.8335e-06 | 519 | 5.4959e-06
692.8915 | 7.7018e-06 | 359.5 | 6.8307e-06 | 519.2 | 5.4979e-06
693.149 | 7.7139e-06 | 359.6 | 6.8286e-06 | 519.4 | 5.5001e-06
693.4064 | 7.7252e-06 | 359.7 | 6.8265e-06 | 519.6 | 5.502e-06
693.6638 | 7.7362e-06 | 359.8 | 6.8244e-06 | 519.8 | 5.5038e-06
693.9211 | 7.747e-06 | 359.9 | 6.8247e-06 | 520 | 5.5055e-06
694.1783 | 7.7563e-06 | 360 | 6.8232e-06 | 520.2 | 5.5072e-06
694.4355 | 7.765e-06 | 360.1 | 6.8217e-06 | 520.4 | 5.5089e-06
694.6926 | 7.773e-06 | 360.2 | 6.8211e-06 | 520.6 | 5.5103e-06
694.9497 | 7.7843e-06 | 360.3 | 6.8204e-06 | 520.8 | 5.5121e-06
695.2067 | 7.7955e-06 | 360.4 | 6.8197e-06 | 521 | 5.5139e-06
695.4637 | 7.8073e-06 | 360.5 | 6.8207e-06 | 521.2 | 5.5159e-06
695.7206 | 7.8196e-06 | 360.6 | 6.8206e-06 | 521.4 | 5.518e-06
695.9774 | 7.8323e-06 | 360.7 | 6.8204e-06 | 521.6 | 5.5201e-06
696.2342 | 7.8455e-06 | 360.8 | 6.8203e-06 | 521.8 | 5.5224e-06
696.4909 | 7.8579e-06 | 360.9 | 6.8213e-06 | 522 | 5.5248e-06
696.7476 | 7.8706e-06 | 361 | 6.8219e-06 | 522.2 | 5.5272e-06
697.0042 | 7.8835e-06 | 361.1 | 6.8225e-06 | 522.4 | 5.5298e-06
697.2607 | 7.8967e-06 | 361.2 | 6.8241e-06 | 522.6 | 5.5322e-06
697.5172 | 7.9116e-06 | 361.3 | 6.825e-06 | 522.8 | 5.5345e-06
697.7736 | 7.9247e-06 | 361.4 | 6.8258e-06 | 523 | 5.5369e-06
698.03 | 7.9362e-06 | 361.5 | 6.8257e-06 | 523.2 | 5.5395e-06
698.2863 | 7.9477e-06 | 361.6 | 6.8264e-06 | 523.4 | 5.5421e-06
698.5426 | 7.9596e-06 | 361.7 | 6.8271e-06 | 523.6 | 5.5445e-06
698.7988 | 7.9714e-06 | 361.8 | 6.8278e-06 | 523.8 | 5.5473e-06
699.0549 | 7.9825e-06 | 361.9 | 6.8246e-06 | 524 | 5.5501e-06
699.311 | 7.9943e-06 | 362 | 6.8238e-06 | 524.2 | 5.5529e-06
699.567 | 8.008e-06 | 362.1 | 6.823e-06 | 524.4 | 5.5556e-06
699.823 | 8.0188e-06 | 362.2 | 6.8197e-06 | 524.6 | 5.5582e-06
700.0789 | 8.0296e-06 | 362.3 | 6.8176e-06 | 524.8 | 5.5607e-06
700.3348 | 8.0405e-06 | 362.4 | 6.8155e-06 | 525 | 5.563e-06
700.5906 | 8.0526e-06 | 362.5 | 6.8161e-06 | 525.2 | 5.5653e-06
700.8463 | 8.0657e-06 | 362.6 | 6.8149e-06 | 525.4 | 5.5675e-06
701.102 | 8.0779e-06 | 362.7 | 6.8138e-06 | 525.6 | 5.5697e-06
701.3576 | 8.0885e-06 | 362.8 | 6.8126e-06 | 525.8 | 5.5721e-06
701.6132 | 8.0995e-06 | 362.9 | 6.816e-06 | 526 | 5.5747e-06
701.8687 | 8.1095e-06 | 363 | 6.8168e-06 | 526.2 | 5.5772e-06
702.1241 | 8.119e-06 | 363.1 | 6.8176e-06 | 526.4 | 5.5796e-06
702.3795 | 8.1273e-06 | 363.2 | 6.8156e-06 | 526.6 | 5.582e-06
702.6349 | 8.1383e-06 | 363.3 | 6.8156e-06 | 526.8 | 5.5848e-06
702.8901 | 8.1473e-06 | 363.4 | 6.8156e-06 | 527 | 5.5872e-06
703.1454 | 8.1565e-06 | 363.5 | 6.8161e-06 | 527.2 | 5.5895e-06
703.4005 | 8.165e-06 | 363.6 | 6.8156e-06 | 527.4 | 5.5916e-06
703.6556 | 8.1738e-06 | 363.7 | 6.8152e-06 | 527.6 | 5.594e-06
703.9107 | 8.1849e-06 | 363.8 | 6.8147e-06 | 527.8 | 5.5966e-06
704.1657 | 8.1982e-06 | 363.9 | 6.8151e-06 | 528 | 5.5992e-06
704.4206 | 8.2094e-06 | 364 | 6.8153e-06 | 528.2 | 5.6015e-06
704.6755 | 8.2228e-06 | 364.1 | 6.8155e-06 | 528.4 | 5.6037e-06
704.9303 | 8.236e-06 | 364.2 | 6.8146e-06 | 528.6 | 5.6059e-06
705.185 | 8.2501e-06 | 364.3 | 6.8142e-06 | 528.8 | 5.608e-06
705.4397 | 8.2634e-06 | 364.4 | 6.8139e-06 | 529 | 5.61e-06
705.6944 | 8.2774e-06 | 364.5 | 6.813e-06 | 529.2 | 5.6116e-06
705.949 | 8.2908e-06 | 364.6 | 6.8123e-06 | 529.4 | 5.6135e-06
706.2035 | 8.3042e-06 | 364.7 | 6.8117e-06 | 529.6 | 5.6156e-06
706.458 | 8.3182e-06 | 364.8 | 6.811e-06 | 529.8 | 5.6178e-06
706.7124 | 8.3294e-06 | 364.9 | 6.8078e-06 | 530 | 5.6203e-06
706.9667 | 8.341e-06 | 365 | 6.8063e-06 | 530.2 | 5.6229e-06
707.221 | 8.3517e-06 | 365.1 | 6.8048e-06 | 530.4 | 5.6255e-06
707.4753 | 8.3616e-06 | 365.2 | 6.8035e-06 | 530.6 | 5.6281e-06
707.7295 | 8.3709e-06 | 365.3 | 6.8018e-06 | 530.8 | 5.6308e-06
707.9836 | 8.3811e-06 | 365.4 | 6.8e-06 | 531 | 5.6335e-06
708.2377 | 8.392e-06 | 365.5 | 6.796e-06 | 531.2 | 5.636e-06
708.4917 | 8.4039e-06 | 365.6 | 6.7936e-06 | 531.4 | 5.6386e-06
708.7456 | 8.4152e-06 | 365.7 | 6.7911e-06 | 531.6 | 5.6413e-06
708.9995 | 8.427e-06 | 365.8 | 6.7887e-06 | 531.8 | 5.644e-06
709.2534 | 8.4383e-06 | 365.9 | 6.7856e-06 | 532 | 5.6467e-06
709.5072 | 8.4521e-06 | 366 | 6.7825e-06 | 532.2 | 5.6494e-06
709.7609 | 8.4665e-06 | 366.1 | 6.7794e-06 | 532.4 | 5.6525e-06
710.0145 | 8.4842e-06 | 366.2 | 6.7781e-06 | 532.6 | 5.6554e-06
710.2682 | 8.5004e-06 | 366.3 | 6.7759e-06 | 532.8 | 5.6583e-06
710.5217 | 8.517e-06 | 366.4 | 6.7738e-06 | 533 | 5.6612e-06
710.7752 | 8.5326e-06 | 366.5 | 6.7661e-06 | 533.2 | 5.6639e-06
711.0287 | 8.5498e-06 | 366.6 | 6.7617e-06 | 533.4 | 5.6667e-06
711.2821 | 8.567e-06 | 366.7 | 6.7573e-06 | 533.6 | 5.6695e-06
711.5354 | 8.5814e-06 | 366.8 | 6.7529e-06 | 533.8 | 5.6723e-06
711.7886 | 8.5969e-06 | 366.9 | 6.75e-06 | 534 | 5.6752e-06
712.0419 | 8.6119e-06 | 367 | 6.7452e-06 | 534.2 | 5.6779e-06
712.295 | 8.6258e-06 | 367.1 | 6.7404e-06 | 534.4 | 5.6808e-06
712.5481 | 8.6362e-06 | 367.2 | 6.7395e-06 | 534.6 | 5.6836e-06
712.8012 | 8.6484e-06 | 367.3 | 6.737e-06 | 534.8 | 5.6863e-06
713.0542 | 8.6623e-06 | 367.4 | 6.7345e-06 | 535 | 5.6888e-06
713.3071 | 8.6771e-06 | 367.5 | 6.7304e-06 | 535.2 | 5.6912e-06
713.5599 | 8.6932e-06 | 367.6 | 6.7273e-06 | 535.4 | 5.6938e-06
713.8128 | 8.7072e-06 | 367.7 | 6.7242e-06 | 535.6 | 5.6965e-06
714.0655 | 8.7229e-06 | 367.8 | 6.7211e-06 | 535.8 | 5.6993e-06
714.3182 | 8.7423e-06 | 367.9 | 6.722e-06 | 536 | 5.7021e-06
714.5709 | 8.7609e-06 | 368 | 6.7199e-06 | 536.2 | 5.7046e-06
714.8235 | 8.7799e-06 | 368.1 | 6.7178e-06 | 536.4 | 5.7074e-06
715.076 | 8.7987e-06 | 368.2 | 6.7181e-06 | 536.6 | 5.7103e-06
715.3285 | 8.8208e-06 | 368.3 | 6.7177e-06 | 536.8 | 5.7132e-06
715.5809 | 8.8412e-06 | 368.4 | 6.7173e-06 | 537 | 5.7164e-06
715.8332 | 8.8613e-06 | 368.5 | 6.7159e-06 | 537.2 | 5.7196e-06
716.0856 | 8.8803e-06 | 368.6 | 6.715e-06 | 537.4 | 5.7227e-06
716.3378 | 8.8993e-06 | 368.7 | 6.7141e-06 | 537.6 | 5.7255e-06
716.59 | 8.918e-06 | 368.8 | 6.7131e-06 | 537.8 | 5.7283e-06
716.8421 | 8.9345e-06 | 368.9 | 6.7126e-06 | 538 | 5.7314e-06
717.0942 | 8.9506e-06 | 369 | 6.7116e-06 | 538.2 | 5.7344e-06
717.3462 | 8.9656e-06 | 369.1 | 6.7107e-06 | 538.4 | 5.7376e-06
717.5982 | 8.9815e-06 | 369.2 | 6.7099e-06 | 538.6 | 5.7408e-06
717.8501 | 8.9984e-06 | 369.3 | 6.7091e-06 | 538.8 | 5.744e-06
718.1019 | 9.0147e-06 | 369.4 | 6.7084e-06 | 539 | 5.7473e-06
718.3537 | 9.0328e-06 | 369.5 | 6.7105e-06 | 539.2 | 5.7507e-06
718.6055 | 9.0498e-06 | 369.6 | 6.7108e-06 | 539.4 | 5.7543e-06
718.8572 | 9.0678e-06 | 369.7 | 6.7112e-06 | 539.6 | 5.7576e-06
719.1088 | 9.0865e-06 | 369.8 | 6.7116e-06 | 539.8 | 5.7603e-06
719.3604 | 9.106e-06 | 369.9 | 6.71e-06 | 540 | 5.7633e-06
719.6119 | 9.127e-06 | 370 | 6.7102e-06 | 540.2 | 5.7661e-06
719.8633 | 9.1498e-06 | 370.1 | 6.7104e-06 | 540.4 | 5.769e-06
720.1147 | 9.1744e-06 | 370.2 | 6.7146e-06 | 540.6 | 5.7718e-06
720.3661 | 9.1995e-06 | 370.3 | 6.7157e-06 | 540.8 | 5.7751e-06
720.6174 | 9.2251e-06 | 370.4 | 6.7167e-06 | 541 | 5.7786e-06
720.8686 | 9.2508e-06 | 370.5 | 6.7186e-06 | 541.2 | 5.7821e-06
721.1198 | 9.2749e-06 | 370.6 | 6.7211e-06 | 541.4 | 5.7856e-06
721.3709 | 9.2999e-06 | 370.7 | 6.7236e-06 | 541.6 | 5.789e-06
721.6219 | 9.3226e-06 | 370.8 | 6.7261e-06 | 541.8 | 5.7923e-06
721.8729 | 9.3443e-06 | 370.9 | 6.7225e-06 | 542 | 5.7954e-06
722.1239 | 9.3666e-06 | 371 | 6.7227e-06 | 542.2 | 5.7985e-06
722.3748 | 9.3896e-06 | 371.1 | 6.7228e-06 | 542.4 | 5.8015e-06
722.6256 | 9.412e-06 | 371.2 | 6.7236e-06 | 542.6 | 5.8045e-06
722.8764 | 9.4307e-06 | 371.3 | 6.7234e-06 | 542.8 | 5.8075e-06
723.1271 | 9.4473e-06 | 371.4 | 6.7233e-06 | 543 | 5.8106e-06
723.3778 | 9.4601e-06 | 371.5 | 6.7228e-06 | 543.2 | 5.8138e-06
723.6284 | 9.4725e-06 | 371.6 | 6.723e-06 | 543.4 | 5.8171e-06
723.879 | 9.4856e-06 | 371.7 | 6.7232e-06 | 543.6 | 5.8207e-06
724.1295 | 9.5013e-06 | 371.8 | 6.7233e-06 | 543.8 | 5.824e-06
724.3799 | 9.5182e-06 | 371.9 | 6.7188e-06 | 544 | 5.8271e-06
724.6303 | 9.5345e-06 | 372 | 6.7173e-06 | 544.2 | 5.8296e-06
724.8806 | 9.5492e-06 | 372.1 | 6.7157e-06 | 544.4 | 5.8324e-06
725.1309 | 9.5617e-06 | 372.2 | 6.7095e-06 | 544.6 | 5.8354e-06
725.3811 | 9.5745e-06 | 372.3 | 6.7058e-06 | 544.8 | 5.8385e-06
725.6313 | 9.5898e-06 | 372.4 | 6.7021e-06 | 545 | 5.8417e-06
725.8814 | 9.6054e-06 | 372.5 | 6.6973e-06 | 545.2 | 5.8451e-06
726.1314 | 9.6234e-06 | 372.6 | 6.6926e-06 | 545.4 | 5.8486e-06
726.3814 | 9.6411e-06 | 372.7 | 6.688e-06 | 545.6 | 5.852e-06
726.6314 | 9.6574e-06 | 372.8 | 6.6834e-06 | 545.8 | 5.8553e-06
726.8813 | 9.669e-06 | 372.9 | 6.681e-06 | 546 | 5.8587e-06
727.1311 | 9.6812e-06 | 373 | 6.6771e-06 | 546.2 | 5.8623e-06
727.3809 | 9.6938e-06 | 373.1 | 6.6732e-06 | 546.4 | 5.866e-06
727.6306 | 9.7082e-06 | 373.2 | 6.6596e-06 | 546.6 | 5.8695e-06
727.8802 | 9.7244e-06 | 373.3 | 6.6526e-06 | 546.8 | 5.8725e-06
728.1298 | 9.7404e-06 | 373.4 | 6.6455e-06 | 547 | 5.8758e-06
728.3794 | 9.7544e-06 | 373.5 | 6.6357e-06 | 547.2 | 5.8792e-06
728.6289 | 9.7673e-06 | 373.6 | 6.6259e-06 | 547.4 | 5.8832e-06
728.8783 | 9.7781e-06 | 373.7 | 6.6161e-06 | 547.6 | 5.8867e-06
729.1277 | 9.7882e-06 | 373.8 | 6.6063e-06 | 547.8 | 5.8904e-06
729.377 | 9.7995e-06 | 373.9 | 6.5995e-06 | 548 | 5.8941e-06
729.6263 | 9.8132e-06 | 374 | 6.591e-06 | 548.2 | 5.8979e-06
729.8755 | 9.8253e-06 | 374.1 | 6.5825e-06 | 548.4 | 5.9016e-06
730.1246 | 9.8378e-06 | 374.2 | 6.5762e-06 | 548.6 | 5.905e-06
730.3737 | 9.8481e-06 | 374.3 | 6.5686e-06 | 548.8 | 5.9083e-06
730.6228 | 9.8561e-06 | 374.4 | 6.561e-06 | 549 | 5.9118e-06
730.8718 | 9.8623e-06 | 374.5 | 6.5532e-06 | 549.2 | 5.9152e-06
731.1207 | 9.8697e-06 | 374.6 | 6.5457e-06 | 549.4 | 5.9187e-06
731.3696 | 9.8804e-06 | 374.7 | 6.5383e-06 | 549.6 | 5.922e-06
731.6184 | 9.8909e-06 | 374.8 | 6.5342e-06 | 549.8 | 5.9253e-06
731.8672 | 9.9026e-06 | 374.9 | 6.5281e-06 | 550 | 5.929e-06
732.1159 | 9.9142e-06 | 375 | 6.5219e-06 | 550.2 | 5.9328e-06
732.3646 | 9.9255e-06 | 375.1 | 6.5157e-06 | 550.4 | 5.9368e-06
732.6132 | 9.9379e-06 | 375.2 | 6.5127e-06 | 550.6 | 5.9405e-06
732.8617 | 9.9502e-06 | 375.3 | 6.5082e-06 | 550.8 | 5.9442e-06
733.1102 | 9.9631e-06 | 375.4 | 6.5037e-06 | 551 | 5.9481e-06
733.3586 | 9.9782e-06 | 375.5 | 6.4995e-06 | 551.2 | 5.9519e-06
733.607 | 9.9947e-06 | 375.6 | 6.4951e-06 | 551.4 | 5.9555e-06
733.8553 | 1.0011e-05 | 375.7 | 6.4907e-06 | 551.6 | 5.9591e-06
734.1036 | 1.0025e-05 | 375.8 | 6.4862e-06 | 551.8 | 5.9627e-06
734.3518 | 1.004e-05 | 375.9 | 6.4818e-06 | 552 | 5.9666e-06
734.6 | 1.0054e-05 | 376 | 6.4774e-06 | 552.2 | 5.9706e-06
734.8481 | 1.0069e-05 | 376.1 | 6.4731e-06 | 552.4 | 5.9746e-06
735.0961 | 1.0081e-05 | 376.2 | 6.4687e-06 | 552.6 | 5.9782e-06
735.3441 | 1.0095e-05 | 376.3 | 6.4643e-06 | 552.8 | 5.9818e-06
735.592 | 1.0108e-05 | 376.4 | 6.4599e-06 | 553 | 5.9854e-06
735.8399 | 1.012e-05 | 376.5 | 6.456e-06 | 553.2 | 5.9891e-06
736.0878 | 1.0133e-05 | 376.6 | 6.4518e-06 | 553.4 | 5.9929e-06
736.3355 | 1.0148e-05 | 376.7 | 6.4476e-06 | 553.6 | 5.997e-06
736.5832 | 1.0165e-05 | 376.8 | 6.4434e-06 | 553.8 | 6.0013e-06
736.8309 | 1.0182e-05 | 376.9 | 6.4393e-06 | 554 | 6.0058e-06
737.0785 | 1.0199e-05 | 377 | 6.4351e-06 | 554.2 | 6.0098e-06
737.3261 | 1.0215e-05 | 377.1 | 6.431e-06 | 554.4 | 6.0137e-06
737.5735 | 1.0233e-05 | 377.2 | 6.4271e-06 | 554.6 | 6.0177e-06
737.821 | 1.0253e-05 | 377.3 | 6.4231e-06 | 554.8 | 6.0218e-06
738.0684 | 1.0273e-05 | 377.4 | 6.419e-06 | 555 | 6.0259e-06
738.3157 | 1.029e-05 | 377.5 | 6.4143e-06 | 555.2 | 6.0297e-06
738.563 | 1.0307e-05 | 377.6 | 6.41e-06 | 555.4 | 6.0335e-06
738.8102 | 1.0324e-05 | 377.7 | 6.4057e-06 | 555.6 | 6.0373e-06
739.0574 | 1.0338e-05 | 377.8 | 6.4019e-06 | 555.8 | 6.0414e-06
739.3045 | 1.035e-05 | 377.9 | 6.3976e-06 | 556 | 6.0456e-06
739.5515 | 1.0363e-05 | 378 | 6.3934e-06 | 556.2 | 6.0497e-06
739.7985 | 1.0376e-05 | 378.1 | 6.3891e-06 | 556.4 | 6.0537e-06
740.0454 | 1.0388e-05 | 378.2 | 6.3843e-06 | 556.6 | 6.0575e-06
740.2923 | 1.04e-05 | 378.3 | 6.3799e-06 | 556.8 | 6.0608e-06
740.5392 | 1.0412e-05 | 378.4 | 6.3756e-06 | 557 | 6.0641e-06
740.7859 | 1.0423e-05 | 378.5 | 6.3713e-06 | 557.2 | 6.0676e-06
741.0327 | 1.0435e-05 | 378.6 | 6.3669e-06 | 557.4 | 6.0715e-06
741.2793 | 1.0451e-05 | 378.7 | 6.3624e-06 | 557.6 | 6.0759e-06
741.5259 | 1.0467e-05 | 378.8 | 6.3569e-06 | 557.8 | 6.0805e-06
741.7725 | 1.0485e-05 | 378.9 | 6.3522e-06 | 558 | 6.085e-06
742.019 | 1.0504e-05 | 379 | 6.3475e-06 | 558.2 | 6.0895e-06
742.2655 | 1.0526e-05 | 379.1 | 6.3431e-06 | 558.4 | 6.094e-06
742.5119 | 1.0547e-05 | 379.2 | 6.3382e-06 | 558.6 | 6.098e-06
742.7582 | 1.057e-05 | 379.3 | 6.3334e-06 | 558.8 | 6.1019e-06
743.0045 | 1.0588e-05 | 379.4 | 6.3286e-06 | 559 | 6.1057e-06
743.2507 | 1.0609e-05 | 379.5 | 6.3248e-06 | 559.2 | 6.1095e-06
743.4969 | 1.063e-05 | 379.6 | 6.3205e-06 | 559.4 | 6.1137e-06
743.743 | 1.0652e-05 | 379.7 | 6.3161e-06 | 559.6 | 6.1181e-06
743.9891 | 1.0668e-05 | 379.8 | 6.3112e-06 | 559.8 | 6.1225e-06
744.2351 | 1.0682e-05 | 379.9 | 6.3066e-06 | 560 | 6.1269e-06
744.481 | 1.0696e-05 | 380 | 6.3021e-06 | 560.2 | 6.1312e-06
744.7269 | 1.0711e-05 | 380.1 | 6.2959e-06 | 560.4 | 6.1358e-06
744.9728 | 1.0724e-05 | 380.2 | 6.2905e-06 | 560.6 | 6.14e-06
745.2186 | 1.0736e-05 | 380.3 | 6.2852e-06 | 560.8 | 6.1441e-06
745.4643 | 1.075e-05 | 380.4 | 6.2798e-06 | 561 | 6.1481e-06
745.71 | 1.0768e-05 | 380.5 | 6.2745e-06 | 561.2 | 6.1521e-06
745.9556 | 1.0785e-05 | 380.6 | 6.269e-06 | 561.4 | 6.1561e-06
746.2012 | 1.0802e-05 | 380.7 | 6.2635e-06 | 561.6 | 6.1603e-06
746.4467 | 1.0821e-05 | 380.8 | 6.2582e-06 | 561.8 | 6.1647e-06
746.6922 | 1.0841e-05 | 380.9 | 6.2529e-06 | 562 | 6.1693e-06
746.9376 | 1.0865e-05 | 381 | 6.2476e-06 | 562.2 | 6.174e-06
747.1829 | 1.089e-05 | 381.1 | 6.2429e-06 | 562.4 | 6.1785e-06
747.4282 | 1.0912e-05 | 381.2 | 6.2378e-06 | 562.6 | 6.1826e-06
747.6735 | 1.0936e-05 | 381.3 | 6.2327e-06 | 562.8 | 6.1866e-06
747.9187 | 1.0959e-05 | 381.4 | 6.2276e-06 | 563 | 6.1906e-06
748.1638 | 1.0982e-05 | 381.5 | 6.2222e-06 | 563.2 | 6.1946e-06
748.4089 | 1.1002e-05 | 381.6 | 6.2171e-06 | 563.4 | 6.1988e-06
748.6539 | 1.102e-05 | 381.7 | 6.212e-06 | 563.6 | 6.2028e-06
748.8989 | 1.1039e-05 | 381.8 | 6.2074e-06 | 563.8 | 6.2069e-06
749.1438 | 1.1059e-05 | 381.9 | 6.2024e-06 | 564 | 6.2112e-06
749.3887 | 1.1078e-05 | 382 | 6.1974e-06 | 564.2 | 6.2154e-06
749.6335 | 1.1094e-05 | 382.1 | 6.1927e-06 | 564.4 | 6.2195e-06
749.8782 | 1.1108e-05 | 382.2 | 6.188e-06 | 564.6 | 6.2236e-06
750.1229 | 1.1125e-05 | 382.3 | 6.1833e-06 | 564.8 | 6.2276e-06
750.3676 | 1.1142e-05 | 382.4 | 6.1792e-06 | 565 | 6.2317e-06
750.6122 | 1.1158e-05 | 382.5 | 6.1747e-06 | 565.2 | 6.2357e-06
750.8567 | 1.1175e-05 | 382.6 | 6.1702e-06 | 565.4 | 6.2396e-06
751.1012 | 1.1197e-05 | 382.7 | 6.1657e-06 | 565.6 | 6.2434e-06
751.3456 | 1.1219e-05 | 382.8 | 6.1604e-06 | 565.8 | 6.247e-06
751.59 | 1.1241e-05 | 382.9 | 6.1557e-06 | 566 | 6.2509e-06
751.8343 | 1.1262e-05 | 383 | 6.1511e-06 | 566.2 | 6.2547e-06
752.0786 | 1.1281e-05 | 383.1 | 6.1449e-06 | 566.4 | 6.2588e-06
752.3228 | 1.1302e-05 | 383.2 | 6.1396e-06 | 566.6 | 6.2628e-06
752.567 | 1.1321e-05 | 383.3 | 6.1342e-06 | 566.8 | 6.2668e-06
752.8111 | 1.134e-05 | 383.4 | 6.1286e-06 | 567 | 6.2706e-06
753.0552 | 1.1358e-05 | 383.5 | 6.1231e-06 | 567.2 | 6.2741e-06
753.2991 | 1.1378e-05 | 383.6 | 6.1175e-06 | 567.4 | 6.2778e-06
753.5431 | 1.1396e-05 | 383.7 | 6.1119e-06 | 567.6 | 6.282e-06
753.787 | 1.141e-05 | 383.8 | 6.1087e-06 | 567.8 | 6.2863e-06
754.0308 | 1.1426e-05 | 383.9 | 6.1039e-06 | 568 | 6.2908e-06
754.2746 | 1.1443e-05 | 384 | 6.0991e-06 | 568.2 | 6.2951e-06
754.5183 | 1.1459e-05 | 384.1 | 6.0958e-06 | 568.4 | 6.2993e-06
754.762 | 1.1481e-05 | 384.2 | 6.0918e-06 | 568.6 | 6.3037e-06
755.0056 | 1.1503e-05 | 384.3 | 6.0878e-06 | 568.8 | 6.3081e-06
755.2492 | 1.1529e-05 | 384.4 | 6.0831e-06 | 569 | 6.3128e-06
755.4927 | 1.1555e-05 | 384.5 | 6.0789e-06 | 569.2 | 6.3175e-06
755.7361 | 1.1581e-05 | 384.6 | 6.0747e-06 | 569.4 | 6.3222e-06
755.9796 | 1.1606e-05 | 384.7 | 6.0696e-06 | 569.6 | 6.3271e-06
756.2229 | 1.1629e-05 | 384.8 | 6.0649e-06 | 569.8 | 6.3318e-06
756.4662 | 1.1652e-05 | 384.9 | 6.0601e-06 | 570 | 6.3363e-06
756.7094 | 1.1673e-05 | 385 | 6.0554e-06 | 570.2 | 6.3409e-06
756.9526 | 1.1695e-05 | 385.1 | 6.0503e-06 | 570.4 | 6.3458e-06
757.1958 | 1.1718e-05 | 385.2 | 6.0454e-06 | 570.6 | 6.3505e-06
757.4389 | 1.1739e-05 | 385.3 | 6.0404e-06 | 570.8 | 6.3553e-06
757.6819 | 1.176e-05 | 385.4 | 6.0356e-06 | 571 | 6.3602e-06
757.9248 | 1.178e-05 | 385.5 | 6.0307e-06 | 571.2 | 6.3647e-06
758.1678 | 1.1798e-05 | 385.6 | 6.0259e-06 | 571.4 | 6.3693e-06
758.4106 | 1.1815e-05 | 385.7 | 6.021e-06 | 571.6 | 6.3738e-06
758.6535 | 1.1834e-05 | 385.8 | 6.0161e-06 | 571.8 | 6.3782e-06
758.8962 | 1.1858e-05 | 385.9 | 6.0113e-06 | 572 | 6.3828e-06
759.1389 | 1.1883e-05 | 386 | 6.0064e-06 | 572.2 | 6.3874e-06
759.3816 | 1.191e-05 | 386.1 | 6.0019e-06 | 572.4 | 6.392e-06
759.6242 | 1.1935e-05 | 386.2 | 5.9972e-06 | 572.6 | 6.3964e-06
759.8667 | 1.1959e-05 | 386.3 | 5.9924e-06 | 572.8 | 6.4007e-06
760.1092 | 1.1984e-05 | 386.4 | 5.9886e-06 | 573 | 6.4049e-06
760.3516 | 1.2007e-05 | 386.5 | 5.9843e-06 | 573.2 | 6.4094e-06
760.594 | 1.2031e-05 | 386.6 | 5.9799e-06 | 573.4 | 6.4139e-06
760.8364 | 1.2056e-05 | 386.7 | 5.9761e-06 | 573.6 | 6.4183e-06
761.0786 | 1.2083e-05 | 386.8 | 5.9721e-06 | 573.8 | 6.4226e-06
761.3208 | 1.2108e-05 | 386.9 | 5.9681e-06 | 574 | 6.4269e-06
761.563 | 1.2131e-05 | 387 | 5.9642e-06 | 574.2 | 6.4313e-06
761.8051 | 1.215e-05 | 387.1 | 5.9603e-06 | 574.4 | 6.4359e-06
762.0472 | 1.2169e-05 | 387.2 | 5.9563e-06 | 574.6 | 6.4407e-06
762.2892 | 1.2188e-05 | 387.3 | 5.9524e-06 | 574.8 | 6.4456e-06
762.5312 | 1.2208e-05 | 387.4 | 5.9474e-06 | 575 | 6.4505e-06
762.7731 | 1.2229e-05 | 387.5 | 5.9431e-06 | 575.2 | 6.4554e-06
763.0149 | 1.2249e-05 | 387.6 | 5.9388e-06 | 575.4 | 6.4601e-06
763.2567 | 1.2269e-05 | 387.7 | 5.9339e-06 | 575.6 | 6.4648e-06
763.4985 | 1.2288e-05 | 387.8 | 5.9292e-06 | 575.8 | 6.4692e-06
763.7401 | 1.2308e-05 | 387.9 | 5.9245e-06 | 576 | 6.4738e-06
763.9818 | 1.2332e-05 | 388 | 5.9208e-06 | 576.2 | 6.4784e-06
764.2234 | 1.2355e-05 | 388.1 | 5.9165e-06 | 576.4 | 6.4831e-06
764.4649 | 1.2382e-05 | 388.2 | 5.9122e-06 | 576.6 | 6.4879e-06
764.7064 | 1.2407e-05 | 388.3 | 5.9079e-06 | 576.8 | 6.4928e-06
764.9478 | 1.2434e-05 | 388.4 | 5.9054e-06 | 577 | 6.4979e-06
765.1892 | 1.2459e-05 | 388.5 | 5.9018e-06 | 577.2 | 6.5025e-06
765.4305 | 1.2482e-05 | 388.6 | 5.8983e-06 | 577.4 | 6.5073e-06
765.6717 | 1.2506e-05 | 388.7 | 5.8939e-06 | 577.6 | 6.5121e-06
765.913 | 1.253e-05 | 388.8 | 5.8902e-06 | 577.8 | 6.5171e-06
766.1541 | 1.2555e-05 | 388.9 | 5.8864e-06 | 578 | 6.522e-06
766.3952 | 1.2577e-05 | 389 | 5.8819e-06 | 578.2 | 6.5268e-06
766.6363 | 1.2595e-05 | 389.1 | 5.8776e-06 | 578.4 | 6.5315e-06
766.8773 | 1.2615e-05 | 389.2 | 5.8733e-06 | 578.6 | 6.5365e-06
767.1182 | 1.2634e-05 | 389.3 | 5.8696e-06 | 578.8 | 6.5417e-06
767.3591 | 1.2656e-05 | 389.4 | 5.8656e-06 | 579 | 6.5469e-06
767.5999 | 1.2679e-05 | 389.5 | 5.8616e-06 | 579.2 | 6.5519e-06
767.8407 | 1.2704e-05 | 389.6 | 5.8575e-06 | 579.4 | 6.5567e-06
768.0815 | 1.2732e-05 | 389.7 | 5.8528e-06 | 579.6 | 6.5614e-06
768.3221 | 1.2758e-05 | 389.8 | 5.8487e-06 | 579.8 | 6.5664e-06
768.5628 | 1.2787e-05 | 389.9 | 5.8445e-06 | 580 | 6.5711e-06
768.8033 | 1.2818e-05 | 390 | 5.8398e-06 | 580.2 | 6.576e-06
769.0439 | 1.285e-05 | 390.1 | 5.8352e-06 | 580.4 | 6.581e-06
769.2843 | 1.2882e-05 | 390.2 | 5.8307e-06 | 580.6 | 6.586e-06
769.5247 | 1.291e-05 | 390.3 | 5.8272e-06 | 580.8 | 6.5912e-06
769.7651 | 1.2937e-05 | 390.4 | 5.823e-06 | 581 | 6.5964e-06
770.0054 | 1.2963e-05 | 390.5 | 5.8189e-06 | 581.2 | 6.6014e-06
770.2457 | 1.2985e-05 | 390.6 | 5.8148e-06 | 581.4 | 6.6067e-06
770.4859 | 1.3004e-05 | 390.7 | 5.8123e-06 | 581.6 | 6.6119e-06
770.726 | 1.3025e-05 | 390.8 | 5.8089e-06 | 581.8 | 6.617e-06
770.9661 | 1.3047e-05 | 390.9 | 5.8054e-06 | 582 | 6.6222e-06
771.2062 | 1.3068e-05 | 391 | 5.8017e-06 | 582.2 | 6.6277e-06
771.4462 | 1.3088e-05 | 391.1 | 5.7982e-06 | 582.4 | 6.6339e-06
771.6861 | 1.3106e-05 | 391.2 | 5.7948e-06 | 582.6 | 6.6402e-06
771.926 | 1.3129e-05 | 391.3 | 5.7899e-06 | 582.8 | 6.6462e-06
772.1658 | 1.3155e-05 | 391.4 | 5.7857e-06 | 583 | 6.6519e-06
772.4056 | 1.3183e-05 | 391.5 | 5.7816e-06 | 583.2 | 6.6573e-06
772.6453 | 1.3212e-05 | 391.6 | 5.777e-06 | 583.4 | 6.6629e-06
772.885 | 1.3242e-05 | 391.7 | 5.7725e-06 | 583.6 | 6.6685e-06
773.1246 | 1.3276e-05 | 391.8 | 5.7679e-06 | 583.8 | 6.6745e-06
773.3642 | 1.3304e-05 | 391.9 | 5.7634e-06 | 584 | 6.6802e-06
773.6037 | 1.3332e-05 | 392 | 5.759e-06 | 584.2 | 6.6859e-06
773.8432 | 1.3359e-05 | 392.1 | 5.7546e-06 | 584.4 | 6.6917e-06
774.0826 | 1.3382e-05 | 392.2 | 5.7501e-06 | 584.6 | 6.6975e-06
774.322 | 1.3404e-05 | 392.3 | 5.7466e-06 | 584.8 | 6.7037e-06
774.5613 | 1.3426e-05 | 392.4 | 5.7424e-06 | 585 | 6.71e-06
774.8005 | 1.3447e-05 | 392.5 | 5.7383e-06 | 585.2 | 6.7164e-06
775.0397 | 1.3469e-05 | 392.6 | 5.7349e-06 | 585.4 | 6.7228e-06
775.2789 | 1.3496e-05 | 392.7 | 5.7312e-06 | 585.6 | 6.7291e-06
775.518 | 1.3523e-05 | 392.8 | 5.7275e-06 | 585.8 | 6.7355e-06
775.757 | 1.3548e-05 | 392.9 | 5.7231e-06 | 586 | 6.7423e-06
775.996 | 1.3578e-05 | 393 | 5.7191e-06 | 586.2 | 6.7495e-06
776.235 | 1.3609e-05 | 393.1 | 5.7151e-06 | 586.4 | 6.7571e-06
776.4738 | 1.3641e-05 | 393.2 | 5.7111e-06 | 586.6 | 6.7648e-06
776.7127 | 1.3675e-05 | 393.3 | 5.7078e-06 | 586.8 | 6.7721e-06
776.9515 | 1.3712e-05 | 393.4 | 5.704e-06 | 587 | 6.7794e-06
777.1902 | 1.3747e-05 | 393.5 | 5.7001e-06 | 587.2 | 6.7868e-06
777.4289 | 1.378e-05 | 393.6 | 5.6966e-06 | 587.4 | 6.7944e-06
777.6675 | 1.3806e-05 | 393.7 | 5.693e-06 | 587.6 | 6.8019e-06
777.9061 | 1.3826e-05 | 393.8 | 5.6894e-06 | 587.8 | 6.8096e-06
778.1446 | 1.3845e-05 | 393.9 | 5.6855e-06 | 588 | 6.8176e-06
778.3831 | 1.3865e-05 | 394 | 5.6818e-06 | 588.2 | 6.8256e-06
778.6215 | 1.3888e-05 | 394.1 | 5.678e-06 | 588.4 | 6.8337e-06
778.8599 | 1.3912e-05 | 394.2 | 5.6743e-06 | 588.6 | 6.8415e-06
779.0982 | 1.3942e-05 | 394.3 | 5.6692e-06 | 588.8 | 6.8497e-06
779.3364 | 1.397e-05 | 394.4 | 5.6651e-06 | 589 | 6.8578e-06
779.5746 | 1.3998e-05 | 394.5 | 5.6609e-06 | 589.2 | 6.8659e-06
779.8128 | 1.4024e-05 | 394.6 | 5.6572e-06 | 589.4 | 6.8739e-06
780.0509 | 1.4054e-05 | 394.7 | 5.6531e-06 | 589.6 | 6.8818e-06
780.2889 | 1.409e-05 | 394.8 | 5.649e-06 | 589.8 | 6.8899e-06
780.527 | 1.4122e-05 | 394.9 | 5.6453e-06 | 590 | 6.8982e-06
780.7649 | 1.4154e-05 | 395 | 5.6416e-06 | 590.2 | 6.9062e-06
781.0028 | 1.4182e-05 | 395.1 | 5.6378e-06 | 590.4 | 6.9144e-06
781.2406 | 1.4207e-05 | 395.2 | 5.6343e-06 | 590.6 | 6.9226e-06
781.4784 | 1.4234e-05 | 395.3 | 5.6306e-06 | 590.8 | 6.9307e-06
781.7162 | 1.4261e-05 | 395.4 | 5.627e-06 | 591 | 6.9384e-06
781.9539 | 1.4294e-05 | 395.5 | 5.6233e-06 | 591.2 | 6.9458e-06
782.1915 | 1.4327e-05 | 395.6 | 5.6189e-06 | 591.4 | 6.953e-06
782.4291 | 1.4362e-05 | 395.7 | 5.6151e-06 | 591.6 | 6.9602e-06
782.6666 | 1.4397e-05 | 395.8 | 5.6112e-06 | 591.8 | 6.9669e-06
782.9041 | 1.4425e-05 | 395.9 | 5.6072e-06 | 592 | 6.9734e-06
783.1415 | 1.4454e-05 | 396 | 5.6032e-06 | 592.2 | 6.9798e-06
783.3789 | 1.4485e-05 | 396.1 | 5.5992e-06 | 592.4 | 6.9861e-06
783.6162 | 1.4512e-05 | 396.2 | 5.5957e-06 | 592.6 | 6.9924e-06
783.8535 | 1.4538e-05 | 396.3 | 5.592e-06 | 592.8 | 6.9988e-06
784.0907 | 1.4562e-05 | 396.4 | 5.5882e-06 | 593 | 7.005e-06
784.3279 | 1.4586e-05 | 396.5 | 5.5848e-06 | 593.2 | 7.0111e-06
784.565 | 1.4611e-05 | 396.6 | 5.5813e-06 | 593.4 | 7.0167e-06
784.8021 | 1.4642e-05 | 396.7 | 5.5778e-06 | 593.6 | 7.0217e-06
785.0391 | 1.4672e-05 | 396.8 | 5.5742e-06 | 593.8 | 7.0268e-06
785.2761 | 1.4706e-05 | 396.9 | 5.5703e-06 | 594 | 7.0317e-06
785.513 | 1.474e-05 | 397 | 5.5666e-06 | 594.2 | 7.0365e-06
785.7498 | 1.4782e-05 | 397.1 | 5.563e-06 | 594.4 | 7.0412e-06
785.9866 | 1.4819e-05 | 397.2 | 5.5597e-06 | 594.6 | 7.0456e-06
786.2234 | 1.4857e-05 | 397.3 | 5.5562e-06 | 594.8 | 7.05e-06
786.4601 | 1.4893e-05 | 397.4 | 5.5526e-06 | 595 | 7.0545e-06
786.6967 | 1.4927e-05 | 397.5 | 5.5494e-06 | 595.2 | 7.059e-06
786.9333 | 1.496e-05 | 397.6 | 5.5462e-06 | 595.4 | 7.0634e-06
787.1699 | 1.499e-05 | 397.7 | 5.5429e-06 | 595.6 | 7.0673e-06
787.4064 | 1.502e-05 | 397.8 | 5.5399e-06 | 595.8 | 7.0706e-06
787.6428 | 1.5048e-05 | 397.9 | 5.5367e-06 | 596 | 7.0731e-06
787.8792 | 1.5078e-05 | 398 | 5.5335e-06 | 596.2 | 7.0756e-06
788.1156 | 1.5115e-05 | 398.1 | 5.5304e-06 | 596.4 | 7.0782e-06
788.3519 | 1.5147e-05 | 398.2 | 5.5269e-06 | 596.6 | 7.0804e-06
788.5881 | 1.5183e-05 | 398.3 | 5.5237e-06 | 596.8 | 7.0825e-06
788.8243 | 1.5218e-05 | 398.4 | 5.5205e-06 | 597 | 7.0841e-06
789.0604 | 1.5252e-05 | 398.5 | 5.5178e-06 | 597.2 | 7.0854e-06
789.2965 | 1.5283e-05 | 398.6 | 5.5147e-06 | 597.4 | 7.0864e-06
789.5326 | 1.5312e-05 | 398.7 | 5.5116e-06 | 597.6 | 7.0873e-06
789.7685 | 1.5341e-05 | 398.8 | 5.5083e-06 | 597.8 | 7.0881e-06
790.0045 | 1.5371e-05 | 398.9 | 5.5052e-06 | 598 | 7.0882e-06
790.2403 | 1.5405e-05 | 399 | 5.5021e-06 | 598.2 | 7.0878e-06
790.4762 | 1.5435e-05 | 399.1 | 5.4993e-06 | 598.4 | 7.087e-06
790.712 | 1.5459e-05 | 399.2 | 5.4963e-06 | 598.6 | 7.0859e-06
790.9477 | 1.5488e-05 | 399.3 | 5.4932e-06 | 598.8 | 7.0845e-06
791.1834 | 1.5515e-05 | 399.4 | 5.4901e-06 | 599 | 7.083e-06
791.419 | 1.5546e-05 | 399.5 | 5.4872e-06 | 599.2 | 7.0812e-06
791.6545 | 1.5582e-05 | 399.6 | 5.4843e-06 | 599.4 | 7.0788e-06
791.8901 | 1.5624e-05 | 399.7 | 5.4814e-06 | 599.6 | 7.076e-06
792.1255 | 1.5667e-05 | 399.8 | 5.4775e-06 | 599.8 | 7.0734e-06
792.361 | 1.5705e-05 | 399.9 | 5.4742e-06 | 600 | 7.0704e-06
792.5963 | 1.5738e-05 | 400 | 5.4709e-06 | 600.2 | 7.067e-06
792.8316 | 1.5771e-05 | 400.1 | 5.4674e-06 | 600.4 | 7.063e-06
793.0669 | 1.5804e-05 | 400.2 | 5.4639e-06 | 600.6 | 7.0584e-06
793.3021 | 1.584e-05 | 400.3 | 5.4604e-06 | 600.8 | 7.0535e-06
793.5373 | 1.5871e-05 | 400.4 | 5.4569e-06 | 601 | 7.0482e-06
793.7724 | 1.5907e-05 | 400.5 | 5.4535e-06 | 601.2 | 7.0428e-06
794.0074 | 1.5942e-05 | 400.6 | 5.45e-06 | 601.4 | 7.0369e-06
794.2425 | 1.5982e-05 | 400.7 | 5.4466e-06 | 601.6 | 7.0306e-06
794.4774 | 1.6018e-05 | 400.8 | 5.4432e-06 | 601.8 | 7.0239e-06
794.7123 | 1.6054e-05 | 400.9 | 5.4397e-06 | 602 | 7.0172e-06
794.9472 | 1.6086e-05 | 401 | 5.4363e-06 | 602.2 | 7.0097e-06
795.182 | 1.6115e-05 | 401.1 | 5.4327e-06 | 602.4 | 7.0021e-06
795.4167 | 1.6142e-05 | 401.2 | 5.4292e-06 | 602.6 | 6.9943e-06
795.6514 | 1.6167e-05 | 401.3 | 5.4257e-06 | 602.8 | 6.9861e-06
795.8861 | 1.6192e-05 | 401.4 | 5.4222e-06 | 603 | 6.9772e-06
796.1207 | 1.6227e-05 | 401.5 | 5.4186e-06 | 603.2 | 6.9678e-06
796.3552 | 1.6264e-05 | 401.6 | 5.4151e-06 | 603.4 | 6.958e-06
796.5897 | 1.6305e-05 | 401.7 | 5.4118e-06 | 603.6 | 6.9483e-06
796.8242 | 1.6339e-05 | 401.8 | 5.4084e-06 | 603.8 | 6.9384e-06
797.0586 | 1.637e-05 | 401.9 | 5.405e-06 | 604 | 6.9281e-06
797.2929 | 1.6395e-05 | 402 | 5.4016e-06 | 604.2 | 6.9179e-06
797.5272 | 1.6424e-05 | 402.1 | 5.3982e-06 | 604.4 | 6.9074e-06
797.7615 | 1.6454e-05 | 402.2 | 5.3948e-06 | 604.6 | 6.897e-06
797.9957 | 1.6488e-05 | 402.3 | 5.3915e-06 | 604.8 | 6.8867e-06
798.2298 | 1.653e-05 | 402.4 | 5.3881e-06 | 605 | 6.8763e-06
798.4639 | 1.6569e-05 | 402.5 | 5.3846e-06 | 605.2 | 6.8662e-06
798.698 | 1.6613e-05 | 402.6 | 5.3812e-06 | 605.4 | 6.8561e-06
798.932 | 1.6654e-05 | 402.7 | 5.378e-06 | 605.6 | 6.8458e-06
799.1659 | 1.6687e-05 | 402.8 | 5.3746e-06 | 605.8 | 6.8355e-06
799.3998 | 1.6721e-05 | 402.9 | 5.3713e-06 | 606 | 6.8251e-06
799.6336 | 1.6754e-05 | 403 | 5.3685e-06 | 606.2 | 6.8151e-06
799.8674 | 1.6792e-05 | 403.1 | 5.3654e-06 | 606.4 | 6.8052e-06
800.1011 | 1.6837e-05 | 403.2 | 5.3624e-06 | 606.6 | 6.7949e-06
800.3348 | 1.6882e-05 | 403.3 | 5.3593e-06 | 606.8 | 6.7845e-06
800.5685 | 1.6927e-05 | 403.4 | 5.356e-06 | 607 | 6.7741e-06
800.8021 | 1.6971e-05 | 403.5 | 5.353e-06 | 607.2 | 6.7643e-06
801.0356 | 1.7011e-05 | 403.6 | 5.3499e-06 | 607.4 | 6.7547e-06
801.2691 | 1.7045e-05 | 403.7 | 5.3465e-06 | 607.6 | 6.7452e-06
801.5025 | 1.7076e-05 | 403.8 | 5.3433e-06 | 607.8 | 6.7358e-06
801.7359 | 1.7118e-05 | 403.9 | 5.34e-06 | 608 | 6.7266e-06
801.9692 | 1.7166e-05 | 404 | 5.3365e-06 | 608.2 | 6.7174e-06
802.2025 | 1.7219e-05 | 404.1 | 5.3331e-06 | 608.4 | 6.708e-06
802.4357 | 1.7269e-05 | 404.2 | 5.3297e-06 | 608.6 | 6.6991e-06
802.6689 | 1.7317e-05 | 404.3 | 5.3261e-06 | 608.8 | 6.6907e-06
802.9021 | 1.7365e-05 | 404.4 | 5.3226e-06 | 609 | 6.6826e-06
803.1351 | 1.7413e-05 | 404.5 | 5.319e-06 | 609.2 | 6.6743e-06
803.3682 | 1.746e-05 | 404.6 | 5.3155e-06 | 609.4 | 6.6662e-06
803.6012 | 1.7503e-05 | 404.7 | 5.3119e-06 | 609.6 | 6.6581e-06
803.8341 | 1.7547e-05 | 404.8 | 5.3083e-06 | 609.8 | 6.6504e-06
804.067 | 1.7587e-05 | 404.9 | 5.3047e-06 | 610 | 6.643e-06
804.2998 | 1.7619e-05 | 405 | 5.3015e-06 | 610.2 | 6.6356e-06
804.5326 | 1.7633e-05 | 405.1 | 5.298e-06 | 610.4 | 6.6283e-06
804.7653 | 1.7631e-05 | 405.2 | 5.2946e-06 | 610.6 | 6.6213e-06
804.998 | 1.7624e-05 | 405.3 | 5.2909e-06 | 610.8 | 6.6147e-06
805.2306 | 1.7684e-05 | 405.4 | 5.2874e-06 | 611 | 6.6085e-06
805.4632 | 1.8011e-05 | 405.5 | 5.2839e-06 | 611.2 | 6.6024e-06
805.6957 | 1.8797e-05 | 405.6 | 5.2805e-06 | 611.4 | 6.5964e-06
805.9282 | 2.0808e-05 | 405.7 | 5.277e-06 | 611.6 | 6.5907e-06
806.1606 | 2.3299e-05 | 405.8 | 5.2735e-06 | 611.8 | 6.585e-06
806.393 | 2.6455e-05 | 405.9 | 5.27e-06 | 612 | 6.5793e-06
806.6253 | 3.0614e-05 | 406 | 5.2659e-06 | 612.2 | 6.5741e-06
806.8576 | 3.3022e-05 | 406.1 | 5.2622e-06 | 612.4 | 6.5696e-06
807.0898 | 3.659e-05 | 406.2 | 5.2586e-06 | 612.6 | 6.5655e-06
807.322 | 4.2406e-05 | 406.3 | 5.2546e-06 | 612.8 | 6.5618e-06
807.5541 | 5.3344e-05 | 406.4 | 5.2507e-06 | 613 | 6.5581e-06
807.7862 | 7.8472e-05 | 406.5 | 5.2469e-06 | 613.2 | 6.5545e-06
|  | 406.6 | 5.2427e-06 | 613.4 | 6.5511e-06
|  | 406.7 | 5.2387e-06 | 613.6 | 6.5477e-06
|  | 406.8 | 5.2348e-06 | 613.8 | 6.5442e-06
|  | 406.9 | 5.2312e-06 | 614 | 6.5407e-06
|  | 407 | 5.2274e-06 | 614.2 | 6.5378e-06
|  | 407.1 | 5.2235e-06 | 614.4 | 6.5353e-06
|  | 407.2 | 5.2196e-06 | 614.6 | 6.5331e-06
|  | 407.3 | 5.2167e-06 | 614.8 | 6.5311e-06
|  | 407.4 | 5.2132e-06 | 615 | 6.5291e-06
|  | 407.5 | 5.2098e-06 | 615.2 | 6.5277e-06
|  | 407.6 | 5.2067e-06 | 615.4 | 6.5263e-06
|  | 407.7 | 5.2034e-06 | 615.6 | 6.5247e-06
|  | 407.8 | 5.2001e-06 | 615.8 | 6.5233e-06
|  | 407.9 | 5.1969e-06 | 616 | 6.5221e-06
|  | 408 | 5.1937e-06 | 616.2 | 6.5213e-06
|  | 408.1 | 5.1905e-06 | 616.4 | 6.5207e-06
|  | 408.2 | 5.187e-06 | 616.6 | 6.5206e-06
|  | 408.3 | 5.1836e-06 | 616.8 | 6.5212e-06
|  | 408.4 | 5.1803e-06 | 617 | 6.522e-06
|  | 408.5 | 5.177e-06 | 617.2 | 6.5229e-06
|  | 408.6 | 5.1731e-06 | 617.4 | 6.5239e-06
|  | 408.7 | 5.1696e-06 | 617.6 | 6.525e-06
|  | 408.8 | 5.1661e-06 | 617.8 | 6.5264e-06
|  | 408.9 | 5.1623e-06 | 618 | 6.5281e-06
|  | 409 | 5.1586e-06 | 618.2 | 6.5302e-06
|  | 409.1 | 5.155e-06 | 618.4 | 6.5329e-06
|  | 409.2 | 5.151e-06 | 618.6 | 6.5362e-06
|  | 409.3 | 5.1473e-06 | 618.8 | 6.5402e-06
|  | 409.4 | 5.1435e-06 | 619 | 6.545e-06
|  | 409.5 | 5.1401e-06 | 619.2 | 6.5501e-06
|  | 409.6 | 5.1365e-06 | 619.4 | 6.5558e-06
|  | 409.7 | 5.1328e-06 | 619.6 | 6.5619e-06
|  | 409.8 | 5.1296e-06 | 619.8 | 6.5683e-06
|  | 409.9 | 5.1262e-06 | 620 | 6.5747e-06
|  | 410 | 5.1228e-06 | 620.2 | 6.5812e-06
|  | 410.1 | 5.1194e-06 | 620.4 | 6.5885e-06
|  | 410.2 | 5.1162e-06 | 620.6 | 6.5963e-06
|  | 410.3 | 5.1129e-06 | 620.8 | 6.605e-06
|  | 410.4 | 5.1096e-06 | 621 | 6.6143e-06
|  | 410.5 | 5.1063e-06 | 621.2 | 6.6241e-06
|  | 410.6 | 5.1029e-06 | 621.4 | 6.634e-06
|  | 410.7 | 5.0996e-06 | 621.6 | 6.6442e-06
|  | 410.8 | 5.0965e-06 | 621.8 | 6.6548e-06
|  | 410.9 | 5.0932e-06 | 622 | 6.6658e-06
|  | 411 | 5.0899e-06 | 622.2 | 6.677e-06
|  | 411.1 | 5.0869e-06 | 622.4 | 6.6884e-06
|  | 411.2 | 5.0838e-06 | 622.6 | 6.7001e-06
|  | 411.3 | 5.0806e-06 | 622.8 | 6.7121e-06
|  | 411.4 | 5.0775e-06 | 623 | 6.7245e-06
|  | 411.5 | 5.075e-06 | 623.2 | 6.7374e-06
|  | 411.6 | 5.072e-06 | 623.4 | 6.7502e-06
|  | 411.7 | 5.0691e-06 | 623.6 | 6.7631e-06
|  | 411.8 | 5.0662e-06 | 623.8 | 6.7759e-06
|  | 411.9 | 5.0633e-06 | 624 | 6.7887e-06
|  | 412 | 5.0604e-06 | 624.2 | 6.8017e-06
|  | 412.1 | 5.0572e-06 | 624.4 | 6.8147e-06
|  | 412.2 | 5.0541e-06 | 624.6 | 6.8279e-06
|  | 412.3 | 5.0511e-06 | 624.8 | 6.841e-06
|  | 412.4 | 5.0485e-06 | 625 | 6.8543e-06
|  | 412.5 | 5.0456e-06 | 625.2 | 6.8676e-06
|  | 412.6 | 5.0427e-06 | 625.4 | 6.881e-06
|  | 412.7 | 5.0398e-06 | 625.6 | 6.8947e-06
|  | 412.8 | 5.0366e-06 | 625.8 | 6.9087e-06
|  | 412.9 | 5.0336e-06 | 626 | 6.923e-06
|  | 413 | 5.0307e-06 | 626.2 | 6.9375e-06
|  | 413.1 | 5.0281e-06 | 626.4 | 6.952e-06
|  | 413.2 | 5.0252e-06 | 626.6 | 6.9663e-06
|  | 413.3 | 5.0223e-06 | 626.8 | 6.9803e-06
|  | 413.4 | 5.0193e-06 | 627 | 6.9945e-06
|  | 413.5 | 5.0165e-06 | 627.2 | 7.0087e-06
|  | 413.6 | 5.0137e-06 | 627.4 | 7.0228e-06
|  | 413.7 | 5.0104e-06 | 627.6 | 7.0367e-06
|  | 413.8 | 5.0073e-06 | 627.8 | 7.0507e-06
|  | 413.9 | 5.0042e-06 | 628 | 7.0644e-06
|  | 414 | 5.0016e-06 | 628.2 | 7.0777e-06
|  | 414.1 | 4.9987e-06 | 628.4 | 7.0908e-06
|  | 414.2 | 4.9959e-06 | 628.6 | 7.1036e-06
|  | 414.3 | 4.993e-06 | 628.8 | 7.1158e-06
|  | 414.4 | 4.9898e-06 | 629 | 7.1272e-06
|  | 414.5 | 4.987e-06 | 629.2 | 7.1382e-06
|  | 414.6 | 4.9841e-06 | 629.4 | 7.1491e-06
|  | 414.7 | 4.9807e-06 | 629.6 | 7.16e-06
|  | 414.8 | 4.9775e-06 | 629.8 | 7.171e-06
|  | 414.9 | 4.9743e-06 | 630 | 7.1817e-06
|  | 415 | 4.9708e-06 | 630.2 | 7.1919e-06
|  | 415.1 | 4.9674e-06 | 630.4 | 7.2018e-06
|  | 415.2 | 4.9641e-06 | 630.6 | 7.2114e-06
|  | 415.3 | 4.9614e-06 | 630.8 | 7.2206e-06
|  | 415.4 | 4.9582e-06 | 631 | 7.2295e-06
|  | 415.5 | 4.9551e-06 | 631.2 | 7.2381e-06
|  | 415.6 | 4.9523e-06 | 631.4 | 7.2462e-06
|  | 415.7 | 4.9495e-06 | 631.6 | 7.2538e-06
|  | 415.8 | 4.9467e-06 | 631.8 | 7.2611e-06
|  | 415.9 | 4.9439e-06 | 632 | 7.2684e-06
|  | 416 | 4.9407e-06 | 632.2 | 7.2753e-06
|  | 416.1 | 4.9377e-06 | 632.4 | 7.2815e-06
|  | 416.2 | 4.9348e-06 | 632.6 | 7.2871e-06
|  | 416.3 | 4.9319e-06 | 632.8 | 7.2922e-06
|  | 416.4 | 4.9289e-06 | 633 | 7.2968e-06
|  | 416.5 | 4.9259e-06 | 633.2 | 7.3004e-06
|  | 416.6 | 4.9227e-06 | 633.4 | 7.3035e-06
|  | 416.7 | 4.9196e-06 | 633.6 | 7.3061e-06
|  | 416.8 | 4.9166e-06 | 633.8 | 7.3082e-06
|  | 416.9 | 4.9136e-06 | 634 | 7.3095e-06
|  | 417 | 4.9105e-06 | 634.2 | 7.3106e-06
|  | 417.1 | 4.9074e-06 | 634.4 | 7.3117e-06
|  | 417.2 | 4.9044e-06 | 634.6 | 7.3128e-06
|  | 417.3 | 4.9022e-06 | 634.8 | 7.3143e-06
|  | 417.4 | 4.8995e-06 | 635 | 7.3158e-06
|  | 417.5 | 4.8968e-06 | 635.2 | 7.3173e-06
|  | 417.6 | 4.8942e-06 | 635.4 | 7.3182e-06
|  | 417.7 | 4.8916e-06 | 635.6 | 7.3186e-06
|  | 417.8 | 4.889e-06 | 635.8 | 7.3185e-06
|  | 417.9 | 4.8862e-06 | 636 | 7.3185e-06
|  | 418 | 4.8835e-06 | 636.2 | 7.3189e-06
|  | 418.1 | 4.8808e-06 | 636.4 | 7.3194e-06
|  | 418.2 | 4.8785e-06 | 636.6 | 7.3193e-06
|  | 418.3 | 4.876e-06 | 636.8 | 7.319e-06
|  | 418.4 | 4.8734e-06 | 637 | 7.3182e-06
|  | 418.5 | 4.8712e-06 | 637.2 | 7.3169e-06
|  | 418.6 | 4.8688e-06 | 637.4 | 7.3151e-06
|  | 418.7 | 4.8665e-06 | 637.6 | 7.3124e-06
|  | 418.8 | 4.8641e-06 | 637.8 | 7.3086e-06
|  | 418.9 | 4.8612e-06 | 638 | 7.3044e-06
|  | 419 | 4.8587e-06 | 638.2 | 7.2998e-06
|  | 419.1 | 4.8562e-06 | 638.4 | 7.2952e-06
|  | 419.2 | 4.8539e-06 | 638.6 | 7.2902e-06
|  | 419.3 | 4.8514e-06 | 638.8 | 7.285e-06
|  | 419.4 | 4.8489e-06 | 639 | 7.2791e-06
|  | 419.5 | 4.8466e-06 | 639.2 | 7.2729e-06
|  | 419.6 | 4.8443e-06 | 639.4 | 7.2659e-06
|  | 419.7 | 4.842e-06 | 639.6 | 7.2586e-06
|  | 419.8 | 4.8401e-06 | 639.8 | 7.2509e-06
|  | 419.9 | 4.8379e-06 | 640 | 7.2425e-06
|  | 420 | 4.8357e-06 | 640.2 | 7.2337e-06
|  | 420.1 | 4.834e-06 | 640.4 | 7.2242e-06
|  | 420.2 | 4.8322e-06 | 640.6 | 7.2143e-06
|  | 420.3 | 4.8303e-06 | 640.8 | 7.2043e-06
|  | 420.4 | 4.8285e-06 | 641 | 7.1936e-06
|  | 420.5 | 4.8268e-06 | 641.2 | 7.182e-06
|  | 420.6 | 4.825e-06 | 641.4 | 7.1701e-06
|  | 420.7 | 4.8232e-06 | 641.6 | 7.158e-06
|  | 420.8 | 4.8209e-06 | 641.8 | 7.146e-06
|  | 420.9 | 4.8189e-06 | 642 | 7.1341e-06
|  | 421 | 4.8169e-06 | 642.2 | 7.1224e-06
|  | 421.1 | 4.8146e-06 | 642.4 | 7.1112e-06
|  | 421.2 | 4.8124e-06 | 642.6 | 7.1002e-06
|  | 421.3 | 4.8102e-06 | 642.8 | 7.0898e-06
|  | 421.4 | 4.8081e-06 | 643 | 7.0799e-06
|  | 421.5 | 4.8059e-06 | 643.2 | 7.0706e-06
|  | 421.6 | 4.8037e-06 | 643.4 | 7.0617e-06
|  | 421.7 | 4.8015e-06 | 643.6 | 7.0529e-06
|  | 421.8 | 4.7993e-06 | 643.8 | 7.0441e-06
|  | 421.9 | 4.7972e-06 | 644 | 7.0353e-06
|  | 422 | 4.795e-06 | 644.2 | 7.0265e-06
|  | 422.1 | 4.7933e-06 | 644.4 | 7.0178e-06
|  | 422.2 | 4.7913e-06 | 644.6 | 7.0092e-06
|  | 422.3 | 4.7893e-06 | 644.8 | 7.0006e-06
|  | 422.4 | 4.788e-06 | 645 | 6.9919e-06
|  | 422.5 | 4.7863e-06 | 645.2 | 6.983e-06
|  | 422.6 | 4.7846e-06 | 645.4 | 6.9746e-06
|  | 422.7 | 4.783e-06 | 645.6 | 6.9662e-06
|  | 422.8 | 4.7815e-06 | 645.8 | 6.9577e-06
|  | 422.9 | 4.7799e-06 | 646 | 6.9486e-06
|  | 423 | 4.7808e-06 | 646.2 | 6.9395e-06
|  | 423.1 | 4.7802e-06 | 646.4 | 6.9306e-06
|  | 423.2 | 4.7797e-06 | 646.6 | 6.922e-06
|  | 423.3 | 4.7825e-06 | 646.8 | 6.9138e-06
|  | 423.4 | 4.784e-06 | 647 | 6.9061e-06
|  | 423.5 | 4.7855e-06 | 647.2 | 6.8991e-06
|  | 423.6 | 4.787e-06 | 647.4 | 6.8923e-06
|  | 423.7 | 4.7922e-06 | 647.6 | 6.8859e-06
|  | 423.8 | 4.7953e-06 | 647.8 | 6.8795e-06
|  | 423.9 | 4.7984e-06 | 648 | 6.8732e-06
|  | 424 | 4.8048e-06 | 648.2 | 6.8669e-06
|  | 424.1 | 4.8094e-06 | 648.4 | 6.8605e-06
|  | 424.2 | 4.8139e-06 | 648.6 | 6.8543e-06
|  | 424.3 | 4.8225e-06 | 648.8 | 6.8482e-06
|  | 424.4 | 4.829e-06 | 649 | 6.8423e-06
|  | 424.5 | 4.8354e-06 | 649.2 | 6.8367e-06
|  | 424.6 | 4.8463e-06 | 649.4 | 6.8306e-06
|  | 424.7 | 4.8551e-06 | 649.6 | 6.8242e-06
|  | 424.8 | 4.8639e-06 | 649.8 | 6.8179e-06
|  | 424.9 | 4.878e-06 | 650 | 6.8119e-06
|  | 425 | 4.8899e-06 | 650.2 | 6.8059e-06
|  | 425.1 | 4.9018e-06 | 650.4 | 6.7998e-06
|  | 425.2 | 4.9136e-06 | 650.6 | 6.7938e-06
|  | 425.3 | 4.9303e-06 | 650.8 | 6.7879e-06
|  | 425.4 | 4.9444e-06 | 651 | 6.7818e-06
|  | 425.5 | 4.9585e-06 | 651.2 | 6.7756e-06
|  | 425.6 | 4.9708e-06 | 651.4 | 6.7692e-06
|  | 425.7 | 4.9844e-06 | 651.6 | 6.7629e-06
|  | 425.8 | 4.9981e-06 | 651.8 | 6.7577e-06
|  | 425.9 | 5.0083e-06 | 652 | 6.7532e-06
|  | 426 | 5.0202e-06 | 652.2 | 6.7494e-06
|  | 426.1 | 5.0322e-06 | 652.4 | 6.7458e-06
|  | 426.2 | 5.0424e-06 | 652.6 | 6.7418e-06
|  | 426.3 | 5.0533e-06 | 652.8 | 6.7376e-06
|  | 426.4 | 5.0642e-06 | 653 | 6.7331e-06
|  | 426.5 | 5.0721e-06 | 653.2 | 6.7287e-06
|  | 426.6 | 5.0816e-06 | 653.4 | 6.7245e-06
|  | 426.7 | 5.091e-06 | 653.6 | 6.7208e-06
|  | 426.8 | 5.1004e-06 | 653.8 | 6.7174e-06
|  | 426.9 | 5.103e-06 | 654 | 6.7146e-06
|  | 427 | 5.1097e-06 | 654.2 | 6.7126e-06
|  | 427.1 | 5.1164e-06 | 654.4 | 6.711e-06
|  | 427.2 | 5.1165e-06 | 654.6 | 6.7091e-06
|  | 427.3 | 5.1203e-06 | 654.8 | 6.7065e-06
|  | 427.4 | 5.1241e-06 | 655 | 6.7037e-06
|  | 427.5 | 5.12e-06 | 655.2 | 6.701e-06
|  | 427.6 | 5.1203e-06 | 655.4 | 6.6989e-06
|  | 427.7 | 5.1205e-06 | 655.6 | 6.6975e-06
|  | 427.8 | 5.1126e-06 | 655.8 | 6.6965e-06
|  | 427.9 | 5.1086e-06 | 656 | 6.6959e-06
|  | 428 | 5.1046e-06 | 656.2 | 6.6954e-06
|  | 428.1 | 5.0934e-06 | 656.4 | 6.6953e-06
|  | 428.2 | 5.0852e-06 | 656.6 | 6.6945e-06
|  | 428.3 | 5.077e-06 | 656.8 | 6.6931e-06
|  | 428.4 | 5.0688e-06 | 657 | 6.6911e-06
|  | 428.5 | 5.0505e-06 | 657.2 | 6.6886e-06
|  | 428.6 | 5.0383e-06 | 657.4 | 6.686e-06
|  | 428.7 | 5.026e-06 | 657.6 | 6.6841e-06
|  | 428.8 | 5.0072e-06 | 657.8 | 6.683e-06
|  | 428.9 | 4.9919e-06 | 658 | 6.6826e-06
|  | 429 | 4.9766e-06 | 658.2 | 6.6823e-06
|  | 429.1 | 4.9607e-06 | 658.4 | 6.682e-06
|  | 429.2 | 4.9448e-06 | 658.6 | 6.6817e-06
|  | 429.3 | 4.9288e-06 | 658.8 | 6.6812e-06
|  | 429.4 | 4.9152e-06 | 659 | 6.6806e-06
|  | 429.5 | 4.9004e-06 | 659.2 | 6.6798e-06
|  | 429.6 | 4.8855e-06 | 659.4 | 6.6793e-06
|  | 429.7 | 4.8743e-06 | 659.6 | 6.6791e-06
|  | 429.8 | 4.8613e-06 | 659.8 | 6.6798e-06
|  | 429.9 | 4.8483e-06 | 660 | 6.6814e-06
|  | 430 | 4.841e-06 | 660.2 | 6.6835e-06
|  | 430.1 | 4.8312e-06 | 660.4 | 6.6862e-06
|  | 430.2 | 4.8215e-06 | 660.6 | 6.6886e-06
|  | 430.3 | 4.8117e-06 | 660.8 | 6.6901e-06
|  | 430.4 | 4.8094e-06 | 661 | 6.6912e-06
|  | 430.5 | 4.8028e-06 | 661.2 | 6.692e-06
|  | 430.6 | 4.7963e-06 | 661.4 | 6.6928e-06
|  | 430.7 | 4.7943e-06 | 661.6 | 6.6937e-06
|  | 430.8 | 4.7898e-06 | 661.8 | 6.6949e-06
|  | 430.9 | 4.7853e-06 | 662 | 6.6964e-06
|  | 431 | 4.784e-06 | 662.2 | 6.6979e-06
|  | 431.1 | 4.7811e-06 | 662.4 | 6.6995e-06
|  | 431.2 | 4.7782e-06 | 662.6 | 6.701e-06
|  | 431.3 | 4.779e-06 | 662.8 | 6.7028e-06
|  | 431.4 | 4.7779e-06 | 663 | 6.7049e-06
|  | 431.5 | 4.7769e-06 | 663.2 | 6.7072e-06
|  | 431.6 | 4.7788e-06 | 663.4 | 6.7097e-06
|  | 431.7 | 4.7796e-06 | 663.6 | 6.7125e-06
|  | 431.8 | 4.7804e-06 | 663.8 | 6.7158e-06
|  | 431.9 | 4.7812e-06 | 664 | 6.7191e-06
|  | 432 | 4.7862e-06 | 664.2 | 6.7227e-06
|  | 432.1 | 4.7887e-06 | 664.4 | 6.7262e-06
|  | 432.2 | 4.7911e-06 | 664.6 | 6.7299e-06
|  | 432.3 | 4.7983e-06 | 664.8 | 6.7336e-06
|  | 432.4 | 4.8028e-06 | 665 | 6.7375e-06
|  | 432.5 | 4.8073e-06 | 665.2 | 6.7415e-06
|  | 432.6 | 4.8162e-06 | 665.4 | 6.7456e-06
|  | 432.7 | 4.8228e-06 | 665.6 | 6.7494e-06
|  | 432.8 | 4.8294e-06 | 665.8 | 6.7532e-06
|  | 432.9 | 4.8385e-06 | 666 | 6.7569e-06
|  | 433 | 4.8464e-06 | 666.2 | 6.7605e-06
|  | 433.1 | 4.8544e-06 | 666.4 | 6.7642e-06
|  | 433.2 | 4.8619e-06 | 666.6 | 6.7681e-06
|  | 433.3 | 4.8699e-06 | 666.8 | 6.7724e-06
|  | 433.4 | 4.8779e-06 | 667 | 6.7769e-06
|  | 433.5 | 4.8824e-06 | 667.2 | 6.7816e-06
|  | 433.6 | 4.8886e-06 | 667.4 | 6.7865e-06
|  | 433.7 | 4.8947e-06 | 667.6 | 6.7916e-06
|  | 433.8 | 4.9009e-06 | 667.8 | 6.7967e-06
|  | 433.9 | 4.9026e-06 | 668 | 6.8017e-06
|  | 434 | 4.9067e-06 | 668.2 | 6.8067e-06
|  | 434.1 | 4.9109e-06 | 668.4 | 6.8116e-06
|  | 434.2 | 4.9134e-06 | 668.6 | 6.8162e-06
|  | 434.3 | 4.9167e-06 | 668.8 | 6.8206e-06
|  | 434.4 | 4.92e-06 | 669 | 6.8254e-06
|  | 434.5 | 4.924e-06 | 669.2 | 6.8303e-06
|  | 434.6 | 4.9275e-06 | 669.4 | 6.8352e-06
|  | 434.7 | 4.9309e-06 | 669.6 | 6.84e-06
|  | 434.8 | 4.9351e-06 | 669.8 | 6.845e-06
|  | 434.9 | 4.939e-06 | 670 | 6.8502e-06
|  | 435 | 4.9429e-06 | 670.2 | 6.8555e-06
|  | 435.1 | 4.9466e-06 | 670.4 | 6.8608e-06
|  | 435.2 | 4.9505e-06 | 670.6 | 6.8664e-06
|  | 435.3 | 4.9544e-06 | 670.8 | 6.8723e-06
|  | 435.4 | 4.957e-06 | 671 | 6.8782e-06
|  | 435.5 | 4.9602e-06 | 671.2 | 6.8841e-06
|  | 435.6 | 4.9634e-06 | 671.4 | 6.8908e-06
|  | 435.7 | 4.9666e-06 | 671.6 | 6.8981e-06
|  | 435.8 | 4.9678e-06 | 671.8 | 6.9055e-06
|  | 435.9 | 4.9701e-06 | 672 | 6.9128e-06
|  | 436 | 4.9724e-06 | 672.2 | 6.9204e-06
|  | 436.1 | 4.9733e-06 | 672.4 | 6.9279e-06
|  | 436.2 | 4.9749e-06 | 672.6 | 6.9352e-06
|  | 436.3 | 4.9766e-06 | 672.8 | 6.9421e-06
|  | 436.4 | 4.9766e-06 | 673 | 6.9492e-06
|  | 436.5 | 4.9776e-06 | 673.2 | 6.9558e-06
|  | 436.6 | 4.9785e-06 | 673.4 | 6.962e-06
|  | 436.7 | 4.9783e-06 | 673.6 | 6.9676e-06
|  | 436.8 | 4.9787e-06 | 673.8 | 6.9735e-06
|  | 436.9 | 4.979e-06 | 674 | 6.9799e-06
|  | 437 | 4.9778e-06 | 674.2 | 6.9866e-06
|  | 437.1 | 4.9774e-06 | 674.4 | 6.9931e-06
|  | 437.2 | 4.9769e-06 | 674.6 | 6.9991e-06
|  | 437.3 | 4.9765e-06 | 674.8 | 7.0051e-06
|  | 437.4 | 4.9743e-06 | 675 | 7.0109e-06
|  | 437.5 | 4.9731e-06 | 675.2 | 7.0169e-06
|  | 437.6 | 4.9719e-06 | 675.4 | 7.0231e-06
|  | 437.7 | 4.9698e-06 | 675.6 | 7.0299e-06
|  | 437.8 | 4.9682e-06 | 675.8 | 7.0366e-06
|  | 437.9 | 4.9666e-06 | 676 | 7.0433e-06
|  | 438 | 4.9645e-06 | 676.2 | 7.0506e-06
|  | 438.1 | 4.9626e-06 | 676.4 | 7.0583e-06
|  | 438.2 | 4.9608e-06 | 676.6 | 7.0656e-06
|  | 438.3 | 4.9587e-06 | 676.8 | 7.0725e-06
|  | 438.4 | 4.9567e-06 | 677 | 7.0792e-06
|  | 438.5 | 4.9547e-06 | 677.2 | 7.086e-06
|  | 438.6 | 4.9523e-06 | 677.4 | 7.0927e-06
|  | 438.7 | 4.9502e-06 | 677.6 | 7.0989e-06
|  | 438.8 | 4.948e-06 | 677.8 | 7.1048e-06
|  | 438.9 | 4.9462e-06 | 678 | 7.1105e-06
|  | 439 | 4.9442e-06 | 678.2 | 7.1164e-06
|  | 439.1 | 4.9421e-06 | 678.4 | 7.1225e-06
|  | 439.2 | 4.9401e-06 | 678.6 | 7.1287e-06
|  | 439.3 | 4.9382e-06 | 678.8 | 7.1355e-06
|  | 439.4 | 4.9363e-06 | 679 | 7.1424e-06
|  | 439.5 | 4.9344e-06 | 679.2 | 7.1499e-06
|  | 439.6 | 4.9317e-06 | 679.4 | 7.1579e-06
|  | 439.7 | 4.9295e-06 | 679.6 | 7.1668e-06
|  | 439.8 | 4.9273e-06 | 679.8 | 7.1756e-06
|  | 439.9 | 4.9248e-06 | 680 | 7.1845e-06
|  | 440 | 4.9225e-06 | 680.2 | 7.193e-06
|  | 440.1 | 4.9201e-06 | 680.4 | 7.2011e-06
|  | 440.2 | 4.918e-06 | 680.6 | 7.2092e-06
|  | 440.3 | 4.9157e-06 | 680.8 | 7.2176e-06
|  | 440.4 | 4.9134e-06 | 681 | 7.2258e-06
|  | 440.5 | 4.9113e-06 | 681.2 | 7.234e-06
|  | 440.6 | 4.9091e-06 | 681.4 | 7.2425e-06
|  | 440.7 | 4.9069e-06 | 681.6 | 7.2506e-06
|  | 440.8 | 4.9046e-06 | 681.8 | 7.2588e-06
|  | 440.9 | 4.9024e-06 | 682 | 7.2668e-06
|  | 441 | 4.9001e-06 | 682.2 | 7.2743e-06
|  | 441.1 | 4.8979e-06 | 682.4 | 7.2817e-06
|  | 441.2 | 4.8956e-06 | 682.6 | 7.2889e-06
|  | 441.3 | 4.8934e-06 | 682.8 | 7.2964e-06
|  | 441.4 | 4.8911e-06 | 683 | 7.3047e-06
|  | 441.5 | 4.8895e-06 | 683.2 | 7.3138e-06
|  | 441.6 | 4.8875e-06 | 683.4 | 7.3235e-06
|  | 441.7 | 4.8854e-06 | 683.6 | 7.3334e-06
|  | 441.8 | 4.8835e-06 | 683.8 | 7.343e-06
|  | 441.9 | 4.8816e-06 | 684 | 7.3521e-06
|  | 442 | 4.8797e-06 | 684.2 | 7.3601e-06
|  | 442.1 | 4.8779e-06 | 684.4 | 7.3674e-06
|  | 442.2 | 4.876e-06 | 684.6 | 7.3748e-06
|  | 442.3 | 4.8741e-06 | 684.8 | 7.3823e-06
|  | 442.4 | 4.8722e-06 | 685 | 7.3897e-06
|  | 442.5 | 4.8703e-06 | 685.2 | 7.3965e-06
|  | 442.6 | 4.8684e-06 | 685.4 | 7.4029e-06
|  | 442.7 | 4.8673e-06 | 685.6 | 7.4092e-06
|  | 442.8 | 4.8657e-06 | 685.8 | 7.4154e-06
|  | 442.9 | 4.8641e-06 | 686 | 7.4219e-06
|  | 443 | 4.863e-06 | 686.2 | 7.4286e-06
|  | 443.1 | 4.8617e-06 | 686.4 | 7.4355e-06
|  | 443.2 | 4.8605e-06 | 686.6 | 7.4428e-06
|  | 443.3 | 4.8593e-06 | 686.8 | 7.4507e-06
|  | 443.4 | 4.8583e-06 | 687 | 7.4596e-06
|  | 443.5 | 4.8572e-06 | 687.2 | 7.4692e-06
|  | 443.6 | 4.856e-06 | 687.4 | 7.4789e-06
|  | 443.7 | 4.8547e-06 | 687.6 | 7.4877e-06
|  | 443.8 | 4.8536e-06 | 687.8 | 7.4961e-06
|  | 443.9 | 4.8524e-06 | 688 | 7.5047e-06
|  | 444 | 4.8511e-06 | 688.2 | 7.5135e-06
|  | 444.1 | 4.8499e-06 | 688.4 | 7.5224e-06
|  | 444.2 | 4.8487e-06 | 688.6 | 7.5308e-06
|  | 444.3 | 4.8475e-06 | 688.8 | 7.5388e-06
|  | 444.4 | 4.8462e-06 | 689 | 7.5466e-06
|  | 444.5 | 4.845e-06 | 689.2 | 7.554e-06
|  | 444.6 | 4.8441e-06 | 689.4 | 7.5608e-06
|  | 444.7 | 4.843e-06 | 689.6 | 7.5673e-06
|  | 444.8 | 4.8419e-06 | 689.8 | 7.5736e-06
|  | 444.9 | 4.841e-06 | 690 | 7.5797e-06
|  | 445 | 4.84e-06 | 690.2 | 7.5853e-06
|  | 445.1 | 4.839e-06 | 690.4 | 7.5918e-06
|  | 445.2 | 4.838e-06 | 690.6 | 7.5996e-06
|  | 445.3 | 4.837e-06 | 690.8 | 7.6077e-06
|  | 445.4 | 4.836e-06 | 691 | 7.6156e-06
|  | 445.5 | 4.835e-06 | 691.2 | 7.624e-06
|  | 445.6 | 4.834e-06 | 691.4 | 7.6333e-06
|  | 445.7 | 4.8331e-06 | 691.6 | 7.6431e-06
|  | 445.8 | 4.8321e-06 | 691.8 | 7.6525e-06
|  | 445.9 | 4.8311e-06 | 692 | 7.6615e-06
|  | 446 | 4.8302e-06 | 692.2 | 7.6708e-06
|  | 446.1 | 4.8292e-06 | 692.4 | 7.6801e-06
|  | 446.2 | 4.8272e-06 | 692.6 | 7.6892e-06
|  | 446.3 | 4.8259e-06 | 692.8 | 7.698e-06
|  | 446.4 | 4.8245e-06 | 693 | 7.7071e-06
|  | 446.5 | 4.8229e-06 | 693.2 | 7.716e-06
|  | 446.6 | 4.8213e-06 | 693.4 | 7.7249e-06
|  | 446.7 | 4.8197e-06 | 693.6 | 7.7335e-06
|  | 446.8 | 4.8186e-06 | 693.8 | 7.7419e-06
|  | 446.9 | 4.8172e-06 | 694 | 7.7497e-06
|  | 447 | 4.8158e-06 | 694.2 | 7.7569e-06
|  | 447.1 | 4.8146e-06 | 694.4 | 7.7637e-06
|  | 447.2 | 4.8133e-06 | 694.6 | 7.7705e-06
|  | 447.3 | 4.812e-06 | 694.8 | 7.7781e-06
|  | 447.4 | 4.8111e-06 | 695 | 7.7864e-06
|  | 447.5 | 4.81e-06 | 695.2 | 7.7953e-06
|  | 447.6 | 4.8089e-06 | 695.4 | 7.8045e-06
|  | 447.7 | 4.8079e-06 | 695.6 | 7.8139e-06
|  | 447.8 | 4.8067e-06 | 695.8 | 7.8235e-06
|  | 447.9 | 4.8056e-06 | 696 | 7.8336e-06
|  | 448 | 4.8046e-06 | 696.2 | 7.8436e-06
|  | 448.1 | 4.8033e-06 | 696.4 | 7.8535e-06
|  | 448.2 | 4.8022e-06 | 696.6 | 7.8633e-06
|  | 448.3 | 4.801e-06 | 696.8 | 7.8732e-06
|  | 448.4 | 4.7999e-06 | 697 | 7.8833e-06
|  | 448.5 | 4.7987e-06 | 697.2 | 7.8938e-06
|  | 448.6 | 4.7976e-06 | 697.4 | 7.9049e-06
|  | 448.7 | 4.7966e-06 | 697.6 | 7.9156e-06
|  | 448.8 | 4.7955e-06 | 697.8 | 7.9255e-06
|  | 448.9 | 4.7944e-06 | 698 | 7.9349e-06
|  | 449 | 4.7935e-06 | 698.2 | 7.9439e-06
|  | 449.1 | 4.7925e-06 | 698.4 | 7.953e-06
|  | 449.2 | 4.7915e-06 | 698.6 | 7.9622e-06
|  | 449.3 | 4.7907e-06 | 698.8 | 7.9712e-06
|  | 449.4 | 4.7898e-06 | 699 | 7.9802e-06
|  | 449.5 | 4.7889e-06 | 699.2 | 7.9892e-06
|  | 449.6 | 4.7881e-06 | 699.4 | 7.9991e-06
|  | 449.7 | 4.7873e-06 | 699.6 | 8.0088e-06
|  | 449.8 | 4.7865e-06 | 699.8 | 8.0179e-06
|  | 449.9 | 4.7857e-06 | 700 | 8.0263e-06
|  | 450 | 4.7851e-06 | 700.2 | 8.0348e-06
|  | 450.1 | 4.7844e-06 | 700.4 | 8.0437e-06
|  | 450.2 | 4.7837e-06 | 700.6 | 8.0533e-06
|  | 450.3 | 4.7831e-06 | 700.8 | 8.0632e-06
|  | 450.4 | 4.7824e-06 | 701 | 8.073e-06
|  | 450.5 | 4.7818e-06 | 701.2 | 8.0819e-06
|  | 450.6 | 4.7812e-06 | 701.4 | 8.0904e-06
|  | 450.7 | 4.7806e-06 | 701.6 | 8.0988e-06
|  | 450.8 | 4.78e-06 | 701.8 | 8.1067e-06
|  | 450.9 | 4.7792e-06 | 702 | 8.1144e-06
|  | 451 | 4.7785e-06 | 702.2 | 8.1214e-06
|  | 451.1 | 4.7779e-06 | 702.4 | 8.1288e-06
|  | 451.2 | 4.7779e-06 | 702.6 | 8.1365e-06
|  | 451.3 | 4.7774e-06 | 702.8 | 8.144e-06
|  | 451.4 | 4.777e-06 | 703 | 8.1513e-06
|  | 451.5 | 4.7768e-06 | 703.2 | 8.1582e-06
|  | 451.6 | 4.7767e-06 | 703.4 | 8.165e-06
|  | 451.7 | 4.7765e-06 | 703.6 | 8.1722e-06
|  | 451.8 | 4.7763e-06 | 703.8 | 8.1802e-06
|  | 451.9 | 4.7761e-06 | 704 | 8.1896e-06
|  | 452 | 4.7759e-06 | 704.2 | 8.1993e-06
|  | 452.1 | 4.7757e-06 | 704.4 | 8.2089e-06
|  | 452.2 | 4.7753e-06 | 704.6 | 8.2189e-06
|  | 452.3 | 4.7751e-06 | 704.8 | 8.2292e-06
|  | 452.4 | 4.7748e-06 | 705 | 8.24e-06
|  | 452.5 | 4.7743e-06 | 705.2 | 8.2507e-06
|  | 452.6 | 4.7739e-06 | 705.4 | 8.2614e-06
|  | 452.7 | 4.7735e-06 | 705.6 | 8.2722e-06
|  | 452.8 | 4.7727e-06 | 705.8 | 8.2829e-06
|  | 452.9 | 4.7722e-06 | 706 | 8.2935e-06
|  | 453 | 4.7716e-06 | 706.2 | 8.3041e-06
|  | 453.1 | 4.7714e-06 | 706.4 | 8.3147e-06
|  | 453.2 | 4.771e-06 | 706.6 | 8.3243e-06
|  | 453.3 | 4.7705e-06 | 706.8 | 8.3336e-06
|  | 453.4 | 4.7704e-06 | 707 | 8.3422e-06
|  | 453.5 | 4.7702e-06 | 707.2 | 8.3507e-06
|  | 453.6 | 4.77e-06 | 707.4 | 8.3586e-06
|  | 453.7 | 4.7694e-06 | 707.6 | 8.3661e-06
|  | 453.8 | 4.7691e-06 | 707.8 | 8.3739e-06
|  | 453.9 | 4.7688e-06 | 708 | 8.382e-06
|  | 454 | 4.7686e-06 | 708.2 | 8.3905e-06
|  | 454.1 | 4.7683e-06 | 708.4 | 8.3996e-06
|  | 454.2 | 4.768e-06 | 708.6 | 8.4086e-06
|  | 454.3 | 4.7678e-06 | 708.8 | 8.4178e-06
|  | 454.4 | 4.7676e-06 | 709 | 8.4269e-06
|  | 454.5 | 4.7674e-06 | 709.2 | 8.4363e-06
|  | 454.6 | 4.7672e-06 | 709.4 | 8.4463e-06
|  | 454.7 | 4.7664e-06 | 709.6 | 8.4572e-06
|  | 454.8 | 4.7659e-06 | 709.8 | 8.4698e-06
|  | 454.9 | 4.7655e-06 | 710 | 8.4829e-06
|  | 455 | 4.7652e-06 | 710.2 | 8.4961e-06
|  | 455.1 | 4.7647e-06 | 710.4 | 8.5091e-06
|  | 455.2 | 4.7643e-06 | 710.6 | 8.5217e-06
|  | 455.3 | 4.7642e-06 | 710.8 | 8.5346e-06
|  | 455.4 | 4.7639e-06 | 711 | 8.5478e-06
|  | 455.5 | 4.7637e-06 | 711.2 | 8.5611e-06
|  | 455.6 | 4.7637e-06 | 711.4 | 8.5734e-06
|  | 455.7 | 4.7635e-06 | 711.6 | 8.5856e-06
|  | 455.8 | 4.7633e-06 | 711.8 | 8.5974e-06
|  | 455.9 | 4.763e-06 | 712 | 8.6092e-06
|  | 456 | 4.7628e-06 | 712.2 | 8.6201e-06
|  | 456.1 | 4.7626e-06 | 712.4 | 8.6297e-06
|  | 456.2 | 4.7626e-06 | 712.6 | 8.6391e-06
|  | 456.3 | 4.7624e-06 | 712.8 | 8.6486e-06
|  | 456.4 | 4.7622e-06 | 713 | 8.6595e-06
|  | 456.5 | 4.7623e-06 | 713.2 | 8.6709e-06
|  | 456.6 | 4.7623e-06 | 713.4 | 8.683e-06
|  | 456.7 | 4.7623e-06 | 713.6 | 8.695e-06
|  | 456.8 | 4.7622e-06 | 713.8 | 8.7068e-06
|  | 456.9 | 4.7622e-06 | 714 | 8.7194e-06
|  | 457 | 4.7622e-06 | 714.2 | 8.7335e-06
|  | 457.1 | 4.7622e-06 | 714.4 | 8.7481e-06
|  | 457.2 | 4.7622e-06 | 714.6 | 8.7632e-06
|  | 457.3 | 4.7621e-06 | 714.8 | 8.7781e-06
|  | 457.4 | 4.7621e-06 | 715 | 8.7935e-06
|  | 457.5 | 4.7618e-06 | 715.2 | 8.8098e-06
|  | 457.6 | 4.7616e-06 | 715.4 | 8.8263e-06
|  | 457.7 | 4.7615e-06 | 715.6 | 8.8426e-06
|  | 457.8 | 4.7616e-06 | 715.8 | 8.8585e-06
|  | 457.9 | 4.7615e-06 | 716 | 8.8738e-06
|  | 458 | 4.7614e-06 | 716.2 | 8.8889e-06
|  | 458.1 | 4.7618e-06 | 716.4 | 8.9039e-06
|  | 458.2 | 4.762e-06 | 716.6 | 8.9181e-06
|  | 458.3 | 4.7621e-06 | 716.8 | 8.9317e-06
|  | 458.4 | 4.7624e-06 | 717 | 8.9444e-06
|  | 458.5 | 4.7627e-06 | 717.2 | 8.9567e-06
|  | 458.6 | 4.7629e-06 | 717.4 | 8.9691e-06
|  | 458.7 | 4.7633e-06 | 717.6 | 8.9819e-06
|  | 458.8 | 4.7636e-06 | 717.8 | 8.995e-06
|  | 458.9 | 4.764e-06 | 718 | 9.0081e-06
|  | 459 | 4.7645e-06 | 718.2 | 9.0218e-06
|  | 459.1 | 4.7649e-06 | 718.4 | 9.0357e-06
|  | 459.2 | 4.7654e-06 | 718.6 | 9.0497e-06
|  | 459.3 | 4.7659e-06 | 718.8 | 9.0638e-06
|  | 459.4 | 4.7664e-06 | 719 | 9.0784e-06
|  | 459.5 | 4.7669e-06 | 719.2 | 9.0936e-06
|  | 459.6 | 4.7675e-06 | 719.4 | 9.1096e-06
|  | 459.7 | 4.7681e-06 | 719.6 | 9.1263e-06
|  | 459.8 | 4.7686e-06 | 719.8 | 9.1443e-06
|  | 459.9 | 4.7692e-06 | 720 | 9.1632e-06
|  | 460 | 4.77e-06 | 720.2 | 9.1828e-06
|  | 460.1 | 4.7706e-06 | 720.4 | 9.203e-06
|  | 460.2 | 4.7713e-06 | 720.6 | 9.2233e-06
|  | 460.3 | 4.7703e-06 | 720.8 | 9.2436e-06
|  | 460.4 | 4.7705e-06 | 721 | 9.2633e-06
|  | 460.5 | 4.7706e-06 | 721.2 | 9.283e-06
|  | 460.6 | 4.7714e-06 | 721.4 | 9.302e-06
|  | 460.7 | 4.7715e-06 | 721.6 | 9.3205e-06
|  | 460.8 | 4.7716e-06 | 721.8 | 9.338e-06
|  | 460.9 | 4.7727e-06 | 722 | 9.3556e-06
|  | 461 | 4.7735e-06 | 722.2 | 9.3736e-06
|  | 461.1 | 4.7742e-06 | 722.4 | 9.3917e-06
|  | 461.2 | 4.7751e-06 | 722.6 | 9.4091e-06
|  | 461.3 | 4.7758e-06 | 722.8 | 9.4245e-06
|  | 461.4 | 4.7766e-06 | 723 | 9.4387e-06
|  | 461.5 | 4.7776e-06 | 723.2 | 9.4508e-06
|  | 461.6 | 4.7785e-06 | 723.4 | 9.4612e-06
|  | 461.7 | 4.7793e-06 | 723.6 | 9.4713e-06
|  | 461.8 | 4.7803e-06 | 723.8 | 9.4819e-06
|  | 461.9 | 4.7812e-06 | 724 | 9.4934e-06
|  | 462 | 4.7821e-06 | 724.2 | 9.5061e-06
|  | 462.1 | 4.783e-06 | 724.4 | 9.5194e-06
|  | 462.2 | 4.7839e-06 | 724.6 | 9.5323e-06
|  | 462.3 | 4.7849e-06 | 724.8 | 9.5441e-06
|  | 462.4 | 4.7855e-06 | 725 | 9.555e-06
|  | 462.5 | 4.7863e-06 | 725.2 | 9.5654e-06
|  | 462.6 | 4.7871e-06 | 725.4 | 9.5762e-06
|  | 462.7 | 4.7879e-06 | 725.6 | 9.5879e-06
|  | 462.8 | 4.789e-06 | 725.8 | 9.6006e-06
|  | 462.9 | 4.7899e-06 | 726 | 9.6142e-06
|  | 463 | 4.7908e-06 | 726.2 | 9.6281e-06
|  | 463.1 | 4.7917e-06 | 726.4 | 9.642e-06
|  | 463.2 | 4.7927e-06 | 726.6 | 9.6546e-06
|  | 463.3 | 4.7937e-06 | 726.8 | 9.665e-06
|  | 463.4 | 4.7941e-06 | 727 | 9.6748e-06
|  | 463.5 | 4.7949e-06 | 727.2 | 9.6847e-06
|  | 463.6 | 4.7957e-06 | 727.4 | 9.6954e-06
|  | 463.7 | 4.7977e-06 | 727.6 | 9.7067e-06
|  | 463.8 | 4.7988e-06 | 727.8 | 9.7193e-06
|  | 463.9 | 4.8e-06 | 728 | 9.732e-06
|  | 464 | 4.8009e-06 | 728.2 | 9.7441e-06
|  | 464.1 | 4.8022e-06 | 728.4 | 9.7552e-06
|  | 464.2 | 4.8035e-06 | 728.6 | 9.7655e-06
|  | 464.3 | 4.8045e-06 | 728.8 | 9.7745e-06
|  | 464.4 | 4.8055e-06 | 729 | 9.783e-06
|  | 464.5 | 4.8065e-06 | 729.2 | 9.7917e-06
|  | 464.6 | 4.8074e-06 | 729.4 | 9.8013e-06
|  | 464.7 | 4.8084e-06 | 729.6 | 9.8115e-06
|  | 464.8 | 4.8094e-06 | 729.8 | 9.8216e-06
|  | 464.9 | 4.8101e-06 | 730 | 9.8316e-06
|  | 465 | 4.8109e-06 | 730.2 | 9.8407e-06
|  | 465.1 | 4.8118e-06 | 730.4 | 9.8484e-06
|  | 465.2 | 4.8125e-06 | 730.6 | 9.8551e-06
|  | 465.3 | 4.8132e-06 | 730.8 | 9.8606e-06
|  | 465.4 | 4.814e-06 | 731 | 9.8662e-06
|  | 465.5 | 4.8146e-06 | 731.2 | 9.8733e-06
|  | 465.6 | 4.8153e-06 | 731.4 | 9.8816e-06
|  | 465.7 | 4.816e-06 | 731.6 | 9.8903e-06
|  | 465.8 | 4.8166e-06 | 731.8 | 9.8995e-06
|  | 465.9 | 4.8173e-06 | 732 | 9.9088e-06
|  | 466 | 4.8179e-06 | 732.2 | 9.918e-06
|  | 466.1 | 4.8185e-06 | 732.4 | 9.9275e-06
|  | 466.2 | 4.8188e-06 | 732.6 | 9.9372e-06
|  | 466.3 | 4.8193e-06 | 732.8 | 9.9472e-06
|  | 466.4 | 4.8198e-06 | 733 | 9.9574e-06
|  | 466.5 | 4.8201e-06 | 733.2 | 9.9686e-06
|  | 466.6 | 4.8205e-06 | 733.4 | 9.9812e-06
|  | 466.7 | 4.8208e-06 | 733.6 | 9.9942e-06
|  | 466.8 | 4.8215e-06 | 733.8 | 1.0007e-05
|  | 466.9 | 4.822e-06 | 734 | 1.0019e-05
|  | 467 | 4.8224e-06 | 734.2 | 1.0031e-05
|  | 467.1 | 4.8231e-06 | 734.4 | 1.0042e-05
|  | 467.2 | 4.8237e-06 | 734.6 | 1.0054e-05
|  | 467.3 | 4.8243e-06 | 734.8 | 1.0066e-05
|  | 467.4 | 4.8249e-06 | 735 | 1.0076e-05
|  | 467.5 | 4.8254e-06 | 735.2 | 1.0087e-05
|  | 467.6 | 4.826e-06 | 735.4 | 1.0098e-05
|  | 467.7 | 4.8266e-06 | 735.6 | 1.0108e-05
|  | 467.8 | 4.8271e-06 | 735.8 | 1.0118e-05
|  | 467.9 | 4.8276e-06 | 736 | 1.0129e-05
|  | 468 | 4.8284e-06 | 736.2 | 1.014e-05
|  | 468.1 | 4.829e-06 | 736.4 | 1.0153e-05
|  | 468.2 | 4.8297e-06 | 736.6 | 1.0166e-05
|  | 468.3 | 4.8306e-06 | 736.8 | 1.018e-05
|  | 468.4 | 4.8315e-06 | 737 | 1.0194e-05
|  | 468.5 | 4.8323e-06 | 737.2 | 1.0207e-05
|  | 468.6 | 4.8332e-06 | 737.4 | 1.0221e-05
|  | 468.7 | 4.8341e-06 | 737.6 | 1.0236e-05
|  | 468.8 | 4.8349e-06 | 737.8 | 1.0252e-05
|  | 468.9 | 4.8358e-06 | 738 | 1.0267e-05
|  | 469 | 4.8369e-06 | 738.2 | 1.0282e-05
|  | 469.1 | 4.8379e-06 | 738.4 | 1.0296e-05
|  | 469.2 | 4.8389e-06 | 738.6 | 1.0309e-05
|  | 469.3 | 4.84e-06 | 738.8 | 1.0323e-05
|  | 469.4 | 4.841e-06 | 739 | 1.0335e-05
|  | 469.5 | 4.842e-06 | 739.2 | 1.0345e-05
|  | 469.6 | 4.8432e-06 | 739.4 | 1.0355e-05
|  | 469.7 | 4.8443e-06 | 739.6 | 1.0365e-05
|  | 469.8 | 4.8454e-06 | 739.8 | 1.0375e-05
|  | 469.9 | 4.8469e-06 | 740 | 1.0386e-05
|  | 470 | 4.8482e-06 | 740.2 | 1.0395e-05
|  | 470.1 | 4.8495e-06 | 740.4 | 1.0405e-05
|  | 470.2 | 4.8508e-06 | 740.6 | 1.0414e-05
|  | 470.3 | 4.8521e-06 | 740.8 | 1.0424e-05
|  | 470.4 | 4.8534e-06 | 741 | 1.0434e-05
|  | 470.5 | 4.8548e-06 | 741.2 | 1.0446e-05
|  | 470.6 | 4.8561e-06 | 741.4 | 1.0459e-05
|  | 470.7 | 4.8574e-06 | 741.6 | 1.0473e-05
|  | 470.8 | 4.8587e-06 | 741.8 | 1.0488e-05
|  | 470.9 | 4.8601e-06 | 742 | 1.0503e-05
|  | 471 | 4.8614e-06 | 742.2 | 1.052e-05
|  | 471.1 | 4.8625e-06 | 742.4 | 1.0538e-05
|  | 471.2 | 4.8638e-06 | 742.6 | 1.0555e-05
|  | 471.3 | 4.865e-06 | 742.8 | 1.0572e-05
|  | 471.4 | 4.8663e-06 | 743 | 1.0588e-05
|  | 471.5 | 4.8675e-06 | 743.2 | 1.0605e-05
|  | 471.6 | 4.8687e-06 | 743.4 | 1.0622e-05
|  | 471.7 | 4.8698e-06 | 743.6 | 1.0639e-05
|  | 471.8 | 4.871e-06 | 743.8 | 1.0655e-05
|  | 471.9 | 4.8722e-06 | 744 | 1.0668e-05
|  | 472 | 4.8735e-06 | 744.2 | 1.068e-05
|  | 472.1 | 4.8747e-06 | 744.4 | 1.0691e-05
|  | 472.2 | 4.8759e-06 | 744.6 | 1.0703e-05
|  | 472.3 | 4.8771e-06 | 744.8 | 1.0714e-05
|  | 472.4 | 4.8783e-06 | 745 | 1.0725e-05
|  | 472.5 | 4.8795e-06 | 745.2 | 1.0735e-05
|  | 472.6 | 4.8808e-06 | 745.4 | 1.0747e-05
|  | 472.7 | 4.8822e-06 | 745.6 | 1.076e-05
|  | 472.8 | 4.8834e-06 | 745.8 | 1.0774e-05
|  | 472.9 | 4.8847e-06 | 746 | 1.0788e-05
|  | 473 | 4.8859e-06 | 746.2 | 1.0802e-05
|  | 473.1 | 4.8872e-06 | 746.4 | 1.0818e-05
|  | 473.2 | 4.8885e-06 | 746.6 | 1.0834e-05
|  | 473.3 | 4.8895e-06 | 746.8 | 1.0852e-05
|  | 473.4 | 4.8907e-06 | 747 | 1.0871e-05
|  | 473.5 | 4.8919e-06 | 747.2 | 1.0891e-05
|  | 473.6 | 4.8932e-06 | 747.4 | 1.091e-05
|  | 473.7 | 4.8944e-06 | 747.6 | 1.0929e-05
|  | 473.8 | 4.8956e-06 | 747.8 | 1.0948e-05
|  | 473.9 | 4.897e-06 | 748 | 1.0967e-05
|  | 474 | 4.8984e-06 | 748.2 | 1.0985e-05
|  | 474.1 | 4.8997e-06 | 748.4 | 1.1001e-05
|  | 474.2 | 4.9012e-06 | 748.6 | 1.1016e-05
|  | 474.3 | 4.9026e-06 | 748.8 | 1.1031e-05
|  | 474.4 | 4.904e-06 | 749 | 1.1047e-05
|  | 474.5 | 4.9052e-06 | 749.2 | 1.1063e-05
|  | 474.6 | 4.9066e-06 | 749.4 | 1.1078e-05
|  | 474.7 | 4.908e-06 | 749.6 | 1.1092e-05
|  | 474.8 | 4.9093e-06 | 749.8 | 1.1104e-05
|  | 474.9 | 4.9106e-06 | 750 | 1.1116e-05
|  | 475 | 4.9119e-06 | 750.2 | 1.113e-05
|  | 475.1 | 4.9133e-06 | 750.4 | 1.1144e-05
|  | 475.2 | 4.9147e-06 | 750.6 | 1.1157e-05
|  | 475.3 | 4.9161e-06 | 750.8 | 1.1172e-05
|  | 475.4 | 4.9177e-06 | 751 | 1.1188e-05
|  | 475.5 | 4.9192e-06 | 751.2 | 1.1206e-05
|  | 475.6 | 4.9207e-06 | 751.4 | 1.1224e-05
|  | 475.7 | 4.9219e-06 | 751.6 | 1.1242e-05
|  | 475.8 | 4.9234e-06 | 751.8 | 1.1259e-05
|  | 475.9 | 4.9248e-06 | 752 | 1.1275e-05
|  | 476 | 4.9262e-06 | 752.2 | 1.1291e-05
|  | 476.1 | 4.9276e-06 | 752.4 | 1.1308e-05
|  | 476.2 | 4.929e-06 | 752.6 | 1.1323e-05
|  | 476.3 | 4.9303e-06 | 752.8 | 1.1339e-05
|  | 476.4 | 4.9323e-06 | 753 | 1.1354e-05
|  | 476.5 | 4.9338e-06 | 753.2 | 1.137e-05
|  | 476.6 | 4.9354e-06 | 753.4 | 1.1385e-05
|  | 476.7 | 4.9369e-06 | 753.6 | 1.1399e-05
|  | 476.8 | 4.9385e-06 | 753.8 | 1.1411e-05
|  | 476.9 | 4.94e-06 | 754 | 1.1424e-05
|  | 477 | 4.9415e-06 | 754.2 | 1.1438e-05
|  | 477.1 | 4.943e-06 | 754.4 | 1.1451e-05
|  | 477.2 | 4.9445e-06 | 754.6 | 1.1467e-05
|  | 477.3 | 4.9458e-06 | 754.8 | 1.1484e-05
|  | 477.4 | 4.9473e-06 | 755 | 1.1503e-05
|  | 477.5 | 4.9487e-06 | 755.2 | 1.1524e-05
|  | 477.6 | 4.9502e-06 | 755.4 | 1.1545e-05
|  | 477.7 | 4.9516e-06 | 755.6 | 1.1567e-05
|  | 477.8 | 4.953e-06 | 755.8 | 1.1587e-05
|  | 477.9 | 4.9546e-06 | 756 | 1.1607e-05
|  | 478 | 4.9561e-06 | 756.2 | 1.1627e-05
|  | 478.1 | 4.9576e-06 | 756.4 | 1.1646e-05
|  | 478.2 | 4.9591e-06 | 756.6 | 1.1663e-05
|  | 478.3 | 4.9607e-06 | 756.8 | 1.1681e-05
|  | 478.4 | 4.9622e-06 | 757 | 1.17e-05
|  | 478.5 | 4.9635e-06 | 757.2 | 1.1718e-05
|  | 478.6 | 4.965e-06 | 757.4 | 1.1736e-05
|  | 478.7 | 4.9664e-06 | 757.6 | 1.1753e-05
|  | 478.8 | 4.9678e-06 | 757.8 | 1.1769e-05
|  | 478.9 | 4.9692e-06 | 758 | 1.1785e-05
|  | 479 | 4.9706e-06 | 758.2 | 1.18e-05
|  | 479.1 | 4.9725e-06 | 758.4 | 1.1815e-05
|  | 479.2 | 4.9741e-06 | 758.6 | 1.1831e-05
|  | 479.3 | 4.9758e-06 | 758.8 | 1.1849e-05
|  | 479.4 | 4.9774e-06 | 759 | 1.1869e-05
|  | 479.5 | 4.9792e-06 | 759.2 | 1.189e-05
|  | 479.6 | 4.9809e-06 | 759.4 | 1.1912e-05
|  | 479.7 | 4.9825e-06 | 759.6 | 1.1932e-05
|  | 479.8 | 4.9841e-06 | 759.8 | 1.1952e-05
|  | 479.9 | 4.9858e-06 | 760 | 1.1973e-05
|  | 480 | 4.9874e-06 | 760.2 | 1.1992e-05
|  | 480.1 | 4.9891e-06 | 760.4 | 1.2012e-05
|  | 480.2 | 4.9907e-06 | 760.6 | 1.2032e-05
|  | 480.3 | 4.9922e-06 | 760.8 | 1.2053e-05
|  | 480.4 | 4.9938e-06 | 761 | 1.2074e-05
|  | 480.5 | 4.9953e-06 | 761.2 | 1.2096e-05
|  | 480.6 | 4.9969e-06 | 761.4 | 1.2115e-05
|  | 480.7 | 4.998e-06 | 761.6 | 1.2133e-05
|  | 480.8 | 4.9994e-06 | 761.8 | 1.215e-05
|  | 480.9 | 5.0008e-06 | 762 | 1.2166e-05
|  | 481 | 5.0024e-06 | 762.2 | 1.2181e-05
|  | 481.1 | 5.0038e-06 | 762.4 | 1.2197e-05
|  | 481.2 | 5.0052e-06 | 762.6 | 1.2214e-05
|  | 481.3 | 5.0069e-06 | 762.8 | 1.2231e-05
|  | 481.4 | 5.0085e-06 | 763 | 1.2248e-05
|  | 481.5 | 5.01e-06 | 763.2 | 1.2264e-05
|  | 481.6 | 5.0117e-06 | 763.4 | 1.228e-05
|  | 481.7 | 5.0133e-06 | 763.6 | 1.2297e-05
|  | 481.8 | 5.0149e-06 | 763.8 | 1.2315e-05
|  | 481.9 | 5.0166e-06 | 764 | 1.2334e-05
|  | 482 | 5.0182e-06 | 764.2 | 1.2354e-05
|  | 482.1 | 5.0198e-06 | 764.4 | 1.2375e-05
|  | 482.2 | 5.0212e-06 | 764.6 | 1.2396e-05
|  | 482.3 | 5.0228e-06 | 764.8 | 1.2418e-05
|  | 482.4 | 5.0243e-06 | 765 | 1.2439e-05
|  | 482.5 | 5.0259e-06 | 765.2 | 1.2459e-05
|  | 482.6 | 5.0274e-06 | 765.4 | 1.2479e-05
|  | 482.7 | 5.0289e-06 | 765.6 | 1.2499e-05
|  | 482.8 | 5.0307e-06 | 765.8 | 1.2519e-05
|  | 482.9 | 5.0323e-06 | 766 | 1.2539e-05
|  | 483 | 5.034e-06 | 766.2 | 1.2559e-05
|  | 483.1 | 5.0356e-06 | 766.4 | 1.2576e-05
|  | 483.2 | 5.0373e-06 | 766.6 | 1.2592e-05
|  | 483.3 | 5.0389e-06 | 766.8 | 1.2608e-05
|  | 483.4 | 5.0408e-06 | 767 | 1.2624e-05
|  | 483.5 | 5.0425e-06 | 767.2 | 1.2641e-05
|  | 483.6 | 5.0442e-06 | 767.4 | 1.266e-05
|  | 483.7 | 5.0459e-06 | 767.6 | 1.2679e-05
|  | 483.8 | 5.0476e-06 | 767.8 | 1.27e-05
|  | 483.9 | 5.0494e-06 | 768 | 1.2722e-05
|  | 484 | 5.0513e-06 | 768.2 | 1.2744e-05
|  | 484.1 | 5.0531e-06 | 768.4 | 1.2768e-05
|  | 484.2 | 5.0549e-06 | 768.6 | 1.2792e-05
|  | 484.3 | 5.0565e-06 | 768.8 | 1.2818e-05
|  | 484.4 | 5.0583e-06 | 769 | 1.2844e-05
|  | 484.5 | 5.06e-06 | 769.2 | 1.287e-05
|  | 484.6 | 5.0615e-06 | 769.4 | 1.2895e-05
|  | 484.7 | 5.0631e-06 | 769.6 | 1.2919e-05
|  | 484.8 | 5.0647e-06 | 769.8 | 1.2941e-05
|  | 484.9 | 5.0662e-06 | 770 | 1.2962e-05
|  | 485 | 5.0677e-06 | 770.2 | 1.2981e-05
|  | 485.1 | 5.0693e-06 | 770.4 | 1.2998e-05
|  | 485.2 | 5.0711e-06 | 770.6 | 1.3014e-05
|  | 485.3 | 5.0728e-06 | 770.8 | 1.3032e-05
|  | 485.4 | 5.0744e-06 | 771 | 1.3049e-05
|  | 485.5 | 5.0761e-06 | 771.2 | 1.3067e-05
|  | 485.6 | 5.0779e-06 | 771.4 | 1.3083e-05
|  | 485.7 | 5.0796e-06 | 771.6 | 1.31e-05
|  | 485.8 | 5.0814e-06 | 771.8 | 1.3117e-05
|  | 485.9 | 5.0831e-06 | 772 | 1.3137e-05
|  | 486 | 5.0848e-06 | 772.2 | 1.3159e-05
|  | 486.1 | 5.0865e-06 | 772.4 | 1.3183e-05
|  | 486.2 | 5.0882e-06 | 772.6 | 1.3207e-05
|  | 486.3 | 5.0899e-06 | 772.8 | 1.3232e-05
|  | 486.4 | 5.0916e-06 | 773 | 1.3258e-05
|  | 486.5 | 5.093e-06 | 773.2 | 1.3284e-05
|  | 486.6 | 5.0946e-06 | 773.4 | 1.3308e-05
|  | 486.7 | 5.0962e-06 | 773.6 | 1.3332e-05
|  | 486.8 | 5.0972e-06 | 773.8 | 1.3353e-05
|  | 486.9 | 5.0985e-06 | 774 | 1.3373e-05
|  | 487 | 5.0999e-06 | 774.2 | 1.3393e-05
|  | 487.1 | 5.1013e-06 | 774.4 | 1.3411e-05
|  | 487.2 | 5.1026e-06 | 774.6 | 1.3429e-05
|  | 487.3 | 5.1039e-06 | 774.8 | 1.3447e-05
|  | 487.4 | 5.1054e-06 | 775 | 1.3466e-05
|  | 487.5 | 5.1068e-06 | 775.2 | 1.3487e-05
|  | 487.6 | 5.1082e-06 | 775.4 | 1.351e-05
|  | 487.7 | 5.1097e-06 | 775.6 | 1.3532e-05
|  | 487.8 | 5.1112e-06 | 775.8 | 1.3555e-05
|  | 487.9 | 5.1126e-06 | 776 | 1.3579e-05
|  | 488 | 5.1142e-06 | 776.2 | 1.3605e-05
|  | 488.1 | 5.1157e-06 | 776.4 | 1.3631e-05
|  | 488.2 | 5.1172e-06 | 776.6 | 1.3659e-05
|  | 488.3 | 5.119e-06 | 776.8 | 1.3689e-05
|  | 488.4 | 5.1206e-06 | 777 | 1.3719e-05
|  | 488.5 | 5.1223e-06 | 777.2 | 1.3748e-05
|  | 488.6 | 5.1236e-06 | 777.4 | 1.3775e-05
|  | 488.7 | 5.1252e-06 | 777.6 | 1.3798e-05
|  | 488.8 | 5.1267e-06 | 777.8 | 1.3817e-05
|  | 488.9 | 5.1284e-06 | 778 | 1.3834e-05
|  | 489 | 5.13e-06 | 778.2 | 1.385e-05
|  | 489.1 | 5.1315e-06 | 778.4 | 1.3867e-05
|  | 489.2 | 5.1333e-06 | 778.6 | 1.3886e-05
|  | 489.3 | 5.135e-06 | 778.8 | 1.3907e-05
|  | 489.4 | 5.1367e-06 | 779 | 1.393e-05
|  | 489.5 | 5.1382e-06 | 779.2 | 1.3954e-05
|  | 489.6 | 5.1399e-06 | 779.4 | 1.3977e-05
|  | 489.7 | 5.1415e-06 | 779.6 | 1.4e-05
|  | 489.8 | 5.1428e-06 | 779.8 | 1.4024e-05
|  | 489.9 | 5.1442e-06 | 780 | 1.4049e-05
|  | 490 | 5.1457e-06 | 780.2 | 1.4077e-05
|  | 490.1 | 5.1472e-06 | 780.4 | 1.4104e-05
|  | 490.2 | 5.1486e-06 | 780.6 | 1.4132e-05
|  | 490.3 | 5.15e-06 | 780.8 | 1.4158e-05
|  | 490.4 | 5.1515e-06 | 781 | 1.4182e-05
|  | 490.5 | 5.153e-06 | 781.2 | 1.4203e-05
|  | 490.6 | 5.1545e-06 | 781.4 | 1.4225e-05
|  | 490.7 | 5.1558e-06 | 781.6 | 1.4248e-05
|  | 490.8 | 5.1571e-06 | 781.8 | 1.4273e-05
|  | 490.9 | 5.1585e-06 | 782 | 1.4301e-05
|  | 491 | 5.1598e-06 | 782.2 | 1.4329e-05
|  | 491.1 | 5.161e-06 | 782.4 | 1.4358e-05
|  | 491.2 | 5.1623e-06 | 782.6 | 1.4386e-05
|  | 491.3 | 5.1638e-06 | 782.8 | 1.4413e-05
|  | 491.4 | 5.1651e-06 | 783 | 1.4437e-05
|  | 491.5 | 5.1665e-06 | 783.2 | 1.4462e-05
|  | 491.6 | 5.1679e-06 | 783.4 | 1.4486e-05
|  | 491.7 | 5.1693e-06 | 783.6 | 1.451e-05
|  | 491.8 | 5.1707e-06 | 783.8 | 1.4532e-05
|  | 491.9 | 5.1722e-06 | 784 | 1.4552e-05
|  | 492 | 5.1736e-06 | 784.2 | 1.4573e-05
|  | 492.1 | 5.1751e-06 | 784.4 | 1.4594e-05
|  | 492.2 | 5.1765e-06 | 784.6 | 1.4617e-05
|  | 492.3 | 5.1779e-06 | 784.8 | 1.4641e-05
|  | 492.4 | 5.1794e-06 | 785 | 1.4668e-05
|  | 492.5 | 5.1808e-06 | 785.2 | 1.4695e-05
|  | 492.6 | 5.1821e-06 | 785.4 | 1.4724e-05
|  | 492.7 | 5.1835e-06 | 785.6 | 1.4756e-05
|  | 492.8 | 5.1848e-06 | 785.8 | 1.4789e-05
|  | 492.9 | 5.1864e-06 | 786 | 1.4821e-05
|  | 493 | 5.1878e-06 | 786.2 | 1.4853e-05
|  | 493.1 | 5.1892e-06 | 786.4 | 1.4884e-05
|  | 493.2 | 5.191e-06 | 786.6 | 1.4913e-05
|  | 493.3 | 5.1926e-06 | 786.8 | 1.4941e-05
|  | 493.4 | 5.1942e-06 | 787 | 1.4968e-05
|  | 493.5 | 5.1956e-06 | 787.2 | 1.4994e-05
|  | 493.6 | 5.1972e-06 | 787.4 | 1.5019e-05
|  | 493.7 | 5.1988e-06 | 787.6 | 1.5043e-05
|  | 493.8 | 5.2002e-06 | 787.8 | 1.5069e-05
|  | 493.9 | 5.2017e-06 | 788 | 1.5098e-05
|  | 494 | 5.2032e-06 | 788.2 | 1.5126e-05
|  | 494.1 | 5.2049e-06 | 788.4 | 1.5155e-05
|  | 494.2 | 5.2064e-06 | 788.6 | 1.5185e-05
|  | 494.3 | 5.208e-06 | 788.8 | 1.5214e-05
|  | 494.4 | 5.2094e-06 | 789 | 1.5243e-05
|  | 494.5 | 5.2109e-06 | 789.2 | 1.527e-05
|  | 494.6 | 5.2124e-06 | 789.4 | 1.5295e-05
|  | 494.7 | 5.2138e-06 | 789.6 | 1.532e-05
|  | 494.8 | 5.2153e-06 | 789.8 | 1.5345e-05
|  | 494.9 | 5.2167e-06 | 790 | 1.5371e-05
|  | 495 | 5.2181e-06 | 790.2 | 1.5398e-05
|  | 495.1 | 5.2195e-06 | 790.4 | 1.5424e-05
|  | 495.2 | 5.221e-06 | 790.6 | 1.5447e-05
|  | 495.3 | 5.2225e-06 | 790.8 | 1.547e-05
|  | 495.4 | 5.224e-06 | 791 | 1.5494e-05
|  | 495.5 | 5.2255e-06 | 791.2 | 1.5518e-05
|  | 495.6 | 5.227e-06 | 791.4 | 1.5544e-05
|  | 495.7 | 5.2285e-06 | 791.6 | 1.5574e-05
|  | 495.8 | 5.23e-06 | 791.8 | 1.5609e-05
|  | 495.9 | 5.2309e-06 | 792 | 1.5644e-05
|  | 496 | 5.2321e-06 | 792.2 | 1.5679e-05
|  | 496.1 | 5.2334e-06 | 792.4 | 1.571e-05
|  | 496.2 | 5.2345e-06 | 792.6 | 1.5739e-05
|  | 496.3 | 5.2357e-06 | 792.8 | 1.5767e-05
|  | 496.4 | 5.2368e-06 | 793 | 1.5795e-05
|  | 496.5 | 5.2383e-06 | 793.2 | 1.5824e-05
|  | 496.6 | 5.2395e-06 | 793.4 | 1.5853e-05
|  | 496.7 | 5.2408e-06 | 793.6 | 1.5881e-05
|  | 496.8 | 5.2423e-06 | 793.8 | 1.5911e-05
|  | 496.9 | 5.2437e-06 | 794 | 1.5942e-05
|  | 497 | 5.2451e-06 | 794.2 | 1.5974e-05
|  | 497.1 | 5.2465e-06 | 794.4 | 1.6006e-05
|  | 497.2 | 5.2478e-06 | 794.6 | 1.6037e-05
|  | 497.3 | 5.2492e-06 | 794.8 | 1.6066e-05
|  | 497.4 | 5.2508e-06 | 795 | 1.6092e-05
|  | 497.5 | 5.2522e-06 | 795.2 | 1.6116e-05
|  | 497.6 | 5.2536e-06 | 795.4 | 1.614e-05
|  | 497.7 | 5.2551e-06 | 795.6 | 1.6161e-05
|  | 497.8 | 5.2566e-06 | 795.8 | 1.6184e-05
|  | 497.9 | 5.2581e-06 | 796 | 1.621e-05
|  | 498 | 5.2591e-06 | 796.2 | 1.6239e-05
|  | 498.1 | 5.2604e-06 | 796.4 | 1.6272e-05
|  | 498.2 | 5.2617e-06 | 796.6 | 1.6305e-05
|  | 498.3 | 5.2629e-06 | 796.8 | 1.6335e-05
|  | 498.4 | 5.2641e-06 | 797 | 1.6361e-05
|  | 498.5 | 5.2653e-06 | 797.2 | 1.6385e-05
|  | 498.6 | 5.2664e-06 | 797.4 | 1.6409e-05
|  | 498.7 | 5.2676e-06 | 797.6 | 1.6433e-05
|  | 498.8 | 5.2688e-06 | 797.8 | 1.646e-05
|  | 498.9 | 5.27e-06 | 798 | 1.6491e-05
|  | 499 | 5.2711e-06 | 798.2 | 1.6524e-05
|  | 499.1 | 5.2722e-06 | 798.4 | 1.6559e-05
|  | 499.2 | 5.2739e-06 | 798.6 | 1.6595e-05
|  | 499.3 | 5.2753e-06 | 798.8 | 1.6631e-05
|  | 499.4 | 5.2767e-06 | 799 | 1.6663e-05
|  | 499.5 | 5.2778e-06 | 799.2 | 1.6692e-05
|  | 499.6 | 5.2792e-06 | 799.4 | 1.6721e-05
|  | 499.7 | 5.2806e-06 | 799.6 | 1.675e-05
|  | 499.8 | 5.2819e-06 | 799.8 | 1.6782e-05
|  | 499.9 | 5.2832e-06 | 800 | 1.6818e-05
|  | 500 | 5.2844e-06 | 800.2 | 1.6856e-05
|  | 500.1 | 5.2858e-06 | 800.4 | 1.6894e-05
|  | 500.2 | 5.2872e-06 | 800.6 | 1.6932e-05
|  | 500.3 | 5.2886e-06 | 800.8 | 1.697e-05
|  | 500.4 | 5.2895e-06 | 801 | 1.7004e-05
|  | 500.5 | 5.2906e-06 | 801.2 | 1.7034e-05
|  | 500.6 | 5.2918e-06 | 801.4 | 1.7062e-05
|  | 500.7 | 5.2928e-06 | 801.6 | 1.7094e-05
|  | 500.8 | 5.2938e-06 | 801.8 | 1.7132e-05
|  | 500.9 | 5.2948e-06 | 802 | 1.7174e-05
|  | 501 | 5.2961e-06 | 802.2 | 1.7218e-05
|  | 501.1 | 5.2973e-06 | 802.4 | 1.7261e-05
|  | 501.2 | 5.2985e-06 | 802.6 | 1.7303e-05
|  | 501.3 | 5.3e-06 | 802.8 | 1.7344e-05
|  | 501.4 | 5.3014e-06 | 803 | 1.7385e-05
|  | 501.5 | 5.3027e-06 | 803.2 | 1.7426e-05
|  | 501.6 | 5.3039e-06 | 803.4 | 1.7465e-05
|  | 501.7 | 5.3052e-06 | 803.6 | 1.7503e-05
|  | 501.8 | 5.3065e-06 | 803.8 | 1.754e-05
|  | 501.9 | 5.3077e-06 | 804 | 1.7574e-05
|  | 502 | 5.309e-06 | 804.2 | 1.7603e-05
|  | 502.1 | 5.3102e-06 | 804.4 | 1.7623e-05
|  | 502.2 | 5.3114e-06 | 804.6 | 1.7632e-05
|  | 502.3 | 5.3127e-06 | 804.8 | 1.7629e-05
|  | 502.4 | 5.3139e-06 | 805 | 1.7644e-05
|  | 502.5 | 5.3148e-06 | 805.2 | 1.7719e-05
|  | 502.6 | 5.3159e-06 | 805.4 | 1.7999e-05
|  | 502.7 | 5.317e-06 | 805.6 | 1.8658e-05
|  | 502.8 | 5.3184e-06 | 805.8 | 1.9833e-05
|  | 502.9 | 5.3196e-06 | 806 | 2.1577e-05
|  | 503 | 5.3207e-06 | 806.2 | 2.3949e-05
|  | 503.1 | 5.3215e-06 | 806.4 | 2.6863e-05
|  | 503.2 | 5.3226e-06 | 806.6 | 2.987e-05
|  | 503.3 | 5.3237e-06 | 806.8 | 3.2513e-05
|  | 503.4 | 5.3247e-06 | 807 | 3.558e-05
|  | 503.5 | 5.3257e-06 | 807.2 | 3.9565e-05
|  | 503.6 | 5.3266e-06 | 807.4 | 4.635e-05
|  | 503.7 | 5.3283e-06 | 807.6 | 6.0527e-05
|  | 503.8 | 5.3296e-06
|  | 503.9 | 5.3309e-06
|  | 504 | 5.332e-06
|  | 504.1 | 5.3334e-06
|  | 504.2 | 5.3348e-06
|  | 504.3 | 5.3358e-06
|  | 504.4 | 5.3368e-06
|  | 504.5 | 5.3379e-06
|  | 504.6 | 5.3387e-06
|  | 504.7 | 5.3397e-06
|  | 504.8 | 5.3406e-06
|  | 504.9 | 5.342e-06
|  | 505 | 5.3431e-06
|  | 505.1 | 5.3442e-06
|  | 505.2 | 5.3454e-06
|  | 505.3 | 5.3467e-06
|  | 505.4 | 5.3479e-06
|  | 505.5 | 5.349e-06
|  | 505.6 | 5.3502e-06
|  | 505.7 | 5.3513e-06
|  | 505.8 | 5.3527e-06
|  | 505.9 | 5.354e-06
|  | 506 | 5.3553e-06
|  | 506.1 | 5.3566e-06
|  | 506.2 | 5.3577e-06
|  | 506.3 | 5.359e-06
|  | 506.4 | 5.3603e-06
|  | 506.5 | 5.3617e-06
|  | 506.6 | 5.3629e-06
|  | 506.7 | 5.3642e-06
|  | 506.8 | 5.3656e-06
|  | 506.9 | 5.367e-06
|  | 507 | 5.3683e-06
|  | 507.1 | 5.3691e-06
|  | 507.2 | 5.3703e-06
|  | 507.3 | 5.3715e-06
|  | 507.4 | 5.3726e-06
|  | 507.5 | 5.3737e-06
|  | 507.6 | 5.3748e-06
|  | 507.7 | 5.3759e-06
|  | 507.8 | 5.377e-06
|  | 507.9 | 5.3781e-06
|  | 508 | 5.3798e-06
|  | 508.1 | 5.3811e-06
|  | 508.2 | 5.3824e-06
|  | 508.3 | 5.3834e-06
|  | 508.4 | 5.3847e-06
|  | 508.5 | 5.386e-06
|  | 508.6 | 5.3867e-06
|  | 508.7 | 5.3877e-06
|  | 508.8 | 5.3887e-06
|  | 508.9 | 5.3901e-06
|  | 509 | 5.3912e-06
|  | 509.1 | 5.3924e-06
|  | 509.2 | 5.3933e-06
|  | 509.3 | 5.3945e-06
|  | 509.4 | 5.3956e-06
|  | 509.5 | 5.3965e-06
|  | 509.6 | 5.3975e-06
|  | 509.7 | 5.3985e-06
|  | 509.8 | 5.3999e-06
|  | 509.9 | 5.401e-06
|  | 510 | 5.4021e-06
|  | 510.1 | 5.403e-06
|  | 510.2 | 5.4041e-06
|  | 510.3 | 5.4052e-06
|  | 510.4 | 5.4059e-06
|  | 510.5 | 5.4069e-06
|  | 510.6 | 5.4078e-06
|  | 510.7 | 5.409e-06
|  | 510.8 | 5.41e-06
|  | 510.9 | 5.4111e-06
|  | 511 | 5.4124e-06
|  | 511.1 | 5.4135e-06
|  | 511.2 | 5.4145e-06
|  | 511.3 | 5.4156e-06
|  | 511.4 | 5.4166e-06
|  | 511.5 | 5.4177e-06
|  | 511.6 | 5.4188e-06
|  | 511.7 | 5.4198e-06
|  | 511.8 | 5.4209e-06
|  | 511.9 | 5.422e-06
|  | 512 | 5.4231e-06
|  | 512.1 | 5.424e-06
|  | 512.2 | 5.4251e-06
|  | 512.3 | 5.4261e-06
|  | 512.4 | 5.4271e-06
|  | 512.5 | 5.4281e-06
|  | 512.6 | 5.4291e-06
|  | 512.7 | 5.4303e-06
|  | 512.8 | 5.4314e-06
|  | 512.9 | 5.4325e-06
|  | 513 | 5.4334e-06
|  | 513.1 | 5.4345e-06
|  | 513.2 | 5.4355e-06
|  | 513.3 | 5.4364e-06
|  | 513.4 | 5.4373e-06
|  | 513.5 | 5.4382e-06
|  | 513.6 | 5.4395e-06
|  | 513.7 | 5.4406e-06
|  | 513.8 | 5.4417e-06
|  | 513.9 | 5.4428e-06
|  | 514 | 5.4439e-06
|  | 514.1 | 5.4451e-06
|  | 514.2 | 5.4463e-06
|  | 514.3 | 5.4474e-06
|  | 514.4 | 5.4485e-06
|  | 514.5 | 5.4493e-06
|  | 514.6 | 5.4503e-06
|  | 514.7 | 5.4514e-06
|  | 514.8 | 5.4522e-06
|  | 514.9 | 5.453e-06
|  | 515 | 5.4538e-06
|  | 515.1 | 5.455e-06
|  | 515.2 | 5.456e-06
|  | 515.3 | 5.4569e-06
|  | 515.4 | 5.458e-06
|  | 515.5 | 5.4591e-06
|  | 515.6 | 5.4602e-06
|  | 515.7 | 5.4613e-06
|  | 515.8 | 5.4623e-06
|  | 515.9 | 5.4634e-06
|  | 516 | 5.4642e-06
|  | 516.1 | 5.4652e-06
|  | 516.2 | 5.4662e-06
|  | 516.3 | 5.4673e-06
|  | 516.4 | 5.4683e-06
|  | 516.5 | 5.4693e-06
|  | 516.6 | 5.4707e-06
|  | 516.7 | 5.4719e-06
|  | 516.8 | 5.4732e-06
|  | 516.9 | 5.4741e-06
|  | 517 | 5.4752e-06
|  | 517.1 | 5.4764e-06
|  | 517.2 | 5.4776e-06
|  | 517.3 | 5.4786e-06
|  | 517.4 | 5.4797e-06
|  | 517.5 | 5.4806e-06
|  | 517.6 | 5.4817e-06
|  | 517.7 | 5.4827e-06
|  | 517.8 | 5.4839e-06
|  | 517.9 | 5.485e-06
|  | 518 | 5.486e-06
|  | 518.1 | 5.4871e-06
|  | 518.2 | 5.4883e-06
|  | 518.3 | 5.4894e-06
|  | 518.4 | 5.4903e-06
|  | 518.5 | 5.4913e-06
|  | 518.6 | 5.4923e-06
|  | 518.7 | 5.4931e-06
|  | 518.8 | 5.494e-06
|  | 518.9 | 5.4948e-06
|  | 519 | 5.4959e-06
|  | 519.1 | 5.4968e-06
|  | 519.2 | 5.4978e-06
|  | 519.3 | 5.499e-06
|  | 519.4 | 5.5001e-06
|  | 519.5 | 5.5012e-06
|  | 519.6 | 5.5019e-06
|  | 519.7 | 5.5029e-06
|  | 519.8 | 5.5039e-06
|  | 519.9 | 5.5047e-06
|  | 520 | 5.5055e-06
|  | 520.1 | 5.5063e-06
|  | 520.2 | 5.5072e-06
|  | 520.3 | 5.5081e-06
|  | 520.4 | 5.509e-06
|  | 520.5 | 5.5095e-06
|  | 520.6 | 5.5103e-06
|  | 520.7 | 5.511e-06
|  | 520.8 | 5.5122e-06
|  | 520.9 | 5.513e-06
|  | 521 | 5.5138e-06
|  | 521.1 | 5.5149e-06
|  | 521.2 | 5.5159e-06
|  | 521.3 | 5.5169e-06
|  | 521.4 | 5.518e-06
|  | 521.5 | 5.519e-06
|  | 521.6 | 5.52e-06
|  | 521.7 | 5.5213e-06
|  | 521.8 | 5.5224e-06
|  | 521.9 | 5.5235e-06
|  | 522 | 5.5248e-06
|  | 522.1 | 5.526e-06
|  | 522.2 | 5.5271e-06
|  | 522.3 | 5.5286e-06
|  | 522.4 | 5.5298e-06
|  | 522.5 | 5.5311e-06
|  | 522.6 | 5.5321e-06
|  | 522.7 | 5.5333e-06
|  | 522.8 | 5.5345e-06
|  | 522.9 | 5.5357e-06
|  | 523 | 5.5369e-06
|  | 523.1 | 5.538e-06
|  | 523.2 | 5.5395e-06
|  | 523.3 | 5.5409e-06
|  | 523.4 | 5.5422e-06
|  | 523.5 | 5.5433e-06
|  | 523.6 | 5.5445e-06
|  | 523.7 | 5.5458e-06
|  | 523.8 | 5.5474e-06
|  | 523.9 | 5.5487e-06
|  | 524 | 5.55e-06
|  | 524.1 | 5.5514e-06
|  | 524.2 | 5.5529e-06
|  | 524.3 | 5.5543e-06
|  | 524.4 | 5.5556e-06
|  | 524.5 | 5.5569e-06
|  | 524.6 | 5.5582e-06
|  | 524.7 | 5.5594e-06
|  | 524.8 | 5.5607e-06
|  | 524.9 | 5.5618e-06
|  | 525 | 5.563e-06
|  | 525.1 | 5.5642e-06
|  | 525.2 | 5.5653e-06
|  | 525.3 | 5.5664e-06
|  | 525.4 | 5.5675e-06
|  | 525.5 | 5.5686e-06
|  | 525.6 | 5.5697e-06
|  | 525.7 | 5.5708e-06
|  | 525.8 | 5.5722e-06
|  | 525.9 | 5.5734e-06
|  | 526 | 5.5747e-06
|  | 526.1 | 5.5759e-06
|  | 526.2 | 5.5772e-06
|  | 526.3 | 5.5784e-06
|  | 526.4 | 5.5796e-06
|  | 526.5 | 5.5807e-06
|  | 526.6 | 5.5819e-06
|  | 526.7 | 5.5834e-06
|  | 526.8 | 5.5848e-06
|  | 526.9 | 5.5861e-06
|  | 527 | 5.5871e-06
|  | 527.1 | 5.5883e-06
|  | 527.2 | 5.5896e-06
|  | 527.3 | 5.5906e-06
|  | 527.4 | 5.5916e-06
|  | 527.5 | 5.5927e-06
|  | 527.6 | 5.5941e-06
|  | 527.7 | 5.5953e-06
|  | 527.8 | 5.5965e-06
|  | 527.9 | 5.5979e-06
|  | 528 | 5.5992e-06
|  | 528.1 | 5.6005e-06
|  | 528.2 | 5.6014e-06
|  | 528.3 | 5.6026e-06
|  | 528.4 | 5.6037e-06
|  | 528.5 | 5.6049e-06
|  | 528.6 | 5.6059e-06
|  | 528.7 | 5.607e-06
|  | 528.8 | 5.608e-06
|  | 528.9 | 5.6091e-06
|  | 529 | 5.6102e-06
|  | 529.1 | 5.6107e-06
|  | 529.2 | 5.6116e-06
|  | 529.3 | 5.6124e-06
|  | 529.4 | 5.6136e-06
|  | 529.5 | 5.6145e-06
|  | 529.6 | 5.6155e-06
|  | 529.7 | 5.6167e-06
|  | 529.8 | 5.6178e-06
|  | 529.9 | 5.6189e-06
|  | 530 | 5.6204e-06
|  | 530.1 | 5.6216e-06
|  | 530.2 | 5.6229e-06
|  | 530.3 | 5.6242e-06
|  | 530.4 | 5.6255e-06
|  | 530.5 | 5.6269e-06
|  | 530.6 | 5.6281e-06
|  | 530.7 | 5.6294e-06
|  | 530.8 | 5.6307e-06
|  | 530.9 | 5.6321e-06
|  | 531 | 5.6335e-06
|  | 531.1 | 5.6348e-06
|  | 531.2 | 5.6359e-06
|  | 531.3 | 5.6372e-06
|  | 531.4 | 5.6387e-06
|  | 531.5 | 5.64e-06
|  | 531.6 | 5.6413e-06
|  | 531.7 | 5.6427e-06
|  | 531.8 | 5.644e-06
|  | 531.9 | 5.6454e-06
|  | 532 | 5.6467e-06
|  | 532.1 | 5.648e-06
|  | 532.2 | 5.6493e-06
|  | 532.3 | 5.651e-06
|  | 532.4 | 5.6525e-06
|  | 532.5 | 5.654e-06
|  | 532.6 | 5.6553e-06
|  | 532.7 | 5.6568e-06
|  | 532.8 | 5.6582e-06
|  | 532.9 | 5.6598e-06
|  | 533 | 5.6612e-06
|  | 533.1 | 5.6626e-06
|  | 533.2 | 5.6639e-06
|  | 533.3 | 5.6653e-06
|  | 533.4 | 5.6667e-06
|  | 533.5 | 5.6681e-06
|  | 533.6 | 5.6695e-06
|  | 533.7 | 5.6708e-06
|  | 533.8 | 5.6723e-06
|  | 533.9 | 5.6737e-06
|  | 534 | 5.6752e-06
|  | 534.1 | 5.6765e-06
|  | 534.2 | 5.6779e-06
|  | 534.3 | 5.6793e-06
|  | 534.4 | 5.6809e-06
|  | 534.5 | 5.6823e-06
|  | 534.6 | 5.6837e-06
|  | 534.7 | 5.6849e-06
|  | 534.8 | 5.6863e-06
|  | 534.9 | 5.6877e-06
|  | 535 | 5.6887e-06
|  | 535.1 | 5.6899e-06
|  | 535.2 | 5.6911e-06
|  | 535.3 | 5.6925e-06
|  | 535.4 | 5.6938e-06
|  | 535.5 | 5.6951e-06
|  | 535.6 | 5.6965e-06
|  | 535.7 | 5.6979e-06
|  | 535.8 | 5.6992e-06
|  | 535.9 | 5.7007e-06
|  | 536 | 5.7021e-06
|  | 536.1 | 5.7035e-06
|  | 536.2 | 5.7046e-06
|  | 536.3 | 5.7058e-06
|  | 536.4 | 5.7075e-06
|  | 536.5 | 5.7089e-06
|  | 536.6 | 5.7102e-06
|  | 536.7 | 5.7117e-06
|  | 536.8 | 5.7132e-06
|  | 536.9 | 5.7148e-06
|  | 537 | 5.7165e-06
|  | 537.1 | 5.718e-06
|  | 537.2 | 5.7196e-06
|  | 537.3 | 5.7211e-06
|  | 537.4 | 5.7227e-06
|  | 537.5 | 5.7243e-06
|  | 537.6 | 5.7255e-06
|  | 537.7 | 5.7269e-06
|  | 537.8 | 5.7282e-06
|  | 537.9 | 5.7299e-06
|  | 538 | 5.7314e-06
|  | 538.1 | 5.7328e-06
|  | 538.2 | 5.7344e-06
|  | 538.3 | 5.736e-06
|  | 538.4 | 5.7376e-06
|  | 538.5 | 5.7392e-06
|  | 538.6 | 5.7408e-06
|  | 538.7 | 5.7423e-06
|  | 538.8 | 5.744e-06
|  | 538.9 | 5.7456e-06
|  | 539 | 5.7473e-06
|  | 539.1 | 5.749e-06
|  | 539.2 | 5.7507e-06
|  | 539.3 | 5.7524e-06
|  | 539.4 | 5.7543e-06
|  | 539.5 | 5.7561e-06
|  | 539.6 | 5.7579e-06
|  | 539.7 | 5.7588e-06
|  | 539.8 | 5.7603e-06
|  | 539.9 | 5.7618e-06
|  | 540 | 5.7634e-06
|  | 540.1 | 5.7648e-06
|  | 540.2 | 5.7662e-06
|  | 540.3 | 5.7675e-06
|  | 540.4 | 5.7689e-06
|  | 540.5 | 5.7704e-06
|  | 540.6 | 5.7718e-06
|  | 540.7 | 5.7732e-06
|  | 540.8 | 5.7751e-06
|  | 540.9 | 5.7769e-06
|  | 541 | 5.7786e-06
|  | 541.1 | 5.7803e-06
|  | 541.2 | 5.7821e-06
|  | 541.3 | 5.7839e-06
|  | 541.4 | 5.7856e-06
|  | 541.5 | 5.7873e-06
|  | 541.6 | 5.789e-06
|  | 541.7 | 5.7906e-06
|  | 541.8 | 5.7923e-06
|  | 541.9 | 5.794e-06
|  | 542 | 5.7954e-06
|  | 542.1 | 5.7969e-06
|  | 542.2 | 5.7985e-06
|  | 542.3 | 5.8e-06
|  | 542.4 | 5.8015e-06
|  | 542.5 | 5.803e-06
|  | 542.6 | 5.8045e-06
|  | 542.7 | 5.806e-06
|  | 542.8 | 5.8075e-06
|  | 542.9 | 5.8091e-06
|  | 543 | 5.8106e-06
|  | 543.1 | 5.8122e-06
|  | 543.2 | 5.8138e-06
|  | 543.3 | 5.8154e-06
|  | 543.4 | 5.817e-06
|  | 543.5 | 5.819e-06
|  | 543.6 | 5.8207e-06
|  | 543.7 | 5.8225e-06
|  | 543.8 | 5.824e-06
|  | 543.9 | 5.8257e-06
|  | 544 | 5.8274e-06
|  | 544.1 | 5.8282e-06
|  | 544.2 | 5.8295e-06
|  | 544.3 | 5.8311e-06
|  | 544.4 | 5.8324e-06
|  | 544.5 | 5.8338e-06
|  | 544.6 | 5.8354e-06
|  | 544.7 | 5.8369e-06
|  | 544.8 | 5.8385e-06
|  | 544.9 | 5.8401e-06
|  | 545 | 5.8417e-06
|  | 545.1 | 5.8433e-06
|  | 545.2 | 5.8451e-06
|  | 545.3 | 5.8469e-06
|  | 545.4 | 5.8486e-06
|  | 545.5 | 5.8502e-06
|  | 545.6 | 5.852e-06
|  | 545.7 | 5.8537e-06
|  | 545.8 | 5.8553e-06
|  | 545.9 | 5.8569e-06
|  | 546 | 5.8585e-06
|  | 546.1 | 5.8605e-06
|  | 546.2 | 5.8623e-06
|  | 546.3 | 5.8641e-06
|  | 546.4 | 5.866e-06
|  | 546.5 | 5.8678e-06
|  | 546.6 | 5.8697e-06
|  | 546.7 | 5.8709e-06
|  | 546.8 | 5.8725e-06
|  | 546.9 | 5.8741e-06
|  | 547 | 5.8758e-06
|  | 547.1 | 5.8774e-06
|  | 547.2 | 5.879e-06
|  | 547.3 | 5.8813e-06
|  | 547.4 | 5.8832e-06
|  | 547.5 | 5.8851e-06
|  | 547.6 | 5.8866e-06
|  | 547.7 | 5.8884e-06
|  | 547.8 | 5.8905e-06
|  | 547.9 | 5.8923e-06
|  | 548 | 5.8941e-06
|  | 548.1 | 5.896e-06
|  | 548.2 | 5.8979e-06
|  | 548.3 | 5.8999e-06
|  | 548.4 | 5.9015e-06
|  | 548.5 | 5.9033e-06
|  | 548.6 | 5.9051e-06
|  | 548.7 | 5.9067e-06
|  | 548.8 | 5.9083e-06
|  | 548.9 | 5.91e-06
|  | 549 | 5.9118e-06
|  | 549.1 | 5.9135e-06
|  | 549.2 | 5.9152e-06
|  | 549.3 | 5.917e-06
|  | 549.4 | 5.9187e-06
|  | 549.5 | 5.9205e-06
|  | 549.6 | 5.9219e-06
|  | 549.7 | 5.9235e-06
|  | 549.8 | 5.9252e-06
|  | 549.9 | 5.9272e-06
|  | 550 | 5.929e-06
|  | 550.1 | 5.9307e-06
|  | 550.2 | 5.9329e-06
|  | 550.3 | 5.9348e-06
|  | 550.4 | 5.9368e-06
|  | 550.5 | 5.9386e-06
|  | 550.6 | 5.9406e-06
|  | 550.7 | 5.9424e-06
|  | 550.8 | 5.9442e-06
|  | 550.9 | 5.946e-06
|  | 551 | 5.9481e-06
|  | 551.1 | 5.95e-06
|  | 551.2 | 5.952e-06
|  | 551.3 | 5.9536e-06
|  | 551.4 | 5.9555e-06
|  | 551.5 | 5.9573e-06
|  | 551.6 | 5.9591e-06
|  | 551.7 | 5.9608e-06
|  | 551.8 | 5.9625e-06
|  | 551.9 | 5.9647e-06
|  | 552 | 5.9666e-06
|  | 552.1 | 5.9686e-06
|  | 552.2 | 5.9706e-06
|  | 552.3 | 5.9727e-06
|  | 552.4 | 5.9747e-06
|  | 552.5 | 5.9763e-06
|  | 552.6 | 5.9782e-06
|  | 552.7 | 5.98e-06
|  | 552.8 | 5.9818e-06
|  | 552.9 | 5.9835e-06
|  | 553 | 5.9853e-06
|  | 553.1 | 5.9872e-06
|  | 553.2 | 5.9891e-06
|  | 553.3 | 5.9909e-06
|  | 553.4 | 5.9929e-06
|  | 553.5 | 5.9948e-06
|  | 553.6 | 5.997e-06
|  | 553.7 | 5.9991e-06
|  | 553.8 | 6.0012e-06
|  | 553.9 | 6.0035e-06
|  | 554 | 6.0058e-06
|  | 554.1 | 6.008e-06
|  | 554.2 | 6.0097e-06
|  | 554.3 | 6.0117e-06
|  | 554.4 | 6.0137e-06
|  | 554.5 | 6.0158e-06
|  | 554.6 | 6.0177e-06
|  | 554.7 | 6.0196e-06
|  | 554.8 | 6.0218e-06
|  | 554.9 | 6.0239e-06
|  | 555 | 6.026e-06
|  | 555.1 | 6.0278e-06
|  | 555.2 | 6.0297e-06
|  | 555.3 | 6.0316e-06
|  | 555.4 | 6.0335e-06
|  | 555.5 | 6.0353e-06
|  | 555.6 | 6.0372e-06
|  | 555.7 | 6.0394e-06
|  | 555.8 | 6.0414e-06
|  | 555.9 | 6.0434e-06
|  | 556 | 6.0457e-06
|  | 556.1 | 6.0478e-06
|  | 556.2 | 6.0497e-06
|  | 556.3 | 6.0517e-06
|  | 556.4 | 6.0538e-06
|  | 556.5 | 6.0556e-06
|  | 556.6 | 6.0575e-06
|  | 556.7 | 6.0593e-06
|  | 556.8 | 6.0607e-06
|  | 556.9 | 6.0624e-06
|  | 557 | 6.064e-06
|  | 557.1 | 6.0659e-06
|  | 557.2 | 6.0676e-06
|  | 557.3 | 6.0692e-06
|  | 557.4 | 6.0717e-06
|  | 557.5 | 6.0737e-06
|  | 557.6 | 6.0758e-06
|  | 557.7 | 6.0782e-06
|  | 557.8 | 6.0805e-06
|  | 557.9 | 6.0828e-06
|  | 558 | 6.085e-06
|  | 558.1 | 6.0873e-06
|  | 558.2 | 6.0896e-06
|  | 558.3 | 6.0917e-06
|  | 558.4 | 6.094e-06
|  | 558.5 | 6.0962e-06
|  | 558.6 | 6.098e-06
|  | 558.7 | 6.1e-06
|  | 558.8 | 6.1019e-06
|  | 558.9 | 6.1038e-06
|  | 559 | 6.1056e-06
|  | 559.1 | 6.1076e-06
|  | 559.2 | 6.1095e-06
|  | 559.3 | 6.1114e-06
|  | 559.4 | 6.1138e-06
|  | 559.5 | 6.116e-06
|  | 559.6 | 6.1181e-06
|  | 559.7 | 6.1202e-06
|  | 559.8 | 6.1225e-06
|  | 559.9 | 6.1247e-06
|  | 560 | 6.1269e-06
|  | 560.1 | 6.129e-06
|  | 560.2 | 6.1312e-06
|  | 560.3 | 6.1335e-06
|  | 560.4 | 6.1358e-06
|  | 560.5 | 6.138e-06
|  | 560.6 | 6.1399e-06
|  | 560.7 | 6.1421e-06
|  | 560.8 | 6.1442e-06
|  | 560.9 | 6.1461e-06
|  | 561 | 6.1481e-06
|  | 561.1 | 6.15e-06
|  | 561.2 | 6.1521e-06
|  | 561.3 | 6.1541e-06
|  | 561.4 | 6.1561e-06
|  | 561.5 | 6.1581e-06
|  | 561.6 | 6.1602e-06
|  | 561.7 | 6.1625e-06
|  | 561.8 | 6.1647e-06
|  | 561.9 | 6.1669e-06
|  | 562 | 6.1693e-06
|  | 562.1 | 6.1717e-06
|  | 562.2 | 6.1741e-06
|  | 562.3 | 6.1762e-06
|  | 562.4 | 6.1785e-06
|  | 562.5 | 6.1808e-06
|  | 562.6 | 6.1825e-06
|  | 562.7 | 6.1846e-06
|  | 562.8 | 6.1866e-06
|  | 562.9 | 6.1886e-06
|  | 563 | 6.1906e-06
|  | 563.1 | 6.1925e-06
|  | 563.2 | 6.1947e-06
|  | 563.3 | 6.1967e-06
|  | 563.4 | 6.1988e-06
|  | 563.5 | 6.2007e-06
|  | 563.6 | 6.2027e-06
|  | 563.7 | 6.2048e-06
|  | 563.8 | 6.2069e-06
|  | 563.9 | 6.2089e-06
|  | 564 | 6.2112e-06
|  | 564.1 | 6.2134e-06
|  | 564.2 | 6.2156e-06
|  | 564.3 | 6.2174e-06
|  | 564.4 | 6.2195e-06
|  | 564.5 | 6.2215e-06
|  | 564.6 | 6.2236e-06
|  | 564.7 | 6.2256e-06
|  | 564.8 | 6.2275e-06
|  | 564.9 | 6.2297e-06
|  | 565 | 6.2317e-06
|  | 565.1 | 6.2338e-06
|  | 565.2 | 6.2356e-06
|  | 565.3 | 6.2376e-06
|  | 565.4 | 6.2396e-06
|  | 565.5 | 6.2416e-06
|  | 565.6 | 6.2435e-06
|  | 565.7 | 6.2452e-06
|  | 565.8 | 6.247e-06
|  | 565.9 | 6.2489e-06
|  | 566 | 6.2509e-06
|  | 566.1 | 6.2528e-06
|  | 566.2 | 6.2547e-06
|  | 566.3 | 6.2567e-06
|  | 566.4 | 6.2588e-06
|  | 566.5 | 6.2608e-06
|  | 566.6 | 6.2628e-06
|  | 566.7 | 6.2649e-06
|  | 566.8 | 6.2669e-06
|  | 566.9 | 6.2687e-06
|  | 567 | 6.2706e-06
|  | 567.1 | 6.2725e-06
|  | 567.2 | 6.2741e-06
|  | 567.3 | 6.2758e-06
|  | 567.4 | 6.2776e-06
|  | 567.5 | 6.28e-06
|  | 567.6 | 6.282e-06
|  | 567.7 | 6.284e-06
|  | 567.8 | 6.2863e-06
|  | 567.9 | 6.2885e-06
|  | 568 | 6.2908e-06
|  | 568.1 | 6.293e-06
|  | 568.2 | 6.2952e-06
|  | 568.3 | 6.2972e-06
|  | 568.4 | 6.2993e-06
|  | 568.5 | 6.3014e-06
|  | 568.6 | 6.3037e-06
|  | 568.7 | 6.3058e-06
|  | 568.8 | 6.308e-06
|  | 568.9 | 6.3105e-06
|  | 569 | 6.3128e-06
|  | 569.1 | 6.3152e-06
|  | 569.2 | 6.3175e-06
|  | 569.3 | 6.3198e-06
|  | 569.4 | 6.3222e-06
|  | 569.5 | 6.3247e-06
|  | 569.6 | 6.3271e-06
|  | 569.7 | 6.3295e-06
|  | 569.8 | 6.3317e-06
|  | 569.9 | 6.3341e-06
|  | 570 | 6.3363e-06
|  | 570.1 | 6.3386e-06
|  | 570.2 | 6.3408e-06
|  | 570.3 | 6.3434e-06
|  | 570.4 | 6.3458e-06
|  | 570.5 | 6.3482e-06
|  | 570.6 | 6.3505e-06
|  | 570.7 | 6.3529e-06
|  | 570.8 | 6.3553e-06
|  | 570.9 | 6.3578e-06
|  | 571 | 6.3602e-06
|  | 571.1 | 6.3626e-06
|  | 571.2 | 6.3646e-06
|  | 571.3 | 6.3669e-06
|  | 571.4 | 6.3693e-06
|  | 571.5 | 6.3716e-06
|  | 571.6 | 6.3738e-06
|  | 571.7 | 6.376e-06
|  | 571.8 | 6.3782e-06
|  | 571.9 | 6.3805e-06
|  | 572 | 6.3829e-06
|  | 572.1 | 6.3851e-06
|  | 572.2 | 6.3874e-06
|  | 572.3 | 6.3896e-06
|  | 572.4 | 6.392e-06
|  | 572.5 | 6.3943e-06
|  | 572.6 | 6.3964e-06
|  | 572.7 | 6.3986e-06
|  | 572.8 | 6.4008e-06
|  | 572.9 | 6.4028e-06
|  | 573 | 6.4049e-06
|  | 573.1 | 6.407e-06
|  | 573.2 | 6.4095e-06
|  | 573.3 | 6.4117e-06
|  | 573.4 | 6.414e-06
|  | 573.5 | 6.4161e-06
|  | 573.6 | 6.4184e-06
|  | 573.7 | 6.4205e-06
|  | 573.8 | 6.4226e-06
|  | 573.9 | 6.4247e-06
|  | 574 | 6.4269e-06
|  | 574.1 | 6.429e-06
|  | 574.2 | 6.4312e-06
|  | 574.3 | 6.4336e-06
|  | 574.4 | 6.4359e-06
|  | 574.5 | 6.4382e-06
|  | 574.6 | 6.4407e-06
|  | 574.7 | 6.4431e-06
|  | 574.8 | 6.4455e-06
|  | 574.9 | 6.4481e-06
|  | 575 | 6.4505e-06
|  | 575.1 | 6.453e-06
|  | 575.2 | 6.4553e-06
|  | 575.3 | 6.4577e-06
|  | 575.4 | 6.4602e-06
|  | 575.5 | 6.4625e-06
|  | 575.6 | 6.4648e-06
|  | 575.7 | 6.467e-06
|  | 575.8 | 6.4692e-06
|  | 575.9 | 6.4715e-06
|  | 576 | 6.4739e-06
|  | 576.1 | 6.4761e-06
|  | 576.2 | 6.4784e-06
|  | 576.3 | 6.4808e-06
|  | 576.4 | 6.4831e-06
|  | 576.5 | 6.4855e-06
|  | 576.6 | 6.4879e-06
|  | 576.7 | 6.4903e-06
|  | 576.8 | 6.4927e-06
|  | 576.9 | 6.4954e-06
|  | 577 | 6.4979e-06
|  | 577.1 | 6.5004e-06
|  | 577.2 | 6.5024e-06
|  | 577.3 | 6.5048e-06
|  | 577.4 | 6.5074e-06
|  | 577.5 | 6.5097e-06
|  | 577.6 | 6.5121e-06
|  | 577.7 | 6.5146e-06
|  | 577.8 | 6.5171e-06
|  | 577.9 | 6.5196e-06
|  | 578 | 6.5219e-06
|  | 578.1 | 6.5244e-06
|  | 578.2 | 6.5268e-06
|  | 578.3 | 6.5291e-06
|  | 578.4 | 6.5315e-06
|  | 578.5 | 6.5339e-06
|  | 578.6 | 6.5366e-06
|  | 578.7 | 6.5391e-06
|  | 578.8 | 6.5416e-06
|  | 578.9 | 6.5444e-06
|  | 579 | 6.547e-06
|  | 579.1 | 6.5494e-06
|  | 579.2 | 6.5519e-06
|  | 579.3 | 6.5544e-06
|  | 579.4 | 6.5567e-06
|  | 579.5 | 6.559e-06
|  | 579.6 | 6.5613e-06
|  | 579.7 | 6.564e-06
|  | 579.8 | 6.5664e-06
|  | 579.9 | 6.5688e-06
|  | 580 | 6.571e-06
|  | 580.1 | 6.5735e-06
|  | 580.2 | 6.5759e-06
|  | 580.3 | 6.5786e-06
|  | 580.4 | 6.581e-06
|  | 580.5 | 6.5835e-06
|  | 580.6 | 6.5859e-06
|  | 580.7 | 6.5885e-06
|  | 580.8 | 6.5913e-06
|  | 580.9 | 6.5938e-06
|  | 581 | 6.5964e-06
|  | 581.1 | 6.5988e-06
|  | 581.2 | 6.6014e-06
|  | 581.3 | 6.604e-06
|  | 581.4 | 6.6068e-06
|  | 581.5 | 6.6094e-06
|  | 581.6 | 6.612e-06
|  | 581.7 | 6.6144e-06
|  | 581.8 | 6.617e-06
|  | 581.9 | 6.6196e-06
|  | 582 | 6.6222e-06
|  | 582.1 | 6.6248e-06
|  | 582.2 | 6.6273e-06
|  | 582.3 | 6.6309e-06
|  | 582.4 | 6.6339e-06
|  | 582.5 | 6.637e-06
|  | 582.6 | 6.6402e-06
|  | 582.7 | 6.6433e-06
|  | 582.8 | 6.6461e-06
|  | 582.9 | 6.649e-06
|  | 583 | 6.652e-06
|  | 583.1 | 6.6546e-06
|  | 583.2 | 6.6573e-06
|  | 583.3 | 6.6601e-06
|  | 583.4 | 6.6629e-06
|  | 583.5 | 6.6656e-06
|  | 583.6 | 6.6684e-06
|  | 583.7 | 6.6715e-06
|  | 583.8 | 6.6745e-06
|  | 583.9 | 6.6774e-06
|  | 584 | 6.6801e-06
|  | 584.1 | 6.683e-06
|  | 584.2 | 6.686e-06
|  | 584.3 | 6.6888e-06
|  | 584.4 | 6.6917e-06
|  | 584.5 | 6.6946e-06
|  | 584.6 | 6.6975e-06
|  | 584.7 | 6.7005e-06
|  | 584.8 | 6.7037e-06
|  | 584.9 | 6.7068e-06
|  | 585 | 6.7099e-06
|  | 585.1 | 6.7132e-06
|  | 585.2 | 6.7164e-06
|  | 585.3 | 6.7196e-06
|  | 585.4 | 6.7227e-06
|  | 585.5 | 6.7259e-06
|  | 585.6 | 6.7291e-06
|  | 585.7 | 6.7322e-06
|  | 585.8 | 6.7353e-06
|  | 585.9 | 6.7389e-06
|  | 586 | 6.7423e-06
|  | 586.1 | 6.7456e-06
|  | 586.2 | 6.7496e-06
|  | 586.3 | 6.7533e-06
|  | 586.4 | 6.7571e-06
|  | 586.5 | 6.761e-06
|  | 586.6 | 6.7648e-06
|  | 586.7 | 6.7687e-06
|  | 586.8 | 6.772e-06
|  | 586.9 | 6.7757e-06
|  | 587 | 6.7794e-06
|  | 587.1 | 6.7831e-06
|  | 587.2 | 6.7868e-06
|  | 587.3 | 6.7906e-06
|  | 587.4 | 6.7944e-06
|  | 587.5 | 6.7981e-06
|  | 587.6 | 6.8019e-06
|  | 587.7 | 6.8057e-06
|  | 587.8 | 6.8095e-06
|  | 587.9 | 6.8137e-06
|  | 588 | 6.8176e-06
|  | 588.1 | 6.8216e-06
|  | 588.2 | 6.8256e-06
|  | 588.3 | 6.8297e-06
|  | 588.4 | 6.8337e-06
|  | 588.5 | 6.8376e-06
|  | 588.6 | 6.8415e-06
|  | 588.7 | 6.8455e-06
|  | 588.8 | 6.8498e-06
|  | 588.9 | 6.8538e-06
|  | 589 | 6.8578e-06
|  | 589.1 | 6.8619e-06
|  | 589.2 | 6.866e-06
|  | 589.3 | 6.8699e-06
|  | 589.4 | 6.8739e-06
|  | 589.5 | 6.8779e-06
|  | 589.6 | 6.8818e-06
|  | 589.7 | 6.8858e-06
|  | 589.8 | 6.8898e-06
|  | 589.9 | 6.8941e-06
|  | 590 | 6.8982e-06
|  | 590.1 | 6.9023e-06
|  | 590.2 | 6.9062e-06
|  | 590.3 | 6.9102e-06
|  | 590.4 | 6.9145e-06
|  | 590.5 | 6.9186e-06
|  | 590.6 | 6.9226e-06
|  | 590.7 | 6.9266e-06
|  | 590.8 | 6.9307e-06
|  | 590.9 | 6.9348e-06
|  | 591 | 6.9383e-06
|  | 591.1 | 6.9421e-06
|  | 591.2 | 6.9459e-06
|  | 591.3 | 6.9494e-06
|  | 591.4 | 6.953e-06
|  | 591.5 | 6.9566e-06
|  | 591.6 | 6.9601e-06
|  | 591.7 | 6.9637e-06
|  | 591.8 | 6.9668e-06
|  | 591.9 | 6.9701e-06
|  | 592 | 6.9734e-06
|  | 592.1 | 6.9766e-06
|  | 592.2 | 6.9798e-06
|  | 592.3 | 6.983e-06
|  | 592.4 | 6.9861e-06
|  | 592.5 | 6.9893e-06
|  | 592.6 | 6.9924e-06
|  | 592.7 | 6.9956e-06
|  | 592.8 | 6.9988e-06
|  | 592.9 | 7.0019e-06
|  | 593 | 7.005e-06
|  | 593.1 | 7.0082e-06
|  | 593.2 | 7.011e-06
|  | 593.3 | 7.014e-06
|  | 593.4 | 7.017e-06
|  | 593.5 | 7.0191e-06
|  | 593.6 | 7.0217e-06
|  | 593.7 | 7.0242e-06
|  | 593.8 | 7.0268e-06
|  | 593.9 | 7.0293e-06
|  | 594 | 7.0317e-06
|  | 594.1 | 7.034e-06
|  | 594.2 | 7.0365e-06
|  | 594.3 | 7.0389e-06
|  | 594.4 | 7.0412e-06
|  | 594.5 | 7.0435e-06
|  | 594.6 | 7.0455e-06
|  | 594.7 | 7.0477e-06
|  | 594.8 | 7.0499e-06
|  | 594.9 | 7.0523e-06
|  | 595 | 7.0545e-06
|  | 595.1 | 7.0567e-06
|  | 595.2 | 7.0589e-06
|  | 595.3 | 7.0613e-06
|  | 595.4 | 7.0636e-06
|  | 595.5 | 7.0653e-06
|  | 595.6 | 7.0673e-06
|  | 595.7 | 7.0694e-06
|  | 595.8 | 7.0704e-06
|  | 595.9 | 7.072e-06
|  | 596 | 7.0731e-06
|  | 596.1 | 7.0743e-06
|  | 596.2 | 7.0755e-06
|  | 596.3 | 7.0769e-06
|  | 596.4 | 7.0782e-06
|  | 596.5 | 7.0794e-06
|  | 596.6 | 7.0803e-06
|  | 596.7 | 7.0815e-06
|  | 596.8 | 7.0827e-06
|  | 596.9 | 7.0833e-06
|  | 597 | 7.0841e-06
|  | 597.1 | 7.085e-06
|  | 597.2 | 7.0853e-06
|  | 597.3 | 7.0859e-06
|  | 597.4 | 7.0864e-06
|  | 597.5 | 7.0869e-06
|  | 597.6 | 7.0874e-06
|  | 597.7 | 7.0877e-06
|  | 597.8 | 7.0881e-06
|  | 597.9 | 7.0885e-06
|  | 598 | 7.0881e-06
|  | 598.1 | 7.0881e-06
|  | 598.2 | 7.0881e-06
|  | 598.3 | 7.0874e-06
|  | 598.4 | 7.087e-06
|  | 598.5 | 7.0867e-06
|  | 598.6 | 7.0858e-06
|  | 598.7 | 7.0853e-06
|  | 598.8 | 7.0845e-06
|  | 598.9 | 7.0838e-06
|  | 599 | 7.0831e-06
|  | 599.1 | 7.0821e-06
|  | 599.2 | 7.0812e-06
|  | 599.3 | 7.0804e-06
|  | 599.4 | 7.0786e-06
|  | 599.5 | 7.0773e-06
|  | 599.6 | 7.076e-06
|  | 599.7 | 7.0748e-06
|  | 599.8 | 7.0734e-06
|  | 599.9 | 7.072e-06
|  | 600 | 7.0703e-06
|  | 600.1 | 7.0689e-06
|  | 600.2 | 7.0669e-06
|  | 600.3 | 7.0651e-06
|  | 600.4 | 7.0633e-06
|  | 600.5 | 7.0606e-06
|  | 600.6 | 7.0584e-06
|  | 600.7 | 7.0561e-06
|  | 600.8 | 7.0534e-06
|  | 600.9 | 7.0509e-06
|  | 601 | 7.0483e-06
|  | 601.1 | 7.0455e-06
|  | 601.2 | 7.0429e-06
|  | 601.3 | 7.0398e-06
|  | 601.4 | 7.0369e-06
|  | 601.5 | 7.034e-06
|  | 601.6 | 7.0305e-06
|  | 601.7 | 7.0272e-06
|  | 601.8 | 7.024e-06
|  | 601.9 | 7.0206e-06
|  | 602 | 7.0172e-06
|  | 602.1 | 7.0137e-06
|  | 602.2 | 7.0096e-06
|  | 602.3 | 7.0059e-06
|  | 602.4 | 7.0022e-06
|  | 602.5 | 6.9983e-06
|  | 602.6 | 6.9944e-06
|  | 602.7 | 6.9901e-06
|  | 602.8 | 6.9861e-06
|  | 602.9 | 6.9821e-06
|  | 603 | 6.9771e-06
|  | 603.1 | 6.9725e-06
|  | 603.2 | 6.9679e-06
|  | 603.3 | 6.9629e-06
|  | 603.4 | 6.958e-06
|  | 603.5 | 6.9532e-06
|  | 603.6 | 6.9483e-06
|  | 603.7 | 6.9434e-06
|  | 603.8 | 6.9385e-06
|  | 603.9 | 6.9332e-06
|  | 604 | 6.9282e-06
|  | 604.1 | 6.923e-06
|  | 604.2 | 6.9179e-06
|  | 604.3 | 6.9127e-06
|  | 604.4 | 6.9074e-06
|  | 604.5 | 6.9021e-06
|  | 604.6 | 6.8969e-06
|  | 604.7 | 6.8919e-06
|  | 604.8 | 6.8867e-06
|  | 604.9 | 6.8815e-06
|  | 605 | 6.8762e-06
|  | 605.1 | 6.8711e-06
|  | 605.2 | 6.8663e-06
|  | 605.3 | 6.8613e-06
|  | 605.4 | 6.8562e-06
|  | 605.5 | 6.8509e-06
|  | 605.6 | 6.8458e-06
|  | 605.7 | 6.8407e-06
|  | 605.8 | 6.8355e-06
|  | 605.9 | 6.8302e-06
|  | 606 | 6.8249e-06
|  | 606.1 | 6.8202e-06
|  | 606.2 | 6.8151e-06
|  | 606.3 | 6.8101e-06
|  | 606.4 | 6.8052e-06
|  | 606.5 | 6.8002e-06
|  | 606.6 | 6.7949e-06
|  | 606.7 | 6.7897e-06
|  | 606.8 | 6.7846e-06
|  | 606.9 | 6.7793e-06
|  | 607 | 6.7741e-06
|  | 607.1 | 6.7688e-06
|  | 607.2 | 6.7644e-06
|  | 607.3 | 6.7595e-06
|  | 607.4 | 6.7546e-06
|  | 607.5 | 6.7499e-06
|  | 607.6 | 6.7452e-06
|  | 607.7 | 6.7406e-06
|  | 607.8 | 6.7358e-06
|  | 607.9 | 6.7311e-06
|  | 608 | 6.7266e-06
|  | 608.1 | 6.7221e-06
|  | 608.2 | 6.7175e-06
|  | 608.3 | 6.7127e-06
|  | 608.4 | 6.708e-06
|  | 608.5 | 6.7034e-06
|  | 608.6 | 6.6992e-06
|  | 608.7 | 6.6947e-06
|  | 608.8 | 6.6908e-06
|  | 608.9 | 6.6867e-06
|  | 609 | 6.6826e-06
|  | 609.1 | 6.6784e-06
|  | 609.2 | 6.6743e-06
|  | 609.3 | 6.6702e-06
|  | 609.4 | 6.6662e-06
|  | 609.5 | 6.6621e-06
|  | 609.6 | 6.658e-06
|  | 609.7 | 6.6543e-06
|  | 609.8 | 6.6504e-06
|  | 609.9 | 6.6465e-06
|  | 610 | 6.6431e-06
|  | 610.1 | 6.6395e-06
|  | 610.2 | 6.6356e-06
|  | 610.3 | 6.6318e-06
|  | 610.4 | 6.6281e-06
|  | 610.5 | 6.6249e-06
|  | 610.6 | 6.6213e-06
|  | 610.7 | 6.6178e-06
|  | 610.8 | 6.6148e-06
|  | 610.9 | 6.6116e-06
|  | 611 | 6.6084e-06
|  | 611.1 | 6.6054e-06
|  | 611.2 | 6.6023e-06
|  | 611.3 | 6.5994e-06
|  | 611.4 | 6.5964e-06
|  | 611.5 | 6.5934e-06
|  | 611.6 | 6.5907e-06
|  | 611.7 | 6.5879e-06
|  | 611.8 | 6.5851e-06
|  | 611.9 | 6.5821e-06
|  | 612 | 6.5793e-06
|  | 612.1 | 6.5765e-06
|  | 612.2 | 6.5742e-06
|  | 612.3 | 6.5715e-06
|  | 612.4 | 6.5696e-06
|  | 612.5 | 6.5675e-06
|  | 612.6 | 6.5654e-06
|  | 612.7 | 6.5637e-06
|  | 612.8 | 6.5618e-06
|  | 612.9 | 6.56e-06
|  | 613 | 6.5581e-06
|  | 613.1 | 6.5562e-06
|  | 613.2 | 6.5543e-06
|  | 613.3 | 6.5529e-06
|  | 613.4 | 6.5511e-06
|  | 613.5 | 6.5493e-06
|  | 613.6 | 6.5477e-06
|  | 613.7 | 6.546e-06
|  | 613.8 | 6.5441e-06
|  | 613.9 | 6.5423e-06
|  | 614 | 6.5405e-06
|  | 614.1 | 6.5393e-06
|  | 614.2 | 6.5378e-06
|  | 614.3 | 6.5362e-06
|  | 614.4 | 6.5354e-06
|  | 614.5 | 6.5342e-06
|  | 614.6 | 6.5332e-06
|  | 614.7 | 6.5321e-06
|  | 614.8 | 6.531e-06
|  | 614.9 | 6.5301e-06
|  | 615 | 6.5291e-06
|  | 615.1 | 6.5281e-06
|  | 615.2 | 6.5278e-06
|  | 615.3 | 6.5271e-06
|  | 615.4 | 6.5264e-06
|  | 615.5 | 6.5254e-06
|  | 615.6 | 6.5247e-06
|  | 615.7 | 6.5241e-06
|  | 615.8 | 6.5233e-06
|  | 615.9 | 6.5225e-06
|  | 616 | 6.5222e-06
|  | 616.1 | 6.5216e-06
|  | 616.2 | 6.5211e-06
|  | 616.3 | 6.521e-06
|  | 616.4 | 6.5207e-06
|  | 616.5 | 6.5204e-06
|  | 616.6 | 6.5208e-06
|  | 616.7 | 6.5207e-06
|  | 616.8 | 6.5213e-06
|  | 616.9 | 6.5216e-06
|  | 617 | 6.522e-06
|  | 617.1 | 6.5224e-06
|  | 617.2 | 6.5229e-06
|  | 617.3 | 6.5233e-06
|  | 617.4 | 6.5239e-06
|  | 617.5 | 6.5244e-06
|  | 617.6 | 6.5249e-06
|  | 617.7 | 6.5256e-06
|  | 617.8 | 6.5262e-06
|  | 617.9 | 6.5272e-06
|  | 618 | 6.5281e-06
|  | 618.1 | 6.5289e-06
|  | 618.2 | 6.5303e-06
|  | 618.3 | 6.5314e-06
|  | 618.4 | 6.5326e-06
|  | 618.5 | 6.5346e-06
|  | 618.6 | 6.5362e-06
|  | 618.7 | 6.5378e-06
|  | 618.8 | 6.5403e-06
|  | 618.9 | 6.5424e-06
|  | 619 | 6.545e-06
|  | 619.1 | 6.5474e-06
|  | 619.2 | 6.5499e-06
|  | 619.3 | 6.553e-06
|  | 619.4 | 6.5558e-06
|  | 619.5 | 6.5586e-06
|  | 619.6 | 6.562e-06
|  | 619.7 | 6.5651e-06
|  | 619.8 | 6.5682e-06
|  | 619.9 | 6.5715e-06
|  | 620 | 6.5747e-06
|  | 620.1 | 6.578e-06
|  | 620.2 | 6.5812e-06
|  | 620.3 | 6.5845e-06
|  | 620.4 | 6.5886e-06
|  | 620.5 | 6.5923e-06
|  | 620.6 | 6.596e-06
|  | 620.7 | 6.6008e-06
|  | 620.8 | 6.605e-06
|  | 620.9 | 6.6093e-06
|  | 621 | 6.6145e-06
|  | 621.1 | 6.6192e-06
|  | 621.2 | 6.6241e-06
|  | 621.3 | 6.6289e-06
|  | 621.4 | 6.6338e-06
|  | 621.5 | 6.6391e-06
|  | 621.6 | 6.6442e-06
|  | 621.7 | 6.6493e-06
|  | 621.8 | 6.6549e-06
|  | 621.9 | 6.6603e-06
|  | 622 | 6.6657e-06
|  | 622.1 | 6.6714e-06
|  | 622.2 | 6.6769e-06
|  | 622.3 | 6.6827e-06
|  | 622.4 | 6.6884e-06
|  | 622.5 | 6.694e-06
|  | 622.6 | 6.7002e-06
|  | 622.7 | 6.7061e-06
|  | 622.8 | 6.712e-06
|  | 622.9 | 6.7183e-06
|  | 623 | 6.7244e-06
|  | 623.1 | 6.731e-06
|  | 623.2 | 6.7374e-06
|  | 623.3 | 6.7438e-06
|  | 623.4 | 6.7502e-06
|  | 623.5 | 6.7567e-06
|  | 623.6 | 6.7632e-06
|  | 623.7 | 6.7695e-06
|  | 623.8 | 6.7759e-06
|  | 623.9 | 6.7823e-06
|  | 624 | 6.7887e-06
|  | 624.1 | 6.7951e-06
|  | 624.2 | 6.8017e-06
|  | 624.3 | 6.8082e-06
|  | 624.4 | 6.8147e-06
|  | 624.5 | 6.8213e-06
|  | 624.6 | 6.8279e-06
|  | 624.7 | 6.8344e-06
|  | 624.8 | 6.841e-06
|  | 624.9 | 6.8476e-06
|  | 625 | 6.8542e-06
|  | 625.1 | 6.8609e-06
|  | 625.2 | 6.8676e-06
|  | 625.3 | 6.8743e-06
|  | 625.4 | 6.881e-06
|  | 625.5 | 6.8877e-06
|  | 625.6 | 6.8948e-06
|  | 625.7 | 6.9017e-06
|  | 625.8 | 6.9085e-06
|  | 625.9 | 6.9159e-06
|  | 626 | 6.923e-06
|  | 626.1 | 6.9302e-06
|  | 626.2 | 6.9375e-06
|  | 626.3 | 6.9448e-06
|  | 626.4 | 6.952e-06
|  | 626.5 | 6.9592e-06
|  | 626.6 | 6.9664e-06
|  | 626.7 | 6.9733e-06
|  | 626.8 | 6.9803e-06
|  | 626.9 | 6.9873e-06
|  | 627 | 6.9946e-06
|  | 627.1 | 7.0017e-06
|  | 627.2 | 7.0087e-06
|  | 627.3 | 7.0158e-06
|  | 627.4 | 7.0229e-06
|  | 627.5 | 7.0298e-06
|  | 627.6 | 7.0367e-06
|  | 627.7 | 7.0437e-06
|  | 627.8 | 7.0507e-06
|  | 627.9 | 7.0577e-06
|  | 628 | 7.0646e-06
|  | 628.1 | 7.0711e-06
|  | 628.2 | 7.0778e-06
|  | 628.3 | 7.0843e-06
|  | 628.4 | 7.0908e-06
|  | 628.5 | 7.0973e-06
|  | 628.6 | 7.1035e-06
|  | 628.7 | 7.1098e-06
|  | 628.8 | 7.1162e-06
|  | 628.9 | 7.1214e-06
|  | 629 | 7.1273e-06
|  | 629.1 | 7.1327e-06
|  | 629.2 | 7.1382e-06
|  | 629.3 | 7.1436e-06
|  | 629.4 | 7.1491e-06
|  | 629.5 | 7.1545e-06
|  | 629.6 | 7.16e-06
|  | 629.7 | 7.1655e-06
|  | 629.8 | 7.171e-06
|  | 629.9 | 7.1765e-06
|  | 630 | 7.1816e-06
|  | 630.1 | 7.1869e-06
|  | 630.2 | 7.1919e-06
|  | 630.3 | 7.1969e-06
|  | 630.4 | 7.202e-06
|  | 630.5 | 7.2066e-06
|  | 630.6 | 7.2114e-06
|  | 630.7 | 7.2162e-06
|  | 630.8 | 7.2205e-06
|  | 630.9 | 7.225e-06
|  | 631 | 7.2296e-06
|  | 631.1 | 7.2339e-06
|  | 631.2 | 7.2383e-06
|  | 631.3 | 7.242e-06
|  | 631.4 | 7.2462e-06
|  | 631.5 | 7.2503e-06
|  | 631.6 | 7.2537e-06
|  | 631.7 | 7.2574e-06
|  | 631.8 | 7.2611e-06
|  | 631.9 | 7.2649e-06
|  | 632 | 7.2686e-06
|  | 632.1 | 7.2718e-06
|  | 632.2 | 7.2753e-06
|  | 632.3 | 7.2788e-06
|  | 632.4 | 7.2814e-06
|  | 632.5 | 7.2844e-06
|  | 632.6 | 7.2873e-06
|  | 632.7 | 7.2896e-06
|  | 632.8 | 7.2922e-06
|  | 632.9 | 7.2947e-06
|  | 633 | 7.2967e-06
|  | 633.1 | 7.299e-06
|  | 633.2 | 7.3002e-06
|  | 633.3 | 7.302e-06
|  | 633.4 | 7.3037e-06
|  | 633.5 | 7.3048e-06
|  | 633.6 | 7.3061e-06
|  | 633.7 | 7.3074e-06
|  | 633.8 | 7.3081e-06
|  | 633.9 | 7.3091e-06
|  | 634 | 7.3095e-06
|  | 634.1 | 7.31e-06
|  | 634.2 | 7.3106e-06
|  | 634.3 | 7.3112e-06
|  | 634.4 | 7.3117e-06
|  | 634.5 | 7.3121e-06
|  | 634.6 | 7.3129e-06
|  | 634.7 | 7.3135e-06
|  | 634.8 | 7.3143e-06
|  | 634.9 | 7.3151e-06
|  | 635 | 7.3158e-06
|  | 635.1 | 7.3166e-06
|  | 635.2 | 7.3173e-06
|  | 635.3 | 7.3181e-06
|  | 635.4 | 7.318e-06
|  | 635.5 | 7.3184e-06
|  | 635.6 | 7.3188e-06
|  | 635.7 | 7.3185e-06
|  | 635.8 | 7.3185e-06
|  | 635.9 | 7.3185e-06
|  | 636 | 7.3185e-06
|  | 636.1 | 7.3185e-06
|  | 636.2 | 7.319e-06
|  | 636.3 | 7.3193e-06
|  | 636.4 | 7.3195e-06
|  | 636.5 | 7.3193e-06
|  | 636.6 | 7.3195e-06
|  | 636.7 | 7.3192e-06
|  | 636.8 | 7.319e-06
|  | 636.9 | 7.3188e-06
|  | 637 | 7.3181e-06
|  | 637.1 | 7.3176e-06
|  | 637.2 | 7.3172e-06
|  | 637.3 | 7.3159e-06
|  | 637.4 | 7.3151e-06
|  | 637.5 | 7.3142e-06
|  | 637.6 | 7.3122e-06
|  | 637.7 | 7.3108e-06
|  | 637.8 | 7.3085e-06
|  | 637.9 | 7.3065e-06
|  | 638 | 7.3045e-06
|  | 638.1 | 7.3021e-06
|  | 638.2 | 7.2998e-06
|  | 638.3 | 7.2975e-06
|  | 638.4 | 7.2951e-06
|  | 638.5 | 7.2928e-06
|  | 638.6 | 7.2902e-06
|  | 638.7 | 7.2877e-06
|  | 638.8 | 7.2852e-06
|  | 638.9 | 7.282e-06
|  | 639 | 7.2791e-06
|  | 639.1 | 7.2762e-06
|  | 639.2 | 7.2728e-06
|  | 639.3 | 7.2696e-06
|  | 639.4 | 7.2658e-06
|  | 639.5 | 7.2623e-06
|  | 639.6 | 7.2587e-06
|  | 639.7 | 7.2547e-06
|  | 639.8 | 7.2509e-06
|  | 639.9 | 7.247e-06
|  | 640 | 7.2424e-06
|  | 640.1 | 7.2382e-06
|  | 640.2 | 7.234e-06
|  | 640.3 | 7.2289e-06
|  | 640.4 | 7.2243e-06
|  | 640.5 | 7.2192e-06
|  | 640.6 | 7.2143e-06
|  | 640.7 | 7.2094e-06
|  | 640.8 | 7.2042e-06
|  | 640.9 | 7.1992e-06
|  | 641 | 7.1941e-06
|  | 641.1 | 7.1877e-06
|  | 641.2 | 7.1821e-06
|  | 641.3 | 7.1761e-06
|  | 641.4 | 7.1701e-06
|  | 641.5 | 7.164e-06
|  | 641.6 | 7.158e-06
|  | 641.7 | 7.152e-06
|  | 641.8 | 7.146e-06
|  | 641.9 | 7.1401e-06
|  | 642 | 7.1341e-06
|  | 642.1 | 7.1283e-06
|  | 642.2 | 7.1224e-06
|  | 642.3 | 7.1166e-06
|  | 642.4 | 7.1113e-06
|  | 642.5 | 7.1057e-06
|  | 642.6 | 7.1001e-06
|  | 642.7 | 7.095e-06
|  | 642.8 | 7.0896e-06
|  | 642.9 | 7.0849e-06
|  | 643 | 7.0799e-06
|  | 643.1 | 7.0749e-06
|  | 643.2 | 7.0706e-06
|  | 643.3 | 7.0661e-06
|  | 643.4 | 7.0616e-06
|  | 643.5 | 7.0573e-06
|  | 643.6 | 7.053e-06
|  | 643.7 | 7.0485e-06
|  | 643.8 | 7.0441e-06
|  | 643.9 | 7.0396e-06
|  | 644 | 7.0353e-06
|  | 644.1 | 7.0309e-06
|  | 644.2 | 7.0265e-06
|  | 644.3 | 7.0221e-06
|  | 644.4 | 7.0178e-06
|  | 644.5 | 7.0134e-06
|  | 644.6 | 7.0092e-06
|  | 644.7 | 7.0049e-06
|  | 644.8 | 7.0006e-06
|  | 644.9 | 6.9963e-06
|  | 645 | 6.9921e-06
|  | 645.1 | 6.9874e-06
|  | 645.2 | 6.983e-06
|  | 645.3 | 6.9786e-06
|  | 645.4 | 6.9747e-06
|  | 645.5 | 6.9704e-06
|  | 645.6 | 6.9662e-06
|  | 645.7 | 6.962e-06
|  | 645.8 | 6.9579e-06
|  | 645.9 | 6.9531e-06
|  | 646 | 6.9486e-06
|  | 646.1 | 6.9441e-06
|  | 646.2 | 6.9395e-06
|  | 646.3 | 6.9349e-06
|  | 646.4 | 6.9307e-06
|  | 646.5 | 6.9263e-06
|  | 646.6 | 6.9219e-06
|  | 646.7 | 6.9179e-06
|  | 646.8 | 6.9138e-06
|  | 646.9 | 6.9096e-06
|  | 647 | 6.9062e-06
|  | 647.1 | 6.9024e-06
|  | 647.2 | 6.8991e-06
|  | 647.3 | 6.8957e-06
|  | 647.4 | 6.8923e-06
|  | 647.5 | 6.8891e-06
|  | 647.6 | 6.8859e-06
|  | 647.7 | 6.8826e-06
|  | 647.8 | 6.8795e-06
|  | 647.9 | 6.8763e-06
|  | 648 | 6.8732e-06
|  | 648.1 | 6.87e-06
|  | 648.2 | 6.8669e-06
|  | 648.3 | 6.8637e-06
|  | 648.4 | 6.8605e-06
|  | 648.5 | 6.8574e-06
|  | 648.6 | 6.8543e-06
|  | 648.7 | 6.8512e-06
|  | 648.8 | 6.8482e-06
|  | 648.9 | 6.8452e-06
|  | 649 | 6.8422e-06
|  | 649.1 | 6.8395e-06
|  | 649.2 | 6.8367e-06
|  | 649.3 | 6.8339e-06
|  | 649.4 | 6.8304e-06
|  | 649.5 | 6.8274e-06
|  | 649.6 | 6.8243e-06
|  | 649.7 | 6.821e-06
|  | 649.8 | 6.8178e-06
|  | 649.9 | 6.8149e-06
|  | 650 | 6.8119e-06
|  | 650.1 | 6.8089e-06
|  | 650.2 | 6.8059e-06
|  | 650.3 | 6.8029e-06
|  | 650.4 | 6.7998e-06
|  | 650.5 | 6.7968e-06
|  | 650.6 | 6.7937e-06
|  | 650.7 | 6.7908e-06
|  | 650.8 | 6.7879e-06
|  | 650.9 | 6.7849e-06
|  | 651 | 6.7818e-06
|  | 651.1 | 6.7788e-06
|  | 651.2 | 6.7755e-06
|  | 651.3 | 6.7724e-06
|  | 651.4 | 6.7692e-06
|  | 651.5 | 6.7661e-06
|  | 651.6 | 6.7629e-06
|  | 651.7 | 6.7597e-06
|  | 651.8 | 6.758e-06
|  | 651.9 | 6.7555e-06
|  | 652 | 6.7529e-06
|  | 652.1 | 6.7512e-06
|  | 652.2 | 6.7492e-06
|  | 652.3 | 6.7476e-06
|  | 652.4 | 6.7458e-06
|  | 652.5 | 6.7439e-06
|  | 652.6 | 6.7417e-06
|  | 652.7 | 6.7397e-06
|  | 652.8 | 6.7378e-06
|  | 652.9 | 6.7354e-06
|  | 653 | 6.7331e-06
|  | 653.1 | 6.7309e-06
|  | 653.2 | 6.7287e-06
|  | 653.3 | 6.7265e-06
|  | 653.4 | 6.7246e-06
|  | 653.5 | 6.7225e-06
|  | 653.6 | 6.7208e-06
|  | 653.7 | 6.719e-06
|  | 653.8 | 6.7172e-06
|  | 653.9 | 6.716e-06
|  | 654 | 6.7146e-06
|  | 654.1 | 6.7132e-06
|  | 654.2 | 6.7127e-06
|  | 654.3 | 6.7117e-06
|  | 654.4 | 6.7109e-06
|  | 654.5 | 6.7102e-06
|  | 654.6 | 6.7094e-06
|  | 654.7 | 6.7077e-06
|  | 654.8 | 6.7065e-06
|  | 654.9 | 6.7053e-06
|  | 655 | 6.7037e-06
|  | 655.1 | 6.7021e-06
|  | 655.2 | 6.701e-06
|  | 655.3 | 6.6998e-06
|  | 655.4 | 6.6986e-06
|  | 655.5 | 6.6982e-06
|  | 655.6 | 6.6975e-06
|  | 655.7 | 6.6968e-06
|  | 655.8 | 6.6966e-06
|  | 655.9 | 6.6962e-06
|  | 656 | 6.6959e-06
|  | 656.1 | 6.6955e-06
|  | 656.2 | 6.6952e-06
|  | 656.3 | 6.6954e-06
|  | 656.4 | 6.6953e-06
|  | 656.5 | 6.6952e-06
|  | 656.6 | 6.6943e-06
|  | 656.7 | 6.694e-06
|  | 656.8 | 6.6931e-06
|  | 656.9 | 6.6922e-06
|  | 657 | 6.6913e-06
|  | 657.1 | 6.6898e-06
|  | 657.2 | 6.6886e-06
|  | 657.3 | 6.6874e-06
|  | 657.4 | 6.686e-06
|  | 657.5 | 6.6847e-06
|  | 657.6 | 6.6843e-06
|  | 657.7 | 6.6835e-06
|  | 657.8 | 6.6827e-06
|  | 657.9 | 6.6828e-06
|  | 658 | 6.6826e-06
|  | 658.1 | 6.6824e-06
|  | 658.2 | 6.6823e-06
|  | 658.3 | 6.6822e-06
|  | 658.4 | 6.682e-06
|  | 658.5 | 6.6818e-06
|  | 658.6 | 6.6817e-06
|  | 658.7 | 6.6814e-06
|  | 658.8 | 6.6812e-06
|  | 658.9 | 6.681e-06
|  | 659 | 6.6805e-06
|  | 659.1 | 6.6802e-06
|  | 659.2 | 6.6798e-06
|  | 659.3 | 6.6795e-06
|  | 659.4 | 6.6791e-06
|  | 659.5 | 6.6792e-06
|  | 659.6 | 6.6791e-06
|  | 659.7 | 6.679e-06
|  | 659.8 | 6.68e-06
|  | 659.9 | 6.6804e-06
|  | 660 | 6.6815e-06
|  | 660.1 | 6.6824e-06
|  | 660.2 | 6.6833e-06
|  | 660.3 | 6.6849e-06
|  | 660.4 | 6.6862e-06
|  | 660.5 | 6.6875e-06
|  | 660.6 | 6.6885e-06
|  | 660.7 | 6.6898e-06
|  | 660.8 | 6.69e-06
|  | 660.9 | 6.6907e-06
|  | 661 | 6.6913e-06
|  | 661.1 | 6.6916e-06
|  | 661.2 | 6.692e-06
|  | 661.3 | 6.6924e-06
|  | 661.4 | 6.6928e-06
|  | 661.5 | 6.6932e-06
|  | 661.6 | 6.6938e-06
|  | 661.7 | 6.6942e-06
|  | 661.8 | 6.6947e-06
|  | 661.9 | 6.6956e-06
|  | 662 | 6.6963e-06
|  | 662.1 | 6.6971e-06
|  | 662.2 | 6.6979e-06
|  | 662.3 | 6.6987e-06
|  | 662.4 | 6.6995e-06
|  | 662.5 | 6.7002e-06
|  | 662.6 | 6.701e-06
|  | 662.7 | 6.7019e-06
|  | 662.8 | 6.7027e-06
|  | 662.9 | 6.7039e-06
|  | 663 | 6.7049e-06
|  | 663.1 | 6.7059e-06
|  | 663.2 | 6.7072e-06
|  | 663.3 | 6.7084e-06
|  | 663.4 | 6.7097e-06
|  | 663.5 | 6.7111e-06
|  | 663.6 | 6.7123e-06
|  | 663.7 | 6.7142e-06
|  | 663.8 | 6.7158e-06
|  | 663.9 | 6.7174e-06
|  | 664 | 6.7191e-06
|  | 664.1 | 6.7209e-06
|  | 664.2 | 6.7227e-06
|  | 664.3 | 6.7245e-06
|  | 664.4 | 6.7262e-06
|  | 664.5 | 6.728e-06
|  | 664.6 | 6.7299e-06
|  | 664.7 | 6.7317e-06
|  | 664.8 | 6.7337e-06
|  | 664.9 | 6.7355e-06
|  | 665 | 6.7374e-06
|  | 665.1 | 6.7395e-06
|  | 665.2 | 6.7415e-06
|  | 665.3 | 6.7436e-06
|  | 665.4 | 6.7456e-06
|  | 665.5 | 6.7477e-06
|  | 665.6 | 6.7494e-06
|  | 665.7 | 6.7513e-06
|  | 665.8 | 6.7532e-06
|  | 665.9 | 6.7551e-06
|  | 666 | 6.757e-06
|  | 666.1 | 6.7587e-06
|  | 666.2 | 6.7605e-06
|  | 666.3 | 6.7623e-06
|  | 666.4 | 6.7642e-06
|  | 666.5 | 6.766e-06
|  | 666.6 | 6.7682e-06
|  | 666.7 | 6.7702e-06
|  | 666.8 | 6.7723e-06
|  | 666.9 | 6.7747e-06
|  | 667 | 6.7769e-06
|  | 667.1 | 6.7792e-06
|  | 667.2 | 6.7816e-06
|  | 667.3 | 6.7839e-06
|  | 667.4 | 6.7865e-06
|  | 667.5 | 6.789e-06
|  | 667.6 | 6.7915e-06
|  | 667.7 | 6.7942e-06
|  | 667.8 | 6.7968e-06
|  | 667.9 | 6.7992e-06
|  | 668 | 6.8017e-06
|  | 668.1 | 6.8042e-06
|  | 668.2 | 6.8067e-06
|  | 668.3 | 6.8092e-06
|  | 668.4 | 6.8117e-06
|  | 668.5 | 6.8139e-06
|  | 668.6 | 6.8163e-06
|  | 668.7 | 6.8184e-06
|  | 668.8 | 6.8206e-06
|  | 668.9 | 6.8228e-06
|  | 669 | 6.8255e-06
|  | 669.1 | 6.8279e-06
|  | 669.2 | 6.8303e-06
|  | 669.3 | 6.8328e-06
|  | 669.4 | 6.8353e-06
|  | 669.5 | 6.8376e-06
|  | 669.6 | 6.84e-06
|  | 669.7 | 6.8424e-06
|  | 669.8 | 6.8451e-06
|  | 669.9 | 6.8476e-06
|  | 670 | 6.8501e-06
|  | 670.1 | 6.8528e-06
|  | 670.2 | 6.8554e-06
|  | 670.3 | 6.8581e-06
|  | 670.4 | 6.8608e-06
|  | 670.5 | 6.8634e-06
|  | 670.6 | 6.8665e-06
|  | 670.7 | 6.8693e-06
|  | 670.8 | 6.8723e-06
|  | 670.9 | 6.8752e-06
|  | 671 | 6.8782e-06
|  | 671.1 | 6.8811e-06
|  | 671.2 | 6.8841e-06
|  | 671.3 | 6.8871e-06
|  | 671.4 | 6.891e-06
|  | 671.5 | 6.8944e-06
|  | 671.6 | 6.8981e-06
|  | 671.7 | 6.9019e-06
|  | 671.8 | 6.9056e-06
|  | 671.9 | 6.9092e-06
|  | 672 | 6.9128e-06
|  | 672.1 | 6.9165e-06
|  | 672.2 | 6.9204e-06
|  | 672.3 | 6.9242e-06
|  | 672.4 | 6.9278e-06
|  | 672.5 | 6.9316e-06
|  | 672.6 | 6.9353e-06
|  | 672.7 | 6.9385e-06
|  | 672.8 | 6.942e-06
|  | 672.9 | 6.9457e-06
|  | 673 | 6.9492e-06
|  | 673.1 | 6.9527e-06
|  | 673.2 | 6.9557e-06
|  | 673.3 | 6.959e-06
|  | 673.4 | 6.9623e-06
|  | 673.5 | 6.9647e-06
|  | 673.6 | 6.9675e-06
|  | 673.7 | 6.9706e-06
|  | 673.8 | 6.9735e-06
|  | 673.9 | 6.9764e-06
|  | 674 | 6.98e-06
|  | 674.1 | 6.9833e-06
|  | 674.2 | 6.9865e-06
|  | 674.3 | 6.9899e-06
|  | 674.4 | 6.9932e-06
|  | 674.5 | 6.9961e-06
|  | 674.6 | 6.9991e-06
|  | 674.7 | 7.0022e-06
|  | 674.8 | 7.005e-06
|  | 674.9 | 7.008e-06
|  | 675 | 7.0109e-06
|  | 675.1 | 7.0138e-06
|  | 675.2 | 7.0167e-06
|  | 675.3 | 7.02e-06
|  | 675.4 | 7.0231e-06
|  | 675.5 | 7.0262e-06
|  | 675.6 | 7.0301e-06
|  | 675.7 | 7.0335e-06
|  | 675.8 | 7.0365e-06
|  | 675.9 | 7.0398e-06
|  | 676 | 7.0431e-06
|  | 676.1 | 7.0469e-06
|  | 676.2 | 7.0504e-06
|  | 676.3 | 7.0544e-06
|  | 676.4 | 7.0583e-06
|  | 676.5 | 7.0623e-06
|  | 676.6 | 7.0654e-06
|  | 676.7 | 7.069e-06
|  | 676.8 | 7.0726e-06
|  | 676.9 | 7.0759e-06
|  | 677 | 7.0792e-06
|  | 677.1 | 7.0826e-06
|  | 677.2 | 7.086e-06
|  | 677.3 | 7.0895e-06
|  | 677.4 | 7.0927e-06
|  | 677.5 | 7.096e-06
|  | 677.6 | 7.0988e-06
|  | 677.7 | 7.1018e-06
|  | 677.8 | 7.1049e-06
|  | 677.9 | 7.1077e-06
|  | 678 | 7.1105e-06
|  | 678.1 | 7.1133e-06
|  | 678.2 | 7.1165e-06
|  | 678.3 | 7.1195e-06
|  | 678.4 | 7.1224e-06
|  | 678.5 | 7.1254e-06
|  | 678.6 | 7.1284e-06
|  | 678.7 | 7.1322e-06
|  | 678.8 | 7.1355e-06
|  | 678.9 | 7.1388e-06
|  | 679 | 7.1425e-06
|  | 679.1 | 7.146e-06
|  | 679.2 | 7.15e-06
|  | 679.3 | 7.1538e-06
|  | 679.4 | 7.1575e-06
|  | 679.5 | 7.1624e-06
|  | 679.6 | 7.1667e-06
|  | 679.7 | 7.1711e-06
|  | 679.8 | 7.1756e-06
|  | 679.9 | 7.1801e-06
|  | 680 | 7.1845e-06
|  | 680.1 | 7.1889e-06
|  | 680.2 | 7.1933e-06
|  | 680.3 | 7.1969e-06
|  | 680.4 | 7.2011e-06
|  | 680.5 | 7.2052e-06
|  | 680.6 | 7.2092e-06
|  | 680.7 | 7.2133e-06
|  | 680.8 | 7.2177e-06
|  | 680.9 | 7.2219e-06
|  | 681 | 7.2258e-06
|  | 681.1 | 7.2298e-06
|  | 681.2 | 7.2339e-06
|  | 681.3 | 7.2383e-06
|  | 681.4 | 7.2425e-06
|  | 681.5 | 7.2466e-06
|  | 681.6 | 7.2505e-06
|  | 681.7 | 7.2547e-06
|  | 681.8 | 7.2589e-06
|  | 681.9 | 7.263e-06
|  | 682 | 7.267e-06
|  | 682.1 | 7.2705e-06
|  | 682.2 | 7.2743e-06
|  | 682.3 | 7.2781e-06
|  | 682.4 | 7.2817e-06
|  | 682.5 | 7.2854e-06
|  | 682.6 | 7.2889e-06
|  | 682.7 | 7.2925e-06
|  | 682.8 | 7.2962e-06
|  | 682.9 | 7.3005e-06
|  | 683 | 7.3043e-06
|  | 683.1 | 7.3093e-06
|  | 683.2 | 7.3138e-06
|  | 683.3 | 7.3184e-06
|  | 683.4 | 7.3235e-06
|  | 683.5 | 7.3285e-06
|  | 683.6 | 7.3334e-06
|  | 683.7 | 7.3383e-06
|  | 683.8 | 7.3432e-06
|  | 683.9 | 7.3475e-06
|  | 684 | 7.3521e-06
|  | 684.1 | 7.3568e-06
|  | 684.2 | 7.3598e-06
|  | 684.3 | 7.3637e-06
|  | 684.4 | 7.3675e-06
|  | 684.5 | 7.3711e-06
|  | 684.6 | 7.3748e-06
|  | 684.7 | 7.3786e-06
|  | 684.8 | 7.3823e-06
|  | 684.9 | 7.386e-06
|  | 685 | 7.3897e-06
|  | 685.1 | 7.3934e-06
|  | 685.2 | 7.3963e-06
|  | 685.3 | 7.3997e-06
|  | 685.4 | 7.4029e-06
|  | 685.5 | 7.4061e-06
|  | 685.6 | 7.4092e-06
|  | 685.7 | 7.4123e-06
|  | 685.8 | 7.4154e-06
|  | 685.9 | 7.4185e-06
|  | 686 | 7.4219e-06
|  | 686.1 | 7.4251e-06
|  | 686.2 | 7.4286e-06
|  | 686.3 | 7.432e-06
|  | 686.4 | 7.4354e-06
|  | 686.5 | 7.439e-06
|  | 686.6 | 7.4425e-06
|  | 686.7 | 7.4468e-06
|  | 686.8 | 7.4507e-06
|  | 686.9 | 7.4547e-06
|  | 687 | 7.4598e-06
|  | 687.1 | 7.4644e-06
|  | 687.2 | 7.469e-06
|  | 687.3 | 7.4743e-06
|  | 687.4 | 7.4792e-06
|  | 687.5 | 7.4832e-06
|  | 687.6 | 7.4877e-06
|  | 687.7 | 7.4922e-06
|  | 687.8 | 7.496e-06
|  | 687.9 | 7.5001e-06
|  | 688 | 7.5047e-06
|  | 688.1 | 7.5091e-06
|  | 688.2 | 7.5135e-06
|  | 688.3 | 7.5179e-06
|  | 688.4 | 7.5224e-06
|  | 688.5 | 7.5269e-06
|  | 688.6 | 7.5306e-06
|  | 688.7 | 7.5348e-06
|  | 688.8 | 7.5387e-06
|  | 688.9 | 7.5427e-06
|  | 689 | 7.5467e-06
|  | 689.1 | 7.5504e-06
|  | 689.2 | 7.5543e-06
|  | 689.3 | 7.5574e-06
|  | 689.4 | 7.5608e-06
|  | 689.5 | 7.5642e-06
|  | 689.6 | 7.5673e-06
|  | 689.7 | 7.5704e-06
|  | 689.8 | 7.5736e-06
|  | 689.9 | 7.5768e-06
|  | 690 | 7.58e-06
|  | 690.1 | 7.5824e-06
|  | 690.2 | 7.5853e-06
|  | 690.3 | 7.5882e-06
|  | 690.4 | 7.5921e-06
|  | 690.5 | 7.5953e-06
|  | 690.6 | 7.5997e-06
|  | 690.7 | 7.6037e-06
|  | 690.8 | 7.6078e-06
|  | 690.9 | 7.6114e-06
|  | 691 | 7.6154e-06
|  | 691.1 | 7.6199e-06
|  | 691.2 | 7.624e-06
|  | 691.3 | 7.6282e-06
|  | 691.4 | 7.6334e-06
|  | 691.5 | 7.6383e-06
|  | 691.6 | 7.6431e-06
|  | 691.7 | 7.6479e-06
|  | 691.8 | 7.6528e-06
|  | 691.9 | 7.657e-06
|  | 692 | 7.6615e-06
|  | 692.1 | 7.6659e-06
|  | 692.2 | 7.6709e-06
|  | 692.3 | 7.6755e-06
|  | 692.4 | 7.68e-06
|  | 692.5 | 7.6847e-06
|  | 692.6 | 7.6894e-06
|  | 692.7 | 7.6935e-06
|  | 692.8 | 7.6979e-06
|  | 692.9 | 7.7026e-06
|  | 693 | 7.7071e-06
|  | 693.1 | 7.7115e-06
|  | 693.2 | 7.716e-06
|  | 693.3 | 7.7205e-06
|  | 693.4 | 7.7251e-06
|  | 693.5 | 7.7292e-06
|  | 693.6 | 7.7335e-06
|  | 693.7 | 7.7377e-06
|  | 693.8 | 7.7419e-06
|  | 693.9 | 7.7461e-06
|  | 694 | 7.7496e-06
|  | 694.1 | 7.7535e-06
|  | 694.2 | 7.7569e-06
|  | 694.3 | 7.7604e-06
|  | 694.4 | 7.7639e-06
|  | 694.5 | 7.7669e-06
|  | 694.6 | 7.7701e-06
|  | 694.7 | 7.7744e-06
|  | 694.8 | 7.7781e-06
|  | 694.9 | 7.7819e-06
|  | 695 | 7.7865e-06
|  | 695.1 | 7.7909e-06
|  | 695.2 | 7.7952e-06
|  | 695.3 | 7.7999e-06
|  | 695.4 | 7.8044e-06
|  | 695.5 | 7.8092e-06
|  | 695.6 | 7.8139e-06
|  | 695.7 | 7.8185e-06
|  | 695.8 | 7.8236e-06
|  | 695.9 | 7.8285e-06
|  | 696 | 7.8336e-06
|  | 696.1 | 7.8387e-06
|  | 696.2 | 7.8437e-06
|  | 696.3 | 7.8485e-06
|  | 696.4 | 7.8535e-06
|  | 696.5 | 7.8584e-06
|  | 696.6 | 7.8633e-06
|  | 696.7 | 7.8682e-06
|  | 696.8 | 7.8733e-06
|  | 696.9 | 7.8782e-06
|  | 697 | 7.8832e-06
|  | 697.1 | 7.8884e-06
|  | 697.2 | 7.8935e-06
|  | 697.3 | 7.8994e-06
|  | 697.4 | 7.9049e-06
|  | 697.5 | 7.9104e-06
|  | 697.6 | 7.9155e-06
|  | 697.7 | 7.921e-06
|  | 697.8 | 7.9254e-06
|  | 697.9 | 7.9302e-06
|  | 698 | 7.935e-06
|  | 698.1 | 7.9393e-06
|  | 698.2 | 7.9438e-06
|  | 698.3 | 7.9484e-06
|  | 698.4 | 7.953e-06
|  | 698.5 | 7.9576e-06
|  | 698.6 | 7.9622e-06
|  | 698.7 | 7.9668e-06
|  | 698.8 | 7.9712e-06
|  | 698.9 | 7.9757e-06
|  | 699 | 7.9801e-06
|  | 699.1 | 7.9847e-06
|  | 699.2 | 7.9892e-06
|  | 699.3 | 7.9937e-06
|  | 699.4 | 7.9994e-06
|  | 699.5 | 8.0043e-06
|  | 699.6 | 8.0086e-06
|  | 699.7 | 8.0134e-06
|  | 699.8 | 8.0182e-06
|  | 699.9 | 8.0221e-06
|  | 700 | 8.0263e-06
|  | 700.1 | 8.0305e-06
|  | 700.2 | 8.0348e-06
|  | 700.3 | 8.039e-06
|  | 700.4 | 8.0438e-06
|  | 700.5 | 8.0483e-06
|  | 700.6 | 8.0534e-06
|  | 700.7 | 8.0583e-06
|  | 700.8 | 8.0632e-06
|  | 700.9 | 8.068e-06
|  | 701 | 8.073e-06
|  | 701.1 | 8.0779e-06
|  | 701.2 | 8.0817e-06
|  | 701.3 | 8.0862e-06
|  | 701.4 | 8.0904e-06
|  | 701.5 | 8.0947e-06
|  | 701.6 | 8.0989e-06
|  | 701.7 | 8.1027e-06
|  | 701.8 | 8.1068e-06
|  | 701.9 | 8.1105e-06
|  | 702 | 8.1144e-06
|  | 702.1 | 8.1182e-06
|  | 702.2 | 8.1213e-06
|  | 702.3 | 8.1248e-06
|  | 702.4 | 8.129e-06
|  | 702.5 | 8.1328e-06
|  | 702.6 | 8.1365e-06
|  | 702.7 | 8.1402e-06
|  | 702.8 | 8.1441e-06
|  | 702.9 | 8.1477e-06
|  | 703 | 8.1513e-06
|  | 703.1 | 8.1548e-06
|  | 703.2 | 8.1581e-06
|  | 703.3 | 8.1616e-06
|  | 703.4 | 8.1651e-06
|  | 703.5 | 8.1685e-06
|  | 703.6 | 8.1719e-06
|  | 703.7 | 8.1763e-06
|  | 703.8 | 8.1802e-06
|  | 703.9 | 8.1841e-06
|  | 704 | 8.1899e-06
|  | 704.1 | 8.1947e-06
|  | 704.2 | 8.1992e-06
|  | 704.3 | 8.204e-06
|  | 704.4 | 8.2088e-06
|  | 704.5 | 8.214e-06
|  | 704.6 | 8.2188e-06
|  | 704.7 | 8.224e-06
|  | 704.8 | 8.2292e-06
|  | 704.9 | 8.2344e-06
|  | 705 | 8.24e-06
|  | 705.1 | 8.2454e-06
|  | 705.2 | 8.2507e-06
|  | 705.3 | 8.256e-06
|  | 705.4 | 8.2614e-06
|  | 705.5 | 8.2669e-06
|  | 705.6 | 8.2722e-06
|  | 705.7 | 8.2775e-06
|  | 705.8 | 8.2829e-06
|  | 705.9 | 8.2883e-06
|  | 706 | 8.2935e-06
|  | 706.1 | 8.2988e-06
|  | 706.2 | 8.304e-06
|  | 706.3 | 8.3096e-06
|  | 706.4 | 8.315e-06
|  | 706.5 | 8.3194e-06
|  | 706.6 | 8.3243e-06
|  | 706.7 | 8.3293e-06
|  | 706.8 | 8.3335e-06
|  | 706.9 | 8.3379e-06
|  | 707 | 8.3422e-06
|  | 707.1 | 8.3466e-06
|  | 707.2 | 8.351e-06
|  | 707.3 | 8.3546e-06
|  | 707.4 | 8.3587e-06
|  | 707.5 | 8.3624e-06
|  | 707.6 | 8.3661e-06
|  | 707.7 | 8.3699e-06
|  | 707.8 | 8.3739e-06
|  | 707.9 | 8.3778e-06
|  | 708 | 8.382e-06
|  | 708.1 | 8.3862e-06
|  | 708.2 | 8.3903e-06
|  | 708.3 | 8.3951e-06
|  | 708.4 | 8.3996e-06
|  | 708.5 | 8.4041e-06
|  | 708.6 | 8.4086e-06
|  | 708.7 | 8.4132e-06
|  | 708.8 | 8.4179e-06
|  | 708.9 | 8.4224e-06
|  | 709 | 8.4269e-06
|  | 709.1 | 8.4314e-06
|  | 709.2 | 8.4359e-06
|  | 709.3 | 8.4414e-06
|  | 709.4 | 8.4463e-06
|  | 709.5 | 8.4513e-06
|  | 709.6 | 8.4574e-06
|  | 709.7 | 8.463e-06
|  | 709.8 | 8.4701e-06
|  | 709.9 | 8.4764e-06
|  | 710 | 8.4827e-06
|  | 710.1 | 8.4894e-06
|  | 710.2 | 8.4961e-06
|  | 710.3 | 8.5026e-06
|  | 710.4 | 8.5091e-06
|  | 710.5 | 8.5155e-06
|  | 710.6 | 8.5216e-06
|  | 710.7 | 8.528e-06
|  | 710.8 | 8.5347e-06
|  | 710.9 | 8.5412e-06
|  | 711 | 8.5477e-06
|  | 711.1 | 8.5546e-06
|  | 711.2 | 8.5614e-06
|  | 711.3 | 8.5672e-06
|  | 711.4 | 8.5734e-06
|  | 711.5 | 8.5796e-06
|  | 711.6 | 8.5856e-06
|  | 711.7 | 8.5915e-06
|  | 711.8 | 8.5974e-06
|  | 711.9 | 8.6034e-06
|  | 712 | 8.6095e-06
|  | 712.1 | 8.6149e-06
|  | 712.2 | 8.6206e-06
|  | 712.3 | 8.6249e-06
|  | 712.4 | 8.6297e-06
|  | 712.5 | 8.6345e-06
|  | 712.6 | 8.6391e-06
|  | 712.7 | 8.6436e-06
|  | 712.8 | 8.648e-06
|  | 712.9 | 8.6541e-06
|  | 713 | 8.6592e-06
|  | 713.1 | 8.6652e-06
|  | 713.2 | 8.6709e-06
|  | 713.3 | 8.6765e-06
|  | 713.4 | 8.6832e-06
|  | 713.5 | 8.6893e-06
|  | 713.6 | 8.6949e-06
|  | 713.7 | 8.7008e-06
|  | 713.8 | 8.7068e-06
|  | 713.9 | 8.7129e-06
|  | 714 | 8.7187e-06
|  | 714.1 | 8.7265e-06
|  | 714.2 | 8.7335e-06
|  | 714.3 | 8.7405e-06
|  | 714.4 | 8.7482e-06
|  | 714.5 | 8.7557e-06
|  | 714.6 | 8.7632e-06
|  | 714.7 | 8.7706e-06
|  | 714.8 | 8.7781e-06
|  | 714.9 | 8.7856e-06
|  | 715 | 8.7931e-06
|  | 715.1 | 8.8017e-06
|  | 715.2 | 8.8098e-06
|  | 715.3 | 8.8179e-06
|  | 715.4 | 8.8262e-06
|  | 715.5 | 8.8347e-06
|  | 715.6 | 8.8426e-06
|  | 715.7 | 8.8506e-06
|  | 715.8 | 8.8587e-06
|  | 715.9 | 8.8661e-06
|  | 716 | 8.8738e-06
|  | 716.1 | 8.8814e-06
|  | 716.2 | 8.8889e-06
|  | 716.3 | 8.8965e-06
|  | 716.4 | 8.9038e-06
|  | 716.5 | 8.9113e-06
|  | 716.6 | 8.918e-06
|  | 716.7 | 8.9249e-06
|  | 716.8 | 8.9319e-06
|  | 716.9 | 8.9381e-06
|  | 717 | 8.9446e-06
|  | 717.1 | 8.9506e-06
|  | 717.2 | 8.9567e-06
|  | 717.3 | 8.9629e-06
|  | 717.4 | 8.9692e-06
|  | 717.5 | 8.9753e-06
|  | 717.6 | 8.9819e-06
|  | 717.7 | 8.9885e-06
|  | 717.8 | 8.995e-06
|  | 717.9 | 9.0015e-06
|  | 718 | 9.0081e-06
|  | 718.1 | 9.0147e-06
|  | 718.2 | 9.022e-06
|  | 718.3 | 9.0288e-06
|  | 718.4 | 9.0357e-06
|  | 718.5 | 9.0426e-06
|  | 718.6 | 9.0496e-06
|  | 718.7 | 9.0567e-06
|  | 718.8 | 9.0637e-06
|  | 718.9 | 9.0711e-06
|  | 719 | 9.0784e-06
|  | 719.1 | 9.0857e-06
|  | 719.2 | 9.0937e-06
|  | 719.3 | 9.1013e-06
|  | 719.4 | 9.1097e-06
|  | 719.5 | 9.1177e-06
|  | 719.6 | 9.1258e-06
|  | 719.7 | 9.1352e-06
|  | 719.8 | 9.1439e-06
|  | 719.9 | 9.1538e-06
|  | 720 | 9.1632e-06
|  | 720.1 | 9.1727e-06
|  | 720.2 | 9.183e-06
|  | 720.3 | 9.1929e-06
|  | 720.4 | 9.2031e-06
|  | 720.5 | 9.2132e-06
|  | 720.6 | 9.2233e-06
|  | 720.7 | 9.2336e-06
|  | 720.8 | 9.2438e-06
|  | 720.9 | 9.2534e-06
|  | 721 | 9.2633e-06
|  | 721.1 | 9.2732e-06
|  | 721.2 | 9.2831e-06
|  | 721.3 | 9.2928e-06
|  | 721.4 | 9.3019e-06
|  | 721.5 | 9.3114e-06
|  | 721.6 | 9.3209e-06
|  | 721.7 | 9.3292e-06
|  | 721.8 | 9.338e-06
|  | 721.9 | 9.3469e-06
|  | 722 | 9.3556e-06
|  | 722.1 | 9.3644e-06
|  | 722.2 | 9.3737e-06
|  | 722.3 | 9.3827e-06
|  | 722.4 | 9.3917e-06
|  | 722.5 | 9.4008e-06
|  | 722.6 | 9.4098e-06
|  | 722.7 | 9.4169e-06
|  | 722.8 | 9.425e-06
|  | 722.9 | 9.4317e-06
|  | 723 | 9.4387e-06
|  | 723.1 | 9.4458e-06
|  | 723.2 | 9.4503e-06
|  | 723.3 | 9.4562e-06
|  | 723.4 | 9.4611e-06
|  | 723.5 | 9.4662e-06
|  | 723.6 | 9.4712e-06
|  | 723.7 | 9.4764e-06
|  | 723.8 | 9.4815e-06
|  | 723.9 | 9.4877e-06
|  | 724 | 9.4934e-06
|  | 724.1 | 9.4992e-06
|  | 724.2 | 9.5063e-06
|  | 724.3 | 9.5128e-06
|  | 724.4 | 9.5193e-06
|  | 724.5 | 9.526e-06
|  | 724.6 | 9.5326e-06
|  | 724.7 | 9.5383e-06
|  | 724.8 | 9.5445e-06
|  | 724.9 | 9.5495e-06
|  | 725 | 9.555e-06
|  | 725.1 | 9.5604e-06
|  | 725.2 | 9.5653e-06
|  | 725.3 | 9.5703e-06
|  | 725.4 | 9.5764e-06
|  | 725.5 | 9.582e-06
|  | 725.6 | 9.5876e-06
|  | 725.7 | 9.5941e-06
|  | 725.8 | 9.6003e-06
|  | 725.9 | 9.6074e-06
|  | 726 | 9.6142e-06
|  | 726.1 | 9.6209e-06
|  | 726.2 | 9.6282e-06
|  | 726.3 | 9.6353e-06
|  | 726.4 | 9.6419e-06
|  | 726.5 | 9.6487e-06
|  | 726.6 | 9.6555e-06
|  | 726.7 | 9.6596e-06
|  | 726.8 | 9.6652e-06
|  | 726.9 | 9.6701e-06
|  | 727 | 9.6748e-06
|  | 727.1 | 9.6796e-06
|  | 727.2 | 9.6848e-06
|  | 727.3 | 9.6897e-06
|  | 727.4 | 9.6955e-06
|  | 727.5 | 9.7009e-06
|  | 727.6 | 9.7063e-06
|  | 727.7 | 9.7131e-06
|  | 727.8 | 9.7192e-06
|  | 727.9 | 9.7256e-06
|  | 728 | 9.732e-06
|  | 728.1 | 9.7385e-06
|  | 728.2 | 9.7439e-06
|  | 728.3 | 9.75e-06
|  | 728.4 | 9.7551e-06
|  | 728.5 | 9.7605e-06
|  | 728.6 | 9.7659e-06
|  | 728.7 | 9.77e-06
|  | 728.8 | 9.7747e-06
|  | 728.9 | 9.7788e-06
|  | 729 | 9.783e-06
|  | 729.1 | 9.7872e-06
|  | 729.2 | 9.7917e-06
|  | 729.3 | 9.796e-06
|  | 729.4 | 9.8015e-06
|  | 729.5 | 9.8065e-06
|  | 729.6 | 9.8115e-06
|  | 729.7 | 9.8165e-06
|  | 729.8 | 9.8217e-06
|  | 729.9 | 9.8267e-06
|  | 730 | 9.8316e-06
|  | 730.1 | 9.8365e-06
|  | 730.2 | 9.8405e-06
|  | 730.3 | 9.8451e-06
|  | 730.4 | 9.8483e-06
|  | 730.5 | 9.8519e-06
|  | 730.6 | 9.8556e-06
|  | 730.7 | 9.8577e-06
|  | 730.8 | 9.8606e-06
|  | 730.9 | 9.8635e-06
|  | 731 | 9.8662e-06
|  | 731.1 | 9.869e-06
|  | 731.2 | 9.8737e-06
|  | 731.3 | 9.8773e-06
|  | 731.4 | 9.8816e-06
|  | 731.5 | 9.8859e-06
|  | 731.6 | 9.8901e-06
|  | 731.7 | 9.8949e-06
|  | 731.8 | 9.8994e-06
|  | 731.9 | 9.9041e-06
|  | 732 | 9.9088e-06
|  | 732.1 | 9.9135e-06
|  | 732.2 | 9.918e-06
|  | 732.3 | 9.9226e-06
|  | 732.4 | 9.9276e-06
|  | 732.5 | 9.9324e-06
|  | 732.6 | 9.9371e-06
|  | 732.7 | 9.9422e-06
|  | 732.8 | 9.9472e-06
|  | 732.9 | 9.9523e-06
|  | 733 | 9.9574e-06
|  | 733.1 | 9.9625e-06
|  | 733.2 | 9.9689e-06
|  | 733.3 | 9.9745e-06
|  | 733.4 | 9.9813e-06
|  | 733.5 | 9.9877e-06
|  | 733.6 | 9.994e-06
|  | 733.7 | 1.0001e-05
|  | 733.8 | 1.0007e-05
|  | 733.9 | 1.0013e-05
|  | 734 | 1.0019e-05
|  | 734.1 | 1.0025e-05
|  | 734.2 | 1.0031e-05
|  | 734.3 | 1.0037e-05
|  | 734.4 | 1.0042e-05
|  | 734.5 | 1.0048e-05
|  | 734.6 | 1.0054e-05
|  | 734.7 | 1.006e-05
|  | 734.8 | 1.0066e-05
|  | 734.9 | 1.0071e-05
|  | 735 | 1.0076e-05
|  | 735.1 | 1.0082e-05
|  | 735.2 | 1.0087e-05
|  | 735.3 | 1.0092e-05
|  | 735.4 | 1.0098e-05
|  | 735.5 | 1.0103e-05
|  | 735.6 | 1.0108e-05
|  | 735.7 | 1.0113e-05
|  | 735.8 | 1.0118e-05
|  | 735.9 | 1.0123e-05
|  | 736 | 1.0129e-05
|  | 736.1 | 1.0135e-05
|  | 736.2 | 1.014e-05
|  | 736.3 | 1.0146e-05
|  | 736.4 | 1.0153e-05
|  | 736.5 | 1.0159e-05
|  | 736.6 | 1.0166e-05
|  | 736.7 | 1.0173e-05
|  | 736.8 | 1.018e-05
|  | 736.9 | 1.0187e-05
|  | 737 | 1.0194e-05
|  | 737.1 | 1.02e-05
|  | 737.2 | 1.0207e-05
|  | 737.3 | 1.0213e-05
|  | 737.4 | 1.0221e-05
|  | 737.5 | 1.0228e-05
|  | 737.6 | 1.0236e-05
|  | 737.7 | 1.0244e-05
|  | 737.8 | 1.0252e-05
|  | 737.9 | 1.026e-05
|  | 738 | 1.0268e-05
|  | 738.1 | 1.0274e-05
|  | 738.2 | 1.0282e-05
|  | 738.3 | 1.0289e-05
|  | 738.4 | 1.0296e-05
|  | 738.5 | 1.0303e-05
|  | 738.6 | 1.0309e-05
|  | 738.7 | 1.0316e-05
|  | 738.8 | 1.0323e-05
|  | 738.9 | 1.0329e-05
|  | 739 | 1.0335e-05
|  | 739.1 | 1.034e-05
|  | 739.2 | 1.0345e-05
|  | 739.3 | 1.0351e-05
|  | 739.4 | 1.0355e-05
|  | 739.5 | 1.036e-05
|  | 739.6 | 1.0365e-05
|  | 739.7 | 1.037e-05
|  | 739.8 | 1.0375e-05
|  | 739.9 | 1.038e-05
|  | 740 | 1.0386e-05
|  | 740.1 | 1.0391e-05
|  | 740.2 | 1.0396e-05
|  | 740.3 | 1.04e-05
|  | 740.4 | 1.0405e-05
|  | 740.5 | 1.041e-05
|  | 740.6 | 1.0414e-05
|  | 740.7 | 1.0419e-05
|  | 740.8 | 1.0424e-05
|  | 740.9 | 1.0428e-05
|  | 741 | 1.0433e-05
|  | 741.1 | 1.044e-05
|  | 741.2 | 1.0446e-05
|  | 741.3 | 1.0452e-05
|  | 741.4 | 1.0459e-05
|  | 741.5 | 1.0466e-05
|  | 741.6 | 1.0473e-05
|  | 741.7 | 1.048e-05
|  | 741.8 | 1.0488e-05
|  | 741.9 | 1.0495e-05
|  | 742 | 1.0503e-05
|  | 742.1 | 1.0512e-05
|  | 742.2 | 1.052e-05
|  | 742.3 | 1.0529e-05
|  | 742.4 | 1.0538e-05
|  | 742.5 | 1.0546e-05
|  | 742.6 | 1.0555e-05
|  | 742.7 | 1.0564e-05
|  | 742.8 | 1.0572e-05
|  | 742.9 | 1.058e-05
|  | 743 | 1.0588e-05
|  | 743.1 | 1.0597e-05
|  | 743.2 | 1.0605e-05
|  | 743.3 | 1.0613e-05
|  | 743.4 | 1.0622e-05
|  | 743.5 | 1.0631e-05
|  | 743.6 | 1.0639e-05
|  | 743.7 | 1.0648e-05
|  | 743.8 | 1.0655e-05
|  | 743.9 | 1.0662e-05
|  | 744 | 1.0668e-05
|  | 744.1 | 1.0674e-05
|  | 744.2 | 1.068e-05
|  | 744.3 | 1.0686e-05
|  | 744.4 | 1.0691e-05
|  | 744.5 | 1.0697e-05
|  | 744.6 | 1.0703e-05
|  | 744.7 | 1.0709e-05
|  | 744.8 | 1.0714e-05
|  | 744.9 | 1.072e-05
|  | 745 | 1.0725e-05
|  | 745.1 | 1.073e-05
|  | 745.2 | 1.0735e-05
|  | 745.3 | 1.0741e-05
|  | 745.4 | 1.0746e-05
|  | 745.5 | 1.0753e-05
|  | 745.6 | 1.076e-05
|  | 745.7 | 1.0767e-05
|  | 745.8 | 1.0774e-05
|  | 745.9 | 1.0781e-05
|  | 746 | 1.0788e-05
|  | 746.1 | 1.0795e-05
|  | 746.2 | 1.0802e-05
|  | 746.3 | 1.081e-05
|  | 746.4 | 1.0817e-05
|  | 746.5 | 1.0826e-05
|  | 746.6 | 1.0834e-05
|  | 746.7 | 1.0843e-05
|  | 746.8 | 1.0852e-05
|  | 746.9 | 1.0861e-05
|  | 747 | 1.0871e-05
|  | 747.1 | 1.0881e-05
|  | 747.2 | 1.089e-05
|  | 747.3 | 1.09e-05
|  | 747.4 | 1.091e-05
|  | 747.5 | 1.0919e-05
|  | 747.6 | 1.0928e-05
|  | 747.7 | 1.0938e-05
|  | 747.8 | 1.0948e-05
|  | 747.9 | 1.0958e-05
|  | 748 | 1.0967e-05
|  | 748.1 | 1.0977e-05
|  | 748.2 | 1.0984e-05
|  | 748.3 | 1.0993e-05
|  | 748.4 | 1.1001e-05
|  | 748.5 | 1.1008e-05
|  | 748.6 | 1.1016e-05
|  | 748.7 | 1.1024e-05
|  | 748.8 | 1.1031e-05
|  | 748.9 | 1.1039e-05
|  | 749 | 1.1047e-05
|  | 749.1 | 1.1055e-05
|  | 749.2 | 1.1063e-05
|  | 749.3 | 1.1071e-05
|  | 749.4 | 1.1078e-05
|  | 749.5 | 1.1085e-05
|  | 749.6 | 1.1092e-05
|  | 749.7 | 1.1097e-05
|  | 749.8 | 1.1103e-05
|  | 749.9 | 1.111e-05
|  | 750 | 1.1116e-05
|  | 750.1 | 1.1123e-05
|  | 750.2 | 1.113e-05
|  | 750.3 | 1.1137e-05
|  | 750.4 | 1.1144e-05
|  | 750.5 | 1.1151e-05
|  | 750.6 | 1.1157e-05
|  | 750.7 | 1.1164e-05
|  | 750.8 | 1.1171e-05
|  | 750.9 | 1.118e-05
|  | 751 | 1.1188e-05
|  | 751.1 | 1.1196e-05
|  | 751.2 | 1.1206e-05
|  | 751.3 | 1.1215e-05
|  | 751.4 | 1.1224e-05
|  | 751.5 | 1.1233e-05
|  | 751.6 | 1.1241e-05
|  | 751.7 | 1.125e-05
|  | 751.8 | 1.1259e-05
|  | 751.9 | 1.1267e-05
|  | 752 | 1.1275e-05
|  | 752.1 | 1.1283e-05
|  | 752.2 | 1.1291e-05
|  | 752.3 | 1.13e-05
|  | 752.4 | 1.1307e-05
|  | 752.5 | 1.1316e-05
|  | 752.6 | 1.1323e-05
|  | 752.7 | 1.1331e-05
|  | 752.8 | 1.1339e-05
|  | 752.9 | 1.1346e-05
|  | 753 | 1.1354e-05
|  | 753.1 | 1.1362e-05
|  | 753.2 | 1.137e-05
|  | 753.3 | 1.1378e-05
|  | 753.4 | 1.1385e-05
|  | 753.5 | 1.1393e-05
|  | 753.6 | 1.1398e-05
|  | 753.7 | 1.1405e-05
|  | 753.8 | 1.1411e-05
|  | 753.9 | 1.1417e-05
|  | 754 | 1.1424e-05
|  | 754.1 | 1.1431e-05
|  | 754.2 | 1.1438e-05
|  | 754.3 | 1.1444e-05
|  | 754.4 | 1.1451e-05
|  | 754.5 | 1.1458e-05
|  | 754.6 | 1.1467e-05
|  | 754.7 | 1.1475e-05
|  | 754.8 | 1.1485e-05
|  | 754.9 | 1.1494e-05
|  | 755 | 1.1503e-05
|  | 755.1 | 1.1514e-05
|  | 755.2 | 1.1524e-05
|  | 755.3 | 1.1535e-05
|  | 755.4 | 1.1545e-05
|  | 755.5 | 1.1556e-05
|  | 755.6 | 1.1567e-05
|  | 755.7 | 1.1577e-05
|  | 755.8 | 1.1587e-05
|  | 755.9 | 1.1598e-05
|  | 756 | 1.1607e-05
|  | 756.1 | 1.1617e-05
|  | 756.2 | 1.1627e-05
|  | 756.3 | 1.1637e-05
|  | 756.4 | 1.1646e-05
|  | 756.5 | 1.1654e-05
|  | 756.6 | 1.1663e-05
|  | 756.7 | 1.1672e-05
|  | 756.8 | 1.1681e-05
|  | 756.9 | 1.169e-05
|  | 757 | 1.17e-05
|  | 757.1 | 1.1709e-05
|  | 757.2 | 1.1718e-05
|  | 757.3 | 1.1727e-05
|  | 757.4 | 1.1736e-05
|  | 757.5 | 1.1744e-05
|  | 757.6 | 1.1753e-05
|  | 757.7 | 1.1761e-05
|  | 757.8 | 1.1769e-05
|  | 757.9 | 1.1778e-05
|  | 758 | 1.1785e-05
|  | 758.1 | 1.1793e-05
|  | 758.2 | 1.18e-05
|  | 758.3 | 1.1807e-05
|  | 758.4 | 1.1815e-05
|  | 758.5 | 1.1823e-05
|  | 758.6 | 1.183e-05
|  | 758.7 | 1.184e-05
|  | 758.8 | 1.1849e-05
|  | 758.9 | 1.1859e-05
|  | 759 | 1.1869e-05
|  | 759.1 | 1.1879e-05
|  | 759.2 | 1.189e-05
|  | 759.3 | 1.1901e-05
|  | 759.4 | 1.1911e-05
|  | 759.5 | 1.1922e-05
|  | 759.6 | 1.1933e-05
|  | 759.7 | 1.1942e-05
|  | 759.8 | 1.1952e-05
|  | 759.9 | 1.1963e-05
|  | 760 | 1.1973e-05
|  | 760.1 | 1.1983e-05
|  | 760.2 | 1.1992e-05
|  | 760.3 | 1.2002e-05
|  | 760.4 | 1.2012e-05
|  | 760.5 | 1.2022e-05
|  | 760.6 | 1.2032e-05
|  | 760.7 | 1.2042e-05
|  | 760.8 | 1.2053e-05
|  | 760.9 | 1.2064e-05
|  | 761 | 1.2075e-05
|  | 761.1 | 1.2085e-05
|  | 761.2 | 1.2096e-05
|  | 761.3 | 1.2106e-05
|  | 761.4 | 1.2115e-05
|  | 761.5 | 1.2125e-05
|  | 761.6 | 1.2133e-05
|  | 761.7 | 1.2142e-05
|  | 761.8 | 1.2151e-05
|  | 761.9 | 1.2158e-05
|  | 762 | 1.2166e-05
|  | 762.1 | 1.2173e-05
|  | 762.2 | 1.2181e-05
|  | 762.3 | 1.2189e-05
|  | 762.4 | 1.2197e-05
|  | 762.5 | 1.2205e-05
|  | 762.6 | 1.2214e-05
|  | 762.7 | 1.2223e-05
|  | 762.8 | 1.2231e-05
|  | 762.9 | 1.224e-05
|  | 763 | 1.2248e-05
|  | 763.1 | 1.2256e-05
|  | 763.2 | 1.2264e-05
|  | 763.3 | 1.2272e-05
|  | 763.4 | 1.228e-05
|  | 763.5 | 1.2288e-05
|  | 763.6 | 1.2297e-05
|  | 763.7 | 1.2305e-05
|  | 763.8 | 1.2315e-05
|  | 763.9 | 1.2324e-05
|  | 764 | 1.2334e-05
|  | 764.1 | 1.2344e-05
|  | 764.2 | 1.2353e-05
|  | 764.3 | 1.2365e-05
|  | 764.4 | 1.2375e-05
|  | 764.5 | 1.2385e-05
|  | 764.6 | 1.2396e-05
|  | 764.7 | 1.2407e-05
|  | 764.8 | 1.2418e-05
|  | 764.9 | 1.2429e-05
|  | 765 | 1.2439e-05
|  | 765.1 | 1.245e-05
|  | 765.2 | 1.2459e-05
|  | 765.3 | 1.2469e-05
|  | 765.4 | 1.2479e-05
|  | 765.5 | 1.2489e-05
|  | 765.6 | 1.2499e-05
|  | 765.7 | 1.2509e-05
|  | 765.8 | 1.2519e-05
|  | 765.9 | 1.2528e-05
|  | 766 | 1.2539e-05
|  | 766.1 | 1.2549e-05
|  | 766.2 | 1.2558e-05
|  | 766.3 | 1.2568e-05
|  | 766.4 | 1.2576e-05
|  | 766.5 | 1.2584e-05
|  | 766.6 | 1.2592e-05
|  | 766.7 | 1.2601e-05
|  | 766.8 | 1.2608e-05
|  | 766.9 | 1.2616e-05
|  | 767 | 1.2624e-05
|  | 767.1 | 1.2632e-05
|  | 767.2 | 1.2642e-05
|  | 767.3 | 1.265e-05
|  | 767.4 | 1.266e-05
|  | 767.5 | 1.2669e-05
|  | 767.6 | 1.268e-05
|  | 767.7 | 1.269e-05
|  | 767.8 | 1.2699e-05
|  | 767.9 | 1.2711e-05
|  | 768 | 1.2722e-05
|  | 768.1 | 1.2733e-05
|  | 768.2 | 1.2744e-05
|  | 768.3 | 1.2756e-05
|  | 768.4 | 1.2768e-05
|  | 768.5 | 1.2779e-05
|  | 768.6 | 1.2792e-05
|  | 768.7 | 1.2805e-05
|  | 768.8 | 1.2817e-05
|  | 768.9 | 1.2831e-05
|  | 769 | 1.2844e-05
|  | 769.1 | 1.2858e-05
|  | 769.2 | 1.2871e-05
|  | 769.3 | 1.2883e-05
|  | 769.4 | 1.2895e-05
|  | 769.5 | 1.2908e-05
|  | 769.6 | 1.2918e-05
|  | 769.7 | 1.293e-05
|  | 769.8 | 1.2941e-05
|  | 769.9 | 1.2952e-05
|  | 770 | 1.2963e-05
|  | 770.1 | 1.2971e-05
|  | 770.2 | 1.2982e-05
|  | 770.3 | 1.2989e-05
|  | 770.4 | 1.2998e-05
|  | 770.5 | 1.3006e-05
|  | 770.6 | 1.3014e-05
|  | 770.7 | 1.3023e-05
|  | 770.8 | 1.3032e-05
|  | 770.9 | 1.3041e-05
|  | 771 | 1.3049e-05
|  | 771.1 | 1.3058e-05
|  | 771.2 | 1.3067e-05
|  | 771.3 | 1.3075e-05
|  | 771.4 | 1.3084e-05
|  | 771.5 | 1.3091e-05
|  | 771.6 | 1.3099e-05
|  | 771.7 | 1.3109e-05
|  | 771.8 | 1.3117e-05
|  | 771.9 | 1.3126e-05
|  | 772 | 1.3137e-05
|  | 772.1 | 1.3147e-05
|  | 772.2 | 1.316e-05
|  | 772.3 | 1.3171e-05
|  | 772.4 | 1.3182e-05
|  | 772.5 | 1.3194e-05
|  | 772.6 | 1.3206e-05
|  | 772.7 | 1.3219e-05
|  | 772.8 | 1.3231e-05
|  | 772.9 | 1.3245e-05
|  | 773 | 1.3258e-05
|  | 773.1 | 1.3272e-05
|  | 773.2 | 1.3284e-05
|  | 773.3 | 1.3297e-05
|  | 773.4 | 1.3308e-05
|  | 773.5 | 1.332e-05
|  | 773.6 | 1.3332e-05
|  | 773.7 | 1.3343e-05
|  | 773.8 | 1.3354e-05
|  | 773.9 | 1.3363e-05
|  | 774 | 1.3374e-05
|  | 774.1 | 1.3383e-05
|  | 774.2 | 1.3393e-05
|  | 774.3 | 1.3402e-05
|  | 774.4 | 1.3411e-05
|  | 774.5 | 1.342e-05
|  | 774.6 | 1.3429e-05
|  | 774.7 | 1.3438e-05
|  | 774.8 | 1.3447e-05
|  | 774.9 | 1.3456e-05
|  | 775 | 1.3465e-05
|  | 775.1 | 1.3477e-05
|  | 775.2 | 1.3487e-05
|  | 775.3 | 1.3498e-05
|  | 775.4 | 1.351e-05
|  | 775.5 | 1.3521e-05
|  | 775.6 | 1.3531e-05
|  | 775.7 | 1.3542e-05
|  | 775.8 | 1.3555e-05
|  | 775.9 | 1.3566e-05
|  | 776 | 1.3579e-05
|  | 776.1 | 1.3592e-05
|  | 776.2 | 1.3604e-05
|  | 776.3 | 1.3618e-05
|  | 776.4 | 1.3631e-05
|  | 776.5 | 1.3645e-05
|  | 776.6 | 1.3659e-05
|  | 776.7 | 1.3673e-05
|  | 776.8 | 1.3689e-05
|  | 776.9 | 1.3704e-05
|  | 777 | 1.3719e-05
|  | 777.1 | 1.3734e-05
|  | 777.2 | 1.3748e-05
|  | 777.3 | 1.3762e-05
|  | 777.4 | 1.3777e-05
|  | 777.5 | 1.3787e-05
|  | 777.6 | 1.3799e-05
|  | 777.7 | 1.3807e-05
|  | 777.8 | 1.3817e-05
|  | 777.9 | 1.3827e-05
|  | 778 | 1.3833e-05
|  | 778.1 | 1.3841e-05
|  | 778.2 | 1.385e-05
|  | 778.3 | 1.3858e-05
|  | 778.4 | 1.3867e-05
|  | 778.5 | 1.3876e-05
|  | 778.6 | 1.3886e-05
|  | 778.7 | 1.3896e-05
|  | 778.8 | 1.3906e-05
|  | 778.9 | 1.3919e-05
|  | 779 | 1.393e-05
|  | 779.1 | 1.3942e-05
|  | 779.2 | 1.3954e-05
|  | 779.3 | 1.3966e-05
|  | 779.4 | 1.3977e-05
|  | 779.5 | 1.3989e-05
|  | 779.6 | 1.4e-05
|  | 779.7 | 1.4012e-05
|  | 779.8 | 1.4023e-05
|  | 779.9 | 1.4036e-05
|  | 780 | 1.4048e-05
|  | 780.1 | 1.4063e-05
|  | 780.2 | 1.4077e-05
|  | 780.3 | 1.409e-05
|  | 780.4 | 1.4104e-05
|  | 780.5 | 1.4119e-05
|  | 780.6 | 1.4132e-05
|  | 780.7 | 1.4145e-05
|  | 780.8 | 1.4157e-05
|  | 780.9 | 1.417e-05
|  | 781 | 1.4183e-05
|  | 781.1 | 1.4192e-05
|  | 781.2 | 1.4203e-05
|  | 781.3 | 1.4214e-05
|  | 781.4 | 1.4225e-05
|  | 781.5 | 1.4237e-05
|  | 781.6 | 1.4248e-05
|  | 781.7 | 1.4259e-05
|  | 781.8 | 1.4274e-05
|  | 781.9 | 1.4286e-05
|  | 782 | 1.4301e-05
|  | 782.1 | 1.4315e-05
|  | 782.2 | 1.4329e-05
|  | 782.3 | 1.4343e-05
|  | 782.4 | 1.4358e-05
|  | 782.5 | 1.4372e-05
|  | 782.6 | 1.4387e-05
|  | 782.7 | 1.4399e-05
|  | 782.8 | 1.4413e-05
|  | 782.9 | 1.4426e-05
|  | 783 | 1.4437e-05
|  | 783.1 | 1.4449e-05
|  | 783.2 | 1.4462e-05
|  | 783.3 | 1.4475e-05
|  | 783.4 | 1.4486e-05
|  | 783.5 | 1.4498e-05
|  | 783.6 | 1.4511e-05
|  | 783.7 | 1.4521e-05
|  | 783.8 | 1.4532e-05
|  | 783.9 | 1.4542e-05
|  | 784 | 1.4552e-05
|  | 784.1 | 1.4563e-05
|  | 784.2 | 1.4573e-05
|  | 784.3 | 1.4583e-05
|  | 784.4 | 1.4594e-05
|  | 784.5 | 1.4604e-05
|  | 784.6 | 1.4617e-05
|  | 784.7 | 1.4629e-05
|  | 784.8 | 1.464e-05
|  | 784.9 | 1.4654e-05
|  | 785 | 1.4667e-05
|  | 785.1 | 1.4681e-05
|  | 785.2 | 1.4695e-05
|  | 785.3 | 1.4709e-05
|  | 785.4 | 1.4724e-05
|  | 785.5 | 1.4738e-05
|  | 785.6 | 1.4756e-05
|  | 785.7 | 1.4772e-05
|  | 785.8 | 1.4789e-05
|  | 785.9 | 1.4805e-05
|  | 786 | 1.4821e-05
|  | 786.1 | 1.4837e-05
|  | 786.2 | 1.4853e-05
|  | 786.3 | 1.4869e-05
|  | 786.4 | 1.4884e-05
|  | 786.5 | 1.4898e-05
|  | 786.6 | 1.4913e-05
|  | 786.7 | 1.4927e-05
|  | 786.8 | 1.4941e-05
|  | 786.9 | 1.4955e-05
|  | 787 | 1.4968e-05
|  | 787.1 | 1.4981e-05
|  | 787.2 | 1.4994e-05
|  | 787.3 | 1.5006e-05
|  | 787.4 | 1.5019e-05
|  | 787.5 | 1.5031e-05
|  | 787.6 | 1.5043e-05
|  | 787.7 | 1.5055e-05
|  | 787.8 | 1.5068e-05
|  | 787.9 | 1.5083e-05
|  | 788 | 1.5098e-05
|  | 788.1 | 1.5112e-05
|  | 788.2 | 1.5126e-05
|  | 788.3 | 1.5141e-05
|  | 788.4 | 1.5155e-05
|  | 788.5 | 1.517e-05
|  | 788.6 | 1.5185e-05
|  | 788.7 | 1.52e-05
|  | 788.8 | 1.5215e-05
|  | 788.9 | 1.5229e-05
|  | 789 | 1.5243e-05
|  | 789.1 | 1.5256e-05
|  | 789.2 | 1.527e-05
|  | 789.3 | 1.5283e-05
|  | 789.4 | 1.5295e-05
|  | 789.5 | 1.5308e-05
|  | 789.6 | 1.532e-05
|  | 789.7 | 1.5332e-05
|  | 789.8 | 1.5345e-05
|  | 789.9 | 1.5357e-05
|  | 790 | 1.537e-05
|  | 790.1 | 1.5385e-05
|  | 790.2 | 1.5398e-05
|  | 790.3 | 1.5412e-05
|  | 790.4 | 1.5425e-05
|  | 790.5 | 1.5436e-05
|  | 790.6 | 1.5447e-05
|  | 790.7 | 1.5459e-05
|  | 790.8 | 1.547e-05
|  | 790.9 | 1.5482e-05
|  | 791 | 1.5494e-05
|  | 791.1 | 1.5505e-05
|  | 791.2 | 1.5518e-05
|  | 791.3 | 1.5531e-05
|  | 791.4 | 1.5543e-05
|  | 791.5 | 1.5559e-05
|  | 791.6 | 1.5573e-05
|  | 791.7 | 1.5591e-05
|  | 791.8 | 1.5608e-05
|  | 791.9 | 1.5626e-05
|  | 792 | 1.5644e-05
|  | 792.1 | 1.5662e-05
|  | 792.2 | 1.5678e-05
|  | 792.3 | 1.5696e-05
|  | 792.4 | 1.5709e-05
|  | 792.5 | 1.5724e-05
|  | 792.6 | 1.5739e-05
|  | 792.7 | 1.5753e-05
|  | 792.8 | 1.5767e-05
|  | 792.9 | 1.5781e-05
|  | 793 | 1.5795e-05
|  | 793.1 | 1.581e-05
|  | 793.2 | 1.5824e-05
|  | 793.3 | 1.5839e-05
|  | 793.4 | 1.5852e-05
|  | 793.5 | 1.5867e-05
|  | 793.6 | 1.5882e-05
|  | 793.7 | 1.5896e-05
|  | 793.8 | 1.5911e-05
|  | 793.9 | 1.5926e-05
|  | 794 | 1.5941e-05
|  | 794.1 | 1.5958e-05
|  | 794.2 | 1.5975e-05
|  | 794.3 | 1.599e-05
|  | 794.4 | 1.6006e-05
|  | 794.5 | 1.6022e-05
|  | 794.6 | 1.6037e-05
|  | 794.7 | 1.6052e-05
|  | 794.8 | 1.6065e-05
|  | 794.9 | 1.608e-05
|  | 795 | 1.6092e-05
|  | 795.1 | 1.6105e-05
|  | 795.2 | 1.6116e-05
|  | 795.3 | 1.6128e-05
|  | 795.4 | 1.614e-05
|  | 795.5 | 1.615e-05
|  | 795.6 | 1.6161e-05
|  | 795.7 | 1.6172e-05
|  | 795.8 | 1.6183e-05
|  | 795.9 | 1.6197e-05
|  | 796 | 1.621e-05
|  | 796.1 | 1.6223e-05
|  | 796.2 | 1.624e-05
|  | 796.3 | 1.6255e-05
|  | 796.4 | 1.6273e-05
|  | 796.5 | 1.6289e-05
|  | 796.6 | 1.6304e-05
|  | 796.7 | 1.632e-05
|  | 796.8 | 1.6336e-05
|  | 796.9 | 1.6348e-05
|  | 797 | 1.6362e-05
|  | 797.1 | 1.6373e-05
|  | 797.2 | 1.6385e-05
|  | 797.3 | 1.6397e-05
|  | 797.4 | 1.6409e-05
|  | 797.5 | 1.642e-05
|  | 797.6 | 1.6434e-05
|  | 797.7 | 1.6446e-05
|  | 797.8 | 1.6461e-05
|  | 797.9 | 1.6474e-05
|  | 798 | 1.6491e-05
|  | 798.1 | 1.6507e-05
|  | 798.2 | 1.6523e-05
|  | 798.3 | 1.6541e-05
|  | 798.4 | 1.6558e-05
|  | 798.5 | 1.6577e-05
|  | 798.6 | 1.6595e-05
|  | 798.7 | 1.6612e-05
|  | 798.8 | 1.6631e-05
|  | 798.9 | 1.6649e-05
|  | 799 | 1.6662e-05
|  | 799.1 | 1.6678e-05
|  | 799.2 | 1.6692e-05
|  | 799.3 | 1.6707e-05
|  | 799.4 | 1.6721e-05
|  | 799.5 | 1.6735e-05
|  | 799.6 | 1.6749e-05
|  | 799.7 | 1.6765e-05
|  | 799.8 | 1.678e-05
|  | 799.9 | 1.68e-05
|  | 800 | 1.6818e-05
|  | 800.1 | 1.6836e-05
|  | 800.2 | 1.6856e-05
|  | 800.3 | 1.6875e-05
|  | 800.4 | 1.6894e-05
|  | 800.5 | 1.6914e-05
|  | 800.6 | 1.6932e-05
|  | 800.7 | 1.6951e-05
|  | 800.8 | 1.697e-05
|  | 800.9 | 1.6987e-05
|  | 801 | 1.7005e-05
|  | 801.1 | 1.7019e-05
|  | 801.2 | 1.7035e-05
|  | 801.3 | 1.7049e-05
|  | 801.4 | 1.7062e-05
|  | 801.5 | 1.7076e-05
|  | 801.6 | 1.7095e-05
|  | 801.7 | 1.7111e-05
|  | 801.8 | 1.7133e-05
|  | 801.9 | 1.7152e-05
|  | 802 | 1.7175e-05
|  | 802.1 | 1.7196e-05
|  | 802.2 | 1.7218e-05
|  | 802.3 | 1.724e-05
|  | 802.4 | 1.7262e-05
|  | 802.5 | 1.7282e-05
|  | 802.6 | 1.7303e-05
|  | 802.7 | 1.7323e-05
|  | 802.8 | 1.7344e-05
|  | 802.9 | 1.7365e-05
|  | 803 | 1.7385e-05
|  | 803.1 | 1.7406e-05
|  | 803.2 | 1.7426e-05
|  | 803.3 | 1.7446e-05
|  | 803.4 | 1.7465e-05
|  | 803.5 | 1.7484e-05
|  | 803.6 | 1.7503e-05
|  | 803.7 | 1.7522e-05
|  | 803.8 | 1.754e-05
|  | 803.9 | 1.7558e-05
|  | 804 | 1.7576e-05
|  | 804.1 | 1.7589e-05
|  | 804.2 | 1.7605e-05
|  | 804.3 | 1.7613e-05
|  | 804.4 | 1.7623e-05
|  | 804.5 | 1.7633e-05
|  | 804.6 | 1.763e-05
|  | 804.7 | 1.7632e-05
|  | 804.8 | 1.7629e-05
|  | 804.9 | 1.7627e-05
|  | 805 | 1.7647e-05
|  | 805.1 | 1.7658e-05
|  | 805.2 | 1.767e-05
|  | 805.3 | 1.7831e-05
|  | 805.4 | 1.7914e-05
|  | 805.5 | 1.8252e-05
|  | 805.6 | 1.8492e-05
|  | 805.7 | 1.9231e-05
|  | 805.8 | 1.9833e-05
|  | 805.9 | 2.0435e-05
|  | 806 | 2.1664e-05
|  | 806.1 | 2.2632e-05
|  | 806.2 | 2.3999e-05
|  | 806.3 | 2.5214e-05
|  | 806.4 | 2.69e-05
|  | 806.5 | 2.8474e-05
|  | 806.6 | 3.0048e-05
|  | 806.7 | 3.1086e-05
|  | 806.8 | 3.25e-05
|  | 806.9 | 3.3954e-05
|  | 807 | 3.5241e-05
|  | 807.1 | 3.7545e-05
|  | 807.2 | 3.9565e-05
|  | 807.3 | 4.1586e-05
|  | 807.4 | 4.6928e-05
|  | 807.5 | 5.0536e-05
|  | 807.6 | 6.1638e-05
|  | 807.7 | 6.9407e-05

## Weighting spectra interpolation
wavelength | raw blanked output (averaged) | new wavelength | forecast over 1, 1/2 nm increments | new wavelength | forecast over 1, 0.2 nm increments | new wavelength | forecast over 1, 0.5 nm increments | new wavelength | forecast over 1, 1 nm increments | 29.5059 | RLE is shown in column k
299.9184 | 28.8166 | 300 | 29.586 | 300.1 | 28.6353 | 300.3 | 27.754 | 300 | 29.57 | 1.0022
299.9224 | 28.936 | 300.1 | 28.2157 | 300.3 | 27.8952 | 300.8 | 24.9581 | 301 | 23.6128 | 0.80027
299.9276 | 29.0895 | 300.2 | 28.1043 | 300.5 | 26.7314 | 301.3 | 21.7367 | 302 | 17.9916 | 0.60976
299.9303 | 29.8872 | 300.3 | 27.8952 | 300.7 | 25.2442 | 301.8 | 18.6086 | 303 | 15.2041 | 0.51529
299.9328 | 29.243 | 300.4 | 27.6861 | 300.9 | 24.5413 | 302.3 | 16.6696 | 304 | 13.4501 | 0.45584
299.938 | 29.3966 | 300.5 | 26.9124 | 301.1 | 23.1929 | 302.8 | 15.336 | 305 | 11.5991 | 0.39311
299.9489 | 29.7207 | 300.6 | 25.5957 | 301.3 | 21.5468 | 303.3 | 14.7553 | 306 | 10.3682 | 0.35139
300.0099 | 29.5671 | 300.7 | 25.2442 | 301.5 | 20.4705 | 303.8 | 13.6634 | 307 | 8.9979 | 0.30495
300.0393 | 28.4754 | 300.8 | 24.8927 | 301.7 | 19.3147 | 304.3 | 12.9905 | 308 | 8.134 | 0.27567
300.0518 | 28.6513 | 300.9 | 24.5413 | 301.9 | 17.9392 | 304.8 | 11.8315 | 309 | 7.8623 | 0.26646
300.1516 | 27.8783 | 301 | 24.1898 | 302.1 | 17.1817 | 305.3 | 11.106 | 310 | 7.7372 | 0.26223
300.1585 | 28.083 | 301.1 | 23.304 | 302.3 | 16.683 | 305.8 | 10.6357 | 311 | 7.695 | 0.26079
300.1654 | 28.2877 | 301.2 | 22.0849 | 302.5 | 16.1442 | 306.3 | 10.0166 | 312 | 7.7418 | 0.26238
300.4154 | 27.6509 | 301.3 | 21.5468 | 302.7 | 15.6054 | 306.8 | 9.324 | 313 | 7.8553 | 0.26623
300.5213 | 26.5627 | 301.4 | 21.0086 | 302.9 | 15.0666 | 307.3 | 8.5394 | 314 | 8.0028 | 0.27123
300.5615 | 24.9886 | 301.5 | 20.4705 | 303.1 | 15.0557 | 307.8 | 8.1703 | 315 | 8.1888 | 0.27753
301.094 | 23.9114 | 301.6 | 19.9323 | 303.3 | 14.7502 | 308.3 | 8.01 | 316 | 8.4197 | 0.28536
301.0969 | 22.8163 | 301.7 | 19.5817 | 303.5 | 14.4599 | 308.8 | 7.8956 | 317 | 8.6989 | 0.29482
301.1963 | 21.8852 | 301.8 | 18.43 | 303.7 | 13.9383 | 309.3 | 7.8045 | 318 | 8.98 | 0.30435
301.6055 | 19.9457 | 301.9 | 18.0166 | 303.9 | 13.3925 | 309.8 | 7.7456 | 319 | 9.2921 | 0.31492
301.6176 | 20.3039 | 302 | 17.3709 | 304.1 | 13.4555 | 310.3 | 7.7152 | 320 | 9.6333 | 0.32649
301.7206 | 19.4339 | 302.1 | 17.2218 | 304.3 | 12.9662 | 310.8 | 7.6948 | 321 | 9.9557 | 0.33741
301.8037 | 18.3031 | 302.2 | 16.9524 | 304.5 | 12.5499 | 311.3 | 7.69 | 322 | 10.2776 | 0.34832
302.086 | 16.8919 | 302.3 | 16.683 | 304.7 | 12.1134 | 311.8 | 7.7265 | 323 | 10.5962 | 0.35912
303.0291 | 14.8034 | 302.4 | 16.4136 | 304.9 | 11.5678 | 312.3 | 7.7644 | 324 | 10.9168 | 0.36999
303.0622 | 15.7816 | 302.5 | 16.1442 | 305.1 | 11.0358 | 312.8 | 7.8274 | 325 | 11.2442 | 0.38108
303.4898 | 14.3212 | 302.6 | 15.8748 | 305.3 | 11.2939 | 313.3 | 7.8955 | 326 | 11.5719 | 0.39219
303.9889 | 13.2378 | 302.7 | 15.6054 | 305.5 | 10.9883 | 313.8 | 7.9734 | 327 | 11.8535 | 0.40173
304.0132 | 13.9552 | 302.8 | 15.336 | 305.7 | 10.7532 | 314.3 | 8.048 | 328 | 12.1499 | 0.41178
304.571 | 12.3857 | 302.9 | 15.0666 | 305.9 | 10.5181 | 314.8 | 8.1259 | 329 | 12.453 | 0.42205
305.127 | 10.9924 | 303 | 14.7972 | 306.1 | 10.283 | 315.3 | 8.2634 | 330 | 12.7473 | 0.43202
306.445 | 10.0063 | 303.1 | 15.1613 | 306.3 | 10.0479 | 315.8 | 8.3744 | 331 | 13.0331 | 0.44171
306.8882 | 9.1282 | 303.2 | 14.9558 | 306.5 | 9.7189 | 316.3 | 8.4871 | 332 | 13.3234 | 0.45155
307.3659 | 8.2802 | 303.3 | 14.7502 | 306.7 | 9.5026 | 316.8 | 8.6454 | 333 | 13.6139 | 0.4614
308.0896 | 8.0198 | 303.4 | 14.5446 | 306.9 | 9.1656 | 317.3 | 8.7853 | 334 | 13.9295 | 0.47209
309.1427 | 7.8319 | 303.5 | 14.4841 | 307.1 | 8.7631 | 317.8 | 8.9244 | 335 | 14.2329 | 0.48238
309.8014 | 7.7293 | 303.6 | 14.2112 | 307.3 | 8.3884 | 318.3 | 9.0634 | 336 | 14.5334 | 0.49256
310.7872 | 7.6848 | 303.7 | 13.9383 | 307.5 | 8.4666 | 318.8 | 9.2025 | 337 | 14.9018 | 0.50504
311.5219 | 7.6946 | 303.8 | 13.6654 | 307.7 | 8.255 | 319.3 | 9.4052 | 338 | 15.2483 | 0.51679
312.3189 | 7.7626 | 303.9 | 13.3925 | 307.9 | 8.0796 | 319.8 | 9.5694 | 339 | 15.588 | 0.5283
313.2499 | 7.8817 | 304 | 13.6068 | 308.1 | 8.0411 | 320.3 | 9.7303 | 340 | 15.9467 | 0.54046
314.1821 | 8.0349 | 304.1 | 13.3824 | 308.3 | 8.0191 | 320.8 | 9.8913 | 341 | 16.3049 | 0.5526
314.9564 | 8.1473 | 304.2 | 13.1743 | 308.5 | 7.9697 | 321.3 | 10.0523 | 342 | 16.6699 | 0.56497
316.5838 | 8.5629 | 304.3 | 12.9662 | 308.7 | 7.9203 | 321.8 | 10.2132 | 343 | 17.0424 | 0.57759
319.0337 | 9.2768 | 304.4 | 12.758 | 308.9 | 7.8709 | 322.3 | 10.3742 | 344 | 17.4065 | 0.58993
323.1772 | 10.6736 | 304.5 | 12.5499 | 309.1 | 7.8215 | 322.8 | 10.5351 | 345 | 17.7775 | 0.6025
326.6276 | 11.7629 | 304.6 | 12.3657 | 309.3 | 7.815 | 323.3 | 10.6899 | 346 | 18.1553 | 0.61531
330.075 | 12.7643 | 304.7 | 12.0997 | 309.5 | 7.777 | 323.8 | 10.8509 | 347 | 18.5304 | 0.62802
333.5225 | 13.7657 | 304.8 | 11.8338 | 309.7 | 7.7429 | 324.3 | 11.0148 | 348 | 18.8775 | 0.63979
336.6275 | 14.7342 | 304.9 | 11.5678 | 309.9 | 7.7427 | 324.8 | 11.1786 | 349 | 19.2325 | 0.65182
339.3925 | 15.7446 | 305 | 11.3018 | 310.1 | 7.7323 | 325.3 | 11.3425 | 350 | 19.5915 | 0.66398
342.156 | 16.711 | 305.1 | 11.0358 | 310.3 | 7.7152 | 325.8 | 11.5063 | 351 | 19.9392 | 0.67577
344.9221 | 17.7544 | 305.2 | 11.3409 | 310.5 | 7.698 | 326.3 | 11.6702 | 352 | 20.2458 | 0.68616
347.6879 | 18.7867 | 305.3 | 11.2234 | 310.7 | 7.6808 | 326.8 | 11.7921 | 353 | 20.5848 | 0.69765
351.3113 | 20.0512 | 305.4 | 11.1059 | 310.9 | 7.7024 | 327.3 | 11.9377 | 354 | 20.9749 | 0.71087
353.4188 | 20.6917 | 305.5 | 10.9883 | 311.1 | 7.6944 | 327.8 | 12.0893 | 355 | 21.3518 | 0.72365
355.4675 | 21.5482 | 305.6 | 10.8708 | 311.3 | 7.69 | 328.3 | 12.2408 | 356 | 21.819 | 0.73948
358.476 | 22.9491 | 305.7 | 10.7532 | 311.5 | 7.6857 | 328.8 | 12.3924 | 357 | 22.2727 | 0.75485
361.5334 | 24.4432 | 305.8 | 10.6357 | 311.7 | 7.7264 | 329.3 | 12.5439 | 358 | 22.7263 | 0.77023
363.9497 | 25.2376 | 305.9 | 10.5181 | 311.9 | 7.7323 | 329.8 | 12.6954 | 359 | 23.2193 | 0.78694
365.3384 | 25.926 | 306 | 10.4006 | 312.1 | 7.7426 | 330.3 | 12.8305 | 360 | 23.6998 | 0.80322
368.4055 | 27.0965 | 306.1 | 10.283 | 312.3 | 7.7528 | 330.8 | 12.9749 | 361 | 24.1771 | 0.8194
370.5969 | 27.5653 | 306.2 | 10.1655 | 312.5 | 7.7978 | 331.3 | 13.1201 | 362 | 24.5135 | 0.8308
373.2585 | 28.0527 | 306.3 | 10.0479 | 312.7 | 7.8162 | 331.8 | 13.2653 | 363 | 24.9176 | 0.8445
375.2632 | 28.485 | 306.4 | 9.9304 | 312.9 | 7.838 | 332.3 | 13.4106 | 364 | 25.3406 | 0.85883
376.4637 | 28.7405 | 306.5 | 9.7001 | 313.1 | 7.8598 | 332.8 | 13.5558 | 365 | 25.7347 | 0.87219
377.5284 | 28.8936 | 306.6 | 9.6013 | 313.3 | 7.8972 | 333.3 | 13.701 | 366 | 26.1297 | 0.88557
378.5579 | 29.0347 | 306.7 | 9.5026 | 313.5 | 7.9296 | 333.8 | 13.8718 | 367 | 26.54 | 0.89948
379.6594 | 29.2512 | 306.8 | 9.4038 | 313.7 | 7.9588 | 334.3 | 14.0226 | 368 | 26.9258 | 0.91256
380.856 | 29.3872 | 306.9 | 9.1377 | 313.9 | 7.988 | 334.8 | 14.1728 | 369 | 27.1578 | 0.92042
382.5124 | 29.494 | 307 | 8.9504 | 314.1 | 8.0173 | 335.3 | 14.3231 | 370 | 27.4592 | 0.93063
383.3768 | 29.5229 | 307.1 | 8.7631 | 314.3 | 8.0478 | 335.8 | 14.4733 | 371 | 27.6304 | 0.93644
384.4357 | 29.5055 | 307.2 | 8.5758 | 314.5 | 8.0791 | 336.3 | 14.6235 | 372 | 27.8165 | 0.94274
386.0687 | 29.3624 | 307.3 | 8.3884 | 314.7 | 8.1103 | 336.8 | 14.8357 | 373 | 28.0174 | 0.94955
387.3406 | 29.2487 | 307.4 | 8.5181 | 314.9 | 8.1415 | 337.3 | 15.0127 | 374 | 28.2222 | 0.95649
388.6604 | 29.1118 | 307.5 | 8.4304 | 315.1 | 8.2217 | 337.8 | 15.181 | 375 | 28.4243 | 0.96334
389.5818 | 28.9273 | 307.6 | 8.3427 | 315.3 | 8.2617 | 338.3 | 15.3493 | 376 | 28.6371 | 0.97055
391.707 | 28.4016 | 307.7 | 8.255 | 315.5 | 8.3068 | 338.8 | 15.5176 | 377 | 28.8136 | 0.97654
393.5994 | 27.9477 | 307.8 | 8.1673 | 315.7 | 8.3519 | 339.3 | 15.6916 | 378 | 28.9585 | 0.98145
395.0835 | 27.3013 | 307.9 | 8.0796 | 315.9 | 8.3969 | 339.8 | 15.876 | 379 | 29.1282 | 0.9872
396.5955 | 26.7076 | 308 | 7.9919 | 316.1 | 8.442 | 340.3 | 16.0546 | 380 | 29.2746 | 0.99216
397.0452 | 26.4138 | 308.1 | 8.0685 | 316.3 | 8.4871 | 340.8 | 16.2334 | 381 | 29.3813 | 0.99578
397.7721 | 26.0473 | 308.2 | 8.0438 | 316.5 | 8.5321 | 341.3 | 16.4122 | 382 | 29.4605 | 0.99846
399.0831 | 25.4025 | 308.3 | 8.0191 | 316.7 | 8.6235 | 341.8 | 16.5909 | 383 | 29.5059 | 1
400.1215 | 24.8594 | 308.4 | 7.9944 | 316.9 | 8.674 | 342.3 | 16.7837 | 384 | 29.5052 | 0.99998
401.071 | 24.2931 | 308.5 | 7.9697 | 317.1 | 8.7297 | 342.8 | 16.9704 | 385 | 29.4429 | 0.99786
402.1291 | 23.6569 | 308.6 | 7.945 | 317.3 | 8.7853 | 343.3 | 17.1521 | 386 | 29.3689 | 0.99535
403.0567 | 23.1186 | 308.7 | 7.9203 | 317.5 | 8.8409 | 343.8 | 17.3338 | 387 | 29.2777 | 0.99227
404.1417 | 22.601 | 308.8 | 7.8956 | 317.7 | 8.8965 | 344.3 | 17.5156 | 388 | 29.1791 | 0.98892
405.553 | 22.0743 | 308.9 | 7.8709 | 317.9 | 8.9522 | 344.8 | 17.6973 | 389 | 29.0339 | 0.984
406.1562 | 21.6037 | 309 | 7.8462 | 318.1 | 9.0078 | 345.3 | 17.893 | 390 | 28.8116 | 0.97647
408.5079 | 20.4817 | 309.1 | 7.8215 | 318.3 | 9.0634 | 345.8 | 18.0801 | 391 | 28.574 | 0.96842
410.8561 | 19.2571 | 309.2 | 7.8281 | 318.5 | 9.1191 | 346.3 | 18.2678 | 392 | 28.3349 | 0.96031
412.8691 | 18.2159 | 309.3 | 7.8111 | 318.7 | 9.1747 | 346.8 | 18.4554 | 393 | 28.0913 | 0.95206
414.8792 | 17.0867 | 309.4 | 7.794 | 318.9 | 9.2303 | 347.3 | 18.643 | 394 | 27.7226 | 0.93956
416.8897 | 15.9721 | 309.5 | 7.777 | 319.1 | 9.3344 | 347.8 | 18.8052 | 395 | 27.3585 | 0.92722
418.5602 | 14.8943 | 309.6 | 7.7599 | 319.3 | 9.4084 | 348.3 | 18.9812 | 396 | 26.9438 | 0.91316
420.2273 | 13.7172 | 309.7 | 7.7429 | 319.5 | 9.4728 | 348.8 | 19.1607 | 397 | 26.4577 | 0.89669
421.8981 | 12.6508 | 309.8 | 7.7258 | 319.7 | 9.5372 | 349.3 | 19.3402 | 398 | 25.9394 | 0.87913
423.2279 | 11.5951 | 309.9 | 7.7495 | 319.9 | 9.6016 | 349.8 | 19.5197 | 399 | 25.4359 | 0.86206
424.2186 | 10.6055 | 310 | 7.7409 | 320.1 | 9.666 | 350.3 | 19.6991 | 400 | 24.9104 | 0.84425
426.2252 | 9.3737 | 310.1 | 7.7323 | 320.3 | 9.7303 | 350.8 | 19.8786 | 401 | 24.3342 | 0.82472
427.5525 | 8.2447 | 310.2 | 7.7237 | 320.5 | 9.7947 | 351.3 | 20.0371 | 402 | 23.7382 | 0.80452
429.2225 | 7.1523 | 310.3 | 7.7152 | 320.7 | 9.8591 | 351.8 | 20.1747 | 403 | 23.1762 | 0.78548
430.8934 | 6.0892 | 310.4 | 7.7066 | 320.9 | 9.9235 | 352.3 | 20.3425 | 404 | 22.691 | 0.76903
432.9055 | 5.0186 | 310.5 | 7.698 | 321.1 | 9.9879 | 352.8 | 20.5096 | 405 | 22.2841 | 0.75524
434.5885 | 4.3145 | 310.6 | 7.6894 | 321.3 | 10.0523 | 353.3 | 20.6924 | 406 | 21.7648 | 0.73764
436.3243 | 3.6729 | 310.7 | 7.6808 | 321.5 | 10.1166 | 353.8 | 20.9071 | 407 | 21.2499 | 0.72019
437.9722 | 3.1764 | 310.8 | 7.7008 | 321.7 | 10.181 | 354.3 | 21.0879 | 408 | 20.7289 | 0.70253
440.3331 | 2.4325 | 310.9 | 7.6987 | 321.9 | 10.2454 | 354.8 | 21.2678 | 409 | 20.2043 | 0.68475
441.6531 | 2.1935 | 311 | 7.6965 | 322.1 | 10.3098 | 355.3 | 21.4703 | 410 | 19.7019 | 0.66773
443.7259 | 1.9616 | 311.1 | 7.6944 | 322.3 | 10.3742 | 355.8 | 21.7342 | 411 | 19.1877 | 0.6503
445.7425 | 1.7146 | 311.2 | 7.6922 | 322.5 | 10.4385 | 356.3 | 21.9592 | 412 | 18.6658 | 0.63261
448.4034 | 1.5168 | 311.3 | 7.69 | 322.7 | 10.5029 | 356.8 | 22.1831 | 413 | 18.1219 | 0.61418
450.8953 | 1.3546 | 311.4 | 7.6879 | 322.9 | 10.5673 | 357.3 | 22.407 | 414 | 17.5759 | 0.59567
453.2563 | 1.1387 | 311.5 | 7.6857 | 323.1 | 10.6317 | 357.8 | 22.6309 | 415 | 17.0253 | 0.57701
457.5001 | 1.0062 | 311.6 | 7.7169 | 323.3 | 10.6854 | 358.3 | 22.8658 | 416 | 16.4662 | 0.55806
461.6758 | 0.90844 | 311.7 | 7.7221 | 323.5 | 10.7526 | 358.8 | 23.1265 | 417 | 15.8635 | 0.53764
467.1132 | 0.79333 | 311.8 | 7.7272 | 323.7 | 10.8181 | 359.3 | 23.3657 | 418 | 15.2573 | 0.51709
472.172 | 0.66125 | 311.9 | 7.7323 | 323.9 | 10.8837 | 359.8 | 23.6044 | 419 | 14.5696 | 0.49378
477.2309 | 0.52916 | 312 | 7.7374 | 324.1 | 10.9492 | 360.3 | 23.843 | 420 | 13.8947 | 0.47091
481.7834 | 0.39724 | 312.1 | 7.7426 | 324.3 | 11.0148 | 360.8 | 24.0817 | 421 | 13.2345 | 0.44854
485.8273 | 0.20027 | 312.2 | 7.7477 | 324.5 | 11.0803 | 361.3 | 24.3203 | 422 | 12.5165 | 0.4242
489.8735 | 0.068519 | 312.3 | 7.7528 | 324.7 | 11.1459 | 361.8 | 24.4159 | 423 | 11.7491 | 0.39819
494.4304 | 0.067033 | 312.4 | 7.7836 | 324.9 | 11.2114 | 362.3 | 24.6228 | 424 | 10.8952 | 0.36925
500.5063 | 0.065052 | 312.5 | 7.7945 | 325.1 | 11.2769 | 362.8 | 24.8334 | 425 | 10.1993 | 0.34567
|  | 312.6 | 7.8054 | 325.3 | 11.3425 | 363.3 | 25.044 | 426 | 9.4607 | 0.32064
|  | 312.7 | 7.8162 | 325.5 | 11.408 | 363.8 | 25.2545 | 427 | 8.7104 | 0.29521
|  | 312.8 | 7.8271 | 325.7 | 11.4736 | 364.3 | 25.4685 | 428 | 8.0068 | 0.27136
|  | 312.9 | 7.838 | 325.9 | 11.5391 | 364.8 | 25.6589 | 429 | 7.2965 | 0.24729
|  | 313 | 7.8489 | 326.1 | 11.6047 | 365.3 | 25.8485 | 430 | 6.6583 | 0.22566
|  | 313.1 | 7.8598 | 326.3 | 11.6702 | 365.8 | 26.0463 | 431 | 6.0775 | 0.20598
|  | 313.2 | 7.8706 | 326.5 | 11.7357 | 366.3 | 26.2521 | 432 | 5.513 | 0.18685
|  | 313.3 | 7.9003 | 326.7 | 11.7634 | 366.8 | 26.4577 | 433 | 5.0106 | 0.16982
|  | 313.4 | 7.9149 | 326.9 | 11.8165 | 367.3 | 26.6634 | 434 | 4.5611 | 0.15458
|  | 313.5 | 7.9296 | 327.1 | 11.8771 | 367.8 | 26.8691 | 435 | 4.1745 | 0.14148
|  | 313.6 | 7.9442 | 327.3 | 11.9377 | 368.3 | 27.0269 | 436 | 3.7994 | 0.12877
|  | 313.7 | 7.9588 | 327.5 | 11.9983 | 368.8 | 27.0821 | 437 | 3.4806 | 0.11796
|  | 313.8 | 7.9734 | 327.7 | 12.059 | 369.3 | 27.2378 | 438 | 3.1581 | 0.10703
|  | 313.9 | 7.988 | 327.9 | 12.1196 | 369.8 | 27.3959 | 439 | 2.8492 | 0.096564
|  | 314 | 8.0026 | 328.1 | 12.1802 | 370.3 | 27.554 | 440 | 2.5549 | 0.08659
|  | 314.1 | 8.0173 | 328.3 | 12.2408 | 370.8 | 27.5833 | 441 | 2.3194 | 0.078607
|  | 314.2 | 8.0323 | 328.5 | 12.3014 | 371.3 | 27.6789 | 442 | 2.1746 | 0.073701
|  | 314.3 | 8.0479 | 328.7 | 12.362 | 371.8 | 27.7772 | 443 | 2.0464 | 0.069356
|  | 314.4 | 8.0635 | 328.9 | 12.4227 | 372.3 | 27.8755 | 444 | 1.9219 | 0.065135
|  | 314.5 | 8.0791 | 329.1 | 12.4833 | 372.8 | 27.9737 | 445 | 1.8052 | 0.06118
|  | 314.6 | 8.0947 | 329.3 | 12.5439 | 373.3 | 28.0802 | 446 | 1.7207 | 0.058316
|  | 314.7 | 8.1103 | 329.5 | 12.6045 | 373.8 | 28.1837 | 447 | 1.633 | 0.055344
|  | 314.8 | 8.1259 | 329.7 | 12.6651 | 374.3 | 28.2818 | 448 | 1.5428 | 0.052288
|  | 314.9 | 8.1415 | 329.9 | 12.7258 | 374.8 | 28.38 | 449 | 1.4811 | 0.050195
|  | 315 | 8.1941 | 330.1 | 12.774 | 375.3 | 28.4875 | 450 | 1.4131 | 0.047894
|  | 315.1 | 8.2167 | 330.3 | 12.8296 | 375.8 | 28.5988 | 451 | 1.3295 | 0.04506
|  | 315.2 | 8.2392 | 330.5 | 12.8877 | 376.3 | 28.6984 | 452 | 1.2474 | 0.042275
|  | 315.3 | 8.2617 | 330.7 | 12.9458 | 376.8 | 28.7758 | 453 | 1.1867 | 0.040219
|  | 315.4 | 8.2843 | 330.9 | 13.0039 | 377.3 | 28.866 | 454 | 1.1583 | 0.039258
|  | 315.5 | 8.3068 | 331.1 | 13.062 | 377.8 | 28.9298 | 455 | 1.1105 | 0.037635
|  | 315.6 | 8.3293 | 331.3 | 13.1201 | 378.3 | 28.9996 | 456 | 1.0603 | 0.035934
|  | 315.7 | 8.3519 | 331.5 | 13.1782 | 378.8 | 29.096 | 457 | 1.0101 | 0.034232
|  | 315.8 | 8.3744 | 331.7 | 13.2363 | 379.3 | 29.1805 | 458 | 0.99948 | 0.033874
|  | 315.9 | 8.3969 | 331.9 | 13.2944 | 379.8 | 29.2438 | 459 | 0.9761 | 0.033082
|  | 316 | 8.4195 | 332.1 | 13.3525 | 380.3 | 29.3174 | 460 | 0.94875 | 0.032155
|  | 316.1 | 8.442 | 332.3 | 13.4106 | 380.8 | 29.374 | 461 | 0.92139 | 0.031227
|  | 316.2 | 8.4645 | 332.5 | 13.4687 | 381.3 | 29.4018 | 462 | 0.90344 | 0.030619
|  | 316.3 | 8.4871 | 332.7 | 13.5268 | 381.8 | 29.4438 | 463 | 0.88268 | 0.029915
|  | 316.4 | 8.5096 | 332.9 | 13.5849 | 382.3 | 29.4857 | 464 | 0.86059 | 0.029167
|  | 316.5 | 8.5321 | 333.1 | 13.643 | 382.8 | 29.4991 | 465 | 0.8385 | 0.028418
|  | 316.6 | 8.5906 | 333.3 | 13.701 | 383.3 | 29.5152 | 466 | 0.81641 | 0.027669
|  | 316.7 | 8.6184 | 333.5 | 13.7591 | 383.8 | 29.509 | 467 | 0.79143 | 0.026823
|  | 316.8 | 8.6462 | 333.7 | 13.8462 | 384.3 | 29.5004 | 468 | 0.76411 | 0.025897
|  | 316.9 | 8.674 | 333.9 | 13.9024 | 384.8 | 29.4534 | 469 | 0.74033 | 0.025091
|  | 317 | 8.7018 | 334.1 | 13.9625 | 385.3 | 29.4219 | 470 | 0.7168 | 0.024294
|  | 317.1 | 8.7297 | 334.3 | 14.0226 | 385.8 | 29.3909 | 471 | 0.69328 | 0.023496
|  | 317.2 | 8.7575 | 334.5 | 14.0827 | 386.3 | 29.3418 | 472 | 0.66784 | 0.022634
|  | 317.3 | 8.7853 | 334.7 | 14.1428 | 386.8 | 29.2969 | 473 | 0.63979 | 0.021683
|  | 317.4 | 8.8131 | 334.9 | 14.2029 | 387.3 | 29.2497 | 474 | 0.61352 | 0.020793
|  | 317.5 | 8.8409 | 335.1 | 14.263 | 387.8 | 29.1979 | 475 | 0.58741 | 0.019908
|  | 317.6 | 8.8687 | 335.3 | 14.3231 | 388.3 | 29.1497 | 476 | 0.5613 | 0.019023
|  | 317.7 | 8.8965 | 335.5 | 14.3832 | 388.8 | 29.0623 | 477 | 0.53362 | 0.018085
|  | 317.8 | 8.9244 | 335.7 | 14.4432 | 389.3 | 28.9873 | 478 | 0.50371 | 0.017071
|  | 317.9 | 8.9522 | 335.9 | 14.5033 | 389.8 | 28.8576 | 479 | 0.47604 | 0.016134
|  | 318 | 8.98 | 336.1 | 14.5634 | 390.3 | 28.7389 | 480 | 0.44859 | 0.015204
|  | 318.1 | 9.0078 | 336.3 | 14.6235 | 390.8 | 28.6211 | 481 | 0.42115 | 0.014273
|  | 318.2 | 9.0356 | 336.5 | 14.6836 | 391.3 | 28.5033 | 482 | 0.36735 | 0.01245
|  | 318.3 | 9.0634 | 336.7 | 14.7995 | 391.8 | 28.3841 | 483 | 0.3228 | 0.01094
|  | 318.4 | 9.0913 | 336.9 | 14.878 | 392.3 | 28.262 | 484 | 0.28473 | 0.00965
|  | 318.5 | 9.1191 | 337.1 | 14.9453 | 392.8 | 28.1401 | 485 | 0.24667 | 0.0083599
|  | 318.6 | 9.1469 | 337.3 | 15.0127 | 393.3 | 28.0181 | 486 | 0.2127 | 0.0072087
|  | 318.7 | 9.1747 | 337.5 | 15.08 | 393.8 | 27.7778 | 487 | 0.17439 | 0.0059105
|  | 318.8 | 9.2025 | 337.7 | 15.1473 | 394.3 | 27.6141 | 488 | 0.13376 | 0.0045334
|  | 318.9 | 9.2303 | 337.9 | 15.2146 | 394.8 | 27.4531 | 489 | 0.093131 | 0.0031563
|  | 319 | 9.2581 | 338.1 | 15.282 | 395.3 | 27.2354 | 490 | 0.097957 | 0.0033199
|  | 319.1 | 9.344 | 338.3 | 15.3493 | 395.8 | 27.0262 | 491 | 0.097436 | 0.0033022
|  | 319.2 | 9.3762 | 338.5 | 15.4166 | 396.3 | 26.8193 | 492 | 0.082266 | 0.0027881
|  | 319.3 | 9.4084 | 338.7 | 15.4839 | 396.8 | 26.5645 | 493 | 0.067096 | 0.002274
|  | 319.4 | 9.4406 | 338.9 | 15.5512 | 397.3 | 26.3005 | 494 | 0.055284 | 0.0018737
|  | 319.5 | 9.4728 | 339.1 | 15.6186 | 397.8 | 26.0351 | 495 | 0.064955 | 0.0022014
|  | 319.6 | 9.505 | 339.3 | 15.6859 | 398.3 | 25.7893 | 496 | 0.066521 | 0.0022545
|  | 319.7 | 9.5372 | 339.5 | 15.7704 | 398.8 | 25.5415 | 497 | 0.066195 | 0.0022434
|  | 319.8 | 9.5694 | 339.7 | 15.8401 | 399.3 | 25.2815 | 498 | 0.065869 | 0.0022324
|  | 319.9 | 9.6016 | 339.9 | 15.9116 | 399.8 | 25.0284 | 499 | 0.065543 | 0.0022213
|  | 320 | 9.6338 | 340.1 | 15.9831 | 400.3 | 24.7439 | 500 | 0.065217 | 0.0022103
|  | 320.1 | 9.666 | 340.3 | 16.0546 | 400.8 | 24.4569
|  | 320.2 | 9.6982 | 340.5 | 16.1261 | 401.3 | 24.1552
|  | 320.3 | 9.7303 | 340.7 | 16.1976 | 401.8 | 23.8548
|  | 320.4 | 9.7625 | 340.9 | 16.2691 | 402.3 | 23.561
|  | 320.5 | 9.7947 | 341.1 | 16.3406 | 402.8 | 23.2668
|  | 320.6 | 9.8269 | 341.3 | 16.4122 | 403.3 | 23.0233
|  | 320.7 | 9.8591 | 341.5 | 16.4837 | 403.8 | 22.7639
|  | 320.8 | 9.8913 | 341.7 | 16.5552 | 404.3 | 22.5624
|  | 320.9 | 9.9235 | 341.9 | 16.6267 | 404.8 | 22.3693
|  | 321 | 9.9557 | 342.1 | 16.6982 | 405.3 | 22.1612
|  | 321.1 | 9.9879 | 342.3 | 16.7916 | 405.8 | 21.851
|  | 321.2 | 10.0201 | 342.5 | 16.8614 | 406.3 | 21.6118
|  | 321.3 | 10.0523 | 342.7 | 16.934 | 406.8 | 21.3547
|  | 321.4 | 10.0844 | 342.9 | 17.0067 | 407.3 | 21.094
|  | 321.5 | 10.1166 | 343.1 | 17.0794 | 407.8 | 20.8332
|  | 321.6 | 10.1488 | 343.3 | 17.1521 | 408.3 | 20.5724
|  | 321.7 | 10.181 | 343.5 | 17.2248 | 408.8 | 20.3015
|  | 321.8 | 10.2132 | 343.7 | 17.2975 | 409.3 | 20.0514
|  | 321.9 | 10.2454 | 343.9 | 17.3702 | 409.8 | 19.8018
|  | 322 | 10.2776 | 344.1 | 17.4429 | 410.3 | 19.5522
|  | 322.1 | 10.3098 | 344.3 | 17.5156 | 410.8 | 19.2951
|  | 322.2 | 10.342 | 344.5 | 17.5883 | 411.3 | 19.0295
|  | 322.3 | 10.3742 | 344.7 | 17.661 | 411.8 | 18.7697
|  | 322.4 | 10.4064 | 344.9 | 17.7336 | 412.3 | 18.5099
|  | 322.5 | 10.4385 | 345.1 | 17.8189 | 412.8 | 18.2361
|  | 322.6 | 10.4707 | 345.3 | 17.8925 | 413.3 | 17.9536
|  | 322.7 | 10.5029 | 345.5 | 17.9676 | 413.8 | 17.6838
|  | 322.8 | 10.5351 | 345.7 | 18.0426 | 414.3 | 17.4141
|  | 322.9 | 10.5673 | 345.9 | 18.1177 | 414.8 | 17.1393
|  | 323 | 10.5995 | 346.1 | 18.1927 | 415.3 | 16.8568
|  | 323.1 | 10.6317 | 346.3 | 18.2678 | 415.8 | 16.5778
|  | 323.2 | 10.6543 | 346.5 | 18.3428 | 416.3 | 16.2987
|  | 323.3 | 10.6871 | 346.7 | 18.4178 | 416.8 | 15.9939
|  | 323.4 | 10.7198 | 346.9 | 18.4929 | 417.3 | 15.6732
|  | 323.5 | 10.7526 | 347.1 | 18.5679 | 417.8 | 15.3761
|  | 323.6 | 10.7854 | 347.3 | 18.643 | 418.3 | 15.079
|  | 323.7 | 10.8181 | 347.5 | 18.718 | 418.8 | 14.7001
|  | 323.8 | 10.8509 | 347.7 | 18.7704 | 419.3 | 14.3606
|  | 323.9 | 10.8837 | 347.9 | 18.8376 | 419.8 | 14.0228
|  | 324 | 10.9165 | 348.1 | 18.9094 | 420.3 | 13.6981
|  | 324.1 | 10.9492 | 348.3 | 18.9812 | 420.8 | 13.37
|  | 324.2 | 10.982 | 348.5 | 19.053 | 421.3 | 13.0339
|  | 324.3 | 11.0148 | 348.7 | 19.1248 | 421.8 | 12.6698
|  | 324.4 | 11.0475 | 348.9 | 19.1966 | 422.3 | 12.2913
|  | 324.5 | 11.0803 | 349.1 | 19.2684 | 422.8 | 11.9391
|  | 324.6 | 11.1131 | 349.3 | 19.3402 | 423.3 | 11.4947
|  | 324.7 | 11.1459 | 349.5 | 19.412 | 423.8 | 11.0166
|  | 324.8 | 11.1786 | 349.7 | 19.4838 | 424.3 | 10.6694
|  | 324.9 | 11.2114 | 349.9 | 19.5556 | 424.8 | 10.3494
|  | 325 | 11.2442 | 350.1 | 19.6273 | 425.3 | 9.9881
|  | 325.1 | 11.2769 | 350.3 | 19.6991 | 425.8 | 9.6268
|  | 325.2 | 11.3097 | 350.5 | 19.7709 | 426.3 | 9.2251
|  | 325.3 | 11.3425 | 350.7 | 19.8427 | 426.8 | 8.8467
|  | 325.4 | 11.3753 | 350.9 | 19.9145 | 427.3 | 8.4965
|  | 325.5 | 11.408 | 351.1 | 19.9863 | 427.8 | 8.158
|  | 325.6 | 11.4408 | 351.3 | 20.0581 | 428.3 | 7.7899
|  | 325.7 | 11.4736 | 351.5 | 20.0668 | 428.8 | 7.4212
|  | 325.8 | 11.5063 | 351.7 | 20.142 | 429.3 | 7.0951
|  | 325.9 | 11.5391 | 351.9 | 20.2088 | 429.8 | 6.7896
|  | 326 | 11.5719 | 352.1 | 20.2757 | 430.3 | 6.467
|  | 326.1 | 11.6047 | 352.3 | 20.3425 | 430.8 | 6.1778
|  | 326.2 | 11.6374 | 352.5 | 20.4093 | 431.3 | 5.9175
|  | 326.3 | 11.6702 | 352.7 | 20.4762 | 431.8 | 5.6286
|  | 326.4 | 11.703 | 352.9 | 20.543 | 432.3 | 5.3397
|  | 326.5 | 11.7357 | 353.1 | 20.6099 | 432.8 | 5.0508
|  | 326.6 | 11.7685 | 353.3 | 20.6767 | 433.3 | 4.9032
|  | 326.7 | 11.7559 | 353.5 | 20.7906 | 433.8 | 4.6568
|  | 326.8 | 11.7862 | 353.7 | 20.8719 | 434.3 | 4.4158
|  | 326.9 | 11.8165 | 353.9 | 20.9439 | 434.8 | 4.2575
|  | 327 | 11.8468 | 354.1 | 21.0159 | 435.3 | 4.0623
|  | 327.1 | 11.8771 | 354.3 | 21.0879 | 435.8 | 3.8656
|  | 327.2 | 11.9074 | 354.5 | 21.1599 | 436.3 | 3.6924
|  | 327.3 | 11.9377 | 354.7 | 21.2318 | 436.8 | 3.552
|  | 327.4 | 11.968 | 354.9 | 21.3038 | 437.3 | 3.383
|  | 327.5 | 11.9983 | 355.1 | 21.3758 | 437.8 | 3.2147
|  | 327.6 | 12.0286 | 355.3 | 21.4478 | 438.3 | 3.0665
|  | 327.7 | 12.059 | 355.5 | 21.5874 | 438.8 | 2.9111
|  | 327.8 | 12.0893 | 355.7 | 21.6905 | 439.3 | 2.7562
|  | 327.9 | 12.1196 | 355.9 | 21.7801 | 439.8 | 2.6013
|  | 328 | 12.1499 | 356.1 | 21.8696 | 440.3 | 2.4756
|  | 328.1 | 12.1802 | 356.3 | 21.9592 | 440.8 | 2.3798
|  | 328.2 | 12.2105 | 356.5 | 22.0488 | 441.3 | 2.2425
|  | 328.3 | 12.2408 | 356.7 | 22.1383 | 441.8 | 2.1998
|  | 328.4 | 12.2711 | 356.9 | 22.2279 | 442.3 | 2.1419
|  | 328.5 | 12.3014 | 357.1 | 22.3175 | 442.8 | 2.0737
|  | 328.6 | 12.3317 | 357.3 | 22.407 | 443.3 | 2.0055
|  | 328.7 | 12.362 | 357.5 | 22.4966 | 443.8 | 1.944
|  | 328.8 | 12.3924 | 357.7 | 22.5861 | 444.3 | 1.8872
|  | 328.9 | 12.4227 | 357.9 | 22.6757 | 444.8 | 1.8286
|  | 329 | 12.453 | 358.1 | 22.7653 | 445.3 | 1.77
|  | 329.1 | 12.4833 | 358.3 | 22.8548 | 445.8 | 1.736
|  | 329.2 | 12.5136 | 358.5 | 22.9773 | 446.3 | 1.6988
|  | 329.3 | 12.5439 | 358.7 | 23.0793 | 446.8 | 1.6518
|  | 329.4 | 12.5742 | 358.9 | 23.1748 | 447.3 | 1.6048
|  | 329.5 | 12.6045 | 359.1 | 23.2702 | 447.8 | 1.5578
|  | 329.6 | 12.6348 | 359.3 | 23.3657 | 448.3 | 1.518
|  | 329.7 | 12.6651 | 359.5 | 23.4612 | 448.8 | 1.4967
|  | 329.8 | 12.6954 | 359.7 | 23.5566 | 449.3 | 1.4621
|  | 329.9 | 12.7258 | 359.9 | 23.6521 | 449.8 | 1.4271
|  | 330 | 12.7561 | 360.1 | 23.7475 | 450.3 | 1.3922
|  | 330.1 | 12.7715 | 360.3 | 23.843 | 450.8 | 1.3488
|  | 330.2 | 12.8006 | 360.5 | 23.9385 | 451.3 | 1.3018
|  | 330.3 | 12.8296 | 360.7 | 24.0339 | 451.8 | 1.2629
|  | 330.4 | 12.8587 | 360.9 | 24.1294 | 452.3 | 1.224
|  | 330.5 | 12.8877 | 361.1 | 24.2248 | 452.8 | 1.1851
|  | 330.6 | 12.9168 | 361.3 | 24.3203 | 453.3 | 1.1784
|  | 330.7 | 12.9458 | 361.5 | 24.4158 | 453.8 | 1.1707
|  | 330.8 | 12.9749 | 361.7 | 24.3485 | 454.3 | 1.1456
|  | 330.9 | 13.0039 | 361.9 | 24.4544 | 454.8 | 1.1205
|  | 331 | 13.033 | 362.1 | 24.5386 | 455.3 | 1.0954
|  | 331.1 | 13.062 | 362.3 | 24.6228 | 455.8 | 1.0703
|  | 331.2 | 13.0911 | 362.5 | 24.707 | 456.3 | 1.0452
|  | 331.3 | 13.1201 | 362.7 | 24.7913 | 456.8 | 1.0201
|  | 331.4 | 13.1492 | 362.9 | 24.8755 | 457.3 | 0.995
|  | 331.5 | 13.1782 | 363.1 | 24.9597 | 457.8 | 1.008
|  | 331.6 | 13.2073 | 363.3 | 25.044 | 458.3 | 0.99525
|  | 331.7 | 13.2363 | 363.5 | 25.1282 | 458.8 | 0.98158
|  | 331.8 | 13.2653 | 363.7 | 25.2124 | 459.3 | 0.9679
|  | 331.9 | 13.2944 | 363.9 | 25.2966 | 459.8 | 0.95422
|  | 332 | 13.3234 | 364.1 | 25.3939 | 460.3 | 0.94054
|  | 332.1 | 13.3525 | 364.3 | 25.4675 | 460.8 | 0.92687
|  | 332.2 | 13.3815 | 364.5 | 25.5441 | 461.3 | 0.91319
|  | 332.3 | 13.4106 | 364.7 | 25.6207 | 461.8 | 0.90801
|  | 332.4 | 13.4396 | 364.9 | 25.6972 | 462.3 | 0.89815
|  | 332.5 | 13.4687 | 365.1 | 25.7738 | 462.8 | 0.8871
|  | 332.6 | 13.4977 | 365.3 | 25.8504 | 463.3 | 0.87605
|  | 332.7 | 13.5268 | 365.5 | 25.9214 | 463.8 | 0.86501
|  | 332.8 | 13.5558 | 365.7 | 26.0053 | 464.3 | 0.85396
|  | 332.9 | 13.5849 | 365.9 | 26.0875 | 464.8 | 0.84292
|  | 333 | 13.6139 | 366.1 | 26.1698 | 465.3 | 0.83187
|  | 333.1 | 13.643 | 366.3 | 26.2521 | 465.8 | 0.82082
|  | 333.2 | 13.672 | 366.5 | 26.3343 | 466.3 | 0.80978
|  | 333.3 | 13.701 | 366.7 | 26.4166 | 466.8 | 0.79873
|  | 333.4 | 13.7301 | 366.9 | 26.4989 | 467.3 | 0.78228
|  | 333.5 | 13.7591 | 367.1 | 26.5812 | 467.8 | 0.76855
|  | 333.6 | 13.8123 | 367.3 | 26.6634 | 468.3 | 0.75679
|  | 333.7 | 13.8424 | 367.5 | 26.7457 | 468.8 | 0.74503
|  | 333.8 | 13.8724 | 367.7 | 26.828 | 469.3 | 0.73327
|  | 333.9 | 13.9024 | 367.9 | 26.9102 | 469.8 | 0.72151
|  | 334 | 13.9325 | 368.1 | 26.9925 | 470.3 | 0.70975
|  | 334.1 | 13.9625 | 368.3 | 27.0748 | 470.8 | 0.69799
|  | 334.2 | 13.9926 | 368.5 | 27.0135 | 471.3 | 0.68623
|  | 334.3 | 14.0226 | 368.7 | 27.0481 | 471.8 | 0.67447
|  | 334.4 | 14.0527 | 368.9 | 27.1113 | 472.3 | 0.6591
|  | 334.5 | 14.0827 | 369.1 | 27.1746 | 472.8 | 0.64485
|  | 334.6 | 14.1128 | 369.3 | 27.2378 | 473.3 | 0.6318
|  | 334.7 | 14.1428 | 369.5 | 27.301 | 473.8 | 0.61874
|  | 334.8 | 14.1728 | 369.7 | 27.3643 | 474.3 | 0.60569
|  | 334.9 | 14.2029 | 369.9 | 27.4275 | 474.8 | 0.59263
|  | 335 | 14.2329 | 370.1 | 27.4908 | 475.3 | 0.57958
|  | 335.1 | 14.263 | 370.3 | 27.554 | 475.8 | 0.56652
|  | 335.2 | 14.293 | 370.5 | 27.6173 | 476.3 | 0.55347
|  | 335.3 | 14.3231 | 370.7 | 27.5451 | 476.8 | 0.54041
|  | 335.4 | 14.3531 | 370.9 | 27.6003 | 477.3 | 0.52442
|  | 335.5 | 14.3832 | 371.1 | 27.6396 | 477.8 | 0.50897
|  | 335.6 | 14.4132 | 371.3 | 27.6789 | 478.3 | 0.49525
|  | 335.7 | 14.4432 | 371.5 | 27.7182 | 478.8 | 0.48152
|  | 335.8 | 14.4733 | 371.7 | 27.7575 | 479.3 | 0.4678
|  | 335.9 | 14.5033 | 371.9 | 27.7968 | 479.8 | 0.45408
|  | 336 | 14.5334 | 372.1 | 27.8362 | 480.3 | 0.44036
|  | 336.1 | 14.5634 | 372.3 | 27.8755 | 480.8 | 0.42664
|  | 336.2 | 14.5935 | 372.5 | 27.9148 | 481.3 | 0.41292
|  | 336.3 | 14.6235 | 372.7 | 27.9541 | 481.8 | 0.37777
|  | 336.4 | 14.6536 | 372.9 | 27.9934 | 482.3 | 0.34944
|  | 336.5 | 14.6836 | 373.1 | 28.0327 | 482.8 | 0.33041
|  | 336.6 | 14.7137 | 373.3 | 28.0832 | 483.3 | 0.31138
|  | 336.7 | 14.8107 | 373.5 | 28.1248 | 483.8 | 0.29234
|  | 336.8 | 14.8444 | 373.7 | 28.164 | 484.3 | 0.27331
|  | 336.9 | 14.878 | 373.9 | 28.2033 | 484.8 | 0.25428
|  | 337 | 14.9117 | 374.1 | 28.2426 | 485.3 | 0.23525
|  | 337.1 | 14.9453 | 374.3 | 28.2818 | 485.8 | 0.21946
|  | 337.2 | 14.979 | 374.5 | 28.3211 | 486.3 | 0.20284
|  | 337.3 | 15.0127 | 374.7 | 28.3604 | 486.8 | 0.18252
|  | 337.4 | 15.0463 | 374.9 | 28.3996 | 487.3 | 0.1622
|  | 337.5 | 15.08 | 375.1 | 28.4389 | 487.8 | 0.14189
|  | 337.6 | 15.1137 | 375.3 | 28.4893 | 488.3 | 0.12157
|  | 337.7 | 15.1473 | 375.5 | 28.5344 | 488.8 | 0.10126
|  | 337.8 | 15.181 | 375.7 | 28.5774 | 489.3 | 0.080941
|  | 337.9 | 15.2146 | 375.9 | 28.6203 | 489.8 | 0.088606
|  | 338 | 15.2483 | 376.1 | 28.6632 | 490.3 | 0.10805
|  | 338.1 | 15.282 | 376.3 | 28.7062 | 490.8 | 0.10047
|  | 338.2 | 15.3156 | 376.5 | 28.7258 | 491.3 | 0.092885
|  | 338.3 | 15.3493 | 376.7 | 28.7573 | 491.8 | 0.0853
|  | 338.4 | 15.3829 | 376.9 | 28.7936 | 492.3 | 0.077715
|  | 338.5 | 15.4166 | 377.1 | 28.8298 | 492.8 | 0.07013
|  | 338.6 | 15.4503 | 377.3 | 28.866 | 493.3 | 0.062545
|  | 338.7 | 15.4839 | 377.5 | 28.9022 | 493.8 | 0.05496
|  | 338.8 | 15.5176 | 377.7 | 28.9128 | 494.3 | 0.053672
|  | 338.9 | 15.5512 | 377.9 | 28.9434 | 494.8 | 0.066597
|  | 339 | 15.5849 | 378.1 | 28.9715 | 495.3 | 0.066749
|  | 339.1 | 15.6186 | 378.3 | 28.9996 | 495.8 | 0.066586
|  | 339.2 | 15.6522 | 378.5 | 29.0277 | 496.3 | 0.066423
|  | 339.3 | 15.6859 | 378.7 | 29.0827 | 496.8 | 0.06626
|  | 339.4 | 15.7328 | 378.9 | 29.1133 | 497.3 | 0.066097
|  | 339.5 | 15.7685 | 379.1 | 29.1469 | 497.8 | 0.065934
|  | 339.6 | 15.8043 | 379.3 | 29.1805 | 498.3 | 0.065771
|  | 339.7 | 15.8401 | 379.5 | 29.2142 | 498.8 | 0.065608
|  | 339.8 | 15.8758 | 379.7 | 29.2294 | 499.3 | 0.065445
|  | 339.9 | 15.9116 | 379.9 | 29.2563 | 499.8 | 0.065282
|  | 340 | 15.9473 | 380.1 | 29.2869 | 500.3 | 0.065119
|  | 340.1 | 15.9831 | 380.3 | 29.3174
|  | 340.2 | 16.0188 | 380.5 | 29.348
|  | 340.3 | 16.0546 | 380.7 | 29.3786
|  | 340.4 | 16.0903 | 380.9 | 29.3751
|  | 340.5 | 16.1261 | 381.1 | 29.3851
|  | 340.6 | 16.1619 | 381.3 | 29.4018
|  | 340.7 | 16.1976 | 381.5 | 29.4186
|  | 340.8 | 16.2334 | 381.7 | 29.4354
|  | 340.9 | 16.2691 | 381.9 | 29.4521
|  | 341 | 16.3049 | 382.1 | 29.4689
|  | 341.1 | 16.3406 | 382.3 | 29.4857
|  | 341.2 | 16.3764 | 382.5 | 29.5025
|  | 341.3 | 16.4122 | 382.7 | 29.4896
|  | 341.4 | 16.4479 | 382.9 | 29.504
|  | 341.5 | 16.4837 | 383.1 | 29.5151
|  | 341.6 | 16.5194 | 383.3 | 29.5261
|  | 341.7 | 16.5552 | 383.5 | 29.5045
|  | 341.8 | 16.5909 | 383.7 | 29.5088
|  | 341.9 | 16.6267 | 383.9 | 29.5098
|  | 342 | 16.6625 | 384.1 | 29.5108
|  | 342.1 | 16.6982 | 384.3 | 29.5119
|  | 342.2 | 16.7523 | 384.5 | 29.4784
|  | 342.3 | 16.7887 | 384.7 | 29.4591
|  | 342.4 | 16.825 | 384.9 | 29.4467
|  | 342.5 | 16.8614 | 385.1 | 29.4343
|  | 342.6 | 16.8977 | 385.3 | 29.4219
|  | 342.7 | 16.934 | 385.5 | 29.4095
|  | 342.8 | 16.9704 | 385.7 | 29.3971
|  | 342.9 | 17.0067 | 385.9 | 29.3847
|  | 343 | 17.0431 | 386.1 | 29.361
|  | 343.1 | 17.0794 | 386.3 | 29.3411
|  | 343.2 | 17.1158 | 386.5 | 29.3234
|  | 343.3 | 17.1521 | 386.7 | 29.3057
|  | 343.4 | 17.1885 | 386.9 | 29.2881
|  | 343.5 | 17.2248 | 387.1 | 29.2704
|  | 343.6 | 17.2612 | 387.3 | 29.2527
|  | 343.7 | 17.2975 | 387.5 | 29.226
|  | 343.8 | 17.3338 | 387.7 | 29.2077
|  | 343.9 | 17.3702 | 387.9 | 29.1884
|  | 344 | 17.4065 | 388.1 | 29.169
|  | 344.1 | 17.4429 | 388.3 | 29.1497
|  | 344.2 | 17.4792 | 388.5 | 29.1303
|  | 344.3 | 17.5156 | 388.7 | 29.0782
|  | 344.4 | 17.5519 | 388.9 | 29.0435
|  | 344.5 | 17.5883 | 389.1 | 29.0154
|  | 344.6 | 17.6246 | 389.3 | 28.9873
|  | 344.7 | 17.661 | 389.5 | 28.9591
|  | 344.8 | 17.6973 | 389.7 | 28.875
|  | 344.9 | 17.7336 | 389.9 | 28.8331
|  | 345 | 17.78 | 390.1 | 28.786
|  | 345.1 | 17.8175 | 390.3 | 28.7389
|  | 345.2 | 17.855 | 390.5 | 28.6918
|  | 345.3 | 17.8925 | 390.7 | 28.6447
|  | 345.4 | 17.93 | 390.9 | 28.5976
|  | 345.5 | 17.9676 | 391.1 | 28.5504
|  | 345.6 | 18.0051 | 391.3 | 28.5033
|  | 345.7 | 18.0426 | 391.5 | 28.4562
|  | 345.8 | 18.0801 | 391.7 | 28.4091
|  | 345.9 | 18.1177 | 391.9 | 28.3594
|  | 346 | 18.1552 | 392.1 | 28.3108
|  | 346.1 | 18.1927 | 392.3 | 28.262
|  | 346.2 | 18.2302 | 392.5 | 28.2132
|  | 346.3 | 18.2678 | 392.7 | 28.1644
|  | 346.4 | 18.3053 | 392.9 | 28.1157
|  | 346.5 | 18.3428 | 393.1 | 28.0669
|  | 346.6 | 18.3803 | 393.3 | 28.0181
|  | 346.7 | 18.4178 | 393.5 | 27.9693
|  | 346.8 | 18.4554 | 393.7 | 27.7911
|  | 346.9 | 18.4929 | 393.9 | 27.7429
|  | 347 | 18.5304 | 394.1 | 27.6785
|  | 347.1 | 18.5679 | 394.3 | 27.6141
|  | 347.2 | 18.6055 | 394.5 | 27.5497
|  | 347.3 | 18.643 | 394.7 | 27.4853
|  | 347.4 | 18.6805 | 394.9 | 27.4209
|  | 347.5 | 18.718 | 395.1 | 27.3227
|  | 347.6 | 18.7556 | 395.3 | 27.2331
|  | 347.7 | 18.7658 | 395.5 | 27.1504
|  | 347.8 | 18.8017 | 395.7 | 27.0676
|  | 347.9 | 18.8376 | 395.9 | 26.9848
|  | 348 | 18.8735 | 396.1 | 26.9021
|  | 348.1 | 18.9094 | 396.3 | 26.8193
|  | 348.2 | 18.9453 | 396.5 | 26.7365
|  | 348.3 | 18.9812 | 396.7 | 26.6001
|  | 348.4 | 19.0171 | 396.9 | 26.5196
|  | 348.5 | 19.053 | 397.1 | 26.4141
|  | 348.6 | 19.0889 | 397.3 | 26.2993
|  | 348.7 | 19.1248 | 397.5 | 26.1882
|  | 348.8 | 19.1607 | 397.7 | 26.077
|  | 348.9 | 19.1966 | 397.9 | 25.9893
|  | 349 | 19.2325 | 398.1 | 25.8885
|  | 349.1 | 19.2684 | 398.3 | 25.7893
|  | 349.2 | 19.3043 | 398.5 | 25.6902
|  | 349.3 | 19.3402 | 398.7 | 25.5911
|  | 349.4 | 19.3761 | 398.9 | 25.4919
|  | 349.5 | 19.412 | 399.1 | 25.3838
|  | 349.6 | 19.4479 | 399.3 | 25.2809
|  | 349.7 | 19.4838 | 399.5 | 25.1799
|  | 349.8 | 19.5197 | 399.7 | 25.0789
|  | 349.9 | 19.5556 | 399.9 | 24.9779
|  | 350 | 19.5915 | 400.1 | 24.8769
|  | 350.1 | 19.6273 | 400.3 | 24.7307
|  | 350.2 | 19.6632 | 400.5 | 24.6241
|  | 350.3 | 19.6991 | 400.7 | 24.5126
|  | 350.4 | 19.735 | 400.9 | 24.4011
|  | 350.5 | 19.7709 | 401.1 | 24.2767
|  | 350.6 | 19.8068 | 401.3 | 24.1543
|  | 350.7 | 19.8427 | 401.5 | 24.0345
|  | 350.8 | 19.8786 | 401.7 | 23.9147
|  | 350.9 | 19.9145 | 401.9 | 23.7949
|  | 351 | 19.9504 | 402.1 | 23.6751
|  | 351.1 | 19.9863 | 402.3 | 23.5636
|  | 351.2 | 20.0222 | 402.5 | 23.4443
|  | 351.3 | 20.0581 | 402.7 | 23.3259
|  | 351.4 | 20.0417 | 402.9 | 23.2076
|  | 351.5 | 20.0751 | 403.1 | 23.1234
|  | 351.6 | 20.1086 | 403.3 | 23.0256
|  | 351.7 | 20.142 | 403.5 | 22.9209
|  | 351.8 | 20.1754 | 403.7 | 22.8163
|  | 351.9 | 20.2088 | 403.9 | 22.7116
|  | 352 | 20.2422 | 404.1 | 22.6069
|  | 352.1 | 20.2757 | 404.3 | 22.5863
|  | 352.2 | 20.3091 | 404.5 | 22.4941
|  | 352.3 | 20.3425 | 404.7 | 22.4109
|  | 352.4 | 20.3759 | 404.9 | 22.3277
|  | 352.5 | 20.4093 | 405.1 | 22.2444
|  | 352.6 | 20.4428 | 405.3 | 22.1612
|  | 352.7 | 20.4762 | 405.5 | 22.078
|  | 352.8 | 20.5096 | 405.7 | 21.8813
|  | 352.9 | 20.543 | 405.9 | 21.8012
|  | 353 | 20.5764 | 406.1 | 21.7065
|  | 353.1 | 20.6099 | 406.3 | 21.6178
|  | 353.2 | 20.6433 | 406.5 | 21.5112
|  | 353.3 | 20.6767 | 406.7 | 21.4069
|  | 353.4 | 20.7101 | 406.9 | 21.3026
|  | 353.5 | 20.8 | 407.1 | 21.1983
|  | 353.6 | 20.836 | 407.3 | 21.094
|  | 353.7 | 20.8719 | 407.5 | 20.9896
|  | 353.8 | 20.9079 | 407.7 | 20.8853
|  | 353.9 | 20.9439 | 407.9 | 20.781
|  | 354 | 20.9799 | 408.1 | 20.6767
|  | 354.1 | 21.0159 | 408.3 | 20.5724
|  | 354.2 | 21.0519 | 408.5 | 20.4681
|  | 354.3 | 21.0879 | 408.7 | 20.3481
|  | 354.4 | 21.1239 | 408.9 | 20.2511
|  | 354.5 | 21.1599 | 409.1 | 20.1513
|  | 354.6 | 21.1959 | 409.3 | 20.0514
|  | 354.7 | 21.2318 | 409.5 | 19.9516
|  | 354.8 | 21.2678 | 409.7 | 19.8517
|  | 354.9 | 21.3038 | 409.9 | 19.7519
|  | 355 | 21.3398 | 410.1 | 19.652
|  | 355.1 | 21.3758 | 410.3 | 19.5522
|  | 355.2 | 21.4118 | 410.5 | 19.4523
|  | 355.3 | 21.4478 | 410.7 | 19.3524
|  | 355.4 | 21.4838 | 410.9 | 19.2399
|  | 355.5 | 21.6009 | 411.1 | 19.1334
|  | 355.6 | 21.6457 | 411.3 | 19.0295
|  | 355.7 | 21.6905 | 411.5 | 18.9256
|  | 355.8 | 21.7353 | 411.7 | 18.8217
|  | 355.9 | 21.7801 | 411.9 | 18.7178
|  | 356 | 21.8248 | 412.1 | 18.6138
|  | 356.1 | 21.8696 | 412.3 | 18.5099
|  | 356.2 | 21.9144 | 412.5 | 18.406
|  | 356.3 | 21.9592 | 412.7 | 18.3021
|  | 356.4 | 22.004 | 412.9 | 18.1742
|  | 356.5 | 22.0488 | 413.1 | 18.0614
|  | 356.6 | 22.0935 | 413.3 | 17.9536
|  | 356.7 | 22.1383 | 413.5 | 17.8457
|  | 356.8 | 22.1831 | 413.7 | 17.7378
|  | 356.9 | 22.2279 | 413.9 | 17.6299
|  | 357 | 22.2727 | 414.1 | 17.522
|  | 357.1 | 22.3175 | 414.3 | 17.4141
|  | 357.2 | 22.3622 | 414.5 | 17.3062
|  | 357.3 | 22.407 | 414.7 | 17.1983
|  | 357.4 | 22.4518 | 414.9 | 17.0817
|  | 357.5 | 22.4966 | 415.1 | 16.9684
|  | 357.6 | 22.5414 | 415.3 | 16.8568
|  | 357.7 | 22.5861 | 415.5 | 16.7452
|  | 357.8 | 22.6309 | 415.7 | 16.6336
|  | 357.9 | 22.6757 | 415.9 | 16.522
|  | 358 | 22.7205 | 416.1 | 16.4104
|  | 358.1 | 22.7653 | 416.3 | 16.2987
|  | 358.2 | 22.8101 | 416.5 | 16.1871
|  | 358.3 | 22.8548 | 416.7 | 16.0755
|  | 358.4 | 22.8996 | 416.9 | 15.9197
|  | 358.5 | 22.9839 | 417.1 | 15.7921
|  | 358.6 | 23.0316 | 417.3 | 15.6732
|  | 358.7 | 23.0793 | 417.5 | 15.5544
|  | 358.8 | 23.1271 | 417.7 | 15.4356
|  | 358.9 | 23.1748 | 417.9 | 15.3167
|  | 359 | 23.2225 | 418.1 | 15.1979
|  | 359.1 | 23.2702 | 418.3 | 15.079
|  | 359.2 | 23.318 | 418.5 | 14.9602
|  | 359.3 | 23.3657 | 418.7 | 14.7562
|  | 359.4 | 23.4134 | 418.9 | 14.6309
|  | 359.5 | 23.4612 | 419.1 | 14.4958
|  | 359.6 | 23.5089 | 419.3 | 14.3606
|  | 359.7 | 23.5566 | 419.5 | 14.2255
|  | 359.8 | 23.6044 | 419.7 | 14.0904
|  | 359.9 | 23.6521 | 419.9 | 13.9553
|  | 360 | 23.6998 | 420.1 | 13.8201
|  | 360.1 | 23.7475 | 420.3 | 13.7026
|  | 360.2 | 23.7953 | 420.5 | 13.5716
|  | 360.3 | 23.843 | 420.7 | 13.4372
|  | 360.4 | 23.8907 | 420.9 | 13.3028
|  | 360.5 | 23.9385 | 421.1 | 13.1684
|  | 360.6 | 23.9862 | 421.3 | 13.0339
|  | 360.7 | 24.0339 | 421.5 | 12.8995
|  | 360.8 | 24.0817 | 421.7 | 12.7651
|  | 360.9 | 24.1294 | 421.9 | 12.5826
|  | 361 | 24.1771 | 422.1 | 12.4321
|  | 361.1 | 24.2248 | 422.3 | 12.2913
|  | 361.2 | 24.2726 | 422.5 | 12.1504
|  | 361.3 | 24.3203 | 422.7 | 12.0095
|  | 361.4 | 24.368 | 422.9 | 11.8687
|  | 361.5 | 24.4158 | 423.1 | 11.7278
|  | 361.6 | 24.328 | 423.3 | 11.4769
|  | 361.7 | 24.3701 | 423.5 | 11.2795
|  | 361.8 | 24.4122 | 423.7 | 11.1042
|  | 361.9 | 24.4544 | 423.9 | 10.9289
|  | 362 | 24.4965 | 424.1 | 10.7536
|  | 362.1 | 24.5386 | 424.3 | 10.6886
|  | 362.2 | 24.5807 | 424.5 | 10.5661
|  | 362.3 | 24.6228 | 424.7 | 10.4216
|  | 362.4 | 24.6649 | 424.9 | 10.2771
|  | 362.5 | 24.707 | 425.1 | 10.1326
|  | 362.6 | 24.7492 | 425.3 | 9.9881
|  | 362.7 | 24.7913 | 425.5 | 9.8435
|  | 362.8 | 24.8334 | 425.7 | 9.699
|  | 362.9 | 24.8755 | 425.9 | 9.5545
|  | 363 | 24.9176 | 426.1 | 9.41
|  | 363.1 | 24.9597 | 426.3 | 9.2084
|  | 363.2 | 25.0018 | 426.5 | 9.0569
|  | 363.3 | 25.044 | 426.7 | 8.9168
|  | 363.4 | 25.0861 | 426.9 | 8.7767
|  | 363.5 | 25.1282 | 427.1 | 8.6366
|  | 363.6 | 25.1703 | 427.3 | 8.4965
|  | 363.7 | 25.2124 | 427.5 | 8.3564
|  | 363.8 | 25.2545 | 427.7 | 8.2363
|  | 363.9 | 25.2966 | 427.9 | 8.0849
|  | 364 | 25.3526 | 428.1 | 7.9374
|  | 364.1 | 25.3909 | 428.3 | 7.7899
|  | 364.2 | 25.4292 | 428.5 | 7.6424
|  | 364.3 | 25.4675 | 428.7 | 7.4949
|  | 364.4 | 25.5058 | 428.9 | 7.3474
|  | 364.5 | 25.5441 | 429.1 | 7.1999
|  | 364.6 | 25.5824 | 429.3 | 7.1023
|  | 364.7 | 25.6207 | 429.5 | 6.9832
|  | 364.8 | 25.6589 | 429.7 | 6.8542
|  | 364.9 | 25.6972 | 429.9 | 6.7251
|  | 365 | 25.7355 | 430.1 | 6.5961
|  | 365.1 | 25.7738 | 430.3 | 6.467
|  | 365.2 | 25.8121 | 430.5 | 6.338
|  | 365.3 | 25.8504 | 430.7 | 6.209
|  | 365.4 | 25.8819 | 430.9 | 6.1371
|  | 365.5 | 25.923 | 431.1 | 6.033
|  | 365.6 | 25.9641 | 431.3 | 5.9175
|  | 365.7 | 26.0053 | 431.5 | 5.8019
|  | 365.8 | 26.0464 | 431.7 | 5.6864
|  | 365.9 | 26.0875 | 431.9 | 5.5708
|  | 366 | 26.1287 | 432.1 | 5.4553
|  | 366.1 | 26.1698 | 432.3 | 5.3397
|  | 366.2 | 26.2109 | 432.5 | 5.2242
|  | 366.3 | 26.2521 | 432.7 | 5.1086
|  | 366.4 | 26.2932 | 432.9 | 4.9931
|  | 366.5 | 26.3343 | 433.1 | 5.0104
|  | 366.6 | 26.3755 | 433.3 | 4.8978
|  | 366.7 | 26.4166 | 433.5 | 4.8014
|  | 366.8 | 26.4577 | 433.7 | 4.705
|  | 366.9 | 26.4989 | 433.9 | 4.6086
|  | 367 | 26.54 | 434.1 | 4.5122
|  | 367.1 | 26.5812 | 434.3 | 4.4158
|  | 367.2 | 26.6223 | 434.5 | 4.3194
|  | 367.3 | 26.6634 | 434.7 | 4.308
|  | 367.4 | 26.7046 | 434.9 | 4.2197
|  | 367.5 | 26.7457 | 435.1 | 4.141
|  | 367.6 | 26.7868 | 435.3 | 4.0623
|  | 367.7 | 26.828 | 435.5 | 3.9836
|  | 367.8 | 26.8691 | 435.7 | 3.9049
|  | 367.9 | 26.9102 | 435.9 | 3.8262
|  | 368 | 26.9514 | 436.1 | 3.7475
|  | 368.1 | 26.9925 | 436.3 | 3.6688
|  | 368.2 | 27.0336 | 436.5 | 3.6608
|  | 368.3 | 27.0748 | 436.7 | 3.585
|  | 368.4 | 27.1159 | 436.9 | 3.5176
|  | 368.5 | 26.9848 | 437.1 | 3.4503
|  | 368.6 | 27.0164 | 437.3 | 3.383
|  | 368.7 | 27.0481 | 437.5 | 3.3157
|  | 368.8 | 27.0797 | 437.7 | 3.2483
|  | 368.9 | 27.1113 | 437.9 | 3.181
|  | 369 | 27.1429 | 438.1 | 3.1295
|  | 369.1 | 27.1746 | 438.3 | 3.066
|  | 369.2 | 27.2062 | 438.5 | 3.0041
|  | 369.3 | 27.2378 | 438.7 | 2.9421
|  | 369.4 | 27.2694 | 438.9 | 2.8801
|  | 369.5 | 27.301 | 439.1 | 2.8182
|  | 369.6 | 27.3327 | 439.3 | 2.7562
|  | 369.7 | 27.3643 | 439.5 | 2.6942
|  | 369.8 | 27.3959 | 439.7 | 2.6323
|  | 369.9 | 27.4275 | 439.9 | 2.5703
|  | 370 | 27.4592 | 440.1 | 2.5083
|  | 370.1 | 27.4908 | 440.3 | 2.4464
|  | 370.2 | 27.5224 | 440.5 | 2.4722
|  | 370.3 | 27.554 | 440.7 | 2.4062
|  | 370.4 | 27.5857 | 440.9 | 2.3516
|  | 370.5 | 27.6173 | 441.1 | 2.2971
|  | 370.6 | 27.5413 | 441.3 | 2.2425
|  | 370.7 | 27.561 | 441.5 | 2.188
|  | 370.8 | 27.5807 | 441.7 | 2.2087
|  | 370.9 | 27.6003 | 441.9 | 2.1964
|  | 371 | 27.62 | 442.1 | 2.1692
|  | 371.1 | 27.6396 | 442.3 | 2.1419
|  | 371.2 | 27.6593 | 442.5 | 2.1146
|  | 371.3 | 27.6789 | 442.7 | 2.0873
|  | 371.4 | 27.6986 | 442.9 | 2.0601
|  | 371.5 | 27.7182 | 443.1 | 2.0328
|  | 371.6 | 27.7379 | 443.3 | 2.0055
|  | 371.7 | 27.7575 | 443.5 | 1.9782
|  | 371.8 | 27.7772 | 443.7 | 1.9509
|  | 371.9 | 27.7968 | 443.9 | 1.9351
|  | 372 | 27.8165 | 444.1 | 1.9106
|  | 372.1 | 27.8362 | 444.3 | 1.8872
|  | 372.2 | 27.8558 | 444.5 | 1.8637
|  | 372.3 | 27.8755 | 444.7 | 1.8403
|  | 372.4 | 27.8951 | 444.9 | 1.8169
|  | 372.5 | 27.9148 | 445.1 | 1.7935
|  | 372.6 | 27.9344 | 445.3 | 1.77
|  | 372.7 | 27.9541 | 445.5 | 1.7466
|  | 372.8 | 27.9737 | 445.7 | 1.7232
|  | 372.9 | 27.9934 | 445.9 | 1.7417
|  | 373 | 28.013 | 446.1 | 1.7176
|  | 373.1 | 28.0327 | 446.3 | 1.6988
|  | 373.2 | 28.0524 | 446.5 | 1.68
|  | 373.3 | 28.0855 | 446.7 | 1.6612
|  | 373.4 | 28.1051 | 446.9 | 1.6424
|  | 373.5 | 28.1248 | 447.1 | 1.6236
|  | 373.6 | 28.1444 | 447.3 | 1.6048
|  | 373.7 | 28.164 | 447.5 | 1.586
|  | 373.8 | 28.1837 | 447.7 | 1.5672
|  | 373.9 | 28.2033 | 447.9 | 1.5484
|  | 374 | 28.2229 | 448.1 | 1.5296
|  | 374.1 | 28.2426 | 448.3 | 1.5107
|  | 374.2 | 28.2622 | 448.5 | 1.5136
|  | 374.3 | 28.2818 | 448.7 | 1.504
|  | 374.4 | 28.3015 | 448.9 | 1.49
|  | 374.5 | 28.3211 | 449.1 | 1.476
|  | 374.6 | 28.3407 | 449.3 | 1.4621
|  | 374.7 | 28.3604 | 449.5 | 1.4481
|  | 374.8 | 28.38 | 449.7 | 1.4341
|  | 374.9 | 28.3996 | 449.9 | 1.4201
|  | 375 | 28.4193 | 450.1 | 1.4062
|  | 375.1 | 28.4389 | 450.3 | 1.3922
|  | 375.2 | 28.4585 | 450.5 | 1.3782
|  | 375.3 | 28.4915 | 450.7 | 1.3642
|  | 375.4 | 28.513 | 450.9 | 1.3358
|  | 375.5 | 28.5344 | 451.1 | 1.3174
|  | 375.6 | 28.5559 | 451.3 | 1.3018
|  | 375.7 | 28.5774 | 451.5 | 1.2863
|  | 375.8 | 28.5988 | 451.7 | 1.2707
|  | 375.9 | 28.6203 | 451.9 | 1.2551
|  | 376 | 28.6418 | 452.1 | 1.2396
|  | 376.1 | 28.6632 | 452.3 | 1.224
|  | 376.2 | 28.6847 | 452.5 | 1.2085
|  | 376.3 | 28.7062 | 452.7 | 1.1929
|  | 376.4 | 28.7276 | 452.9 | 1.1773
|  | 376.5 | 28.7211 | 453.1 | 1.1618
|  | 376.6 | 28.7392 | 453.3 | 1.1875
|  | 376.7 | 28.7573 | 453.5 | 1.1858
|  | 376.8 | 28.7755 | 453.7 | 1.1757
|  | 376.9 | 28.7936 | 453.9 | 1.1657
|  | 377 | 28.8117 | 454.1 | 1.1556
|  | 377.1 | 28.8298 | 454.3 | 1.1456
|  | 377.2 | 28.8479 | 454.5 | 1.1356
|  | 377.3 | 28.866 | 454.7 | 1.1255
|  | 377.4 | 28.8841 | 454.9 | 1.1155
|  | 377.5 | 28.9022 | 455.1 | 1.1054
|  | 377.6 | 28.9013 | 455.3 | 1.0954
|  | 377.7 | 28.9153 | 455.5 | 1.0854
|  | 377.8 | 28.9294 | 455.7 | 1.0753
|  | 377.9 | 28.9434 | 455.9 | 1.0653
|  | 378 | 28.9575 | 456.1 | 1.0552
|  | 378.1 | 28.9715 | 456.3 | 1.0452
|  | 378.2 | 28.9856 | 456.5 | 1.0352
|  | 378.3 | 28.9996 | 456.7 | 1.0251
|  | 378.4 | 29.0137 | 456.9 | 1.0151
|  | 378.5 | 29.0277 | 457.1 | 1.005
|  | 378.6 | 29.0629 | 457.3 | 0.995
|  | 378.7 | 29.0797 | 457.5 | 0.98496
|  | 378.8 | 29.0965 | 457.7 | 1.017
|  | 378.9 | 29.1133 | 457.9 | 1.0062
|  | 379 | 29.1301 | 458.1 | 1.0007
|  | 379.1 | 29.1469 | 458.3 | 0.99525
|  | 379.2 | 29.1637 | 458.5 | 0.98978
|  | 379.3 | 29.1805 | 458.7 | 0.98431
|  | 379.4 | 29.1973 | 458.9 | 0.97884
|  | 379.5 | 29.2142 | 459.1 | 0.97337
|  | 379.6 | 29.231 | 459.3 | 0.9679
|  | 379.7 | 29.2257 | 459.5 | 0.96243
|  | 379.8 | 29.241 | 459.7 | 0.95696
|  | 379.9 | 29.2563 | 459.9 | 0.95148
|  | 380 | 29.2716 | 460.1 | 0.94601
|  | 380.1 | 29.2869 | 460.3 | 0.94054
|  | 380.2 | 29.3021 | 460.5 | 0.93507
|  | 380.3 | 29.3174 | 460.7 | 0.9296
|  | 380.4 | 29.3327 | 460.9 | 0.92413
|  | 380.5 | 29.348 | 461.1 | 0.91866
|  | 380.6 | 29.3633 | 461.3 | 0.91319
|  | 380.7 | 29.3786 | 461.5 | 0.90772
|  | 380.8 | 29.3939 | 461.7 | 0.90987
|  | 380.9 | 29.3683 | 461.9 | 0.90698
|  | 381 | 29.3767 | 462.1 | 0.90256
|  | 381.1 | 29.3851 | 462.3 | 0.89815
|  | 381.2 | 29.3935 | 462.5 | 0.89373
|  | 381.3 | 29.4018 | 462.7 | 0.88931
|  | 381.4 | 29.4102 | 462.9 | 0.88489
|  | 381.5 | 29.4186 | 463.1 | 0.88047
|  | 381.6 | 29.427 | 463.3 | 0.87605
|  | 381.7 | 29.4354 | 463.5 | 0.87164
|  | 381.8 | 29.4438 | 463.7 | 0.86722
|  | 381.9 | 29.4521 | 463.9 | 0.8628
|  | 382 | 29.4605 | 464.1 | 0.85838
|  | 382.1 | 29.4689 | 464.3 | 0.85396
|  | 382.2 | 29.4773 | 464.5 | 0.84954
|  | 382.3 | 29.4857 | 464.7 | 0.84512
|  | 382.4 | 29.4941 | 464.9 | 0.84071
|  | 382.5 | 29.5025 | 465.1 | 0.83629
|  | 382.6 | 29.4875 | 465.3 | 0.83187
|  | 382.7 | 29.493 | 465.5 | 0.82745
|  | 382.8 | 29.4985 | 465.7 | 0.82303
|  | 382.9 | 29.504 | 465.9 | 0.81861
|  | 383 | 29.5096 | 466.1 | 0.8142
|  | 383.1 | 29.5151 | 466.3 | 0.80978
|  | 383.2 | 29.5206 | 466.5 | 0.80536
|  | 383.3 | 29.5261 | 466.7 | 0.80094
|  | 383.4 | 29.5072 | 466.9 | 0.79652
|  | 383.5 | 29.5078 | 467.1 | 0.7921
|  | 383.6 | 29.5083 | 467.3 | 0.77913
|  | 383.7 | 29.5088 | 467.5 | 0.77561
|  | 383.8 | 29.5093 | 467.7 | 0.7709
|  | 383.9 | 29.5098 | 467.9 | 0.7662
|  | 384 | 29.5103 | 468.1 | 0.76149
|  | 384.1 | 29.5108 | 468.3 | 0.75679
|  | 384.2 | 29.5114 | 468.5 | 0.75209
|  | 384.3 | 29.5119 | 468.7 | 0.74738
|  | 384.4 | 29.5124 | 468.9 | 0.74268
|  | 384.5 | 29.4715 | 469.1 | 0.73797
|  | 384.6 | 29.4653 | 469.3 | 0.73327
|  | 384.7 | 29.4591 | 469.5 | 0.72856
|  | 384.8 | 29.4529 | 469.7 | 0.72386
|  | 384.9 | 29.4467 | 469.9 | 0.71916
|  | 385 | 29.4405 | 470.1 | 0.71445
|  | 385.1 | 29.4343 | 470.3 | 0.70975
|  | 385.2 | 29.4281 | 470.5 | 0.70504
|  | 385.3 | 29.4219 | 470.7 | 0.70034
|  | 385.4 | 29.4157 | 470.9 | 0.69564
|  | 385.5 | 29.4095 | 471.1 | 0.69093
|  | 385.6 | 29.4033 | 471.3 | 0.68623
|  | 385.7 | 29.3971 | 471.5 | 0.68152
|  | 385.8 | 29.3909 | 471.7 | 0.67682
|  | 385.9 | 29.3847 | 471.9 | 0.67211
|  | 386 | 29.3785 | 472.1 | 0.66741
|  | 386.1 | 29.3588 | 472.3 | 0.65719
|  | 386.2 | 29.3499 | 472.5 | 0.65268
|  | 386.3 | 29.3411 | 472.7 | 0.64746
|  | 386.4 | 29.3323 | 472.9 | 0.64224
|  | 386.5 | 29.3234 | 473.1 | 0.63702
|  | 386.6 | 29.3146 | 473.3 | 0.6318
|  | 386.7 | 29.3057 | 473.5 | 0.62657
|  | 386.8 | 29.2969 | 473.7 | 0.62135
|  | 386.9 | 29.2881 | 473.9 | 0.61613
|  | 387 | 29.2792 | 474.1 | 0.61091
|  | 387.1 | 29.2704 | 474.3 | 0.60569
|  | 387.2 | 29.2616 | 474.5 | 0.60046
|  | 387.3 | 29.2527 | 474.7 | 0.59524
|  | 387.4 | 29.2367 | 474.9 | 0.59002
|  | 387.5 | 29.2271 | 475.1 | 0.5848
|  | 387.6 | 29.2174 | 475.3 | 0.57958
|  | 387.7 | 29.2077 | 475.5 | 0.57436
|  | 387.8 | 29.198 | 475.7 | 0.56913
|  | 387.9 | 29.1884 | 475.9 | 0.56391
|  | 388 | 29.1787 | 476.1 | 0.55869
|  | 388.1 | 29.169 | 476.3 | 0.55347
|  | 388.2 | 29.1593 | 476.5 | 0.54825
|  | 388.3 | 29.1497 | 476.7 | 0.54302
|  | 388.4 | 29.14 | 476.9 | 0.5378
|  | 388.5 | 29.1303 | 477.1 | 0.53258
|  | 388.6 | 29.1206 | 477.3 | 0.52347
|  | 388.7 | 29.0717 | 477.5 | 0.5172
|  | 388.8 | 29.0576 | 477.7 | 0.51171
|  | 388.9 | 29.0435 | 477.9 | 0.50622
|  | 389 | 29.0295 | 478.1 | 0.50073
|  | 389.1 | 29.0154 | 478.3 | 0.49525
|  | 389.2 | 29.0013 | 478.5 | 0.48976
|  | 389.3 | 28.9873 | 478.7 | 0.48427
|  | 389.4 | 28.9732 | 478.9 | 0.47878
|  | 389.5 | 28.9591 | 479.1 | 0.47329
|  | 389.6 | 28.9038 | 479.3 | 0.4678
|  | 389.7 | 28.8803 | 479.5 | 0.46231
|  | 389.8 | 28.8567 | 479.7 | 0.45683
|  | 389.9 | 28.8331 | 479.9 | 0.45134
|  | 390 | 28.8096 | 480.1 | 0.44585
|  | 390.1 | 28.786 | 480.3 | 0.44036
|  | 390.2 | 28.7625 | 480.5 | 0.43487
|  | 390.3 | 28.7389 | 480.7 | 0.42938
|  | 390.4 | 28.7153 | 480.9 | 0.42389
|  | 390.5 | 28.6918 | 481.1 | 0.41841
|  | 390.6 | 28.6682 | 481.3 | 0.41292
|  | 390.7 | 28.6447 | 481.5 | 0.40743
|  | 390.8 | 28.6211 | 481.7 | 0.40194
|  | 390.9 | 28.5976 | 481.9 | 0.35972
|  | 391 | 28.574 | 482.1 | 0.35705
|  | 391.1 | 28.5504 | 482.3 | 0.34944
|  | 391.2 | 28.5269 | 482.5 | 0.34183
|  | 391.3 | 28.5033 | 482.7 | 0.33422
|  | 391.4 | 28.4798 | 482.9 | 0.3266
|  | 391.5 | 28.4562 | 483.1 | 0.31899
|  | 391.6 | 28.4326 | 483.3 | 0.31138
|  | 391.7 | 28.4091 | 483.5 | 0.30376
|  | 391.8 | 28.384 | 483.7 | 0.29615
|  | 391.9 | 28.3596 | 483.9 | 0.28854
|  | 392 | 28.3352 | 484.1 | 0.28092
|  | 392.1 | 28.3108 | 484.3 | 0.27331
|  | 392.2 | 28.2864 | 484.5 | 0.2657
|  | 392.3 | 28.262 | 484.7 | 0.25809
|  | 392.4 | 28.2376 | 484.9 | 0.25047
|  | 392.5 | 28.2132 | 485.1 | 0.24286
|  | 392.6 | 28.1888 | 485.3 | 0.23525
|  | 392.7 | 28.1644 | 485.5 | 0.22763
|  | 392.8 | 28.1401 | 485.7 | 0.22002
|  | 392.9 | 28.1157 | 485.9 | 0.21798
|  | 393 | 28.0913 | 486.1 | 0.21096
|  | 393.1 | 28.0669 | 486.3 | 0.20284
|  | 393.2 | 28.0425 | 486.5 | 0.19471
|  | 393.3 | 28.0181 | 486.7 | 0.18658
|  | 393.4 | 27.9937 | 486.9 | 0.17846
|  | 393.5 | 27.9693 | 487.1 | 0.17033
|  | 393.6 | 27.8395 | 487.3 | 0.1622
|  | 393.7 | 27.8073 | 487.5 | 0.15408
|  | 393.8 | 27.7751 | 487.7 | 0.14595
|  | 393.9 | 27.7429 | 487.9 | 0.13783
|  | 394 | 27.7107 | 488.1 | 0.1297
|  | 394.1 | 27.6785 | 488.3 | 0.12157
|  | 394.2 | 27.6463 | 488.5 | 0.11345
|  | 394.3 | 27.6141 | 488.7 | 0.10532
|  | 394.4 | 27.5819 | 488.9 | 0.097194
|  | 394.5 | 27.5497 | 489.1 | 0.089067
|  | 394.6 | 27.5175 | 489.3 | 0.080941
|  | 394.7 | 27.4853 | 489.5 | 0.072815
|  | 394.8 | 27.4531 | 489.7 | 0.064688
|  | 394.9 | 27.4209 | 489.9 | 0.10453
|  | 395 | 27.3887 | 490.1 | 0.11109
|  | 395.1 | 27.3159 | 490.3 | 0.10805
|  | 395.2 | 27.2745 | 490.5 | 0.10502
|  | 395.3 | 27.2331 | 490.7 | 0.10199
|  | 395.4 | 27.1918 | 490.9 | 0.098953
|  | 395.5 | 27.1504 | 491.1 | 0.095919
|  | 395.6 | 27.109 | 491.3 | 0.092885
|  | 395.7 | 27.0676 | 491.5 | 0.089851
|  | 395.8 | 27.0262 | 491.7 | 0.086817
|  | 395.9 | 26.9848 | 491.9 | 0.083783
|  | 396 | 26.9435 | 492.1 | 0.080749
|  | 396.1 | 26.9021 | 492.3 | 0.077715
|  | 396.2 | 26.8607 | 492.5 | 0.074681
|  | 396.3 | 26.8193 | 492.7 | 0.071647
|  | 396.4 | 26.7779 | 492.9 | 0.068613
|  | 396.5 | 26.7365 | 493.1 | 0.065579
|  | 396.6 | 26.6508 | 493.3 | 0.062545
|  | 396.7 | 26.6071 | 493.5 | 0.059511
|  | 396.8 | 26.5633 | 493.7 | 0.056477
|  | 396.9 | 26.5196 | 493.9 | 0.053443
|  | 397 | 26.4759 | 494.1 | 0.050409
|  | 397.1 | 26.4105 | 494.3 | 0.047375
|  | 397.2 | 26.3549 | 494.5 | 0.063232
|  | 397.3 | 26.2993 | 494.7 | 0.066945
|  | 397.4 | 26.2437 | 494.9 | 0.06688
|  | 397.5 | 26.1882 | 495.1 | 0.066814
|  | 397.6 | 26.1326 | 495.3 | 0.066749
|  | 397.7 | 26.077 | 495.5 | 0.066684
|  | 397.8 | 26.0372 | 495.7 | 0.066619
|  | 397.9 | 25.9876 | 495.9 | 0.066554
|  | 398 | 25.9381 | 496.1 | 0.066488
|  | 398.1 | 25.8885 | 496.3 | 0.066423
|  | 398.2 | 25.8389 | 496.5 | 0.066358
|  | 398.3 | 25.7893 | 496.7 | 0.066293
|  | 398.4 | 25.7398 | 496.9 | 0.066228
|  | 398.5 | 25.6902 | 497.1 | 0.066162
|  | 398.6 | 25.6406 | 497.3 | 0.066097
|  | 398.7 | 25.5911 | 497.5 | 0.066032
|  | 398.8 | 25.5415 | 497.7 | 0.065967
|  | 398.9 | 25.4919 | 497.9 | 0.065901
|  | 399 | 25.4423 | 498.1 | 0.065836
|  | 399.1 | 25.382 | 498.3 | 0.065771
|  | 399.2 | 25.3315 | 498.5 | 0.065706
|  | 399.3 | 25.2809 | 498.7 | 0.065641
|  | 399.4 | 25.2304 | 498.9 | 0.065575
|  | 399.5 | 25.1799 | 499.1 | 0.06551
|  | 399.6 | 25.1294 | 499.3 | 0.065445
|  | 399.7 | 25.0789 | 499.5 | 0.06538
|  | 399.8 | 25.0284 | 499.7 | 0.065314
|  | 399.9 | 24.9779 | 499.9 | 0.065249
|  | 400 | 24.9274 | 500.1 | 0.065184
|  | 400.1 | 24.8769 | 500.3 | 0.065119
|  | 400.2 | 24.7914 | 500.5 | 0.065054
|  | 400.3 | 24.7356
|  | 400.4 | 24.6799
|  | 400.5 | 24.6241
|  | 400.6 | 24.5684
|  | 400.7 | 24.5126
|  | 400.8 | 24.4569
|  | 400.9 | 24.4011
|  | 401 | 24.3454
|  | 401.1 | 24.2741
|  | 401.2 | 24.2142
|  | 401.3 | 24.1543
|  | 401.4 | 24.0944
|  | 401.5 | 24.0345
|  | 401.6 | 23.9746
|  | 401.7 | 23.9147
|  | 401.8 | 23.8548
|  | 401.9 | 23.7949
|  | 402 | 23.735
|  | 402.1 | 23.6751
|  | 402.2 | 23.6218
|  | 402.3 | 23.5626
|  | 402.4 | 23.5035
|  | 402.5 | 23.4443
|  | 402.6 | 23.3851
|  | 402.7 | 23.3259
|  | 402.8 | 23.2668
|  | 402.9 | 23.2076
|  | 403 | 23.1484
|  | 403.1 | 23.1303
|  | 403.2 | 23.0779
|  | 403.3 | 23.0256
|  | 403.4 | 22.9733
|  | 403.5 | 22.9209
|  | 403.6 | 22.8686
|  | 403.7 | 22.8163
|  | 403.8 | 22.7639
|  | 403.9 | 22.7116
|  | 404 | 22.6593
|  | 404.1 | 22.6069
|  | 404.2 | 22.6189
|  | 404.3 | 22.5773
|  | 404.4 | 22.5357
|  | 404.5 | 22.4941
|  | 404.6 | 22.4525
|  | 404.7 | 22.4109
|  | 404.8 | 22.3693
|  | 404.9 | 22.3277
|  | 405 | 22.286
|  | 405.1 | 22.2444
|  | 405.2 | 22.2028
|  | 405.3 | 22.1612
|  | 405.4 | 22.1196
|  | 405.5 | 22.078
|  | 405.6 | 21.9432
|  | 405.7 | 21.8959
|  | 405.8 | 21.8485
|  | 405.9 | 21.8012
|  | 406 | 21.7539
|  | 406.1 | 21.7065
|  | 406.2 | 21.6677
|  | 406.3 | 21.6155
|  | 406.4 | 21.5634
|  | 406.5 | 21.5112
|  | 406.6 | 21.4591
|  | 406.7 | 21.4069
|  | 406.8 | 21.3547
|  | 406.9 | 21.3026
|  | 407 | 21.2504
|  | 407.1 | 21.1983
|  | 407.2 | 21.1461
|  | 407.3 | 21.094
|  | 407.4 | 21.0418
|  | 407.5 | 20.9896
|  | 407.6 | 20.9375
|  | 407.7 | 20.8853
|  | 407.8 | 20.8332
|  | 407.9 | 20.781
|  | 408 | 20.7289
|  | 408.1 | 20.6767
|  | 408.2 | 20.6245
|  | 408.3 | 20.5724
|  | 408.4 | 20.5202
|  | 408.5 | 20.4681
|  | 408.6 | 20.4009
|  | 408.7 | 20.351
|  | 408.8 | 20.3011
|  | 408.9 | 20.2511
|  | 409 | 20.2012
|  | 409.1 | 20.1513
|  | 409.2 | 20.1014
|  | 409.3 | 20.0514
|  | 409.4 | 20.0015
|  | 409.5 | 19.9516
|  | 409.6 | 19.9016
|  | 409.7 | 19.8517
|  | 409.8 | 19.8018
|  | 409.9 | 19.7519
|  | 410 | 19.7019
|  | 410.1 | 19.652
|  | 410.2 | 19.6021
|  | 410.3 | 19.5522
|  | 410.4 | 19.5022
|  | 410.5 | 19.4523
|  | 410.6 | 19.4024
|  | 410.7 | 19.3524
|  | 410.8 | 19.3025
|  | 410.9 | 19.2373
|  | 411 | 19.1854
|  | 411.1 | 19.1334
|  | 411.2 | 19.0815
|  | 411.3 | 19.0295
|  | 411.4 | 18.9776
|  | 411.5 | 18.9256
|  | 411.6 | 18.8736
|  | 411.7 | 18.8217
|  | 411.8 | 18.7697
|  | 411.9 | 18.7178
|  | 412 | 18.6658
|  | 412.1 | 18.6138
|  | 412.2 | 18.5619
|  | 412.3 | 18.5099
|  | 412.4 | 18.458
|  | 412.5 | 18.406
|  | 412.6 | 18.354
|  | 412.7 | 18.3021
|  | 412.8 | 18.2501
|  | 412.9 | 18.1693
|  | 413 | 18.1154
|  | 413.1 | 18.0614
|  | 413.2 | 18.0075
|  | 413.3 | 17.9536
|  | 413.4 | 17.8996
|  | 413.5 | 17.8457
|  | 413.6 | 17.7917
|  | 413.7 | 17.7378
|  | 413.8 | 17.6838
|  | 413.9 | 17.6299
|  | 414 | 17.5759
|  | 414.1 | 17.522
|  | 414.2 | 17.468
|  | 414.3 | 17.4141
|  | 414.4 | 17.3601
|  | 414.5 | 17.3062
|  | 414.6 | 17.2522
|  | 414.7 | 17.1983
|  | 414.8 | 17.1443
|  | 414.9 | 17.08
|  | 415 | 17.0242
|  | 415.1 | 16.9684
|  | 415.2 | 16.9126
|  | 415.3 | 16.8568
|  | 415.4 | 16.801
|  | 415.5 | 16.7452
|  | 415.6 | 16.6894
|  | 415.7 | 16.6336
|  | 415.8 | 16.5778
|  | 415.9 | 16.522
|  | 416 | 16.4662
|  | 416.1 | 16.4104
|  | 416.2 | 16.3545
|  | 416.3 | 16.2987
|  | 416.4 | 16.2429
|  | 416.5 | 16.1871
|  | 416.6 | 16.1313
|  | 416.7 | 16.0755
|  | 416.8 | 16.0197
|  | 416.9 | 15.9109
|  | 417 | 15.8515
|  | 417.1 | 15.7921
|  | 417.2 | 15.7326
|  | 417.3 | 15.6732
|  | 417.4 | 15.6138
|  | 417.5 | 15.5544
|  | 417.6 | 15.495
|  | 417.7 | 15.4356
|  | 417.8 | 15.3761
|  | 417.9 | 15.3167
|  | 418 | 15.2573
|  | 418.1 | 15.1979
|  | 418.2 | 15.1385
|  | 418.3 | 15.079
|  | 418.4 | 15.0196
|  | 418.5 | 14.9602
|  | 418.6 | 14.8336
|  | 418.7 | 14.766
|  | 418.8 | 14.6984
|  | 418.9 | 14.6309
|  | 419 | 14.5633
|  | 419.1 | 14.4958
|  | 419.2 | 14.4282
|  | 419.3 | 14.3606
|  | 419.4 | 14.2931
|  | 419.5 | 14.2255
|  | 419.6 | 14.1579
|  | 419.7 | 14.0904
|  | 419.8 | 14.0228
|  | 419.9 | 13.9553
|  | 420 | 13.8877
|  | 420.1 | 13.8201
|  | 420.2 | 13.7526
|  | 420.3 | 13.7061
|  | 420.4 | 13.6389
|  | 420.5 | 13.5716
|  | 420.6 | 13.5044
|  | 420.7 | 13.4372
|  | 420.8 | 13.37
|  | 420.9 | 13.3028
|  | 421 | 13.2356
|  | 421.1 | 13.1684
|  | 421.2 | 13.1012
|  | 421.3 | 13.0339
|  | 421.4 | 12.9667
|  | 421.5 | 12.8995
|  | 421.6 | 12.8323
|  | 421.7 | 12.7651
|  | 421.8 | 12.6979
|  | 421.9 | 12.573
|  | 422 | 12.5025
|  | 422.1 | 12.4321
|  | 422.2 | 12.3617
|  | 422.3 | 12.2913
|  | 422.4 | 12.2208
|  | 422.5 | 12.1504
|  | 422.6 | 12.08
|  | 422.7 | 12.0095
|  | 422.8 | 11.9391
|  | 422.9 | 11.8687
|  | 423 | 11.7982
|  | 423.1 | 11.7278
|  | 423.2 | 11.6574
|  | 423.3 | 11.4549
|  | 423.4 | 11.3672
|  | 423.5 | 11.2795
|  | 423.6 | 11.1919
|  | 423.7 | 11.1042
|  | 423.8 | 11.0166
|  | 423.9 | 10.9289
|  | 424 | 10.8413
|  | 424.1 | 10.7536
|  | 424.2 | 10.666
|  | 424.3 | 10.7106
|  | 424.4 | 10.6384
|  | 424.5 | 10.5661
|  | 424.6 | 10.4939
|  | 424.7 | 10.4216
|  | 424.8 | 10.3494
|  | 424.9 | 10.2771
|  | 425 | 10.2048
|  | 425.1 | 10.1326
|  | 425.2 | 10.0603
|  | 425.3 | 9.9881
|  | 425.4 | 9.9158
|  | 425.5 | 9.8435
|  | 425.6 | 9.7713
|  | 425.7 | 9.699
|  | 425.8 | 9.6268
|  | 425.9 | 9.5545
|  | 426 | 9.4822
|  | 426.1 | 9.41
|  | 426.2 | 9.3377
|  | 426.3 | 9.197
|  | 426.4 | 9.1269
|  | 426.5 | 9.0569
|  | 426.6 | 8.9868
|  | 426.7 | 8.9168
|  | 426.8 | 8.8467
|  | 426.9 | 8.7767
|  | 427 | 8.7066
|  | 427.1 | 8.6366
|  | 427.2 | 8.5665
|  | 427.3 | 8.4965
|  | 427.4 | 8.4264
|  | 427.5 | 8.3564
|  | 427.6 | 8.3061
|  | 427.7 | 8.2324
|  | 427.8 | 8.1586
|  | 427.9 | 8.0849
|  | 428 | 8.0111
|  | 428.1 | 7.9374
|  | 428.2 | 7.8636
|  | 428.3 | 7.7899
|  | 428.4 | 7.7162
|  | 428.5 | 7.6424
|  | 428.6 | 7.5687
|  | 428.7 | 7.4949
|  | 428.8 | 7.4212
|  | 428.9 | 7.3474
|  | 429 | 7.2737
|  | 429.1 | 7.1999
|  | 429.2 | 7.1262
|  | 429.3 | 7.1123
|  | 429.4 | 7.0477
|  | 429.5 | 6.9832
|  | 429.6 | 6.9187
|  | 429.7 | 6.8542
|  | 429.8 | 6.7896
|  | 429.9 | 6.7251
|  | 430 | 6.6606
|  | 430.1 | 6.5961
|  | 430.2 | 6.5316
|  | 430.3 | 6.467
|  | 430.4 | 6.4025
|  | 430.5 | 6.338
|  | 430.6 | 6.2735
|  | 430.7 | 6.209
|  | 430.8 | 6.1444
|  | 430.9 | 6.1486
|  | 431 | 6.0908
|  | 431.1 | 6.033
|  | 431.2 | 5.9752
|  | 431.3 | 5.9175
|  | 431.4 | 5.8597
|  | 431.5 | 5.8019
|  | 431.6 | 5.7441
|  | 431.7 | 5.6864
|  | 431.8 | 5.6286
|  | 431.9 | 5.5708
|  | 432 | 5.513
|  | 432.1 | 5.4553
|  | 432.2 | 5.3975
|  | 432.3 | 5.3397
|  | 432.4 | 5.2819
|  | 432.5 | 5.2242
|  | 432.6 | 5.1664
|  | 432.7 | 5.1086
|  | 432.8 | 5.0508
|  | 432.9 | 4.9931
|  | 433 | 5.0423
|  | 433.1 | 4.9941
|  | 433.2 | 4.946
|  | 433.3 | 4.8978
|  | 433.4 | 4.8496
|  | 433.5 | 4.8014
|  | 433.6 | 4.7532
|  | 433.7 | 4.705
|  | 433.8 | 4.6568
|  | 433.9 | 4.6086
|  | 434 | 4.5604
|  | 434.1 | 4.5122
|  | 434.2 | 4.464
|  | 434.3 | 4.4158
|  | 434.4 | 4.3676
|  | 434.5 | 4.3194
|  | 434.6 | 4.3377
|  | 434.7 | 4.2984
|  | 434.8 | 4.2591
|  | 434.9 | 4.2197
|  | 435 | 4.1804
|  | 435.1 | 4.141
|  | 435.2 | 4.1017
|  | 435.3 | 4.0623
|  | 435.4 | 4.023
|  | 435.5 | 3.9836
|  | 435.6 | 3.9443
|  | 435.7 | 3.9049
|  | 435.8 | 3.8656
|  | 435.9 | 3.8262
|  | 436 | 3.7869
|  | 436.1 | 3.7475
|  | 436.2 | 3.7082
|  | 436.3 | 3.6688
|  | 436.4 | 3.686
|  | 436.5 | 3.6523
|  | 436.6 | 3.6186
|  | 436.7 | 3.585
|  | 436.8 | 3.5513
|  | 436.9 | 3.5176
|  | 437 | 3.484
|  | 437.1 | 3.4503
|  | 437.2 | 3.4166
|  | 437.3 | 3.383
|  | 437.4 | 3.3493
|  | 437.5 | 3.3157
|  | 437.6 | 3.282
|  | 437.7 | 3.2483
|  | 437.8 | 3.2147
|  | 437.9 | 3.181
|  | 438 | 3.159
|  | 438.1 | 3.128
|  | 438.2 | 3.097
|  | 438.3 | 3.066
|  | 438.4 | 3.0351
|  | 438.5 | 3.0041
|  | 438.6 | 2.9731
|  | 438.7 | 2.9421
|  | 438.8 | 2.9111
|  | 438.9 | 2.8801
|  | 439 | 2.8492
|  | 439.1 | 2.8182
|  | 439.2 | 2.7872
|  | 439.3 | 2.7562
|  | 439.4 | 2.7252
|  | 439.5 | 2.6942
|  | 439.6 | 2.6633
|  | 439.7 | 2.6323
|  | 439.8 | 2.6013
|  | 439.9 | 2.5703
|  | 440 | 2.5393
|  | 440.1 | 2.5083
|  | 440.2 | 2.4774
|  | 440.3 | 2.4464
|  | 440.4 | 2.488
|  | 440.5 | 2.4607
|  | 440.6 | 2.4334
|  | 440.7 | 2.4062
|  | 440.8 | 2.3789
|  | 440.9 | 2.3516
|  | 441 | 2.3243
|  | 441.1 | 2.2971
|  | 441.2 | 2.2698
|  | 441.3 | 2.2425
|  | 441.4 | 2.2153
|  | 441.5 | 2.188
|  | 441.6 | 2.1607
|  | 441.7 | 2.2237
|  | 441.8 | 2.2101
|  | 441.9 | 2.1964
|  | 442 | 2.1828
|  | 442.1 | 2.1692
|  | 442.2 | 2.1555
|  | 442.3 | 2.1419
|  | 442.4 | 2.1282
|  | 442.5 | 2.1146
|  | 442.6 | 2.101
|  | 442.7 | 2.0873
|  | 442.8 | 2.0737
|  | 442.9 | 2.0601
|  | 443 | 2.0464
|  | 443.1 | 2.0328
|  | 443.2 | 2.0191
|  | 443.3 | 2.0055
|  | 443.4 | 1.9919
|  | 443.5 | 1.9782
|  | 443.6 | 1.9646
|  | 443.7 | 1.9509
|  | 443.8 | 1.9457
|  | 443.9 | 1.934
|  | 444 | 1.9223
|  | 444.1 | 1.9106
|  | 444.2 | 1.8989
|  | 444.3 | 1.8872
|  | 444.4 | 1.8754
|  | 444.5 | 1.8637
|  | 444.6 | 1.852
|  | 444.7 | 1.8403
|  | 444.8 | 1.8286
|  | 444.9 | 1.8169
|  | 445 | 1.8052
|  | 445.1 | 1.7935
|  | 445.2 | 1.7818
|  | 445.3 | 1.77
|  | 445.4 | 1.7583
|  | 445.5 | 1.7466
|  | 445.6 | 1.7349
|  | 445.7 | 1.7232
|  | 445.8 | 1.7458
|  | 445.9 | 1.7364
|  | 446 | 1.727
|  | 446.1 | 1.7176
|  | 446.2 | 1.7082
|  | 446.3 | 1.6988
|  | 446.4 | 1.6894
|  | 446.5 | 1.68
|  | 446.6 | 1.6706
|  | 446.7 | 1.6612
|  | 446.8 | 1.6518
|  | 446.9 | 1.6424
|  | 447 | 1.633
|  | 447.1 | 1.6236
|  | 447.2 | 1.6142
|  | 447.3 | 1.6048
|  | 447.4 | 1.5954
|  | 447.5 | 1.586
|  | 447.6 | 1.5766
|  | 447.7 | 1.5672
|  | 447.8 | 1.5578
|  | 447.9 | 1.5484
|  | 448 | 1.539
|  | 448.1 | 1.5296
|  | 448.2 | 1.5202
|  | 448.3 | 1.5107
|  | 448.4 | 1.5013
|  | 448.5 | 1.518
|  | 448.6 | 1.511
|  | 448.7 | 1.504
|  | 448.8 | 1.497
|  | 448.9 | 1.49
|  | 449 | 1.483
|  | 449.1 | 1.476
|  | 449.2 | 1.4691
|  | 449.3 | 1.4621
|  | 449.4 | 1.4551
|  | 449.5 | 1.4481
|  | 449.6 | 1.4411
|  | 449.7 | 1.4341
|  | 449.8 | 1.4271
|  | 449.9 | 1.4201
|  | 450 | 1.4131
|  | 450.1 | 1.4062
|  | 450.2 | 1.3992
|  | 450.3 | 1.3922
|  | 450.4 | 1.3852
|  | 450.5 | 1.3782
|  | 450.6 | 1.3712
|  | 450.7 | 1.3642
|  | 450.8 | 1.3572
|  | 450.9 | 1.333
|  | 451 | 1.3252
|  | 451.1 | 1.3174
|  | 451.2 | 1.3096
|  | 451.3 | 1.3018
|  | 451.4 | 1.2941
|  | 451.5 | 1.2863
|  | 451.6 | 1.2785
|  | 451.7 | 1.2707
|  | 451.8 | 1.2629
|  | 451.9 | 1.2551
|  | 452 | 1.2474
|  | 452.1 | 1.2396
|  | 452.2 | 1.2318
|  | 452.3 | 1.224
|  | 452.4 | 1.2162
|  | 452.5 | 1.2085
|  | 452.6 | 1.2007
|  | 452.7 | 1.1929
|  | 452.8 | 1.1851
|  | 452.9 | 1.1773
|  | 453 | 1.1696
|  | 453.1 | 1.1618
|  | 453.2 | 1.154
|  | 453.3 | 1.1958
|  | 453.4 | 1.1908
|  | 453.5 | 1.1858
|  | 453.6 | 1.1807
|  | 453.7 | 1.1757
|  | 453.8 | 1.1707
|  | 453.9 | 1.1657
|  | 454 | 1.1607
|  | 454.1 | 1.1556
|  | 454.2 | 1.1506
|  | 454.3 | 1.1456
|  | 454.4 | 1.1406
|  | 454.5 | 1.1356
|  | 454.6 | 1.1305
|  | 454.7 | 1.1255
|  | 454.8 | 1.1205
|  | 454.9 | 1.1155
|  | 455 | 1.1105
|  | 455.1 | 1.1054
|  | 455.2 | 1.1004
|  | 455.3 | 1.0954
|  | 455.4 | 1.0904
|  | 455.5 | 1.0854
|  | 455.6 | 1.0803
|  | 455.7 | 1.0753
|  | 455.8 | 1.0703
|  | 455.9 | 1.0653
|  | 456 | 1.0603
|  | 456.1 | 1.0552
|  | 456.2 | 1.0502
|  | 456.3 | 1.0452
|  | 456.4 | 1.0402
|  | 456.5 | 1.0352
|  | 456.6 | 1.0301
|  | 456.7 | 1.0251
|  | 456.8 | 1.0201
|  | 456.9 | 1.0151
|  | 457 | 1.0101
|  | 457.1 | 1.005
|  | 457.2 | 1
|  | 457.3 | 0.995
|  | 457.4 | 0.98998
|  | 457.5 | 0.98496
|  | 457.6 | 1.0144
|  | 457.7 | 1.0117
|  | 457.8 | 1.0089
|  | 457.9 | 1.0062
|  | 458 | 1.0035
|  | 458.1 | 1.0007
|  | 458.2 | 0.99799
|  | 458.3 | 0.99525
|  | 458.4 | 0.99252
|  | 458.5 | 0.98978
|  | 458.6 | 0.98705
|  | 458.7 | 0.98431
|  | 458.8 | 0.98158
|  | 458.9 | 0.97884
|  | 459 | 0.9761
|  | 459.1 | 0.97337
|  | 459.2 | 0.97063
|  | 459.3 | 0.9679
|  | 459.4 | 0.96516
|  | 459.5 | 0.96243
|  | 459.6 | 0.95969
|  | 459.7 | 0.95696
|  | 459.8 | 0.95422
|  | 459.9 | 0.95148
|  | 460 | 0.94875
|  | 460.1 | 0.94601
|  | 460.2 | 0.94328
|  | 460.3 | 0.94054
|  | 460.4 | 0.93781
|  | 460.5 | 0.93507
|  | 460.6 | 0.93234
|  | 460.7 | 0.9296
|  | 460.8 | 0.92687
|  | 460.9 | 0.92413
|  | 461 | 0.92139
|  | 461.1 | 0.91866
|  | 461.2 | 0.91592
|  | 461.3 | 0.91319
|  | 461.4 | 0.91045
|  | 461.5 | 0.90772
|  | 461.6 | 0.90498
|  | 461.7 | 0.9114
|  | 461.8 | 0.90919
|  | 461.9 | 0.90698
|  | 462 | 0.90477
|  | 462.1 | 0.90256
|  | 462.2 | 0.90035
|  | 462.3 | 0.89815
|  | 462.4 | 0.89594
|  | 462.5 | 0.89373
|  | 462.6 | 0.89152
|  | 462.7 | 0.88931
|  | 462.8 | 0.8871
|  | 462.9 | 0.88489
|  | 463 | 0.88268
|  | 463.1 | 0.88047
|  | 463.2 | 0.87826
|  | 463.3 | 0.87605
|  | 463.4 | 0.87384
|  | 463.5 | 0.87164
|  | 463.6 | 0.86943
|  | 463.7 | 0.86722
|  | 463.8 | 0.86501
|  | 463.9 | 0.8628
|  | 464 | 0.86059
|  | 464.1 | 0.85838
|  | 464.2 | 0.85617
|  | 464.3 | 0.85396
|  | 464.4 | 0.85175
|  | 464.5 | 0.84954
|  | 464.6 | 0.84733
|  | 464.7 | 0.84512
|  | 464.8 | 0.84292
|  | 464.9 | 0.84071
|  | 465 | 0.8385
|  | 465.1 | 0.83629
|  | 465.2 | 0.83408
|  | 465.3 | 0.83187
|  | 465.4 | 0.82966
|  | 465.5 | 0.82745
|  | 465.6 | 0.82524
|  | 465.7 | 0.82303
|  | 465.8 | 0.82082
|  | 465.9 | 0.81861
|  | 466 | 0.81641
|  | 466.1 | 0.8142
|  | 466.2 | 0.81199
|  | 466.3 | 0.80978
|  | 466.4 | 0.80757
|  | 466.5 | 0.80536
|  | 466.6 | 0.80315
|  | 466.7 | 0.80094
|  | 466.8 | 0.79873
|  | 466.9 | 0.79652
|  | 467 | 0.79431
|  | 467.1 | 0.7921
|  | 467.2 | 0.78266
|  | 467.3 | 0.78031
|  | 467.4 | 0.77796
|  | 467.5 | 0.77561
|  | 467.6 | 0.77325
|  | 467.7 | 0.7709
|  | 467.8 | 0.76855
|  | 467.9 | 0.7662
|  | 468 | 0.76385
|  | 468.1 | 0.76149
|  | 468.2 | 0.75914
|  | 468.3 | 0.75679
|  | 468.4 | 0.75444
|  | 468.5 | 0.75209
|  | 468.6 | 0.74973
|  | 468.7 | 0.74738
|  | 468.8 | 0.74503
|  | 468.9 | 0.74268
|  | 469 | 0.74033
|  | 469.1 | 0.73797
|  | 469.2 | 0.73562
|  | 469.3 | 0.73327
|  | 469.4 | 0.73092
|  | 469.5 | 0.72856
|  | 469.6 | 0.72621
|  | 469.7 | 0.72386
|  | 469.8 | 0.72151
|  | 469.9 | 0.71916
|  | 470 | 0.7168
|  | 470.1 | 0.71445
|  | 470.2 | 0.7121
|  | 470.3 | 0.70975
|  | 470.4 | 0.7074
|  | 470.5 | 0.70504
|  | 470.6 | 0.70269
|  | 470.7 | 0.70034
|  | 470.8 | 0.69799
|  | 470.9 | 0.69564
|  | 471 | 0.69328
|  | 471.1 | 0.69093
|  | 471.2 | 0.68858
|  | 471.3 | 0.68623
|  | 471.4 | 0.68388
|  | 471.5 | 0.68152
|  | 471.6 | 0.67917
|  | 471.7 | 0.67682
|  | 471.8 | 0.67447
|  | 471.9 | 0.67211
|  | 472 | 0.66976
|  | 472.1 | 0.66741
|  | 472.2 | 0.66052
|  | 472.3 | 0.65791
|  | 472.4 | 0.6553
|  | 472.5 | 0.65268
|  | 472.6 | 0.65007
|  | 472.7 | 0.64746
|  | 472.8 | 0.64485
|  | 472.9 | 0.64224
|  | 473 | 0.63963
|  | 473.1 | 0.63702
|  | 473.2 | 0.63441
|  | 473.3 | 0.6318
|  | 473.4 | 0.62919
|  | 473.5 | 0.62657
|  | 473.6 | 0.62396
|  | 473.7 | 0.62135
|  | 473.8 | 0.61874
|  | 473.9 | 0.61613
|  | 474 | 0.61352
|  | 474.1 | 0.61091
|  | 474.2 | 0.6083
|  | 474.3 | 0.60569
|  | 474.4 | 0.60308
|  | 474.5 | 0.60046
|  | 474.6 | 0.59785
|  | 474.7 | 0.59524
|  | 474.8 | 0.59263
|  | 474.9 | 0.59002
|  | 475 | 0.58741
|  | 475.1 | 0.5848
|  | 475.2 | 0.58219
|  | 475.3 | 0.57958
|  | 475.4 | 0.57697
|  | 475.5 | 0.57436
|  | 475.6 | 0.57174
|  | 475.7 | 0.56913
|  | 475.8 | 0.56652
|  | 475.9 | 0.56391
|  | 476 | 0.5613
|  | 476.1 | 0.55869
|  | 476.2 | 0.55608
|  | 476.3 | 0.55347
|  | 476.4 | 0.55086
|  | 476.5 | 0.54825
|  | 476.6 | 0.54563
|  | 476.7 | 0.54302
|  | 476.8 | 0.54041
|  | 476.9 | 0.5378
|  | 477 | 0.53519
|  | 477.1 | 0.53258
|  | 477.2 | 0.52997
|  | 477.3 | 0.52269
|  | 477.4 | 0.51994
|  | 477.5 | 0.5172
|  | 477.6 | 0.51446
|  | 477.7 | 0.51171
|  | 477.8 | 0.50897
|  | 477.9 | 0.50622
|  | 478 | 0.50348
|  | 478.1 | 0.50073
|  | 478.2 | 0.49799
|  | 478.3 | 0.49525
|  | 478.4 | 0.4925
|  | 478.5 | 0.48976
|  | 478.6 | 0.48701
|  | 478.7 | 0.48427
|  | 478.8 | 0.48152
|  | 478.9 | 0.47878
|  | 479 | 0.47604
|  | 479.1 | 0.47329
|  | 479.2 | 0.47055
|  | 479.3 | 0.4678
|  | 479.4 | 0.46506
|  | 479.5 | 0.46231
|  | 479.6 | 0.45957
|  | 479.7 | 0.45683
|  | 479.8 | 0.45408
|  | 479.9 | 0.45134
|  | 480 | 0.44859
|  | 480.1 | 0.44585
|  | 480.2 | 0.4431
|  | 480.3 | 0.44036
|  | 480.4 | 0.43762
|  | 480.5 | 0.43487
|  | 480.6 | 0.43213
|  | 480.7 | 0.42938
|  | 480.8 | 0.42664
|  | 480.9 | 0.42389
|  | 481 | 0.42115
|  | 481.1 | 0.41841
|  | 481.2 | 0.41566
|  | 481.3 | 0.41292
|  | 481.4 | 0.41017
|  | 481.5 | 0.40743
|  | 481.6 | 0.40468
|  | 481.7 | 0.40194
|  | 481.8 | 0.36847
|  | 481.9 | 0.36467
|  | 482 | 0.36086
|  | 482.1 | 0.35705
|  | 482.2 | 0.35325
|  | 482.3 | 0.34944
|  | 482.4 | 0.34564
|  | 482.5 | 0.34183
|  | 482.6 | 0.33802
|  | 482.7 | 0.33422
|  | 482.8 | 0.33041
|  | 482.9 | 0.3266
|  | 483 | 0.3228
|  | 483.1 | 0.31899
|  | 483.2 | 0.31518
|  | 483.3 | 0.31138
|  | 483.4 | 0.30757
|  | 483.5 | 0.30376
|  | 483.6 | 0.29996
|  | 483.7 | 0.29615
|  | 483.8 | 0.29234
|  | 483.9 | 0.28854
|  | 484 | 0.28473
|  | 484.1 | 0.28092
|  | 484.2 | 0.27712
|  | 484.3 | 0.27331
|  | 484.4 | 0.26951
|  | 484.5 | 0.2657
|  | 484.6 | 0.26189
|  | 484.7 | 0.25809
|  | 484.8 | 0.25428
|  | 484.9 | 0.25047
|  | 485 | 0.24667
|  | 485.1 | 0.24286
|  | 485.2 | 0.23905
|  | 485.3 | 0.23525
|  | 485.4 | 0.23144
|  | 485.5 | 0.22763
|  | 485.6 | 0.22383
|  | 485.7 | 0.22002
|  | 485.8 | 0.21621
|  | 485.9 | 0.21909
|  | 486 | 0.21503
|  | 486.1 | 0.21096
|  | 486.2 | 0.2069
|  | 486.3 | 0.20284
|  | 486.4 | 0.19877
|  | 486.5 | 0.19471
|  | 486.6 | 0.19065
|  | 486.7 | 0.18658
|  | 486.8 | 0.18252
|  | 486.9 | 0.17846
|  | 487 | 0.17439
|  | 487.1 | 0.17033
|  | 487.2 | 0.16627
|  | 487.3 | 0.1622
|  | 487.4 | 0.15814
|  | 487.5 | 0.15408
|  | 487.6 | 0.15002
|  | 487.7 | 0.14595
|  | 487.8 | 0.14189
|  | 487.9 | 0.13783
|  | 488 | 0.13376
|  | 488.1 | 0.1297
|  | 488.2 | 0.12564
|  | 488.3 | 0.12157
|  | 488.4 | 0.11751
|  | 488.5 | 0.11345
|  | 488.6 | 0.10938
|  | 488.7 | 0.10532
|  | 488.8 | 0.10126
|  | 488.9 | 0.097194
|  | 489 | 0.093131
|  | 489.1 | 0.089067
|  | 489.2 | 0.085004
|  | 489.3 | 0.080941
|  | 489.4 | 0.076878
|  | 489.5 | 0.072815
|  | 489.6 | 0.068751
|  | 489.7 | 0.064688
|  | 489.8 | 0.060625
|  | 489.9 | 0.11412
|  | 490 | 0.11261
|  | 490.1 | 0.11109
|  | 490.2 | 0.10957
|  | 490.3 | 0.10805
|  | 490.4 | 0.10654
|  | 490.5 | 0.10502
|  | 490.6 | 0.1035
|  | 490.7 | 0.10199
|  | 490.8 | 0.10047
|  | 490.9 | 0.098953
|  | 491 | 0.097436
|  | 491.1 | 0.095919
|  | 491.2 | 0.094402
|  | 491.3 | 0.092885
|  | 491.4 | 0.091368
|  | 491.5 | 0.089851
|  | 491.6 | 0.088334
|  | 491.7 | 0.086817
|  | 491.8 | 0.0853
|  | 491.9 | 0.083783
|  | 492 | 0.082266
|  | 492.1 | 0.080749
|  | 492.2 | 0.079232
|  | 492.3 | 0.077715
|  | 492.4 | 0.076198
|  | 492.5 | 0.074681
|  | 492.6 | 0.073164
|  | 492.7 | 0.071647
|  | 492.8 | 0.07013
|  | 492.9 | 0.068613
|  | 493 | 0.067096
|  | 493.1 | 0.065579
|  | 493.2 | 0.064062
|  | 493.3 | 0.062545
|  | 493.4 | 0.061028
|  | 493.5 | 0.059511
|  | 493.6 | 0.057994
|  | 493.7 | 0.056477
|  | 493.8 | 0.05496
|  | 493.9 | 0.053443
|  | 494 | 0.051926
|  | 494.1 | 0.050409
|  | 494.2 | 0.048892
|  | 494.3 | 0.047375
|  | 494.4 | 0.045858
|  | 494.5 | 0.06701
|  | 494.6 | 0.066978
|  | 494.7 | 0.066945
|  | 494.8 | 0.066912
|  | 494.9 | 0.06688
|  | 495 | 0.066847
|  | 495.1 | 0.066814
|  | 495.2 | 0.066782
|  | 495.3 | 0.066749
|  | 495.4 | 0.066717
|  | 495.5 | 0.066684
|  | 495.6 | 0.066651
|  | 495.7 | 0.066619
|  | 495.8 | 0.066586
|  | 495.9 | 0.066554
|  | 496 | 0.066521
|  | 496.1 | 0.066488
|  | 496.2 | 0.066456
|  | 496.3 | 0.066423
|  | 496.4 | 0.066391
|  | 496.5 | 0.066358
|  | 496.6 | 0.066325
|  | 496.7 | 0.066293
|  | 496.8 | 0.06626
|  | 496.9 | 0.066228
|  | 497 | 0.066195
|  | 497.1 | 0.066162
|  | 497.2 | 0.06613
|  | 497.3 | 0.066097
|  | 497.4 | 0.066064
|  | 497.5 | 0.066032
|  | 497.6 | 0.065999
|  | 497.7 | 0.065967
|  | 497.8 | 0.065934
|  | 497.9 | 0.065901
|  | 498 | 0.065869
|  | 498.1 | 0.065836
|  | 498.2 | 0.065804
|  | 498.3 | 0.065771
|  | 498.4 | 0.065738
|  | 498.5 | 0.065706
|  | 498.6 | 0.065673
|  | 498.7 | 0.065641
|  | 498.8 | 0.065608
|  | 498.9 | 0.065575
|  | 499 | 0.065543
|  | 499.1 | 0.06551
|  | 499.2 | 0.065478
|  | 499.3 | 0.065445
|  | 499.4 | 0.065412
|  | 499.5 | 0.06538
|  | 499.6 | 0.065347
|  | 499.7 | 0.065314
|  | 499.8 | 0.065282
|  | 499.9 | 0.065249
|  | 500 | 0.065217
|  | 500.1 | 0.065184
|  | 500.2 | 0.065151
|  | 500.3 | 0.065119
|  | 500.4 | 0.065086
|  | 500.5 | 0.065054