10x v3.1 HT Genomics Sample Processing

- Scope and Applicability: Allows for rapid generation of 3’ transcriptomic-NGS-ready- single-cell-libraries from pools of cells. Doubling the capacity of the original 10xv3.1 protocol, allowing for a capture of ~20,0000 cells instead of ~10,000 cells per port.

- Materials:

  - Chromium Next GEM Single Cell 3’ HT Kit v3.1 48 rxns (-20°C) (10x Genomics 1000351)

    - RT Reagent B (2000435)

    - RT Enzyme C (2000436)

    - Template Switch Oligo (3000228)

    - Reducing Agent B (2000087)

    - Cleanup Buffer (2000438)

    - cDNA Primers (2000439)

    - Amp Mix (2000440)

  - Library Construction Kit (-20°C) (10x Genomics 1000352)

    - Fragmentation Enzyme (2000090)

    - Fragmentation Buffer (2000091)

    - Ligation Buffer (2000092)

    - DNA Ligase (220110)

    - Adaptor Oligos (2000094)

    - Amplification Mix (2000047)

  - DynaBeads MyOne Silane (10x Genomics 2000048) (Stored at 4°C)

  - Chromium Next GEM Single Cell 3’ HT Gel Bead Kit v3.1 (10x Genomics 2000443) (-80°C)

  - Chromium Next GEM Chip M Single Cell Kit (10x Genomics 1000349) (RT)

    - Next GEM Chip M Single Cell (2000417)

    - Chip Gasket HT (3000614)

    - Partitioning Oil (220088)

    - Recovery Agent (2000434)

  - Dual Index Kit TT Set A (10x Genomics 1000215) (Stored at -20°C)

  - TempAssure PCR 8-tube strip (USA Scientific 1402-4700)

  - 0.2 mL TempAssure PCR Tube with frosted caps (USA Scientific 1402-8100)

  - 1.5 mL DNA LoBind Tubes (Eppendorf 022431021)

  - 2.0 mL DNA LoBind Tubes (Eppendorf 022431048)

  - SPRIselect Reagent Kit 60 mL (Beckman Coulter B23318)

  - 50% Glycerol (v/v) aqueous solution (Ricca Chemical Company 3290-32)

  - Nuclease Free Water (Thermo AM9937)

  - Low TE Buffer (Teknova T0221)

  - Tween 20 (Millipore 655204-100mL)

  - 100% Ethanol (AAPER/Pharmco E200G-P)

  - Twin.Tek 96-well PCR Plate, Semi-Skirted (Eppendorf 0030129326)

  - Twin.Tek 96-well PCR Plate, Full-Skirted (Eppendorf 0030129300)

  - BioRad Microseal ‘B’ (BioRad MSB1001)

  - High Sensitivity NGS Kit (Agilent DNF-474-33)

  - Quant-it PicoGreen Reagent Kit (Thermo P7589)

  - 50 mL conical polypropylene Falcon tubes (Corning 352098)

  - 15 mL conical polypropylene Falcon tube (Corning 430790)

  - 10 mL serological pipette tip (Greiner 607180)

  - 25 mL serological pipette tip (Greiner 357525)

  - Two 5 mL divided reservoirs (VWR 41428-958)

  - Qiagen Buffer EB (Qiagen19086)

- Equipment:

  - PCR Hood with UV for chip loading and processing

  - 10x Genomics Chromium Controller

  - BioRad C1000 Thermocycler with a deep-well block

  - Rainin P1000, P20, P200, P10 single channel pipettes

  - Rainin P200, P20, and P10 8-channel pipettes

  - Rainin Pasteur P1000 and P200 low retention filtered tips

  - Rainin 200 and 20 µL low retention filtered tips

  - Rainin 1000, 200 and 20 µL tips filtered tips

  - Eppendorf ThermoMixer C with 1.5 mL SmartBlock and heated lid (Eppendorf 2231000574)

  - Water bath

  - Vortex Mixer

  - Ice Bucket

  - Integra aluminum PCR block

  - Aluminum 2 mL microtube rack

  - MicroCentrifuge for both 1.5-2 mL tubes and 8-strips

  - Plate Sealer

  - Plastic spatula for sealing plates.

  - 10x Vortex Adapter (10x Genomics 330002)

  - Chromium X Chip Holder (10x Genomics 3000598)

  - 10x Magnetic Separator HT (10x Genomics 2000431)

  - Fragment Analyzer Instrument

  - Post-amplification PCR Hood

  - Post-amplification Vortex Mixer

  - Post-amplification Microcentrifuge for both 1.5-2 mL tubes and 8-strips

  - iPad for photo-documenting GEMs

  - Serological pipette

  - Lab timer with clock

- Safety:

  - Nitrile Gloves

  - Eye protection

  - Lab coat

  - Disposable Lab Sleeves

Warning: Personal Protective Equipment (PPE) should be used at all times while operating this protocol. If you are unsure what PPE you should be using, see your immediate supervisor.

- Output:

  - Massively multiplexed single-cell NGS libraries.

- Reference Documents:

  - Chromium Next GEM Single Cell 3ʹ HT Reagent Kits v3.1 User Guide

  - MB0073 Fragment Analyzer Operation for PCR Products

    - dx.doi.org/10.17504/protocols.io.5jyl8mxwrg2w/v2

  - MB0194 384-Well PicoGreen

    - dx.doi.org/10.17504/protocols.io.36wgq5ekxgk5/v2

- Setup (At Least One Hour Before Starting):

  - Aliquot out 1 mL of 50% Glycerol into DNA LoBind tubes

    - Store the aliquots at -20°C until needed.

  - Prepare 50X 1 mL aliquots of 10% Tween 20 by adding 5 mL of Tween 20 to 45 mL of NFdH2O in a 50 mL falcon tube.

    - Invert 20X to mix and let stand at room temperature for 5 minutes.

    - Aliquot out 1 mL of the 10% Tween 20 into 1.5 mL or 2 mL DNA LoBind Tubes.

    - Store the aliquots at room temperature until needed.

    - Dispose of the aliquots after 1 year.

  - Thaw the 5 mL vials of Cleanup Buffer (2000438) in a waterbath at 65°C for approximately 30 min.

    - Vortex at max speed to make sure no precipitate remains.

    - Repeat every 2 min until solution is clear.

    - Aliquot the 5 mL Cleanup Buffer (2000438) vials into 2 mL tubes for future use.

    - Keep one at room temperature and store the rest at -20°C until needed.

  - Fill the ice bucket with ice and place both the Integra aluminum PCR block and the aluminum 2.0 mL microtube rack onto the ice to cool.

  - If processing fewer than 16 reactions, set aside 4 aliquots of 50% Glycerol and equilibrate to room temperature. This will be used to fill unused wells.

  - Remove Reducing Agent B (2000087) from the kit and allow it to thaw at room temperature for at least 1 hour.

  - Remove the Chromium Next GEM Single Cell 3’ HT Gel Bead Kit v3.1 (2000443) from the -80°C and allow them to equilibrate at room temperature for at least 30 minutes.

  - Resuspend the lyophilized Template Switch Oligo (3000228) in 80 µL of Low TE Buffer (Teknova T0221):

    - Centrifuge the Template Switch Oligo at 18G for 10 min, then use a P200 @ 80 µL to add the low TE, making sure the lyophilized pellet is visible before addition of low TE.

    - Vortex the Template Switch Oligo on max for 30s, centrifuge briefly, and then leave at room temperature for at least 1 hour.

    - Measure the concentration (ng/µL) of the resuspended Template Switch Oligo to ensure adequate resuspension.

      - We found the median concentration of 42 separate Template Switch Oligo tubes quantified in triplicate was 14,757 ng/µL with a coefficient of variation of 4.6%.

| N of TSO tubes quantified in triplicate | Median Concentration | %CV | Minimum Concentration | Maximum Concentration |
| --- | --- | --- | --- | --- |
| 42 | 14,757 ng/µL | 4.6% | 13,308 ng/µL | 16,380 ng/µL |

  - Allow the RT Reagent B (2000435) to thaw and equilibrate to room temperature.

  - Once thawed, vortex the RT Reagent B (2000435), Reducing Agent B (2000087), and re-suspended Template Switch Oligo (3000228) for 15s on max.

    - Verify there are no precipitates before returning to the RT Reagent B (2000435), the Reducing Agent B (2000087) and Template Switch Oligo (3000228) to room temperature.

      - Repeat every 5 minutes as necessary until no precipitates are present.

  - Remove the RT Enzyme C (2000436) from the kit and place it in the aluminum block in the ice bucket.

- Methodology/Procedures:

  - Chip Loading and GEM Generation:

    - Conduct the following steps in the designated PCR clean UV hood to prevent contamination and potential wetting failures.

    - Assemble the Chromium Next GEM chip

      - Close the holder lid and carefully attach the blue HT gasket onto the hooks starting from the left to the right. Make sure to do this before inserting the chip so that the gasket will align properly.

      - Open the chip holder with the gasket still attached. Careful not to touch the smooth side of the gasket.

      - Using care to only handle the edges and avoiding contact with the ports or bottom surface, place a Next GEM Chip M Single Cell (2000417) into the Chromium X chip holder (3000598).

      - Cover both the gasket and the chip each with a pipette tip box lid at all times. Only remove the tip box for the chip when you are actively adding solutions to the chip.

    - Prior to chip loading, if fewer than 16 reactions are planned, add the following volumes of 50% Glycerol to each unused well, in the following order:

      - 130 µL in labeled row 1 unused wells using a P200 @ 130 µL.

      - 140 µL in labeled rows 2A & 2B unused wells using a P200 @ 140 µL.

      - 140 µL in labeled rows 3A & 3B unused wells using a P200 @ 140 µL.

      - DO NOT load glycerol into the corresponding well for odd number of samples (a mock sample mix will need to be loaded instead).

    - Prepare enough RT master mix for each reaction with 10% excess volume by consulting the following table, vortexing each reagent for 5s at max and adding reagents in the order listed to a 1.5 mL Eppendorf LoBind Tube (Eppendorf 022431021):

      - Add 330 µL of vortexed RT Reagent B (2000435) to the 1.5 mL Eppendorf LoBind master mix tube using a P1000 @ 330 µL.

      - Add 41.4 µL of Template Switch Oligo (3000228) to the master mix tube using a P200 @ 41.4 µL.

      - Add 34.3 µL of Reducing Agent B (2000087) to the master mix tube using a P200 @ 34.3 µL.

      - Add 154 µL of RT Enzyme C (2000436) to the master mix tube using a P200 @ 154 µL.

      - If there’s an odd number of samples, label the last well on the 2B TempAssure PCR 8-strip for the mock sample. Still make an even number of rxn (sample + mock-sample) of mastermix.

    - Vortex the master mix tube for 10 seconds at max speed.

    - Briefly spin down the master mix tube to collect.

    - Aliquot 63.6 µL of RT master mix into each well including tube marked for mock sample of a TempAssure PCR 8-strip using a P200 @ 63.6 µL.

      - If there is an odd number of samples, add 86.4 µL of nuclease free water into mock sample tube using a P200 @ 86.4 µL.

    - Briefly spin down this strip and place it in the Integra aluminum block on ice.

    - After the cells or nuclei are sorted, cleaned, and counted, use the targeted cell recovery table to calculate the volume of cells and water needed for addition to the TempAssure 8-strip containing the RT master mix.

    - Under witness, add the water first to the mastermix, if needed, using a P200 and the volume listed in blue text above. If adding water, add appropriate water volumes to all samples in the strip tube. Vortex the strip tube containing mastermix and water then spin tubes.

    - To add the cells, triturate the cell solution 10 times, with the tip opening in the middle of the cell solution, using a P200 with a low-retention (regular for nuclei and wide bore for cells) filtered tip @ the volume listed in red in the targeted cell recovery table.

      - Under witness, immediately add the homogenous cell solution to the correct TempAssure 8-strip containing RT master mix using the same P200 pipette and tips that were used to homogenize the sample.

    - Snap the Next GEM Singe Cell 3’ HT Gel Beads (2000443) into the 10x vortex adapter (330002) and vortex on max speed for 30 seconds.

      - Remove the beads from the vortex adapter, snap off the tube strip holder, remove bead strip, and spin down for a full 5 seconds. Place the bead strip back, and snap into place with the tube strip holder.

      - Pierce the foil on the bead strips using fresh P200 tips.

        - Discard the tips after piercing.

      - Using an 8-channel P200@ 130 µl and low-retention tips, add 130 µL of Next GEM Single Cell 3’ HT Gel Beads (2000443) into the wells in labeled row 1 of the chip.

        - Pipette slowly with the tips in the bottom of the central well depression and taking 5 seconds to depress the pipette plunger while raising the tips at the same rate the solution is filling the wells.

        - Wait 30 sec before loading the samples.

    - Gently pipette mix (10X) the first 8-strip tubes containing the cells and RT master mix using an 8-channel P200 equipped with low-retention regular tips @ 140 µL.

      - Without discarding the tips, transfer 140 µL of the cells + master mix to the wells in labeled row 2A of the chip using the same 8-channel P200 @ 140 µL.

      - Pipette slowly with the tips in the bottom of the central well depression and taking 5 seconds to depress the pipette plunger while raising the tips at the same rate the solution is filling the wells.

      - Take care to not introduce air bubbles.

    - Follow step 8.1.13 again with the second 8-strip tube, this time dispensing into the wells in labeled row 2B of the chip.

      - Always ensure the number of sample use for 2A is equal to 2B.

      - If there is an odd number of samples, add the mock sample to the well next to the last sample well in row labeled 2B of the chip.

    - Allow the chip to sit and gravity prime for 30 seconds before adding partitioning oil.

    - Add 140 µL of partitioning oil (220088) to both labeled row 3A and 3B of the chip using a P200 @ 140 µL and a regular low-retention tip.

      - Pipette the oil slowly just above the bottom of the well, avoiding bubbles.

    - Close the lid with the pre-attached gasket, careful not to touch the smooth side of the gasket.

      - Do not tip the chip holder, keep the assembly horizontal to avoid any wetting of the 10x Gasket with partitioning oil.

      - Do not press downwards on the gasket. This will lead directly to “wetting errors” that produce failed libraries.

    - Place the assembly into the Chromium X Controller.

      - Press the large “Eject” button in the middle of the screen to load the chip.

      - Press the large “Play” button on the screen to begin running the Chromium Controller

    - Runs will take around 18 minutes to finish.

    - Return reagents to their proper storage locations.

      - Note the freeze/thaw cycles of the beads as this may affect performance.

    - Proceed directly to GEM Transfer and RT once the chip has finished running.

  - GEM Transfer and RT:

    - Label 2 semi-skirted Twin.Tek PCR plates (Eppendorf 0030129326) with the capture names according to the paperwork. Distinguish the two transfers with “-1” and “-2” after load name.

      - Place the 2 plates onto the Integra aluminum blocks in the ice bucket to chill.

        - With lab marker, circle the wells associated with the load for both plates. Plate rows A-H = left-right port position on the 10x chip.

    - Start the GEM-RT protocol on the C1000 thermocycler to get the block and lid to temperature.

    - Upon completion of GEM formation on the Chromium Controller, press the large “eject” button on the screen.

    - Remove the chip assembly from the Chromium Controller.

    - Carefully remove and discard the 10x Gasket from the assembly.

    - Open the 10x Chip holder (3000614) and fold the lid backwards until it clicks to create a 45 degree well angle.

    - Slowly aspirate the first 90 µL of GEMs from the lowest point of the recovery wells in the chip (labeled 3A) using a P200 @ 90 µL and low-retention regular tips.

      - GEMs have high viscosity and need to be pipetted slowly.

      - GEMs should appear even; excess partitioning oil suggests a clog occurred during the run.

      - Photo document the GEMs, saving it in the load document. If there is obviously less than 90 µL of GEMs, take a picture of the chip. Note in a red circle on picture if there is leftover sample in the port and/or beads in the port.

      - Under witness, dispense the GEMs into the first labeled (-1) Twin.Tek PCR plate in the first labeled column slowly over the course of 20 seconds.

      - Pipette against the sidewalls of the wells, keeping the tips above liquid level to minimize GEM loss.

      - Repeat all the steps in 8.2.7. for the second aliquot of 90 µL of GEMs from 3A into the second labeled (-2) Twin.Tek PCR plate in the same numbered column as the first plate.

    - Repeat all the steps for 8.2.7. for the recovery wells in the chip labeled 3B.

    - Seal the Twin.Tek PCR plates using BioRad microseal ‘B’ (BioRad MSB1001) and a plastic spatula.

    - If the GEM solution has air gaps, spin down the plate for 10s at 1000 x g.

    - Load the plate onto the thermocycler running the GEM-RT program.

| Lid Temperature | Reaction Volume | Run Time |
| --- | --- | --- |
| 53°C | 125 µl | 55 min |
| Step | Temperature °C | Time (Min) |
| 1 | 53 | Hold |
| 2 | 53 | 45:00:00 |
| 3 | 85 | 5:00 |
| 4 | 4 | Hold |

    - Advance the thermocycler program to begin RT by pressing ‘skip step’ on the program.

      - Run time is around 55 minutes.

    - Re-seal the plate with a PlateLoc aluminum seal when done.

    - Upon completion, plate may be stored at 4°C for up to 72 hours or at -20°C for up to one week.

    - When ready, proceed directly to Post GEM-RT Cleanup and cDNA Amplification.

  - Post GEM-RT Cleanup and cDNA Amplification:

    - Remove the 2 ml aliquoted tube of Cleanup Buffer (2000438) from the kit and place in the Eppendorf ThermoMixer C at 65°C for 10 minutes at max speed (3,000 x RPM).

      - Check for precipitates and crystals. Vortex and return to the ThermoMixer for an additional 2 minutes if crystals are present.

      - Repeat as needed.

      - Allow tube to cool to room temperature before using

    - Conduct the following steps in a Post-AMP PCR clean hood.

    - Fill the ice bucket with ice and place both the Integra aluminum PCR block and 1.5-2.0 mL aluminum microtube block on ice to cool. Leave ice bucket outside, next to the hood

    - From the Chromium Next GEM Single Cell 3’ HT GEM Kit v3.1 (10X Genomics 1000351) stored at -20°C, remove the following items and allow them to thaw.

      - cDNA Primers (2000439)

      - Reducing Agent B (2000435)

    - Once thawed, vortex and briefly centrifuge.

    - Check Reducing Agent B (2000435) for precipitates.

    - Keep the cDNA Primers (2000439) at room temperature until needed.

    - Remove the Dynabeads MyOne Silane beads (2000048) from the 4°C and allow them to equilibrate to room temperature.

    - From the 10x Chromium Next GEM Single Cell 3’ GEM Kit v3.1 (10x Genomics 1000139) stored at -20°C, remove the AMP Mix (2000440) and place it in the aluminum microtube block on ice.

    - Prepare 15 mL of 80% Ethanol in a 15 mL conical Falcon tube (Corning 430790) by combining 12 mL of 100% Ethanol with 3 mL of water using two 10 mL serological pipettes, switching tips between reagents.

    - Retrieve the RT-GEM plates from either -20°C or 4°C storage and let thaw to room temperature.

    - Add 125 µL of Recovery Agent (2000434) to each well using a single-channel P200 @ 125 µL.

      - Use care not to touch the surface of the fluid to avoid complexity loss.

      - Do not pipette mix or vortex this biphasic mixture.

    - After 2 min, transfer the biphasic mixture to a TempAssure PCR 8-strip (USA Scientific 1402-4700) using 2X P200 @ 115 µL and low-retention Pasteur tips. This step is done under witness.

    - Carefully pipette out the organic pink phase from the bottom of the tubes using an 8-channel P200 @ 115 µL and low-retention tips.

      - Remove any significant volumes of organic pink phase using a single channel P20 with low-retention tips.

      - Do not remove all the organic phase since the majority of the cDNA exists at the interphase between layers. Leave between 5-10 µL of pink fluid in the bottom of every well of the sample strip.

        - The strip should resemble the below image when finished.

        - Photo document the tubes after recovery reagent removal and save them in the GEM-RT Cleanup and cDNA Amplification paperwork

    - Vortex the DynaBeads MyOne Silane Beads (2000048) until fully re-suspended. To bring the bead volume down to the bottom of the tube, manually tap the tube on the bench top until the beads go down to the bottom. Do not spin down the tubes.

  - Prepare DynaBeads cleanup mix by adding the following reagents, in order, to a 5.0 mL DNA LoBind Tube (Eppendorf 0030122348).

      - Add 88 µL of NfdH2O to the DynaBeads Cleanup Mix 5 mL Eppendorf LoBind tube by using a P200 @ 88 µL.

      - Vortex and spin down the room-temperature Cleanup Buffer (2000438).

      - Add 3203.2 µL of Cleanup Buffer (2000438) to the 5 mL Eppendorf tube using a P1000 @ 1000 µL (X3) and a P1000 @ 203.2 µL.

      - Add 140.8 µL of homogenized and fully re-suspended DynaBeads MyOne Silane beads (2000048) to the DynaBeads Cleanup Mix 5 mL Eppendorf tube using a P200 @ 140.8 µL.

        - Vortex the DynaBeads MyOne Silane beads (2000048) thoroughly for >30s immediately before addition to the mix. Do not spin down beads.

      - Vortex and add 88 µL of Reducing Agent B (2000087) to the DynaBeads Cleanup Mix 5 mL tube using a P200 @ 88 µL.

    - Vortex the DynaBeads Cleanup Mix 5 mL tube at max speed for 15 seconds.

      - Do not spin down mix.

    - Add 200 µL of DynaBeads Cleanup Mix to the sample strip using a P200 @ 200 µL using low retention tips.

      - Do not touch the sample with the pipette tips or pipette mix the solution.

    - Close the caps and briefly vortex the strip for 5 seconds at max speed.

    - Allow the samples to sit at room temperature for at least 10 minutes.

      - Vortex the strip again for 5 seconds at max speed after the first 5 minutes of the incubation.

    - During the 10-minute incubation, prepare Elution Solution 1 by adding the reagents in the order shown below to a 5.0 mL DNA LoBind Tube (Eppendorf 0030122348):

      - Add 3920 µL of Buffer EB (Qiagen 19086) using a P1000 @ 1000 µL (X3) and P1000 @ 920 µL to the Elution Solution 1 tube.

      - Add 40 µL of 10% Tween 20, from the 1 mL aliquot prepared earlier, to the Elution Solution 1 tube by using a P200 @ 40 µL.

      - Vortex and add 40 µL of Reducing Agent B (2000087) to the Elution Solution 1 tube using a P200 @ 40 µL.

    - Aliquot 40 µL of the prepared Elution Solution 1 into each well of a new TempAssure PCR 8-strip using a P200 @ 40 µL.

      - Label that strip “ES1” and set it aside for now.

    - Label new TempAssure PCR 8-strip tubes for each sample, separating strips by PCR cycles needed, and set aside for now.

    - Once the 10-minute incubation is complete, briefly spin down the tube strip and place it in the 10x Magnetic Separator (120250) in the “High” position.

    - Wait 2 minutes or until the supernatant is clear.

    - Free pour roughly 5 mL of 80% Ethanol into a 25 mL divided reservoir (VWR 41428-958) and set this aside.

    - Once the sample strip is clear, carefully remove the supernatant (waste) using a P200 @ 150 µL (twice) equipped with low retention tips.

      - Discard tips and waste into an appropriate waste container.

    - Add 150 µL of 80% Ethanol to each well of the sample strip using an 8-channel P200 @ 150 µL and low retention tips.

    - Immediately add another 150 µL of 80%, using an 8-channel P200 @ 150 µL and new low retention tips, for a total of 300 µL.

    - Let this solution stand for 30 seconds.

    - Carefully remove and discard the Ethanol wash using an 8-channel P200 @ 150 µL with low retention tips.

      - Repeat with new tips to remove all the residual Ethanol.

    - Carefully add 200 µL of 80% Ethanol to the 8-strip using an 8-channel P200 @ 200 µL with low retention tips.

      - Let solution stand for 30 seconds.

    - Remove the Ethanol wash using an 8-channel P200 @ 200 µL with low retention tips.

    - Start a timer for 1 minute.

    - During the 1 minute wait, Remove the strip from the 10x Magnetic Separator (120250) and briefly spin it down.

    - Place the strip back into the 10x Magnetic Separator (120250) on the “Low” position.

    - Remove any remaining Ethanol from the strips using an 8-channel P20 @ 10 µL with low retention tips.

    - At the end of the 1 minute drying period, remove the strip form the 10x Magnetic Separator (120250) and immediately add 36 µL of Elution Solution 1 from the “ES1” strip using an 8-channel P200 @ 36 µL with low retention tips.

    - Close the strip caps and pulse vortex them at max speed for 15 seconds.

    - Briefly spin the strip to collect the sample and incubate it off magnet at room temperature for 5 minutes.

    - Return the strip to the 10X Magnetic Separator (120250) in the “Low” position.

      - Leave the strip in the separator until it clears (around 1 minute).

    - Transfer 35 µL of purified GEM-RT products to pre-labeled TempAssure PCR 8-strip tubes using a P200 @ 35 µL with low retention tips.

    - Place PCR tubes in the Integra aluminum block in the ice bucket and proceed with cDNA amplification.

    - Prepare the cDNA Amplification Reaction Mix by adding the following reagents in order, to a 1.5 mL Eppendorf low-bind tube (Eppendorf 022431021).

      - Vortex and add 880 µL of AMP Mix (2000440) to the cDNA Amplification Master Mix tube using a P1000 @ 880 µL.

      - Vortex and add 264 µL of cDNA Primers (2000439) to the cDNA Amplification Master Mix tube using a P1000 @ 264 µL.

    - Vortex the cDNA Amplification Reaction Mix for 5s on max.

    - Spin down the master mix tube briefly to collect.

    - Add 65 µL of cDNA Amplification master mix to each 35 µL sample using a P200 @ 65 µL.

      - Take care to not touch the pipette tip to the sample when pipetting. Do not pipette mix the sample.

    - Close the caps on the sample tubes and briefly vortex for 5 seconds on max speed.

    - Spin down the PCR tubes to collect the volume and return the tubes to the aluminum block in the ice bucket.

    - To determine PCR cycling conditions, consult the following table for some basic guidelines surrounding cell number captured during the load.

cDNA AMPLIFICATION 10xV3.1_HT; Mouse Cells; Mouse Nuclei; Human Nuclei (<40% NeuN-negative). Targeted Total Cell or Nuclei Capture; mouse baseline; +1; human baseline. <2000; 14; 15; 16. 2000-6000; 12; 13; 14. 6000-7000; 11; 12; 13. 7000-11000; 10; 11; 12. >11000; 8; 9; 10.

      - There are many layers of factors that go into the cycling decisions, the most important thing is that these decisions be made using empirical data.  A few examples that influence cycling conditions are sample type, species, isolation method, NeuN negative percentage, and region of interest dissected.

    - Load the tubes containing cDNA and master mix onto a thermocycler and run “cDNA PCR NN v3” where NN is equal to the number of cycles for each sample.

      - Upon completion of PCR, samples may either be stored at 4°C for up to 72 hours, -20°C for up to 1 week, or proceed directly with SPRIselect cleanup.

    - To begin cleanup, vortex the SPRIselect reagent on max speed until fully re-suspended.

    - Free pour 5 mL of 80% Ethanol into the 25 mL divided reservoir.

    - Label two TempAssure PCR 8-strips (USA Scientific 1402-4700), one for sample QC, and one for EB storage. Label one 96-well Eppendorf Twin.Tek Full-Skirted PCR plate (Eppendorf 0030129300) for the post-cleanup samples.

    - Fill the 8-strip labeled for EB storage with 50 µL of Qiagen Buffer EB using a P200 @ 50 µL.

    - Place freshly vortexed beads in a 25 mL reservoir. Using a multichannel P200, Transfer 60 µL of the homogenized SPRI to each sample in the sample strip.

    - Pulse vortex the 8-strip with PCR products and SPRI beads 15 times at max speed.

    - Incubate the strip for 5-minutes at room temperature.

    - Once the 5-minute incubation is over, place the 8-strip into the 10X Magnetic Separator (120250) using the “High” position.

      - Leave the strips on the magnet with their lids closed until clear.

        - This step generally takes between 2-5 minutes. Check every 2 minutes until clear.

    - Carefully remove and discard the supernatant using an 8-channel P200 @ 180 µL taking care not to disturb the bead pellet.

    - Carefully wash the beads by slowly adding 200 µL of 80% Ethanol from the reservoir using a P200 @ 200 µL.

      - Let the wash solution stand on the beads for 30 seconds.

    - Carefully remove and discard the wash solution using an 8-channel P200 @ 200 µL.

    - Repeat the wash step by adding a fresh aliquot of 80% Ethanol from the reagent reservoir using an 8-channel P200 @ 200 µL.

      - Let the wash solution stand for 30 seconds.

    - Remove and discard the Ethanol wash using an 8-channel P200 @ 200 µL.

    - Remove the 8-strip from the magnet, cap it, and very briefly (<10s) spin the tube down.

    - Immediately return the 8-strip to the magnet, this time in the “Low” position.

    - Start a lab timer for 2 minutes, uncap the tubes, and remove any residual Ethanol using an 8-channel P20 @ 10 µL.

      - It is very important not to over dry the samples, so small amounts of residual Ethanol are fine.

    - Upon completing the end of the 2-minute drying, remove the strip from the magnet and add 41 µL of Qiagen Buffer EB (Qiagen 19086) from the EB 8-strip using an 8-channel P200 @ 41 µL.

    - Close the caps and pulse vortex the strip 10x at max speed until beads are fully resuspended.

    - Let the strip incubate off magnet for 5 minutes at room temperature.

    - Briefly spin down the strip to collect the volume.

    - Upon completion of the 5-minute room temperature incubation, place the 8-strip back into the 10x Magnetic Separator in the “High” position.

      - Leave the sample on the magnet until it clears (around 2 minutes).

    - Transfer 40 µL of the cleaned post cDNA amplified products to the pre-labeled 96-well Eppendorf Twin.Tek Full-Skirted PCR plate using a P200 @ 40 µL.

      - Be sure to distinguish the samples with each plate (-1 or -2)

    - Under witness, transfer 15 µL of cDNA from each of the two plates to a third plate labeled with Amp name and “-P”.

      - There should be 30 µL in each sample well when done. This plate will be used to generate libraries after QC.

      - Vortex and spin down newly pooled Amp “-P” plate before QC and storage.

    - Add 9 µL of Qiagen Buffer EB (Qiagen 19086) from the EB 8-strip to the QC labeled 8-strip using an 8-channel P10 @ 9 µL.

    - Transfer 1 µL of cleaned post cDNA amplified product to the QC strip using an 8-channel P10 @ 1 µL.

      - This will need to be done for all three plates created above.

      - Close and vortex the QC strip for 5 seconds at max speed then spin down to collect.

      - Store the QC strip and 4°C until ready for Post AMP 10x Quant and QC.

    - Store all three plates of cleaned cDNA samples at 4°C for up to 72 hours, or -20°C for up to one week before library construction.

  - Post AMP 10x Quant and QC

    - Follow MB0194: 384-well PicoGreen to quantify the sample, using 3 µl from the 10x diluted QC strip.

    - Multiply the resulting concentration and yield by 10 to achieve the actual sample concentration and yields.

    - Run 1 µL of the 10X diluted 10x QC sample on the FA using the FA SOP.

      - Check the trace of the amplified cDNA. High quality cDNA should closely resemble the following:

    - In addition, check the trace for:

      - Any extraneous peaks.

      - Significant products smaller than 400 bp.

      - Significant differences in yield when compared to PicoGreen results.

  - Quantify the nanogram input for library based on PicoGreen results. In general, target a cDNA input of 111-500 ng.

    - Nanogram input will vary based on a variety of factors such as sample type, species, isolation method, NeuN negative percentage, region of interest dissected, and sample quality.

    - In general, samples with an input of less than 20 ng for library are considered failed. However, there are instances of samples of known lower quality, or that cannot be re-done that will still pass to library. This is not recommended if the sample can be rescued.

    - Below are the yield ranges we have seen for 265 v3.1 HT cDNA amplifications that have passed on to library construction.

| Passing v3.1 HT cDNA Amplifications |  |  |
| --- | --- | --- |
|  | ng input for lib | PG (ng/µL) |
| Sample Number | 265 | 265 |
| Min | 14.47 | 0.72 |
| Max | 657.90 | 32.90 |
| Average | 128.27 | 6.46 |
| Median | 93.76 | 4.69 |

| ng Input for Libray Range | Passing Sample Number | % of Samples |
| --- | --- | --- |
| <10 ng | 0 | 0% |
| 10-30 ng | 17 | 6% |
| 30-60 ng | 63 | 24% |
| 60-100 ng | 57 | 22% |
| 100-200 ng | 76 | 29% |
| 200-300 ng | 30 | 11% |
| 300-500 ng | 20 | 8% |
| >500 ng | 2 | 1% |
| Total: | 265 |  |

    - Samples which do not pass QC should not proceed to library construction. All other samples should proceed with 10x Library Construction.

  - 10x Library Construction

    - Conduct the following steps in a Post-AMP PCR hood.

    - Fill an ice bucket with ice and place both the Integra aluminum block and 1.5-2 mL microtube block on the ice to chill.

    - Remove the following items from the Chromium Next GEM Single Cell 3’ HT Library Kit v3.1 (10x Genomics 1000352) and allow them to equilibrate to room temperature:

      - Ligation Buffer (2000092)

      - Fragmentation Buffer (2000091)

      - Adaptor Oligos (2000094)

    - Remove the following items from the Chromium Next GEM Single Cell 3’ Library Kit v3.1 (10x Genomics 1000158) and place them directly in the microtube block on ice:

      - Fragmentation Enzyme (2000090)

      - DNA Ligase (220110)

      - AMP Mix (2000047)

    - Remove the Dual Index Plate TT Set A (10x Genomics 1000215) from the -20°C and place it at room temperature to thaw.

    - Fill a 50 mL Falcon tube (Corning 352098) with 16 mL of 100% Ethanol using a 25 mL serological pipette (Greiner 357525).

    - Add 4 mL of NFdH2O (Thermo AM9937) to the 16 mL of Ethanol using a 10 mL serological pipette (Greiner 607180) to make 20 mL of 80% Ethanol.

    - Prepare a thermocycler by initiating the Frag, A-Tail protocol.

    - Remove the cleaned and pooled post cDNA amplified product plate labeled with “-P” from either the 4°C or -20°C and allow it to thaw.

      - Once thawed, vortex, and spin down sample plate.

    - Vortex the Fragmentation Buffer (2000091) and verify there is no precipitate.

      - Vortex again if needed.

    - Prepare Fragmentation master mix in a 1.5 mL Eppendorf LoBind tube (Eppendorf 022431021) by vortexing and adding the following regents in order:

      - Vortex and add 44 µL of Fragmentation Buffer (2000091) to the fragmentation mix tube using a P200 @ 44 µL.

      - Vortex and add 88 µL of fragmentation Enzyme (2000090) to the fragmentation mix tube using a P200 @ 88 µL.

    - Vortex the fragmentation master mix for 10 seconds at max speed. Spin down to collect.

    - At this point, label the wells with assigned index primers to each sample that were checked out during paperwork creation.

    - Together with a witness, label a new TempAssure PCR 8-tube strip with the sample names determined by the index well that will be used and place it on ice.

    - Add 15 µL of fragmentation master mix to each well in the new, labeled TempAssure PCR strip using a P20 @ 15 µL.

    - Add 15 µL of Qiagen Buffer EB to each well using a P20 @ 15 µL, to bring the volume up to 30 µL.

    - Transfer 20 µL of purified cDNA to each well containing the fragmentation master mix using an 8-channel P20 @ 20 µL.  Do this under witness to ensure sample integrity.

      - The remaining cDNA can be stored long term at -20°C for generation of additional 3’ gene expression libraries.

    - Close the caps and pulse vortex 5 times at max speed.

    - Spin down the strip to collect the volume.

    - Transfer the tube strip to the pre-cooled thermocycler running the Frag, A-Tail protocol and press “Skip” to initiate the protocol.

    - Upon completion of the thermocycler protocol, proceed directly with double sided size selection using SPRIselect.

    - Vortex the SPRIselect reagent until fully re-suspended.

    - Fill a TempAssure 8-strip with 60 µL of Qiagen buffer EB (Qiagen 19086) using a P200 @ 60 µL and label this strip “EB”.

    - Add 30 µL (0.6X) of homogenized SPRIselect reagent to each sample using a P200 @ 30 µL.

    - Cap and pulse vortex the strip 5 times at max speed.

    - Allow the strip to incubate at room temperature for 5 minutes.

    - Spin down the strip to collect the volume and place it in the 10x magnetic separator in the “High” position. Let strips sit in magnet for minimum of 2 minutes.

    - Label a new TempAssure 8-strip.

    - Add 10 µL of homogenized SPRIselect reagent to each well of the newly labeled TempAssure strip using a P10 @ 10 µL.

    - Once the sample strip in the magnetic separator clears, transfer all 75 µL of the supernatant, using an 8-channel P200 @ 75 µL, to the new strip containing 10 µL of beads.

      - Do not discard the supernatant as it contains the desired fraction of the sample.

      - Discard the strip sitting on the magnetic separator (now with supernatant removed).

    - Pulse vortex the TempAssure strip containing the sample and beads 5 times at max speed.

    - Allow the strip to incubate at room temperature for 5 minutes.

    - Spin down the strip to collect and place it in the 10x magnetic separator in the “High” position. Let strips sit in magnet for minimum of 2 minutes.

    - Fill the long half of a divided reservoir (VWR  41428-958) with 5 mL of 80% Ethanol by free pouring.

    - Once the solution is clear, remove and discard 80 µL of supernatant using an 8-channel P200 @ 80 µL. There should be several µl of supernatant left in the tube.  If there are beads left in the supernatant, that is okay.

      - As the bead volume is small, it is critical to not disturb or remove any of the beads. Some residual volume is preferable to prevent this from occurring.

    - Add 125 µL of 80% Ethanol from the divided reservoir using an 8-channel P200 @ 125 µL.

    - Allow the Ethanol wash to stand for at least 30 seconds.

    - Carefully remove and discard the Ethanol wash with an 8-channel P200 @ 125 µL.

    - Repeat the wash by adding 125 µL of 80% Ethanol from the divided reservoir using an 8-channel P200 @ 125 µL.

    - Incubate the Ethanol wash for at least 30 seconds.

    - Carefully remove and discard the Ethanol wash with an 8-channel P200 @ 125 µL.

    - Cap the strip and briefly spin the tubes and return to the 10x magnetic separator in the “Low” position.

    - Remove the residual Ethanol with an 8-channel P20 @ 10 µL, being careful to not remove any of the beads.

    - Immediately remove the strip from the magnetic separator and add 51 µL of Qiagen buffer EB to the 8-strip using an 8-channel P200 @ 51 µL.

      - Vortex mix the beads until re-suspended.

      - Due to the low bead volume, it is critical to not over dry the beads, so re-suspend the beads immediately after the Ethanol is removed.

    - Incubate the strip and beads for 5 minutes at room temperature.

    - Spin down the strip briefly to collect the volume and place the strip back into the 10x magnetic separator in the “High” position. Let strips sit in magnet for minimum of 2 minutes.

    - Label a new TempAssure 8-strip with the sample IDs.

    - Transfer 50 µL of the double-sided SPRIselect cleaned products (eluant) into the newly labeled TempAssure 8-strip tube using a P200 @ 50 µL.

    - Dispose of the divided reservoir containing 80% Ethanol.

    - Start the 20C LIGATE protocol on the thermocycler.

    - Prepare adaptor ligation mix in a 1.5 mL Eppendorf LoBind tube (Eppendorf 022431021) by vortexing and adding the following regents in order:

      - Vortex and add 176 µL of ligation buffer (2000092) using a P200 @ 176 µL.

      - Vortex and add 88 µL of DNA Ligase (220110) using a P200 @ 88 µL.

      - Vortex and add 176 µL Adaptor Oligos (2000094) using a P200 @ 176 µL.

    - Vortex the adaptor ligation master mix at max speed for 5 seconds and spin down briefly to collect.

    - Add 50 µL of adaptor ligation master mix to each well of the 8-strip containing double-sided SPRIselect cleaned products.

    - Close the caps and vortex the strip for 5 seconds.

    - Spin down the strip briefly to collect the volume.

    - Transfer the strip to the thermocycler running the 20°C LIGATE and advance the protocol to begin the 15 minute 20°C ligation incubation.

    - Upon completion, return the samples to the Post-AMP hood to begin post-ligation cleanup.

    - Free-pour 5 mL of 80% Ethanol into a 25 mL divided reservoir (VWR 41428-958).

    - Label TempAssure 8-tube strip(s) (USA Scientific 1402-4700) to accommodate the final cycling conditions and library IDs of each sample.

    - Label a TempAssure 8-tube strip (USA Scientific 1402-4700) “EB” and add 40 µL of Qiagen Buffer EB (Qiagen 19086).

    - Vortex the SPRIselect reagent until fully re-suspended.

    - Add 80 µL of SPRIselect reagent to each sample in the 8-strip using a P200 @ 80 µL.

    - Cap the strip and pulse vortex it at max speed for 5 seconds.

    - Incubate the strip for 5 minutes at room temperature.

    - Place the strip in the 10x magnetic separator in the “High” position until the solution clears. Let strips sit in magnet for minimum of 2 minutes.

    - Carefully remove and discard the supernatant using an 8-channel P200 @ 180 µL.

    - Add 200 µL of 80% Ethanol from the reagent reservoir to the strip using an 8-channel P200 @ 200 µL.

    - Let the Ethanol wash sit for 30 seconds.

    - Carefully remove and discard the Ethanol wash using an 8-channel P200 @ 200 µL.

    - Repeat the Ethanol wash a second time by adding 200 µL of 80% Ethanol using an 8-channel P200 @ 200 µL.

    - After 30 seconds carefully remove and discard the Ethanol wash using an 8-channel P200 @ 200 µL.

    - Cap and briefly centrifuge the 8-strip to collect the residual Ethanol.

    - Start a timer for 2 minutes, place the 8-strip into the 10x magnetic separator on the “Low” position, and uncap the tubes.

    - Remove any residual Ethanol using an 8-channel P20 @10 µL.

    - Just before the 2-minute drying time is up, remove the strip from the magnet and add 31 µL of Qiagen Buffer EB (Qiagen 19086) from the “EB” strip using an 8-channel P200 @ 31 µL.

      - Vortex the beads and elution buffer at speed 9 until the beads are homogenous in solution.

    - Incubate the eluted beads in the strip for 5 minutes at room temperature.

    - Briefly centrifuge the strip to collect the volume and place it in the 10X magnetic separator in the “Low” position until the solution clears (around 2 minutes).

    - Transfer 30 µL of the post-ligation cleanup products (eluant) to the labeled TempAssure 8-strip tubes.

    - Vortex and spin down the thawed Dual Index Kit TT Set A (10x Genomics 1000215).

    - Briefly vortex the Sample Index PCR master mix reagents for 5 seconds at max speed.

    - Add 50 µL of Amp Mix (2000047) to each tube containing 30 µL of post ligation sample using a P200 @ 50 µL.

    - Under witness, add 20 µL of individual index from the Dual Index Kit TT Set A (10x Genomics 1000215) to each corresponding well.

      - Record the index IDs associated with each sample for demultiplexing.

    - Cap and vortex the tubes for 10 seconds at max speed. Spin down tubes.

    - Consult the following table for the number of PCR cycles based off the amount of input determined by PicoGreen from the amplification step earlier:

|  | LIBRARY | (KS - 10/26/17) |
| --- | --- | --- |
|  | ng Library Input | Sample Index Cycles |
|  | 1-12 ng | 16 |
|  | 13-25 ng | 15 |
|  | 26-65 ng | 14 |
|  | 66-110 ng | 13 |
| cDNA target | 111-15 0ng | 12 |
| cDNA target | 151-300 ng | 11 |
| cDNA target | 301-500 ng | 10 |
|  | 501-750 ng | 9 |
|  | 751-100 ng | 8 |
|  | 1001-1250 ng | 7 |
|  | 1251-1500 ng | 6 |
|  | >1500 ng | 5 |

    - After determining the PCR cycling conditions, run the following PCR protocol on the thermocycler: INDEX PCR NN where NN is equal to the number of PCR cycles.

    - Store samples at 4°C for up to 72 hours, -20°C for up to one week, or proceed with the final double-sided size selection.

    - Label four TempAssure 8-tube PCR strips, one for QC, one for buffer EB, and two for samples during cleanup.

    - In black ink, label a 96-well Eppendorf Twin.Tek Full-Skirted PCR plate (Eppendorf 0030129300) with the library name for post-cleanup sample storage.

    - Add 50 µL of Qiagen Buffer EB (Qiagen 19086) to the EB labeled strip.

    - Free pour 5 mL of 80% Ethanol into a new divided reservoir (VWR 41428-958).

    - Vortex the SPRIselect reagent (Beckman Coulter B23318) until completely re-suspended.

    - Label a TempAssure 8-tube PCR strip with the sample IDs.

    - If necessary, to condense the samples into single 8-strips, transfer the entire volume (100 µL) of sample from the 8-strip tubes to the corresponding wells of the sample ID labeled 8-strip.

    - Add 60 µL of homogenized SPRIselect reagent (Beckman Coulter B23318) reagent to the sample labeled 8-strip using a P200 @ 60 µL.

    - Cap and vortex the tube strip and then let it stand at room temperature for 5 minutes.

    - Briefly spin down the strip to collect the volume and place it in the 10x magnetic separator in the “High” position. Let strips sit in magnet for minimum of 2 minutes.

    - In the other TempAssure 8-tube PCR strip labeled for cleanup, add 20 µL of homogenized SPRIselect reagent to each well of the new 8-strip using a P20 @ 20 µL.

    - Once clear, transfer 150 µL of the supernatant from the strip in the magnetic separator to the strip containing 20 µL of SPRIselect using an 8-channel P200 @ 150 µL.

      - Take care to not transfer any of the beads at this point.

      - Discard the strip from the 10x magnetic separator containing the large products bound to the beads.

    - Cap and vortex the cleanup strip for 10 seconds at max speed.

    - Allow the strip to incubate off-magnet for 5 minutes at room temperature.

    - Once the incubation is over, place the strip into the 10x magnetic separator in the “High” position until the solution is clear. Let strips sit in magnet for minimum of 2 minutes.

    - Carefully remove and discard the supernatant using an 8-channel P200 @ 165 µL.

      - NOTE: Due to the low volume of beads it is important to not discard any beads with the supernatant, so some small volume will be left behind after aspiration.

    - With the strip still in the magnetic separator, add 200 µL of 80% Ethanol to the beads to wash them using an 8-channel P200 @ 200 µL.

    - Allow the wash to stand for 30 seconds.

    - Remove and discard the Ethanol wash using an 8-channel P200 @ 200 µL.

    - Repeat the wash a second time by adding 200 µL of 80% Ethanol using an 8-channel P200 @ 200 µL.

    - Allow this wash to stand for 30 seconds.

    - Remove and discard the Ethanol wash using an 8-channel P200 @ 200 µL.

    - Cap the tubes and briefly centrifuge them to collect any residual volume.

    - Place the strip in the 10x magnetic separator in the “Low” position.

    - Remove and discard any residual Ethanol wash using an 8-channel P20 @ 10 µL.

    - Remove the strip from the magnet and immediately add 36 µL of Qiagen Buffer EB (Qiagen 19086) from the “EB” labeled strip to the strip containing the beads using an 8-channel P200 @ 36 µL.

    - Vortex the beads on speed 9 until the solution is homogenous.

    - Allow the strip to stand off-magnet at room temperature for 5 minutes.

    - Return the strip to the 10x magnetic separator in the “Low” position until clear. Let strips sit in magnet for minimum of 2 minutes.

    - Transfer 35 µL from the tube strip into the labeled Twin.Tek 96-well PCR plate (Eppendorf 0030129300) using an 8-channel P200 @ 35 µL.

      - The samples should be transferred to the well in the plate that corresponds with their selected index.

    - Add 9 µL of Qiagen Buffer EB (Qiagen 19086) to the “QC” labeled strip using an 8-channel P10 @ 9 µL.

    - Add 1 µL of the cleaned 10x library product in the Twin.Tek 96-well PCR plate to the “QC” strip using an 8-channel P10 @ 1 µL. Cap, vortex and then spin down QC strip.

    - Store samples at 4°C for up to 72 hours, or -20°C long term.

    - Proceed to Library Quant and QC using the 10-fold diluted 10x libraries in the QC strip.

  - Library Quant and QC

    - Follow MB0194: 384-well PicoGreen to quantify the sample, using 3 µL from the 10x diluted QC strip.

    - Multiply the resulting concentration and yield by 10 to achieve the actual sample concentration and yields.

    - Run 1 µL of the 10X diluted 10x QC sample on the FA using the FA SOP.

      - Check the trace of the amplified cDNA. High quality cDNA should closely resemble the following:

      - All products should be between 200-600 bp since those are ideal sizes for sequencing.

      - Check for significant differences between the FA yield and PicoGreen yield.

    - Calculate the molarity of each library using the average size from the FA and the concentration from PicoGreen.

      - In general, we target molarity above 60 nM, and samples below 20 nM would be considered a failure and should be rescued with additional index PCR cycles. However, like the cDNA amplification, there are instances of samples of known lower quality, or that cannot be re-done that will still pass.

      - Below are the molarity ranges we have seen for 265 v3.1 HT libraries that have passed on to sequencing.

| Passing v3.1 HT Libraries |  |
| --- | --- |
|  | Stock nM |
| Sample Number | 265 |
| Min | 25.46 |
| Max | 444.82 |
| Average | 128.28 |
| Median | 85.67 |

| Passing v3.1 HT Libraries |  |
| --- | --- |
|  | Stock nM |
| Sample Number | 265 |
| Min | 25.46 |
| Max | 444.82 |
| Average | 128.28 |
| Median | 85.67 |

- Take Down:

  - Return any unused reagents.

  - Clean the hood and ice bucket.