## Fluence Calculations
Modified from Bolton LP Spreadsheet May 2016 by Natalie Hull | [7 empty] | Excimer Lamp 222 nm (modified by DTM for Autumn 2022, OSU CEGE, Hull Lab) | [7 empty] | Determination of the "Petri Factor"
INSTRUCTIONS AND NOTES: All inputs are shaded yellow
Make sure radiometer calibration data is up to date on Weighting tab | [18 empty] | x | -10 | -9.75 | -9.5 | -9.25 | -9 | -8.75 | -8.5 | -8.25 | -8 | -7.75 | -7.5 | -7.25 | -7 | -6.75 | -6.5 | -6.25 | -6 | -5.75 | -5.5 | -5.25 | -5 | -4.75 | -4.5 | -4.25 | -4 | -3.75 | -3.5 | -3.25 | -3 | -2.75 | -2.5 | -2.25 | -2 | -1.75 | -1.5 | -1.25 | -1 | -0.75 | -0.5 | -0.25 | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.25 | 1.5 | 1.75 | 2 | 2.25 | 2.5 | 2.75 | 3 | 3.25 | 3.5 | 3.75 | 4 | 4.25 | 4.5 | 4.75 | 5 | 5.25 | 5.5 | 5.75 | 6 | 6.25 | 6.5 | 6.75 | 7 | 7.25 | 7.5 | 7.75 | 8 | 8.25 | 8.5 | 8.75 | 9 | 9.25 | 9.5 | 9.75 | 10 |  |  |  | x | y | Ratio | x | y | Ratio
Enter lamp spectra from Spectra tab in Weighting tab. Normalize radiometer factor to peak output wavelength, based on Spectra tab information | [10 empty] | Steps: | [7 empty] | y
| Or normalize to 254 for germicidal comparisons (useful for MP UV). | [9 empty] | 1. Draw a 0.5 cm x 0.5 cm grid and place the center of the grid at the center of the collimated beam | [7 empty] | -10 | [82 empty] | 0 | 0 | 0 | -10 | 0 | -10 | 0 | 0
Set up a 'quasi' collimated beam apparatus. A 'collimating tube' may be used, but it is preferable to use circular 'masks' to define the beam. Make sure that | [10 empty] | and at the height of the sample surface/ radiometer detector plane. | [7 empty] | -9.75 | [82 empty] | 0 | 0 | 0 | -9.75 | 0 | -9.75 | 0 | 0
safety measures are taken to protect workers from exposure to the UV from the lamp. EYE PROTECTION IS AN ABSOLUTE REQUIREMENT. | [10 empty] | 2. Measure the UV irradiance (with a radiometer) every 0.5 cm in the x and y directions and input readings into columns S and X. | [7 empty] | -9.5 | [82 empty] | 0 | 0 | 0 | -9.5 | 0 | -9.5 | 0 | 0
Place the detector head of the UV radiometer on a horizontal surface, containing a 0.5 cm x 0.5 cm grid, such that the 'calibration plane' (see the Calibration | [10 empty] | 3. Take readings out to 0.5 cm beyond the radius of the sample dish (the distance required will be shaded in grey once you input the diameter) | [7 empty] | -9.25 | [82 empty] | 0 | 0 | 0 | -9.25 | 0 | -9.25 | 0 | 0
Sheet provided by the manufacturer of the Radiometer) is at the level of where the top of the solution will be during exposures to the UV. | [10 empty] | 4. The Petri factor is calculated in Cell DL90. | [7 empty] | -9 | [82 empty] | 0 | 0 | 0 | -9 | 0 | -9 | 0 | 0
Zero the radiometer with the cap on, and measure the output at the center of the dish, ensuring to use the radiometer factor for the lamp peak wavelength | [18 empty] | -8.75 | [82 empty] | 0 | 0 | 0 | -8.75 | 0 | -8.75 | 0 | 0
Determine the Petri Factor (See column Q) | [18 empty] | -8.5 | [82 empty] | 0 | 0 | 0 | -8.5 | 0 | -8.5 | 0 | 0
Make sure that the spectrophotometer is balanced with distilled water in the same cuvette. Then measure the absorption coefficient (1 cm absorbance) | [10 empty] | UV Distribution across the Petri Dish for a low pressure UV Lamp | [7 empty] | -8.25 | [82 empty] | 0 | 0 | 0 | -8.25 | 0 | -8.25 | 0 | 0
from 200-400nm for the sample to be irradiated and paste into Weighting tab. | [18 empty] | -8 | [82 empty] | 0 | 0 | 0 | -8 | 0 | -8 | 0 | 0
Insert the solution volume added to the Petri dish | [12 empty] | Meter |  |  |  | Meter |  | -7.75 | [82 empty] | 0 | 0 | 0 | -7.75 | 0 | -7.75 | 0 | 0
Insert the stirrer volume | [10 empty] | x | y | Reading | Ratio | x | y | Reading | Ratio | -7.5 | [82 empty] | 0 | 0 | 0 | -7.5 | 0 | -7.5 | 0 | 0
Insert the distance from the center of the UV lamp to the surface of the water in the Petri Dish | [18 empty] | -7.25 | [82 empty] | 0 | 0 | 0 | -7.25 | 0 | -7.25 | 0 | 0
Insert the desired Fluences (UV Doses) | [10 empty] | 0 | -10 |  | 0 | -10 | 0 |  | 0 | -7 | [82 empty] | 0 | 0 | 0 | -7 | 0 | -7 | 0 | 0
Choose a spectral weighting basis (ie germicidal, DNA absorbance, action spectrum, or normalized to 254 for MP) and apply the formula in the weighting tab | [10 empty] | 0 | -9.5 |  | 0 | -9.5 | 0 |  | 0 | -6.75 | [82 empty] | 0 | 0 | 0 | -6.75 | 0 | -6.75 | 0 | 0
Remove the radiometer detector and place a Petri Dish (or other container), containing the cell suspension, on a stirring motor placed so that the top of | [10 empty] | 0 | -9 |  | 0 | -9 | 0 |  | 0 | -6.5 | [82 empty] | 0 | 0 | 0 | -6.5 | 0 | -6.5 | 0 | 0
the solution is at the same level as that of the calibration plane of the detector head. Add a very small stir bar and make sure that the stirring rate is such that | [10 empty] | 0 | -8.5 |  | 0 | -8.5 | 0 |  | 0 | -6.25 | [82 empty] | 0 | 0 | 0 | -6.25 | 0 | -6.25 | 0 | 0
there is no vortex. | [10 empty] | 0 | -8 |  | 0 | -8 | 0 |  | 0 | -6 | [82 empty] | 0 | 0 | 0 | -6 | 0 | -6 | 0 | 0
Expose samples in the UV beam for the times calculated | [10 empty] | 0 | -7.5 |  | 0 | -7.5 | 0 |  | 0 | -5.75 | [82 empty] | 0 | 0 | 0 | -5.75 | 0 | -5.75 | 0 | 0
For a control, or zero dose determination, a sample solution is added to the Petri dish, set into the collimated beam apparatus and left stirring for a | [10 empty] | 0 | -7 |  | 0 | -7 | 0 |  | 0 | -5.5 | [82 empty] | 0 | 0 | 0 | -5.5 | 0 | -5.5 | 0 | 0
time equal to the average of the exposure times used. The resulting counts are used as the 0 dose concentration | [10 empty] | 0 | -6.5 |  | 0 | -6.5 | 0 |  | 0 | -5.25 | [82 empty] | 0 | 0 | 0 | -5.25 | 0 | -5.25 | 0 | 0
Note: the exposure times should be at least 30 s. If they are calculated to be shorter, arrange the irradiation platform further away from the UV lamp so | [10 empty] | 0 | -6 |  | 0 | -6 | 0 |  | 0 | -5 | [82 empty] | 0 | 0 | 0 | -5 | 0 | -5 | 0 | 0
that the irradiance will be smaller. If this is done, the Petri factor must be re-measured | [10 empty] | 0 | -5.5 |  | 0 | -5.5 | 0 |  | 0 | -4.75 | [82 empty] | 0 | 0 | 0 | -4.75 | 0 | -4.75 | 0 | 0
[11 empty] | 0 | -5 |  | 0 | -5 | 0 |  | 0 | -4.5 | [82 empty] | 0 | 0 | 0 | -4.5 | 0 | -4.5 | 0 | 0
|  |  |  | λ peak = | 254 | from Spectra tab |  |  |  |  | 0 | -4.5 |  | 0 | -4.5 | 0 |  | 0 | -4.25 | [82 empty] | 0 | 0 | 0 | -4.25 | 0 | -4.25 | 0 | 0
|  |  |  | λ avg = | 258.47 | from Spectra tab |  |  |  |  | 0 | -4 |  | 0 | -4 | 0 |  | 0 | -4 | [82 empty] | 0 | 0 | 0 | -4 | 0 | -4 | 0 | 0
[11 empty] | 0 | -3.5 | 0.1868 | 0.90416 | -3.5 | 0 | 0.1872 | 0.90962 | -3.75 | [82 empty] | 0 | 0 | 0 | -3.75 | 0.45208 | -3.75 | 0 | 0.45481
|  |  |  | Divergence Factor = | 0.98553 | =d/(d+l). Target >0.9 |  |  |  |  | 0 | -3 | 0.1909 | 0.92401 | -3 | 0 | 0.1906 | 0.92614 | -3.5 | [82 empty] | 0 | 0 | 0 | -3.5 | 0.90416 | -3.5 | 0 | 0.90962
|  |  |  | Reflection Factor at λpeak = | 0.97505 | from Reflection Factor tab |  |  |  |  | 0 | -2.5 | 0.1946 | 0.94192 | -2.5 | 0 | 0.194 | 0.94266 | -3.25 | [82 empty] | 0 | 0 | 0 | -3.25 | 0.91409 | -3.25 | 0 | 0.91788
|  |  |  | Petri Factor = | 0.98665 | calculated in cell DL90. Target >0.9 |  |  |  |  | 0 | -2 | 0.1986 | 0.96128 | -2 | 0 | 0.1965 | 0.95481 | -3 | [82 empty] | 0 | 0 | 0 | -3 | 0.92401 | -3 | 0 | 0.92614
|  |  |  | Spectral Water Factor = | 0.99628 | = Incident / Average Intensity. Target >0.9 |  |  |  |  | 0 | -1.5 | 0.2016 | 0.9758 | -1.5 | 0 | 0.199 | 0.96696 | -2.75 | [82 empty] | 0 | 0 | 0 | -2.75 | 0.93296 | -2.75 | 0 | 0.9344
[11 empty] | 0 | -1 | 0.2035 | 0.985 | -1 | 0 | 0.2017 | 0.98008 | -2.5 | [82 empty] | 0 | 0 | 0 | -2.5 | 0.94192 | -2.5 | 0 | 0.94266
|  |  |  | solution volume added to Petri dish = | 4 | mL |  |  |  |  | 0 | -0.5 | 0.205 | 0.99226 | -0.5 | 0 | 0.205 | 0.99611 | -2.25 | [82 empty] | 0 | 0 | 0 | -2.25 | 0.9516 | -2.25 | 0 | 0.94874
|  |  |  | stirrer volume = | 0 | mL |  |  |  |  | 0 | 0 | 0.2066 | 1 | 0 | 0 | 0.2058 | 1 | -2 | [82 empty] | 0 | 0 | 0 | -2 | 0.96128 | -2 | 0 | 0.95481
|  |  |  | Petri dish internal diameter = | 3.4 | cm |  |  |  |  | 0 | 0.5 | 0.2062 | 0.99806 | 0.5 | 0 | 0.2069 | 1.0053 | -1.75 | [82 empty] | 0 | 0 | 0 | -1.75 | 0.96854 | -1.75 | 0 | 0.96088
|  |  |  | total volume in Petri dish = | 4 | mL |  |  |  |  | 0 | 1 | 0.2045 | 0.98984 | 1 | 0 | 0.2066 | 1.0039 | -1.5 | [37 empty] | 0.96418 | 0.97201 | 0.9739 | 0.9758 | 0.97841 | 0.98101 | 0.9803 | [38 empty] | 6.8256 | 7 | 0 | -1.5 | 0.9758 | -1.5 | 0 | 0.96696
|  |  |  | water path length = | 0.44057 | cm |  |  |  |  | 0 | 1.5 | 0.203 | 0.98258 | 1.5 | 0 | 0.2065 | 1.0034 | -1.25 | [36 empty] | 0.96087 | 0.96873 | 0.97659 | 0.97849 | 0.9804 | 0.98302 | 0.98564 | 0.98492 | 0.98421 | [37 empty] | 8.8028 | 9 | 0 | -1.25 | 0.9804 | -1.25 | 0 | 0.97352
|  |  |  | distance from UV lamp to top of water surface = | 30 | cm |  |  |  |  | 0 | 2 | 0.204 | 0.98742 | 2 | 0 | 0.2067 | 1.0044 | -1 | [35 empty] | 0.95891 | 0.96537 | 0.97327 | 0.98117 | 0.98308 | 0.985 | 0.98763 | 0.99026 | 0.98954 | 0.98882 | 0.98858 | [36 empty] | 10.7916 | 11 | 0 | -1 | 0.985 | -1 | 0 | 0.98008
|  |  |  | Absorbance scan = | done | use Absorbance tab |  |  |  |  | 0 | 2.5 | 0.2019 | 0.97725 | 2.5 | 0 | 0.2072 | 1.0068 | -0.75 | [34 empty] | 0.95596 | 0.96244 | 0.96893 | 0.97686 | 0.98478 | 0.9867 | 0.98863 | 0.99127 | 0.99391 | 0.99319 | 0.99247 | 0.99223 | 0.99199 | [35 empty] | 12.7794 | 13 | 0 | -0.75 | 0.98863 | -0.75 | 0 | 0.9881
|  |  |  | absorption coefficient at 254 nm = | 0.0076716 | cm-1 |  |  |  |  | 0 | 3 | 0.196 | 0.94869 | 3 | 0 | 0.207 | 1.0058 | -0.5 | [34 empty] | 0.95947 | 0.96598 | 0.97249 | 0.98044 | 0.9884 | 0.99033 | 0.99226 | 0.99491 | 0.99756 | 0.99684 | 0.99611 | 0.99587 | 0.99563 | [35 empty] | 12.8263 | 13 | 0 | -0.5 | 0.99226 | -0.5 | 0 | 0.99611
|  |  |  | UVT % at 254 = | 98.2491 | make sure this isn't <90% for disinfection studies |  |  |  |  | 0 | 3.5 | 0.1907 | 0.92304 | 3.5 | 0 | 0.2047 | 0.99466 | -0.25 | [34 empty] | 0.96321 | 0.96975 | 0.97628 | 0.98427 | 0.99226 | 0.99419 | 0.99613 | 0.99879 | 1.0015 | 1.0007 | 1 | 0.99976 | 0.99952 | [35 empty] | 12.8763 | 13 | 0 | -0.25 | 0.99613 | -0.25 | 0 | 0.99806
[11 empty] | 0 | 4 |  | 0 | 4 | 0 |  | 0 | 0 | [34 empty] | 0.96696 | 0.97352 | 0.98008 | 0.9881 | 0.99611 | 0.99806 | 1 | 1.0027 | 1.0053 | 1.0046 | 1.0039 | 1.0036 | 1.0034 | [35 empty] | 12.9264 | 13 | 0 | 0 | 1 | 0 | 0 | 1
|  |  |  | Radiometer Factor at λpeak = | 0.002796 | [5 empty] | 0 | 4.5 |  | 0 | 4.5 | 0 |  | 0 | 0.25 | [34 empty] | 0.96602 | 0.97258 | 0.97913 | 0.98714 | 0.99515 | 0.99709 | 0.99903 | 1.0017 | 1.0044 | 1.0036 | 1.0029 | 1.0027 | 1.0024 | [35 empty] | 12.9139 | 13 | 0 | 0.25 | 0.99903 | 0.25 | 0 | 1.0027
|  |  |  | Radiometer Reading Center of Dish = | 0.19 | mW/cm2 |  | W/cm2 |  |  | 0 | 5 |  | 0 | 5 | 0 |  | 0 | 0.5 | [34 empty] | 0.96509 | 0.97163 | 0.97818 | 0.98618 | 0.99418 | 0.99612 | 0.99806 | 1.0007 | 1.0034 | 1.0027 | 1.0019 | 1.0017 | 1.0015 | [35 empty] | 12.9014 | 13 | 0 | 0.5 | 0.99806 | 0.5 | 0 | 1.0053
|  |  |  | Incident Irradiance= |  | mW/cm2 |  |  |  |  | 0 | 5.5 |  | 0 | 5.5 | 0 |  | 0 | 0.75 | [34 empty] | 0.96111 | 0.96763 | 0.97415 | 0.98212 | 0.99009 | 0.99202 | 0.99395 | 0.99661 | 0.99926 | 0.99854 | 0.99781 | 0.99757 | 0.99733 | [35 empty] | 12.8482 | 13 | 0 | 0.75 | 0.99395 | 0.75 | 0 | 1.0046
|  |  |  | Average Irradiance= | 0.18529 | mW/cm2 |  |  |  |  | 0 | 6 |  | 0 | 6 | 0 |  | 0 | 1 | [35 empty] | 0.96362 | 0.97012 | 0.97805 | 0.98599 | 0.98791 | 0.98984 | 0.99248 | 0.99513 | 0.9944 | 0.99368 | 0.99344 | [36 empty] | 10.8447 | 11 | 0 | 1 | 0.98984 | 1 | 0 | 1.0039
[11 empty] | 0 | 6.5 |  | 0 | 6.5 | 0 |  | 0 | 1.25 | [36 empty] | 0.96656 | 0.97446 | 0.98237 | 0.98429 | 0.98621 | 0.98884 | 0.99148 | 0.99076 | 0.99004 | [37 empty] | 8.855 | 9 | 0 | 1.25 | 0.98621 | 1.25 | 0 | 1.0036
|  |  |  | Weighting Factor = | NA | choose and apply formula on Weighting tab, record weighting type here |  |  |  |  | 0 | 7 |  | 0 | 7 | 0 |  | 0 | 1.5 | [37 empty] | 0.97088 | 0.97876 | 0.98067 | 0.98258 | 0.9852 | 0.98783 | 0.98711 | [38 empty] | 6.873 | 7 | 0 | 1.5 | 0.98258 | 1.5 | 0 | 1.0034
|  |  |  | Weighted Average Irradiance= |  | mW/cm2 |  |  |  |  | 0 | 7.5 |  | 0 | 7.5 | 0 |  | 0 | 1.75 | [82 empty] | 0 | 0 | 0 | 1.75 | 0.985 | 1.75 | 0 | 1.0039
[11 empty] | 0 | 8 |  | 0 | 8 | 0 |  | 0 | 2 | [82 empty] | 0 | 0 | 0 | 2 | 0.98742 | 2 | 0 | 1.0044
|  | Seconds for a Fluence (UV Dose) of | 1 | mJ/cm2 = | 5.3968 | change formula to use weighted average irradiance, if applicable |  |  |  |  | 0 | 8.5 |  | 0 | 8.5 | 0 |  | 0 | 2.25 | [82 empty] | 0 | 0 | 0 | 2.25 | 0.98233 | 2.25 | 0 | 1.0056
[11 empty] | 0 | 9 |  | 0 | 9 | 0 |  | 0 | 2.5 | [82 empty] | 0 | 0 | 0 | 2.5 | 0.97725 | 2.5 | 0 | 1.0068
[9 empty] | min | seconds | 0 | 9.5 |  | 0 | 9.5 | 0 |  | 0 | 2.75 | [82 empty] | 0 | 0 | 0 | 2.75 | 0.96297 | 2.75 | 0 | 1.0063
|  |  |  | Time for a Fluence (UV Dose) of | 0 | mJ/cm2 = | 0 | s = | 0 | 0 | 0 | 10 |  | 0 | 10 | 0 |  | 0 | 3 | [82 empty] | 0 | 0 | 0 | 3 | 0.94869 | 3 | 0 | 1.0058
|  |  |  | Time for a Fluence (UV Dose) of | 10 | mJ/cm2 = | 53.9684 | s = | 0 | 53.9684 | [8 empty] | 3.25 | [82 empty] | 0 | 0 | 0 | 3.25 | 0.93587 | 3.25 | 0 | 1.0002
|  |  |  | Time for a Fluence (UV Dose) of | 20 | mJ/cm2 = | 107.9368 | s = | 1 | 47.9368 | [8 empty] | 3.5 | [82 empty] | 0 | 0 | 0 | 3.5 | 0.92304 | 3.5 | 0 | 0.99466
|  |  |  | Time for a Fluence (UV Dose) of | 30 | mJ/cm2 = | 161.9052 | s = | 2 | 41.9052 | [8 empty] | 3.75 | [82 empty] | 0 | 0 | 0 | 3.75 | 0.46152 | 3.75 | 0 | 0.49733
|  |  |  | Time for a Fluence (UV Dose) of | 40 | mJ/cm2 = | 215.8736 | s = | 3 | 35.8736 | [8 empty] | 4 | [82 empty] | 0 | 0 | 0 | 4 | 0 | 4 | 0 | 0
|  |  |  | Time for a Fluence (UV Dose) of | 50 | mJ/cm2 = | 269.8419 | s = | 4 | 29.8419 | [8 empty] | 4.25 | [82 empty] | 0 | 0 | 0 | 4.25 | 0 | 4.25 | 0 | 0
|  |  |  | Time for a Fluence (UV Dose) of | 60 | mJ/cm2 = | 323.8103 | s = | 5 | 23.8103 | [8 empty] | 4.5 | [82 empty] | 0 | 0 | 0 | 4.5 | 0 | 4.5 | 0 | 0
|  |  |  | Time for a Fluence (UV Dose) of |  | mJ/cm2 = | 0 | s = | 0 | 0 | [8 empty] | 4.75 | [82 empty] | 0 | 0 | 0 | 4.75 | 0 | 4.75 | 0 | 0
|  |  |  | Time for a Fluence (UV Dose) of |  | mJ/cm2 = | 0 | s = | 0 | 0 | [8 empty] | 5 | [82 empty] | 0 | 0 | 0 | 5 | 0 | 5 | 0 | 0
|  |  |  | Time for a Fluence (UV Dose) of |  | mJ/cm2 = | 0 | s = | 0 | 0 | [8 empty] | 5.25 | [82 empty] | 0 | 0 | 0 | 5.25 | 0 | 5.25 | 0 | 0
|  |  |  | Time for a Fluence (UV Dose) of |  | mJ/cm2 = | 0 | s = | 0 | 0 | [8 empty] | 5.5 | [82 empty] | 0 | 0 | 0 | 5.5 | 0 | 5.5 | 0 | 0
|  |  |  | Time for a Fluence (UV Dose) of |  | mJ/cm2 = | 0 | s = | 0 | 0 | [8 empty] | 5.75 | [82 empty] | 0 | 0 | 0 | 5.75 | 0 | 5.75 | 0 | 0
|  |  |  | Time for a Fluence (UV Dose) of |  | mJ/cm2 = | 0 | s = | 0 | 0 | [8 empty] | 6 | [82 empty] | 0 | 0 | 0 | 6 | 0 | 6 | 0 | 0
[19 empty] | 6.25 | [82 empty] | 0 | 0 | 0 | 6.25 | 0 | 6.25 | 0 | 0
[19 empty] | 6.5 | [82 empty] | 0 | 0 | 0 | 6.5 | 0 | 6.5 | 0 | 0
[19 empty] | 6.75 | [82 empty] | 0 | 0 | 0 | 6.75 | 0 | 6.75 | 0 | 0
[19 empty] | 7 | [82 empty] | 0 | 0 | 0 | 7 | 0 | 7 | 0 | 0
[19 empty] | 7.25 | [82 empty] | 0 | 0 | 0 | 7.25 | 0 | 7.25 | 0 | 0
[19 empty] | 7.5 | [82 empty] | 0 | 0 | 0 | 7.5 | 0 | 7.5 | 0 | 0
[19 empty] | 7.75 | [82 empty] | 0 | 0 | 0 | 7.75 | 0 | 7.75 | 0 | 0
[19 empty] | 8 | [82 empty] | 0 | 0 | 0 | 8 | 0 | 8 | 0 | 0
[19 empty] | 8.25 | [82 empty] | 0 | 0 | 0 | 8.25 | 0 | 8.25 | 0 | 0
[19 empty] | 8.5 | [82 empty] | 0 | 0 | 0 | 8.5 | 0 | 8.5 | 0 | 0
[19 empty] | 8.75 | [82 empty] | 0 | 0 | 0 | 8.75 | 0 | 8.75 | 0 | 0
[19 empty] | 9 | [82 empty] | 0 | 0 | 0 | 9 | 0 | 9 | 0 | 0
[19 empty] | 9.25 | [82 empty] | 0 | 0 | 0 | 9.25 | 0 | 9.25 | 0 | 0
[19 empty] | 9.5 | [82 empty] | 0 | 0 | 0 | 9.5 | 0 | 9.5 | 0 | 0
[19 empty] | 9.75 | [82 empty] | 0 | 0 | 0 | 9.75 | 0 | 9.75 | 0 | 0
[19 empty] | 10 | [82 empty] | 0 | 0 | 0 | 10 | 0 | 10 | 0 | 0
[102 empty] | 143.0645 | 145
[101 empty] | Petri factor = | 0.98665
Note on G38: Jim Bolton: Measure carefully with a plastic mm ruler or better with a pair of calipers.
Note on G40: Jim Bolton: The water path length may be entered directly.

## Absorbance
Take up to triplicate absorbance scans of your sample using a quartz cuvette
Enter the absorbance spectra into the columns Abs 1, 2 and 3 (if you have that many)
Final Absorbance will be calculated as the average of the scans
Final Absorbance gets input into the Weighting Tab
Remember to correct absorbance readings for culture dilutions in DI to back calculate original sample absorbance
Input
Calculated
Final Absorbance |  |  | Input, paste raw data values below (up to 3 replicates)
Wavelength | Absorbance | Wavelength | Abs 1 | Abs 2 | Abs 3 | Average Abs
200 | 2.6329 | 800 | -8.1688e-05 | 0.00063241 | 0.00034175 | 0.00029749
201 | 1.9545 | 799 | -7.0197e-05 | 0.00057704 | 0.00035421 | 0.00028702
202 | 1.4456 | 798 | -0.00010726 | 0.00061906 | 0.00035056 | 0.00028746
203 | 1.0686 | 797 | -6.6574e-05 | 0.00059255 | 0.00037838 | 0.00030145
204 | 0.79096 | 796 | -0.00011124 | 0.00064831 | 0.00039852 | 0.00031186
205 | 0.59011 | 795 | -8.4846e-05 | 0.00060102 | 0.00039292 | 0.00030303
206 | 0.45001 | 794 | -1.0976e-05 | 0.00060846 | 0.00047566 | 0.00035771
207 | 0.35229 | 793 | -8.7072e-05 | 0.00066005 | 0.00042945 | 0.00033414
208 | 0.28227 | 792 | 3.1996e-05 | 0.00068673 | 0.00049678 | 0.00040517
209 | 0.23264 | 791 | -6.7039e-05 | 0.00062718 | 0.00041365 | 0.0003246
210 | 0.19676 | 790 | -6.5745e-05 | 0.0005737 | 0.00037411 | 0.00029402
211 | 0.17034 | 789 | -0.0001644 | 0.00054871 | 0.00033745 | 0.00024059
212 | 0.14993 | 788 | -9.2196e-05 | 0.00060693 | 0.00040147 | 0.0003054
213 | 0.13359 | 787 | 6.9375e-06 | 0.00070817 | 0.00052684 | 0.00041398
214 | 0.12033 | 786 | -2.951e-06 | 0.00067903 | 0.00049305 | 0.00038971
215 | 0.10931 | 785 | 7.3315e-05 | 0.0006593 | 0.00047664 | 0.00040308
216 | 0.099818 | 784 | 1.7344e-06 | 0.00066345 | 0.00045625 | 0.00037381
217 | 0.091121 | 783 | -7.6357e-05 | 0.00061453 | 0.00039784 | 0.000312
218 | 0.083208 | 782 | -7.6409e-05 | 0.00062803 | 0.00040416 | 0.0003186
219 | 0.075766 | 781 | -3.9655e-05 | 0.00065466 | 0.0003816 | 0.0003322
220 | 0.06864 | 780 | 6.5233e-06 | 0.0006158 | 0.00045179 | 0.00035804
221 | 0.062169 | 779 | 8.0765e-06 | 0.00066816 | 0.00046902 | 0.00038175
222 | 0.056064 | 778 | -0.00010633 | 0.00058757 | 0.00035779 | 0.00027968
223 | 0.050643 | 777 | -5.6221e-05 | 0.00072673 | 0.00048867 | 0.00038639
224 | 0.045281 | 776 | -1.5272e-05 | 0.00065178 | 0.00047791 | 0.00037147
225 | 0.040349 | 775 | -3.277e-05 | 0.00065909 | 0.00044591 | 0.00035741
226 | 0.035324 | 774 | 2.6637e-05 | 0.00071657 | 0.00049398 | 0.0004124
227 | 0.030548 | 773 | 5.2681e-05 | 0.00070812 | 0.0005152 | 0.00042533
228 | 0.02633 | 772 | 3.0003e-05 | 0.00070381 | 0.00042704 | 0.00038695
229 | 0.022218 | 771 | -4.7213e-05 | 0.00064875 | 0.00044018 | 0.00034724
230 | 0.018826 | 770 | 4.6157e-05 | 0.00072567 | 0.00053835 | 0.00043673
231 | 0.015839 | 769 | 6.2386e-06 | 0.00072228 | 0.00050204 | 0.00041018
232 | 0.013256 | 768 | 8.8012e-07 | 0.00068787 | 0.00048232 | 0.00039036
233 | 0.011326 | 767 | 0.00010174 | 0.00077914 | 0.00052013 | 0.000467
234 | 0.0096702 | 766 | 3.2307e-05 | 0.00083311 | 0.00051678 | 0.00046073
235 | 0.0084813 | 765 | 6.2416e-05 | 0.00075998 | 0.00049105 | 0.00043781
236 | 0.0074639 | 764 | 6.0577e-05 | 0.0008772 | 0.00059426 | 0.00051068
237 | 0.0066378 | 763 | 8.3231e-05 | 0.00077242 | 0.00052352 | 0.00045973
238 | 0.0062107 | 762 | 6.3839e-05 | 0.00076013 | 0.00055693 | 0.0004603
239 | 0.0059227 | 761 | 9.5064e-05 | 0.00089118 | 0.00063018 | 0.00053881
240 | 0.0056289 | 760 | 3.6242e-05 | 0.00087858 | 0.00059376 | 0.00050286
241 | 0.0056816 | 759 | 0.00012893 | 0.00083472 | 0.00060779 | 0.00052381
242 | 0.0055786 | 758 | 0.00014146 | 0.00090185 | 0.00070545 | 0.00058292
243 | 0.0056838 | 757 | 0.00012163 | 0.00086898 | 0.00067758 | 0.00055606
244 | 0.0057924 | 756 | 8.1315e-05 | 0.00089513 | 0.00062026 | 0.00053223
245 | 0.0060034 | 755 | 0.00012049 | 0.00089162 | 0.00066967 | 0.00056059
246 | 0.0062176 | 754 | 0.00015439 | 0.0009082 | 0.00070438 | 0.00058899
247 | 0.0064445 | 753 | 9.7705e-05 | 0.00090646 | 0.00069127 | 0.00056514
248 | 0.0065917 | 752 | 0.00013385 | 0.0009358 | 0.00072225 | 0.0005973
249 | 0.0068964 | 751 | 0.00015503 | 0.00098351 | 0.00069891 | 0.00061249
250 | 0.0070442 | 750 | 0.00016676 | 0.00087593 | 0.00063563 | 0.00055944
251 | 0.0071719 | 749 | 0.0001273 | 0.00087534 | 0.00073661 | 0.00057975
252 | 0.0074077 | 748 | 7.0001e-05 | 0.00092376 | 0.00068844 | 0.00056074
253 | 0.0074842 | 747 | 0.00018826 | 0.00097902 | 0.0007627 | 0.00064333
254 | 0.0076716 | 746 | 0.00014294 | 0.00095279 | 0.00077242 | 0.00062272
255 | 0.0077713 | 745 | 0.00012153 | 0.00091673 | 0.00073729 | 0.00059185
256 | 0.0079251 | 744 | 0.0001722 | 0.00095604 | 0.00080191 | 0.00064338
257 | 0.00809 | 743 | 8.7141e-05 | 0.00084701 | 0.00064999 | 0.00052805
258 | 0.0081185 | 742 | 0.00015783 | 0.00099275 | 0.00072186 | 0.00062415
259 | 0.0083162 | 741 | 0.00023244 | 0.00093619 | 0.00074074 | 0.00063646
260 | 0.0082852 | 740 | 0.0001764 | 0.00097176 | 0.00078806 | 0.00064541
261 | 0.0083284 | 739 | 5.5063e-05 | 0.00088268 | 0.00069305 | 0.0005436
262 | 0.0083959 | 738 | 0.0001266 | 0.00085534 | 0.00068955 | 0.00055716
263 | 0.0083966 | 737 | 0.00018538 | 0.00096462 | 0.00074424 | 0.00063141
264 | 0.0085184 | 736 | 0.00014985 | 0.00092976 | 0.00074828 | 0.0006093
265 | 0.0085185 | 735 | 8.9264e-05 | 0.00081174 | 0.00065429 | 0.00051843
266 | 0.0084257 | 734 | 9.9957e-05 | 0.00089009 | 0.00069974 | 0.00056327
267 | 0.0084511 | 733 | 8.305e-05 | 0.00083767 | 0.00068123 | 0.00053399
268 | 0.0083631 | 732 | -2.1226e-06 | 0.00078215 | 0.00064984 | 0.00047662
269 | 0.0083403 | 731 | 3.9245e-05 | 0.00077455 | 0.0006901 | 0.0005013
270 | 0.0082375 | 730 | 0.00010892 | 0.00081835 | 0.0006622 | 0.00052982
271 | 0.0082114 | 729 | 3.7122e-05 | 0.00081197 | 0.00065499 | 0.00050136
272 | 0.0080789 | 728 | 5.3225e-05 | 0.00080785 | 0.00064618 | 0.00050242
273 | 0.0079362 | 727 | -3.7276e-06 | 0.00075793 | 0.00059128 | 0.00044849
274 | 0.0078569 | 726 | 5.3199e-05 | 0.00075627 | 0.00063166 | 0.00048038
275 | 0.0077657 | 725 | 0.0001135 | 0.00082502 | 0.00063011 | 0.00052287
276 | 0.0076867 | 724 | 6.0526e-05 | 0.00078412 | 0.00059483 | 0.00047982
277 | 0.0074587 | 723 | -1.1907e-05 | 0.00071258 | 0.00057417 | 0.00042495
278 | 0.007289 | 722 | -4.3488e-06 | 0.00075401 | 0.00056152 | 0.00043706
279 | 0.0071891 | 721 | 2.2055e-05 | 0.00073568 | 0.00060157 | 0.0004531
280 | 0.0070557 | 720 | 6.3037e-05 | 0.00069757 | 0.00053827 | 0.00043296
281 | 0.0068292 | 719 | -1.3046e-05 | 0.00066207 | 0.00054402 | 0.00039768
282 | 0.0065943 | 718 | -3.7119e-05 | 0.00069505 | 0.00051401 | 0.00039065
283 | 0.0062673 | 717 | 7.1995e-05 | 0.00078617 | 0.00063291 | 0.00049702
284 | 0.0059948 | 716 | -4.0328e-05 | 0.00069591 | 0.00052554 | 0.00039371
285 | 0.0056155 | 715 | -1.7913e-05 | 0.0006761 | 0.00054721 | 0.0004018
286 | 0.005357 | 714 | -3.8258e-05 | 0.0006684 | 0.00050556 | 0.00037857
287 | 0.0051804 | 713 | -1.3823e-05 | 0.00065686 | 0.00053534 | 0.00039279
288 | 0.0049243 | 712 | -5.2545e-05 | 0.00066365 | 0.00048553 | 0.00036555
289 | 0.0046842 | 711 | 1.2943e-07 | 0.0007054 | 0.00049903 | 0.00040152
290 | 0.0044139 | 710 | -7.2785e-05 | 0.00059104 | 0.00047975 | 0.00033267
291 | 0.0040187 | 709 | -4.3331e-05 | 0.00067864 | 0.00050543 | 0.00038025
292 | 0.0037011 | 708 | 4.3387e-05 | 0.00071276 | 0.00051697 | 0.00042437
293 | 0.0033761 | 707 | -3.7844e-05 | 0.00062202 | 0.00042865 | 0.00033761
294 | 0.0030782 | 706 | 6.2386e-06 | 0.00067361 | 0.00049294 | 0.00039093
295 | 0.0029073 | 705 | -4.7938e-05 | 0.00059853 | 0.0004621 | 0.00033757
296 | 0.0027085 | 704 | -2.4902e-05 | 0.00066142 | 0.00047612 | 0.00037088
297 | 0.0026422 | 703 | 1.2115e-05 | 0.00066975 | 0.00052718 | 0.00040301
298 | 0.0023456 | 702 | -6.7712e-05 | 0.00064875 | 0.00049727 | 0.00035943
299 | 0.0023566 | 701 | -6.4451e-05 | 0.00064854 | 0.00048569 | 0.00035659
300 | 0.002246 | 700 | -2.428e-05 | 0.00066321 | 0.0004918 | 0.00037691
301 | 0.0023133 | 699 | -5.9689e-05 | 0.00062684 | 0.00043567 | 0.00033427
302 | 0.0020654 | 698 | -3.774e-05 | 0.00067039 | 0.00049317 | 0.00037528
303 | 0.0020675 | 697 | -4.2037e-05 | 0.00065564 | 0.00050836 | 0.00037399
304 | 0.0019218 | 696 | 1.3978e-06 | 0.00067174 | 0.00051951 | 0.00039755
305 | 0.0019577 | 695 | 1.8379e-06 | 0.00074839 | 0.00054327 | 0.00043117
306 | 0.001946 | 694 | -2.0191e-06 | 0.00072756 | 0.00052072 | 0.00041542
307 | 0.0018501 | 693 | -4.1778e-05 | 0.00065499 | 0.00050155 | 0.00037159
308 | 0.0017135 | 692 | 1.2167e-05 | 0.00076651 | 0.00052915 | 0.00043594
309 | 0.0018036 | 691 | 3.8311e-06 | 0.00067312 | 0.00050546 | 0.00039414
310 | 0.0018024 | 690 | -3.1528e-05 | 0.00067662 | 0.00047944 | 0.00037484
311 | 0.0018007 | 689 | -6.7821e-06 | 0.00066863 | 0.00042645 | 0.00036277
312 | 0.0016782 | 688 | -4.4159e-05 | 0.00066863 | 0.00047483 | 0.00036643
313 | 0.0016608 | 687 | -1.0095e-05 | 0.00066936 | 0.00045614 | 0.0003718
314 | 0.0016318 | 686 | -3.0027e-05 | 0.00077121 | 0.00053684 | 0.00042601
315 | 0.0016031 | 685 | 2.19e-05 | 0.0007224 | 0.00052718 | 0.00042383
316 | 0.0015152 | 684 | 3.2514e-05 | 0.00069308 | 0.00051121 | 0.00041227
317 | 0.0015568 | 683 | 3.9219e-05 | 0.00068032 | 0.00049139 | 0.00040364
318 | 0.0015829 | 682 | -1.9673e-05 | 0.00062386 | 0.00045345 | 0.00035254
319 | 0.0015757 | 681 | -4.0121e-05 | 0.00062868 | 0.00052536 | 0.00037131
320 | 0.0014859 | 680 | 1.0018e-05 | 0.00061287 | 0.00048055 | 0.00036781
321 | 0.0015423 | 679 | 1.6127e-05 | 0.00066106 | 0.00048983 | 0.00038901
322 | 0.0015245 | 678 | -2.4384e-05 | 0.00066728 | 0.00050414 | 0.00038235
323 | 0.0014518 | 677 | 5.1772e-06 | 0.00067433 | 0.00050958 | 0.00039636
324 | 0.0014213 | 676 | 2.7828e-05 | 0.00066542 | 0.00050427 | 0.00039917
325 | 0.0015273 | 675 | 2.0968e-05 | 0.00065331 | 0.00046656 | 0.00038028
326 | 0.0014433 | 674 | -2.133e-05 | 0.00062435 | 0.00050704 | 0.00037002
327 | 0.0013595 | 673 | 1.0096e-05 | 0.00065904 | 0.00050222 | 0.00039045
328 | 0.0013031 | 672 | -5.3167e-05 | 0.00060258 | 0.0004422 | 0.00033054
329 | 0.0013448 | 671 | 4.9183e-07 | 0.00063742 | 0.00049517 | 0.00037769
330 | 0.0013985 | 670 | -3.422e-05 | 0.0005848 | 0.00041427 | 0.00032161
331 | 0.0013526 | 669 | 1.8224e-05 | 0.00064317 | 0.00047374 | 0.00037838
332 | 0.0013813 | 668 | -1.6049e-05 | 0.00065416 | 0.00044852 | 0.00036221
333 | 0.0011794 | 667 | -2.0708e-05 | 0.00062282 | 0.00046545 | 0.00035585
334 | 0.0012495 | 666 | 2.5524e-05 | 0.00066651 | 0.00049286 | 0.00039496
335 | 0.0012781 | 665 | -4.6594e-06 | 0.00065069 | 0.00048185 | 0.00037596
336 | 0.0011952 | 664 | 9.6038e-06 | 0.00069264 | 0.00050421 | 0.00040215
337 | 0.0011471 | 663 | -4.09e-06 | 0.0006607 | 0.00048302 | 0.00037987
338 | 0.0011351 | 662 | 7.927e-05 | 0.00066446 | 0.0004655 | 0.00040308
339 | 0.0012439 | 661 | 4.6595e-06 | 0.00065715 | 0.00051458 | 0.00039213
340 | 0.00096916 | 660 | -3.2874e-05 | 0.0006453 | 0.00047426 | 0.00036223
341 | 0.0011623 | 659 | -4.447e-05 | 0.00063418 | 0.00043347 | 0.00034106
342 | 0.0010939 | 658 | 3.342e-05 | 0.00066832 | 0.00040927 | 0.00037034
343 | 0.0011048 | 657 | 4.1937e-05 | 0.00066969 | 0.00048833 | 0.00039999
344 | 0.00094027 | 656 | 4.7141e-05 | 0.00064271 | 0.00047599 | 0.00038861
345 | 0.00099814 | 655 | 1.3331e-05 | 0.00070573 | 0.0004982 | 0.00040576
346 | 0.0010038 | 654 | -5.0992e-05 | 0.00068898 | 0.00040183 | 0.00034661
347 | 0.0009077 | 653 | -2.5212e-05 | 0.00064812 | 0.00044155 | 0.00035482
348 | 0.0011222 | 652 | -8.0763e-06 | 0.00071608 | 0.00050753 | 0.00040518
349 | 0.00084142 | 651 | -1.843e-05 | 0.00069108 | 0.00040429 | 0.00035898
350 | 0.0010446 | 650 | 4.1756e-05 | 0.00073355 | 0.00043834 | 0.00040455
351 | 0.0010236 | 649 | -6.6268e-06 | 0.00068489 | 0.00041303 | 0.00036376
352 | 0.00099329 | 648 | -6.8437e-05 | 0.00064175 | 0.00045967 | 0.00034433
353 | 0.0010458 | 647 | -8.5423e-06 | 0.00070114 | 0.00047809 | 0.00039023
354 | 0.00095677 | 646 | 4.9183e-07 | 0.0007007 | 0.00041904 | 0.00037341
355 | 0.0010009 | 645 | -2.9509e-05 | 0.00067721 | 0.00043119 | 0.00035963
356 | 0.00092927 | 644 | -3.5721e-05 | 0.00068478 | 0.00043502 | 0.00036136
357 | 0.00088763 | 643 | -2.2313e-05 | 0.00068235 | 0.00044907 | 0.0003697
358 | 0.00097059 | 642 | 1.9829e-05 | 0.00069922 | 0.00043505 | 0.0003847
359 | 0.00080683 | 641 | -3.7637e-05 | 0.00066993 | 0.00035644 | 0.00032958
360 | 0.00091058 | 640 | -2.9198e-05 | 0.00067467 | 0.00037883 | 0.00034143
361 | 0.00086687 | 639 | -1.8948e-05 | 0.00067257 | 0.00044391 | 0.00036585
362 | 0.0008648 | 638 | 5.5244e-05 | 0.00076519 | 0.00054617 | 0.00045554
363 | 0.00090633 | 637 | 2.3609e-05 | 0.00071774 | 0.00043344 | 0.0003916
364 | 0.00083712 | 636 | -5.3632e-05 | 0.00068686 | 0.00042653 | 0.00035325
365 | 0.00080704 | 635 | -2.076e-05 | 0.00067911 | 0.00040168 | 0.00035334
366 | 0.00086933 | 634 | 2.5964e-05 | 0.00071745 | 0.00046651 | 0.00040331
367 | 0.00083938 | 633 | -9.6812e-06 | 0.00065085 | 0.00042409 | 0.00035508
368 | 0.00078049 | 632 | -4.1674e-05 | 0.00068328 | 0.00039266 | 0.00034476
369 | 0.00078234 | 631 | -1.4444e-05 | 0.00062604 | 0.00043699 | 0.00034953
370 | 0.00077894 | 630 | 7.9988e-06 | 0.00066381 | 0.00040756 | 0.00035979
371 | 0.00077278 | 629 | -4.2453e-06 | 0.00066949 | 0.0004323 | 0.00036585
372 | 0.00076664 | 628 | -5.3684e-05 | 0.00061611 | 0.00037419 | 0.0003122
373 | 0.00075445 | 627 | -2.9613e-05 | 0.00058625 | 0.00042722 | 0.00032795
374 | 0.00072634 | 626 | -3.277e-05 | 0.00064066 | 0.00038263 | 0.00033017
375 | 0.00068561 | 625 | 2.5731e-05 | 0.00067081 | 0.00045293 | 0.00038316
376 | 0.00074424 | 624 | -3.5255e-05 | 0.00065199 | 0.00043829 | 0.00035167
377 | 0.00067934 | 623 | 1.3357e-05 | 0.00067981 | 0.00038546 | 0.00035954
378 | 0.0006805 | 622 | -7.7657e-06 | 0.00069119 | 0.0004151 | 0.00036617
379 | 0.00070633 | 621 | -1.2943e-06 | 0.00066622 | 0.00044047 | 0.00036846
380 | 0.00065437 | 620 | 5.3067e-06 | 0.0006713 | 0.00047799 | 0.00038487
381 | 0.00067834 | 619 | -1.8068e-05 | 0.00061761 | 0.00038701 | 0.00032885
382 | 0.00066288 | 618 | 2.0813e-05 | 0.00067537 | 0.00043728 | 0.00037782
383 | 0.00060943 | 617 | 1.271e-05 | 0.00064201 | 0.00041642 | 0.00035705
384 | 0.00063289 | 616 | 4.2971e-06 | 0.00069803 | 0.00047602 | 0.00039278
385 | 0.0006262 | 615 | 4.3439e-05 | 0.00064691 | 0.00044158 | 0.00037731
386 | 0.00060363 | 614 | 3.3757e-05 | 0.00068066 | 0.00038673 | 0.00036705
387 | 0.00059613 | 613 | -2.6299e-05 | 0.00062207 | 0.00042938 | 0.00034172
388 | 0.0006371 | 612 | 1.8587e-05 | 0.00065201 | 0.00045762 | 0.00037607
389 | 0.00061753 | 611 | 2.9977e-05 | 0.00069977 | 0.00047084 | 0.00040019
390 | 0.00055253 | 610 | 5.2603e-05 | 0.00067211 | 0.00044308 | 0.00038926
391 | 0.00058807 | 609 | 5.7031e-05 | 0.0006755 | 0.00046311 | 0.00039855
392 | 0.00055449 | 608 | 4.7322e-05 | 0.0007424 | 0.00055462 | 0.00044811
393 | 0.00056846 | 607 | 9.4831e-05 | 0.00076153 | 0.00052941 | 0.00046192
394 | 0.00052676 | 606 | 6.1302e-05 | 0.00064076 | 0.00046482 | 0.00038896
395 | 0.00057925 | 605 | 9.5064e-05 | 0.00068445 | 0.00047294 | 0.00041748
396 | 0.00053536 | 604 | 7.7587e-05 | 0.00065805 | 0.00044694 | 0.00039419
397 | 0.00050619 | 603 | 0.00010503 | 0.00065831 | 0.00047122 | 0.00041152
398 | 0.00055289 | 602 | 0.00010661 | 0.00073379 | 0.00053843 | 0.00045961
399 | 0.00057306 | 601 | 5.3846e-05 | 0.00063895 | 0.00046379 | 0.00038553
400 | 0.0005442 | 600 | 7.1529e-05 | 0.00064175 | 0.00051354 | 0.00040894
401 | 0.00052124 | 599 | 0.00013108 | 0.00065533 | 0.0004555 | 0.00041397
402 | 0.00061468 | 598 | 9.3122e-05 | 0.00067931 | 0.00050123 | 0.00042456
403 | 0.00050916 | 597 | 0.00017303 | 0.00063444 | 0.00051031 | 0.00043926
404 | 0.00055373 | 596 | 9.4132e-05 | 0.00060644 | 0.00049631 | 0.00039896
405 | 0.00058331 | 595 | 0.00011438 | 0.00064268 | 0.00044337 | 0.00040014
406 | 0.000551 | 594 | 8.4396e-05 | 0.00063983 | 0.0004789 | 0.00040104
407 | 0.00052925 | 593 | 5.6461e-05 | 0.00062446 | 0.00046951 | 0.00038348
408 | 0.00052116 | 592 | 0.00018766 | 0.00068994 | 0.00055823 | 0.00047861
409 | 0.00055817 | 591 | 0.00015946 | 0.00062658 | 0.00045775 | 0.0004146
410 | 0.00063309 | 590 | 0.00015107 | 0.00063558 | 0.00043984 | 0.00040883
411 | 0.00051781 | 589 | 0.00014584 | 0.00062534 | 0.00050821 | 0.00042646
412 | 0.00051501 | 588 | 0.00014335 | 0.00065712 | 0.00054027 | 0.00044691
413 | 0.00055124 | 587 | 0.0001035 | 0.00062192 | 0.00042953 | 0.00038498
414 | 0.00057503 | 586 | 0.00012777 | 0.00064473 | 0.00044663 | 0.00040638
415 | 0.00051668 | 585 | 0.00016731 | 0.00066513 | 0.00044412 | 0.00042552
416 | 0.00048552 | 584 | 0.00013416 | 0.00064299 | 0.00056481 | 0.00044732
417 | 0.00061801 | 583 | 0.00012626 | 0.00070571 | 0.00048496 | 0.00043898
418 | 0.00053989 | 582 | 9.8611e-05 | 0.00061839 | 0.00041691 | 0.00037797
419 | 0.00057448 | 581 | 0.00012443 | 0.00069583 | 0.00042513 | 0.00041513
420 | 0.0005144 | 580 | 0.00014708 | 0.00063013 | 0.0004399 | 0.0004057
421 | 0.00055901 | 579 | 0.00014089 | 0.00066643 | 0.00050015 | 0.00043582
422 | 0.00056563 | 578 | 0.00018171 | 0.00066713 | 0.00046607 | 0.0004383
423 | 0.00053231 | 577 | 0.0001491 | 0.00064895 | 0.00048618 | 0.00042808
424 | 0.00049147 | 576 | 0.00016029 | 0.00073558 | 0.00048089 | 0.00045892
425 | 0.00054632 | 575 | 0.00015677 | 0.00066002 | 0.00051409 | 0.00044363
426 | 0.00047028 | 574 | 0.00010428 | 0.00066059 | 0.00043824 | 0.00040104
427 | 0.00053645 | 573 | 0.00016676 | 0.00070765 | 0.00050258 | 0.000459
428 | 0.00051524 | 572 | 0.00019893 | 0.00069033 | 0.00044298 | 0.00044408
429 | 0.00047051 | 571 | 0.00016974 | 0.00073153 | 0.00049836 | 0.00046654
430 | 0.00055252 | 570 | 0.00017192 | 0.00073682 | 0.00048737 | 0.00046537
431 | 0.00048916 | 569 | 0.0001795 | 0.00075054 | 0.00053347 | 0.00048784
432 | 0.00047466 | 568 | 0.00025484 | 0.00078202 | 0.00054058 | 0.00052581
433 | 0.00049778 | 567 | 0.00022418 | 0.00074992 | 0.0005334 | 0.0005025
434 | 0.00051411 | 566 | 0.00025212 | 0.00074299 | 0.0004947 | 0.00049661
435 | 0.00048392 | 565 | 0.00026197 | 0.00071162 | 0.00047542 | 0.000483
436 | 0.0004898 | 564 | 0.00022369 | 0.00073174 | 0.00050175 | 0.00048573
437 | 0.0004826 | 563 | 0.00021667 | 0.00078643 | 0.00053565 | 0.00051292
438 | 0.000501 | 562 | 0.00018176 | 0.00072702 | 0.00051647 | 0.00047508
439 | 0.00051104 | 561 | 0.0002707 | 0.00079263 | 0.00051987 | 0.00052773
440 | 0.00049688 | 560 | 0.00022514 | 0.00072575 | 0.0005343 | 0.00049506
441 | 0.00053501 | 559 | 0.00026028 | 0.00079174 | 0.00055636 | 0.00053613
442 | 0.00051632 | 558 | 0.00024552 | 0.00075399 | 0.00054729 | 0.0005156
443 | 0.00049213 | 557 | 0.00022806 | 0.00075824 | 0.00050764 | 0.00049798
444 | 0.00051183 | 556 | 0.00022265 | 0.00075707 | 0.00050764 | 0.00049579
445 | 0.00053958 | 555 | 0.00024526 | 0.00078817 | 0.00052282 | 0.00051875
446 | 0.00050487 | 554 | 0.00023099 | 0.0007925 | 0.00058731 | 0.00053693
447 | 0.0005335 | 553 | 0.00024327 | 0.00076356 | 0.00052917 | 0.000512
448 | 0.00051222 | 552 | 0.00026124 | 0.00073654 | 0.00046109 | 0.00048629
449 | 0.00051558 | 551 | 0.00026554 | 0.00077958 | 0.00060056 | 0.00054856
450 | 0.00054095 | 550 | 0.00027085 | 0.00074852 | 0.00058902 | 0.00053613
451 | 0.00052587 | 549 | 0.00022625 | 0.00071514 | 0.00054864 | 0.00049668
452 | 0.00049508 | 548 | 0.00025394 | 0.0008029 | 0.00055691 | 0.00053791
453 | 0.00054349 | 547 | 0.00026044 | 0.00072829 | 0.00054454 | 0.00051109
454 | 0.0005233 | 546 | 0.00028238 | 0.00074035 | 0.00055709 | 0.0005266
455 | 0.00051986 | 545 | 0.00032053 | 0.00080469 | 0.00059532 | 0.00057351
456 | 0.00049841 | 544 | 0.00021302 | 0.0007299 | 0.0005454 | 0.0004961
457 | 0.00052277 | 543 | 0.00023203 | 0.00073236 | 0.00054677 | 0.00050372
458 | 0.00057907 | 542 | 0.00026694 | 0.00078651 | 0.00058982 | 0.00054776
459 | 0.00053427 | 541 | 0.00025257 | 0.00071706 | 0.00054128 | 0.00050363
460 | 0.0005692 | 540 | 0.00026528 | 0.00076125 | 0.00055284 | 0.00052646
461 | 0.00056779 | 539 | 0.00028176 | 0.00076794 | 0.00049338 | 0.00051436
462 | 0.00057544 | 538 | 0.00034375 | 0.00082543 | 0.00059973 | 0.00058964
463 | 0.00057663 | 537 | 0.00031372 | 0.00081929 | 0.00053477 | 0.00055593
464 | 0.00057231 | 536 | 0.00033831 | 0.00077893 | 0.00057266 | 0.0005633
465 | 0.00051656 | 535 | 0.00025852 | 0.00069547 | 0.00052505 | 0.00049301
466 | 0.00056058 | 534 | 0.00026502 | 0.00072303 | 0.0005257 | 0.00050458
467 | 0.00060351 | 533 | 0.00026925 | 0.00072816 | 0.00055589 | 0.00051777
468 | 0.0005507 | 532 | 0.000296 | 0.00076506 | 0.00054501 | 0.00053536
469 | 0.00065351 | 531 | 0.00031333 | 0.00080782 | 0.00055748 | 0.00055954
470 | 0.0005723 | 530 | 0.00028238 | 0.00078391 | 0.00058187 | 0.00054939
471 | 0.00060292 | 529 | 0.0002787 | 0.00077113 | 0.00055854 | 0.00053612
472 | 0.00053975 | 528 | 0.00028476 | 0.00073205 | 0.00051538 | 0.00051073
473 | 0.00062903 | 527 | 0.00030082 | 0.00074447 | 0.00056802 | 0.00053777
474 | 0.00060257 | 526 | 0.00028515 | 0.00073288 | 0.00055252 | 0.00052352
475 | 0.00062243 | 525 | 0.00029551 | 0.00075129 | 0.00055716 | 0.00053466
476 | 0.00061843 | 524 | 0.00029634 | 0.00076667 | 0.00056059 | 0.0005412
477 | 0.00063669 | 523 | 0.0003711 | 0.00071084 | 0.00057743 | 0.00055313
478 | 0.00061959 | 522 | 0.00025065 | 0.00077061 | 0.00051837 | 0.00051321
479 | 0.00060556 | 521 | 0.0003689 | 0.00078835 | 0.00058923 | 0.00058216
480 | 0.00060316 | 520 | 0.0002752 | 0.0007384 | 0.00055799 | 0.00052387
481 | 0.00061462 | 519 | 0.00030712 | 0.00070446 | 0.00055307 | 0.00052155
482 | 0.0005614 | 518 | 0.00034224 | 0.00080186 | 0.00057489 | 0.000573
483 | 0.00062181 | 517 | 0.00031502 | 0.00072824 | 0.00049916 | 0.00051414
484 | 0.0005985 | 516 | 0.00032129 | 0.00075261 | 0.00054566 | 0.00053985
485 | 0.00061572 | 515 | 0.0002774 | 0.00071574 | 0.0005143 | 0.00050248
486 | 0.00057133 | 514 | 0.00030815 | 0.00077823 | 0.0005926 | 0.00055966
487 | 0.00059566 | 513 | 0.00031758 | 0.00078477 | 0.00060649 | 0.00056961
488 | 0.00056578 | 512 | 0.00034592 | 0.00074489 | 0.00056795 | 0.00055292
489 | 0.00058939 | 511 | 0.00031155 | 0.00075816 | 0.00057051 | 0.00054674
490 | 0.00060449 | 510 | 0.00028328 | 0.00070433 | 0.00054794 | 0.00051185
491 | 0.00062304 | 509 | 0.00032504 | 0.00074546 | 0.00055973 | 0.00054341
492 | 0.00059458 | 508 | 0.00031111 | 0.0007968 | 0.00054376 | 0.00055056
493 | 0.0006432 | 507 | 0.00033328 | 0.00070553 | 0.00059781 | 0.00054554
494 | 0.00058325 | 506 | 0.00034219 | 0.0008037 | 0.00066466 | 0.00060352
495 | 0.00056948 | 505 | 0.00031724 | 0.00086963 | 0.00063107 | 0.00060598
496 | 0.00061528 | 504 | 0.00036901 | 0.00082844 | 0.0006671 | 0.00062152
497 | 0.00060661 | 503 | 0.00030554 | 0.00081835 | 0.00068434 | 0.00060274
498 | 0.00062095 | 502 | 0.00032051 | 0.0008183 | 0.00067574 | 0.00060485
499 | 0.00058423 | 501 | 0.00034201 | 0.00082154 | 0.00067724 | 0.0006136
500 | 0.00059495 | 500 | 0.0003416 | 0.00079037 | 0.00065289 | 0.00059495
501 | 0.0006136 | 499 | 0.00031421 | 0.00077564 | 0.00066282 | 0.00058423
502 | 0.00060485 | 498 | 0.00033414 | 0.00082219 | 0.00070651 | 0.00062095
503 | 0.00060274 | 497 | 0.00033797 | 0.00082102 | 0.00066083 | 0.00060661
504 | 0.00062152 | 496 | 0.00034333 | 0.00081584 | 0.00068668 | 0.00061528
505 | 0.00060598 | 495 | 0.0002937 | 0.00077045 | 0.00064429 | 0.00056948
506 | 0.00060352 | 494 | 0.00032727 | 0.00078819 | 0.00063428 | 0.00058325
507 | 0.00054554 | 493 | 0.00039865 | 0.00081576 | 0.0007152 | 0.0006432
508 | 0.00055056 | 492 | 0.00032846 | 0.00078715 | 0.00066811 | 0.00059458
509 | 0.00054341 | 491 | 0.00036846 | 0.00077665 | 0.00072401 | 0.00062304
510 | 0.00051185 | 490 | 0.00035548 | 0.00082403 | 0.00063394 | 0.00060449
511 | 0.00054674 | 489 | 0.00034095 | 0.00078806 | 0.00063915 | 0.00058939
512 | 0.00055292 | 488 | 0.00029738 | 0.00076763 | 0.00063234 | 0.00056578
513 | 0.00056961 | 487 | 0.00035789 | 0.00078033 | 0.00064875 | 0.00059566
514 | 0.00055966 | 486 | 0.00030629 | 0.00077001 | 0.0006377 | 0.00057133
515 | 0.00050248 | 485 | 0.00038243 | 0.00078529 | 0.00067944 | 0.00061572
516 | 0.00053985 | 484 | 0.00029466 | 0.00082691 | 0.00067392 | 0.0005985
517 | 0.00051414 | 483 | 0.00036908 | 0.00081132 | 0.00068502 | 0.00062181
518 | 0.000573 | 482 | 0.00029663 | 0.00075925 | 0.00062832 | 0.0005614
519 | 0.00052155 | 481 | 0.00036502 | 0.00083015 | 0.00064869 | 0.00061462
520 | 0.00052387 | 480 | 0.00034924 | 0.00078433 | 0.00067592 | 0.00060316
521 | 0.00058216 | 479 | 0.00034388 | 0.0007988 | 0.000674 | 0.00060556
522 | 0.00051321 | 478 | 0.00034147 | 0.0008599 | 0.00065741 | 0.00061959
523 | 0.00055313 | 477 | 0.00036818 | 0.0008447 | 0.00069718 | 0.00063669
524 | 0.0005412 | 476 | 0.00033678 | 0.00085409 | 0.00066443 | 0.00061843
525 | 0.00053466 | 475 | 0.00033341 | 0.00084543 | 0.00068844 | 0.00062243
526 | 0.00052352 | 474 | 0.00035665 | 0.00080461 | 0.00064644 | 0.00060257
527 | 0.00053777 | 473 | 0.0003737 | 0.00082995 | 0.00068344 | 0.00062903
528 | 0.00051073 | 472 | 0.00026425 | 0.00076597 | 0.00058905 | 0.00053975
529 | 0.00053612 | 471 | 0.00033579 | 0.00084725 | 0.00062573 | 0.00060292
530 | 0.00054939 | 470 | 0.00028344 | 0.00079683 | 0.00063664 | 0.0005723
531 | 0.00055954 | 469 | 0.0003959 | 0.00086032 | 0.00070431 | 0.00065351
532 | 0.00053536 | 468 | 0.00029691 | 0.00078503 | 0.00057018 | 0.0005507
533 | 0.00051777 | 467 | 0.00036937 | 0.00083796 | 0.0006032 | 0.00060351
534 | 0.00050458 | 466 | 0.00032186 | 0.00076612 | 0.00059376 | 0.00056058
535 | 0.00049301 | 465 | 0.00031072 | 0.00067135 | 0.00056761 | 0.00051656
536 | 0.0005633 | 464 | 0.00037201 | 0.00078098 | 0.00056393 | 0.00057231
537 | 0.00055593 | 463 | 0.00035882 | 0.00079851 | 0.00057256 | 0.00057663
538 | 0.00058964 | 462 | 0.00037289 | 0.00078313 | 0.00057028 | 0.00057544
539 | 0.00051436 | 461 | 0.00033263 | 0.00075915 | 0.0006116 | 0.00056779
540 | 0.00052646 | 460 | 0.00038934 | 0.00076006 | 0.0005582 | 0.0005692
541 | 0.00050363 | 459 | 0.00030823 | 0.00075891 | 0.00053568 | 0.00053427
542 | 0.00054776 | 458 | 0.0003653 | 0.00078993 | 0.00058197 | 0.00057907
543 | 0.00050372 | 457 | 0.00033105 | 0.00068976 | 0.0005475 | 0.00052277
544 | 0.0004961 | 456 | 0.00029069 | 0.00068598 | 0.00051855 | 0.00049841
545 | 0.00057351 | 455 | 0.00030595 | 0.00075951 | 0.00049413 | 0.00051986
546 | 0.0005266 | 454 | 0.00030888 | 0.0007257 | 0.00053532 | 0.0005233
547 | 0.00051109 | 453 | 0.00032618 | 0.00073304 | 0.00057124 | 0.00054349
548 | 0.00053791 | 452 | 0.0002752 | 0.00069373 | 0.00051632 | 0.00049508
549 | 0.00049668 | 451 | 0.00029497 | 0.00072476 | 0.00055786 | 0.00052587
550 | 0.00053613 | 450 | 0.00033976 | 0.00076542 | 0.00051767 | 0.00054095
551 | 0.00054856 | 449 | 0.00029036 | 0.00070112 | 0.00055527 | 0.00051558
552 | 0.00048629 | 448 | 0.00027673 | 0.0007342 | 0.00052573 | 0.00051222
553 | 0.000512 | 447 | 0.00032719 | 0.00077222 | 0.0005011 | 0.0005335
554 | 0.00053693 | 446 | 0.0002838 | 0.0007168 | 0.00051401 | 0.00050487
555 | 0.00051875 | 445 | 0.00029396 | 0.00077385 | 0.00055094 | 0.00053958
556 | 0.00049579 | 444 | 0.00028531 | 0.00070459 | 0.0005456 | 0.00051183
557 | 0.00049798 | 443 | 0.00027251 | 0.00072476 | 0.00047913 | 0.00049213
558 | 0.0005156 | 442 | 0.0002479 | 0.00074367 | 0.0005574 | 0.00051632
559 | 0.00053613 | 441 | 0.00031818 | 0.00073485 | 0.00055201 | 0.00053501
560 | 0.00049506 | 440 | 0.00026805 | 0.00070241 | 0.00052018 | 0.00049688
561 | 0.00052773 | 439 | 0.00026407 | 0.00072093 | 0.00054812 | 0.00051104
562 | 0.00047508 | 438 | 0.00028261 | 0.00072788 | 0.00049253 | 0.000501
563 | 0.00051292 | 437 | 0.00024311 | 0.00073381 | 0.00047086 | 0.0004826
564 | 0.00048573 | 436 | 0.00024943 | 0.00071673 | 0.00050323 | 0.0004898
565 | 0.000483 | 435 | 0.00024933 | 0.00070324 | 0.00049919 | 0.00048392
566 | 0.00049661 | 434 | 0.00030699 | 0.00072583 | 0.0005095 | 0.00051411
567 | 0.0005025 | 433 | 0.00030256 | 0.00073495 | 0.00045583 | 0.00049778
568 | 0.00052581 | 432 | 0.00022174 | 0.00068201 | 0.00052023 | 0.00047466
569 | 0.00048784 | 431 | 0.00034338 | 0.00071748 | 0.00040663 | 0.00048916
570 | 0.00046537 | 430 | 0.00030872 | 0.00080492 | 0.00054392 | 0.00055252
571 | 0.00046654 | 429 | 0.00026642 | 0.00074061 | 0.0004045 | 0.00047051
572 | 0.00044408 | 428 | 0.00031421 | 0.00073568 | 0.00049582 | 0.00051524
573 | 0.000459 | 427 | 0.00032421 | 0.00080105 | 0.00048408 | 0.00053645
574 | 0.00040104 | 426 | 0.00026479 | 0.00074574 | 0.0004003 | 0.00047028
575 | 0.00044363 | 425 | 0.0003146 | 0.00080393 | 0.00052044 | 0.00054632
576 | 0.00045892 | 424 | 0.00027057 | 0.00073462 | 0.00046923 | 0.00049147
577 | 0.00042808 | 423 | 0.00032623 | 0.00078202 | 0.00048867 | 0.00053231
578 | 0.0004383 | 422 | 0.00036351 | 0.00081247 | 0.0005209 | 0.00056563
579 | 0.00043582 | 421 | 0.00030017 | 0.00085842 | 0.00051844 | 0.00055901
580 | 0.0004057 | 420 | 0.00023633 | 0.00081112 | 0.00049577 | 0.0005144
581 | 0.00041513 | 419 | 0.00035761 | 0.00082748 | 0.00053835 | 0.00057448
582 | 0.00037797 | 418 | 0.00033965 | 0.00078923 | 0.00049079 | 0.00053989
583 | 0.00043898 | 417 | 0.00040725 | 0.00088314 | 0.00056364 | 0.00061801
584 | 0.00044732 | 416 | 0.00025111 | 0.00074717 | 0.00045827 | 0.00048552
585 | 0.00042552 | 415 | 0.00028554 | 0.00075894 | 0.00050556 | 0.00051668
586 | 0.00040638 | 414 | 0.00039035 | 0.0008066 | 0.00052814 | 0.00057503
587 | 0.00038498 | 413 | 0.00031958 | 0.00079991 | 0.00053423 | 0.00055124
588 | 0.00044691 | 412 | 0.00030712 | 0.00078347 | 0.00045443 | 0.00051501
589 | 0.00042646 | 411 | 0.00029183 | 0.00074647 | 0.00051513 | 0.00051781
590 | 0.00040883 | 410 | 0.00037002 | 0.00089012 | 0.00063913 | 0.00063309
591 | 0.0004146 | 409 | 0.0003438 | 0.00081153 | 0.00051917 | 0.00055817
592 | 0.00047861 | 408 | 0.0002701 | 0.00081563 | 0.00047776 | 0.00052116
593 | 0.00038348 | 407 | 0.00031002 | 0.00076529 | 0.00051243 | 0.00052925
594 | 0.00040104 | 406 | 0.00035426 | 0.00078337 | 0.00051536 | 0.000551
595 | 0.00040014 | 405 | 0.00037678 | 0.00084719 | 0.00052596 | 0.00058331
596 | 0.00039896 | 404 | 0.00030774 | 0.00085848 | 0.00049496 | 0.00055373
597 | 0.00043926 | 403 | 0.00032432 | 0.00075077 | 0.00045239 | 0.00050916
598 | 0.00042456 | 402 | 0.00043751 | 0.00083544 | 0.00057108 | 0.00061468
599 | 0.00041397 | 401 | 0.00030618 | 0.00077056 | 0.00048698 | 0.00052124
600 | 0.00040894 | 400 | 0.00033926 | 0.00081428 | 0.00047905 | 0.0005442
601 | 0.00038553 | 399 | 0.00039061 | 0.00084369 | 0.00048488 | 0.00057306
602 | 0.00045961 | 398 | 0.00039735 | 0.00080811 | 0.00045321 | 0.00055289
603 | 0.00041152 | 397 | 0.00031087 | 0.00077217 | 0.00043554 | 0.00050619
604 | 0.00039419 | 396 | 0.00036968 | 0.00079732 | 0.00043909 | 0.00053536
605 | 0.00041748 | 395 | 0.00039898 | 0.00086263 | 0.00047615 | 0.00057925
606 | 0.00038896 | 394 | 0.00029919 | 0.00078581 | 0.00049527 | 0.00052676
607 | 0.00046192 | 393 | 0.00041147 | 0.00077154 | 0.00052236 | 0.00056846
608 | 0.00044811 | 392 | 0.00040406 | 0.00079958 | 0.00045982 | 0.00055449
609 | 0.00039855 | 391 | 0.00042139 | 0.00083731 | 0.00050551 | 0.00058807
610 | 0.00038926 | 390 | 0.00037963 | 0.00084712 | 0.00043085 | 0.00055253
611 | 0.00040019 | 389 | 0.00043261 | 0.00088454 | 0.00053542 | 0.00061753
612 | 0.00037607 | 388 | 0.00044762 | 0.00091666 | 0.00054703 | 0.0006371
613 | 0.00034172 | 387 | 0.00042409 | 0.00089417 | 0.00047014 | 0.00059613
614 | 0.00036705 | 386 | 0.00041518 | 0.00087985 | 0.00051588 | 0.00060363
615 | 0.00037731 | 385 | 0.0004307 | 0.00089303 | 0.00055488 | 0.0006262
616 | 0.00039278 | 384 | 0.00048131 | 0.00089961 | 0.00051774 | 0.00063289
617 | 0.00035705 | 383 | 0.00044858 | 0.0008827 | 0.00049701 | 0.00060943
618 | 0.00037782 | 382 | 0.00046589 | 0.00096797 | 0.00055478 | 0.00066288
619 | 0.00032885 | 381 | 0.00046389 | 0.00096434 | 0.00060678 | 0.00067834
620 | 0.00038487 | 380 | 0.00052129 | 0.00088989 | 0.00055193 | 0.00065437
621 | 0.00036846 | 379 | 0.00051992 | 0.0010102 | 0.00058889 | 0.00070633
622 | 0.00036617 | 378 | 0.00048053 | 0.00096909 | 0.00059187 | 0.0006805
623 | 0.00035954 | 377 | 0.00053327 | 0.00097085 | 0.00053389 | 0.00067934
624 | 0.00035167 | 376 | 0.00059791 | 0.00099197 | 0.00064284 | 0.00074424
625 | 0.00038316 | 375 | 0.00053591 | 0.00094932 | 0.0005716 | 0.00068561
626 | 0.00033017 | 374 | 0.00054122 | 0.0010334 | 0.00060442 | 0.00072634
627 | 0.00032795 | 373 | 0.0005547 | 0.0010629 | 0.00064574 | 0.00075445
628 | 0.0003122 | 372 | 0.00059498 | 0.0010449 | 0.00066002 | 0.00076664
629 | 0.00036585 | 371 | 0.00061463 | 0.001039 | 0.00066466 | 0.00077278
630 | 0.00035979 | 370 | 0.00058321 | 0.0010527 | 0.00070086 | 0.00077894
631 | 0.00034953 | 369 | 0.00061416 | 0.0010598 | 0.00067309 | 0.00078234
632 | 0.00034476 | 368 | 0.00062324 | 0.0010752 | 0.00064307 | 0.00078049
633 | 0.00035508 | 367 | 0.00065702 | 0.0011122 | 0.00074891 | 0.00083938
634 | 0.00040331 | 366 | 0.00065728 | 0.001158 | 0.00079276 | 0.00086933
635 | 0.00035334 | 365 | 0.00065061 | 0.0010849 | 0.00068556 | 0.00080704
636 | 0.00035325 | 364 | 0.00069075 | 0.0011356 | 0.00068499 | 0.00083712
637 | 0.0003916 | 363 | 0.00073941 | 0.0011863 | 0.00079325 | 0.00090633
638 | 0.00045554 | 362 | 0.00069233 | 0.0011586 | 0.00074351 | 0.0008648
639 | 0.00036585 | 361 | 0.00064846 | 0.001173 | 0.00077912 | 0.00086687
640 | 0.00034143 | 360 | 0.00072526 | 0.0012255 | 0.00078101 | 0.00091058
641 | 0.00032958 | 359 | 0.00057878 | 0.0011484 | 0.00069329 | 0.00080683
642 | 0.0003847 | 358 | 0.00078625 | 0.0012706 | 0.0008549 | 0.00097059
643 | 0.0003697 | 357 | 0.00066547 | 0.0012413 | 0.00075606 | 0.00088763
644 | 0.00036136 | 356 | 0.0007717 | 0.0012007 | 0.0008154 | 0.00092927
645 | 0.00035963 | 355 | 0.00080316 | 0.0013127 | 0.00088688 | 0.0010009
646 | 0.00037341 | 354 | 0.00076952 | 0.0012462 | 0.00085456 | 0.00095677
647 | 0.00039023 | 353 | 0.00085523 | 0.0013634 | 0.00091884 | 0.0010458
648 | 0.00034433 | 352 | 0.0008045 | 0.0012777 | 0.00089769 | 0.00099329
649 | 0.00036376 | 351 | 0.00082349 | 0.0013118 | 0.00093546 | 0.0010236
650 | 0.00040455 | 350 | 0.00088849 | 0.00131 | 0.00093546 | 0.0010446
651 | 0.00035898 | 349 | 0.00076662 | 0.0011577 | 0.00059999 | 0.00084142
652 | 0.00040518 | 348 | 0.00095095 | 0.0013248 | 0.0010908 | 0.0011222
653 | 0.00035482 | 347 | 0.00099897 | 0.0010701 | 0.00065404 | 0.0009077
654 | 0.00034661 | 346 | 0.00077709 | 0.0014283 | 0.00080603 | 0.0010038
655 | 0.00040576 | 345 | 0.0011338 | 0.0010061 | 0.00085453 | 0.00099814
656 | 0.00038861 | 344 | 0.00084097 | 0.001309 | 0.00067086 | 0.00094027
657 | 0.00039999 | 343 | 0.000999 | 0.0013408 | 0.00097453 | 0.0011048
658 | 0.00037034 | 342 | 0.0010644 | 0.0014331 | 0.00078412 | 0.0010939
659 | 0.00034106 | 341 | 0.0010147 | 0.00143 | 0.001042 | 0.0011623
660 | 0.00036223 | 340 | 0.0009468 | 0.0012062 | 0.00075445 | 0.00096916
661 | 0.00039213 | 339 | 0.0012536 | 0.0014017 | 0.0010765 | 0.0012439
662 | 0.00040308 | 338 | 0.0010217 | 0.0014487 | 0.00093497 | 0.0011351
663 | 0.00037987 | 337 | 0.0011549 | 0.0013433 | 0.00094307 | 0.0011471
664 | 0.00040215 | 336 | 0.0011579 | 0.0014624 | 0.00096527 | 0.0011952
665 | 0.00037596 | 335 | 0.0012515 | 0.0014002 | 0.0011825 | 0.0012781
666 | 0.00039496 | 334 | 0.0011836 | 0.0014821 | 0.0010828 | 0.0012495
667 | 0.00035585 | 333 | 0.0010954 | 0.0013802 | 0.0010624 | 0.0011794
668 | 0.00036221 | 332 | 0.0013225 | 0.001605 | 0.0012163 | 0.0013813
669 | 0.00037838 | 331 | 0.0012706 | 0.0015628 | 0.0012244 | 0.0013526
670 | 0.00032161 | 330 | 0.0014008 | 0.0015716 | 0.0012232 | 0.0013985
671 | 0.00037769 | 329 | 0.001218 | 0.0016424 | 0.001174 | 0.0013448
672 | 0.00033054 | 328 | 0.0012687 | 0.0015532 | 0.0010875 | 0.0013031
673 | 0.00039045 | 327 | 0.0013493 | 0.0015923 | 0.0011368 | 0.0013595
674 | 0.00037002 | 326 | 0.0014032 | 0.0016232 | 0.0013035 | 0.0014433
675 | 0.00038028 | 325 | 0.0015125 | 0.0017102 | 0.0013592 | 0.0015273
676 | 0.00039917 | 324 | 0.0013565 | 0.0016853 | 0.0012221 | 0.0014213
677 | 0.00039636 | 323 | 0.001418 | 0.0015936 | 0.0013439 | 0.0014518
678 | 0.00038235 | 322 | 0.0014821 | 0.0017151 | 0.0013763 | 0.0015245
679 | 0.00038901 | 321 | 0.0014856 | 0.0017483 | 0.0013928 | 0.0015423
680 | 0.00036781 | 320 | 0.0014817 | 0.0016551 | 0.0013209 | 0.0014859
681 | 0.00037131 | 319 | 0.001482 | 0.0017983 | 0.0014469 | 0.0015757
682 | 0.00035254 | 318 | 0.0015629 | 0.0017248 | 0.001461 | 0.0015829
683 | 0.00040364 | 317 | 0.0015973 | 0.0017139 | 0.0013593 | 0.0015568
684 | 0.00041227 | 316 | 0.0014786 | 0.0017129 | 0.0013543 | 0.0015152
685 | 0.00042383 | 315 | 0.0016079 | 0.0017606 | 0.0014408 | 0.0016031
686 | 0.00042601 | 314 | 0.0015653 | 0.0018268 | 0.0015035 | 0.0016318
687 | 0.0003718 | 313 | 0.0017019 | 0.0017878 | 0.0014928 | 0.0016608
688 | 0.00036643 | 312 | 0.0016483 | 0.0018068 | 0.0015794 | 0.0016782
689 | 0.00036277 | 311 | 0.0017718 | 0.001965 | 0.0016653 | 0.0018007
690 | 0.00037484 | 310 | 0.0017586 | 0.0019888 | 0.0016598 | 0.0018024
691 | 0.00039414 | 309 | 0.0017602 | 0.0019741 | 0.0016766 | 0.0018036
692 | 0.00043594 | 308 | 0.0016578 | 0.0018653 | 0.0016175 | 0.0017135
693 | 0.00037159 | 307 | 0.0017588 | 0.0020929 | 0.0016986 | 0.0018501
694 | 0.00041542 | 306 | 0.0019048 | 0.0021948 | 0.0017383 | 0.001946
695 | 0.00043117 | 305 | 0.0019075 | 0.0021699 | 0.0017956 | 0.0019577
696 | 0.00039755 | 304 | 0.0018059 | 0.0021804 | 0.0017791 | 0.0019218
697 | 0.00037399 | 303 | 0.0019978 | 0.0023386 | 0.0018661 | 0.0020675
698 | 0.00037528 | 302 | 0.0019817 | 0.0023225 | 0.0018921 | 0.0020654
699 | 0.00033427 | 301 | 0.0022768 | 0.0024731 | 0.00219 | 0.0023133
700 | 0.00037691 | 300 | 0.0021606 | 0.0024625 | 0.0021149 | 0.002246
701 | 0.00035659 | 299 | 0.0022894 | 0.0025861 | 0.0021943 | 0.0023566
702 | 0.00035943 | 298 | 0.0022769 | 0.0026055 | 0.0021543 | 0.0023456
703 | 0.00040301 | 297 | 0.0025343 | 0.0028926 | 0.0024996 | 0.0026422
704 | 0.00037088 | 296 | 0.0026444 | 0.0029349 | 0.0025464 | 0.0027085
705 | 0.00033757 | 295 | 0.002844 | 0.0031446 | 0.0027333 | 0.0029073
706 | 0.00039093 | 294 | 0.0029688 | 0.0033287 | 0.002937 | 0.0030782
707 | 0.00033761 | 293 | 0.0033291 | 0.003564 | 0.0032352 | 0.0033761
708 | 0.00042437 | 292 | 0.0036054 | 0.0039191 | 0.0035789 | 0.0037011
709 | 0.00038025 | 291 | 0.0039252 | 0.0042693 | 0.0038617 | 0.0040187
710 | 0.00033267 | 290 | 0.0043513 | 0.0046441 | 0.0042462 | 0.0044139
711 | 0.00040152 | 289 | 0.0045825 | 0.0049296 | 0.0045404 | 0.0046842
712 | 0.00036555 | 288 | 0.0048775 | 0.0051192 | 0.0047761 | 0.0049243
713 | 0.00039279 | 287 | 0.0051229 | 0.0053834 | 0.0050348 | 0.0051804
714 | 0.00037857 | 286 | 0.0052843 | 0.0055539 | 0.005233 | 0.005357
715 | 0.0004018 | 285 | 0.0055572 | 0.0058187 | 0.0054706 | 0.0056155
716 | 0.00039371 | 284 | 0.0059621 | 0.0061368 | 0.0058855 | 0.0059948
717 | 0.00049702 | 283 | 0.0061796 | 0.0065089 | 0.0061133 | 0.0062673
718 | 0.00039065 | 282 | 0.0064861 | 0.006809 | 0.0064879 | 0.0065943
719 | 0.00039768 | 281 | 0.0067822 | 0.0070061 | 0.0066995 | 0.0068292
720 | 0.00043296 | 280 | 0.0070378 | 0.0072685 | 0.0068608 | 0.0070557
721 | 0.0004531 | 279 | 0.0071392 | 0.0074005 | 0.0070276 | 0.0071891
722 | 0.00043706 | 278 | 0.0072139 | 0.0075114 | 0.0071418 | 0.007289
723 | 0.00042495 | 277 | 0.0074124 | 0.0076469 | 0.0073168 | 0.0074587
724 | 0.00047982 | 276 | 0.0076467 | 0.007922 | 0.0074913 | 0.0076867
725 | 0.00052287 | 275 | 0.0076761 | 0.0080075 | 0.0076136 | 0.0077657
726 | 0.00048038 | 274 | 0.0078221 | 0.0080534 | 0.0076953 | 0.0078569
727 | 0.00044849 | 273 | 0.0078491 | 0.0081657 | 0.0077937 | 0.0079362
728 | 0.00050242 | 272 | 0.0080254 | 0.008311 | 0.0079003 | 0.0080789
729 | 0.00050136 | 271 | 0.0081481 | 0.0084098 | 0.0080762 | 0.0082114
730 | 0.00052982 | 270 | 0.0081538 | 0.0084594 | 0.0080994 | 0.0082375
731 | 0.0005013 | 269 | 0.0082855 | 0.0085294 | 0.008206 | 0.0083403
732 | 0.00047662 | 268 | 0.0082978 | 0.0085641 | 0.0082275 | 0.0083631
733 | 0.00053399 | 267 | 0.0084243 | 0.0086251 | 0.008304 | 0.0084511
734 | 0.00056327 | 266 | 0.0083602 | 0.0086531 | 0.0082637 | 0.0084257
735 | 0.00051843 | 265 | 0.0084618 | 0.0087254 | 0.0083684 | 0.0085185
736 | 0.0006093 | 264 | 0.0084565 | 0.0087235 | 0.008375 | 0.0085184
737 | 0.00063141 | 263 | 0.0084091 | 0.0086045 | 0.0081762 | 0.0083966
738 | 0.00055716 | 262 | 0.0083455 | 0.0086318 | 0.0082104 | 0.0083959
739 | 0.0005436 | 261 | 0.0082852 | 0.0085451 | 0.0081549 | 0.0083284
740 | 0.00064541 | 260 | 0.0082402 | 0.0084713 | 0.008144 | 0.0082852
741 | 0.00063646 | 259 | 0.0082627 | 0.0085095 | 0.0081764 | 0.0083162
742 | 0.00062415 | 258 | 0.0081116 | 0.0083033 | 0.0079405 | 0.0081185
743 | 0.00052805 | 257 | 0.008038 | 0.00826 | 0.007972 | 0.00809
744 | 0.00064338 | 256 | 0.0079149 | 0.0081175 | 0.0077429 | 0.0079251
745 | 0.00059185 | 255 | 0.0077676 | 0.0079283 | 0.007618 | 0.0077713
746 | 0.00062272 | 254 | 0.007637 | 0.0078771 | 0.0075006 | 0.0076716
747 | 0.00064333 | 253 | 0.0074667 | 0.0076576 | 0.0073284 | 0.0074842
748 | 0.00056074 | 252 | 0.0073546 | 0.0075672 | 0.0073012 | 0.0074077
749 | 0.00057975 | 251 | 0.0071421 | 0.0073773 | 0.0069964 | 0.0071719
750 | 0.00055944 | 250 | 0.0070273 | 0.007223 | 0.0068824 | 0.0070442
751 | 0.00061249 | 249 | 0.0068409 | 0.0070956 | 0.0067526 | 0.0068964
752 | 0.0005973 | 248 | 0.006585 | 0.0068106 | 0.0063795 | 0.0065917
753 | 0.00056514 | 247 | 0.0063869 | 0.0066201 | 0.0063265 | 0.0064445
754 | 0.00058899 | 246 | 0.0061757 | 0.0064691 | 0.0060079 | 0.0062176
755 | 0.00056059 | 245 | 0.0060302 | 0.0061833 | 0.0057967 | 0.0060034
756 | 0.00053223 | 244 | 0.0058157 | 0.0059765 | 0.0055851 | 0.0057924
757 | 0.00055606 | 243 | 0.0056264 | 0.0058893 | 0.0055356 | 0.0056838
758 | 0.00058292 | 242 | 0.0055491 | 0.005837 | 0.0053496 | 0.0055786
759 | 0.00052381 | 241 | 0.005707 | 0.0058909 | 0.0054469 | 0.0056816
760 | 0.00050286 | 240 | 0.0056083 | 0.0058306 | 0.0054477 | 0.0056289
761 | 0.00053881 | 239 | 0.005916 | 0.0061283 | 0.0057239 | 0.0059227
762 | 0.0004603 | 238 | 0.0061642 | 0.0064248 | 0.0060431 | 0.0062107
763 | 0.00045973 | 237 | 0.0065958 | 0.0069958 | 0.0063218 | 0.0066378
764 | 0.00051068 | 236 | 0.0074715 | 0.0077693 | 0.0071508 | 0.0074639
765 | 0.00043781 | 235 | 0.0085035 | 0.0087875 | 0.0081529 | 0.0084813
766 | 0.00046073 | 234 | 0.0097113 | 0.010066 | 0.0092334 | 0.0096702
767 | 0.000467 | 233 | 0.011315 | 0.011663 | 0.010998 | 0.011326
768 | 0.00039036 | 232 | 0.013338 | 0.013661 | 0.012769 | 0.013256
769 | 0.00041018 | 231 | 0.015845 | 0.016206 | 0.015467 | 0.015839
770 | 0.00043673 | 230 | 0.018826 | 0.019221 | 0.018432 | 0.018826
771 | 0.00034724 | 229 | 0.02222 | 0.02266 | 0.021774 | 0.022218
772 | 0.00038695 | 228 | 0.026324 | 0.026777 | 0.025889 | 0.02633
773 | 0.00042533 | 227 | 0.030583 | 0.031042 | 0.030021 | 0.030548
774 | 0.0004124 | 226 | 0.035303 | 0.035769 | 0.034899 | 0.035324
775 | 0.00035741 | 225 | 0.040334 | 0.040856 | 0.039859 | 0.040349
776 | 0.00037147 | 224 | 0.045282 | 0.045827 | 0.044735 | 0.045281
777 | 0.00038639 | 223 | 0.050662 | 0.051211 | 0.050056 | 0.050643
778 | 0.00027968 | 222 | 0.05604 | 0.056702 | 0.05545 | 0.056064
779 | 0.00038175 | 221 | 0.062042 | 0.062836 | 0.06163 | 0.062169
780 | 0.00035804 | 220 | 0.068521 | 0.069452 | 0.067945 | 0.06864
781 | 0.0003322 | 219 | 0.075654 | 0.076593 | 0.075051 | 0.075766
782 | 0.0003186 | 218 | 0.083053 | 0.084086 | 0.082483 | 0.083208
783 | 0.000312 | 217 | 0.090864 | 0.092085 | 0.090413 | 0.091121
784 | 0.00037381 | 216 | 0.099575 | 0.10091 | 0.098965 | 0.099818
785 | 0.00040308 | 215 | 0.10904 | 0.11045 | 0.10845 | 0.10931
786 | 0.00038971 | 214 | 0.11989 | 0.12162 | 0.11947 | 0.12033
787 | 0.00041398 | 213 | 0.13294 | 0.13495 | 0.13289 | 0.13359
788 | 0.0003054 | 212 | 0.14905 | 0.1514 | 0.14934 | 0.14993
789 | 0.00024059 | 211 | 0.1691 | 0.17206 | 0.16985 | 0.17034
790 | 0.00029402 | 210 | 0.19515 | 0.19853 | 0.19661 | 0.19676
791 | 0.0003246 | 209 | 0.23014 | 0.23463 | 0.23314 | 0.23264
792 | 0.00040517 | 208 | 0.27872 | 0.2846 | 0.28348 | 0.28227
793 | 0.00033414 | 207 | 0.34727 | 0.35514 | 0.35446 | 0.35229
794 | 0.00035771 | 206 | 0.44288 | 0.45334 | 0.45379 | 0.45001
795 | 0.00030303 | 205 | 0.5796 | 0.59404 | 0.59668 | 0.59011
796 | 0.00031186 | 204 | 0.77611 | 0.79581 | 0.80094 | 0.79096
797 | 0.00030145 | 203 | 1.0485 | 1.0749 | 1.0825 | 1.0686
798 | 0.00028746 | 202 | 1.4192 | 1.4528 | 1.4649 | 1.4456
799 | 0.00028702 | 201 | 1.9192 | 1.9632 | 1.9811 | 1.9545
800 | 0.00029749 | 200 | 2.5867 | 2.6439 | 2.6683 | 2.6329

## Weighting
IR Radiometer Reading |  | 0.19 | mW/cm2 | from fluence calculations sheet
d sample depth |  | 0.44057 | cm | from fluence calculations sheet
fp radiometer lamp correction factor |  | 1.0324 | unitless | =(sumRLE)/(sumRLE*fR)
Incident Intensity |  | 0.18599 | mW/cm2 | =fp*IR*DF*PF*RF
Average Intensity |  | 0.18529 | mW/cm2 | =sum Iavg
Weighted Average Intensity |  | 0 | mW/cm2 | =sum Iavg*Rweight
Rλpeak |  | 0.002796 | unitless | from calibration data below
COLUMN SUM | 2.6221 | 1 | 0.25319 | 90.5554 | 2.5397 | 0.18599 | 12.0719 | 0.18529 | 48.637 | 0 | 46.1582 | 0 | 0 |  | 0
λ | RLE | L(λ) | R | fR | RLE*fR | I(λ) | a | Iavg | ADNA | R DNA | AMS2 | R Action, MS2 | R 254 | R etc | Iavg*Rweight
Wavelength (nm) | SpectralOutput, relative lamp energy, Spectra tab | RLE/ Total RLE | 9/6/2022 Radiometer Calibration Factor (DTM) - Units are (A)(cm2)(W-1) | R / Rλpeak |  | IncidentIntensity*L(λ) | Sample Absorbance (cm-1) | Beer Law using abs over sample depth | DNA abs from 2001 LP/MP sheet | DNAabs / λDNAabs (choose wavelength) | MS2 action from 2001 LP/MP sheet | Action / λAction (choose wavelength) | weighting rel 254 for germicidal |  | select weight and apply formula
200 | 0 | 0 | 0.002097 | 0.75 | 0 | 0 | 2.6329 | 0 | 0.83
201 | 0 | 0 | 0.002137 | 0.76431 | 0 | 0 | 1.9545 | 0 | 0.82
202 | 0 | 0 | 0.002177 | 0.77861 | 0 | 0 | 1.4456 | 0 | 0.81
203 | 0 | 0 | 0.002213 | 0.79149 | 0 | 0 | 1.0686 | 0 | 0.8
204 | 0 | 0 | 0.002249 | 0.80436 | 0 | 0 | 0.79096 | 0 | 0.79
205 | 0 | 0 | 0.0022945 | 0.82064 | 0 | 0 | 0.59011 | 0 | 0.776
206 | 0 | 0 | 0.00234 | 0.83691 | 0 | 0 | 0.45001 | 0 | 0.762
207 | 0 | 0 | 0.0023935 | 0.85604 | 0 | 0 | 0.35229 | 0 | 0.748
208 | 0 | 0 | 0.002447 | 0.87518 | 0 | 0 | 0.28227 | 0 | 0.735
209 | 0 | 0 | 0.0024875 | 0.88966 | 0 | 0 | 0.23264 | 0 | 0.71
210 | 0 | 0 | 0.002528 | 0.90415 | 0 | 0 | 0.19676 | 0 | 0.685
211 | 0 | 0 | 0.0025835 | 0.924 | 0 | 0 | 0.17034 | 0 | 0.66
212 | 0 | 0 | 0.002639 | 0.94385 | 0 | 0 | 0.14993 | 0 | 0.645
213 | 0 | 0 | 0.002683 | 0.95959 | 0 | 0 | 0.13359 | 0 | 0.624
214 | 0 | 0 | 0.002727 | 0.97532 | 0 | 0 | 0.12033 | 0 | 0.603
215 | 0 | 0 | 0.002766 | 0.98927 | 0 | 0 | 0.10931 | 0 | 0.582
216 | 0 | 0 | 0.002805 | 1.0032 | 0 | 0 | 0.099818 | 0 | 0.56
217 | 0 | 0 | 0.002852 | 1.02 | 0 | 0 | 0.091121 | 0 | 0.54
218 | 0 | 0 | 0.002899 | 1.0368 | 0 | 0 | 0.083208 | 0 | 0.52
219 | 0 | 0 | 0.0029355 | 1.0499 | 0 | 0 | 0.075766 | 0 | 0.5
220 | 0 | 0 | 0.002972 | 1.0629 | 0 | 0 | 0.06864 | 0 | 0.48 |  | 1.02
221 | 0 | 0 | 0.002998 | 1.0722 | 0 | 0 | 0.062169 | 0 | 0.453 |  | 1
222 | 0 | 0 | 0.003024 | 1.0815 | 0 | 0 | 0.056064 | 0 | 0.426 |  | 0.99
223 | 0 | 0 | 0.0030445 | 1.0889 | 0 | 0 | 0.050643 | 0 | 0.399 |  | 0.97
224 | 0 | 0 | 0.003065 | 1.0962 | 0 | 0 | 0.045281 | 0 | 0.37 |  | 0.955
225 | 0 | 0 | 0.0030775 | 1.1007 | 0 | 0 | 0.040349 | 0 | 0.35 |  | 0.94
226 | 0 | 0 | 0.00309 | 1.1052 | 0 | 0 | 0.035324 | 0 | 0.337 |  | 0.912
227 | 0 | 0 | 0.003107 | 1.1112 | 0 | 0 | 0.030548 | 0 | 0.324 |  | 0.884
228 | 0 | 0 | 0.003124 | 1.1173 | 0 | 0 | 0.02633 | 0 | 0.31 |  | 0.856
229 | 0 | 0 | 0.003116 | 1.1144 | 0 | 0 | 0.022218 | 0 | 0.305 |  | 0.828
230 | 0.0016841 | 0.00064227 | 0.003108 | 1.1116 | 0.001872 | 0.00011945 | 0.018826 | 0.00011832 | 0.3 |  | 0.80504
231 | 0.0015092 | 0.00057556 | 0.003121 | 1.1162 | 0.0016846 | 0.00010705 | 0.015839 | 0.00010619 | 0.3 |  | 0.76478
232 | 0.001238 | 0.00047215 | 0.003134 | 1.1209 | 0.0013877 | 8.7813e-05 | 0.013256 | 8.7225e-05 | 0.3 |  | 0.72453
233 | 0.0010502 | 0.00040052 | 0.0031385 | 1.1225 | 0.0011788 | 7.4491e-05 | 0.011326 | 7.4065e-05 | 0.305 |  | 0.68427
234 | 0.0013836 | 0.00052768 | 0.003143 | 1.1241 | 0.0015553 | 9.8141e-05 | 0.0096702 | 9.7661e-05 | 0.31 |  | 0.64402
235 | 0.0013284 | 0.00050663 | 0.0031415 | 1.1236 | 0.0014926 | 9.4226e-05 | 0.0084813 | 9.3822e-05 | 0.32 |  | 0.60376
236 | 0.0011563 | 0.000441 | 0.00314 | 1.123 | 0.0012986 | 8.2019e-05 | 0.0074639 | 8.1709e-05 | 0.33 |  | 0.575
237 | 0.0011401 | 0.00043481 | 0.0031355 | 1.1214 | 0.0012785 | 8.0868e-05 | 0.0066378 | 8.0596e-05 | 0.35 |  | 0.55
238 | 0.0014046 | 0.0005357 | 0.003131 | 1.1198 | 0.0015729 | 9.9632e-05 | 0.0062107 | 9.9319e-05 | 0.37 |  | 0.535
239 | 0.0013012 | 0.00049625 | 0.003127 | 1.1184 | 0.0014552 | 9.2295e-05 | 0.0059227 | 9.2019e-05 | 0.39 |  | 0.53
240 | 0.001468 | 0.00055985 | 0.003123 | 1.117 | 0.0016396 | 0.00010412 | 0.0056289 | 0.00010383 | 0.41 |  | 0.53
241 | 0.0015839 | 0.00060406 | 0.003109 | 1.1119 | 0.0017612 | 0.00011235 | 0.0056816 | 0.00011202 | 0.433 |  | 0.54
242 | 0.0015535 | 0.00059248 | 0.003095 | 1.1069 | 0.0017197 | 0.00011019 | 0.0055786 | 0.00010988 | 0.456 |  | 0.55
243 | 0.0017261 | 0.00065831 | 0.0030785 | 1.101 | 0.0019005 | 0.00012244 | 0.0056838 | 0.00012208 | 0.479 |  | 0.563
244 | 0.0021687 | 0.00082711 | 0.003062 | 1.0951 | 0.002375 | 0.00015383 | 0.0057924 | 0.00015338 | 0.5 |  | 0.582
245 | 0.0027368 | 0.0010437 | 0.0030435 | 1.0885 | 0.002979 | 0.00019412 | 0.0060034 | 0.00019353 | 0.52 |  | 0.60376
246 | 0.0033488 | 0.0012771 | 0.003025 | 1.0819 | 0.003623 | 0.00023753 | 0.0062176 | 0.00023678 | 0.54 |  | 0.62389
247 | 0.0041234 | 0.0015726 | 0.003 | 1.073 | 0.0044242 | 0.00029247 | 0.0064445 | 0.00029152 | 0.56 |  | 0.64402
248 | 0.0057771 | 0.0022032 | 0.002975 | 1.064 | 0.0061469 | 0.00040977 | 0.0065917 | 0.00040841 | 0.59 |  | 0.66414
249 | 0.0071633 | 0.0027319 | 0.002948 | 1.0544 | 0.0075527 | 0.0005081 | 0.0068964 | 0.00050633 | 0.61 |  | 0.68427
250 | 0.010401 | 0.0039666 | 0.002921 | 1.0447 | 0.010866 | 0.00073773 | 0.0070442 | 0.0007351 | 0.63 |  | 0.7044
251 | 0.048274 | 0.018411 | 0.0028895 | 1.0334 | 0.049888 | 0.0034241 | 0.0071719 | 0.0034117 | 0.65 |  | 0.71952
252 | 0.27607 | 0.10529 | 0.002858 | 1.0222 | 0.28219 | 0.019582 | 0.0074077 | 0.019508 | 0.665 |  | 0.73464
253 | 0.90668 | 0.34579 | 0.002827 | 1.0111 | 0.91673 | 0.064312 | 0.0074842 | 0.064068 | 0.68 |  | 0.74976
254 | 1 | 0.38138 | 0.002796 | 1 | 1 | 0.070931 | 0.0076716 | 0.070656 | 0.695 |  | 0.76488
255 | 0.10573 | 0.040323 | 0.0027625 | 0.98802 | 0.10446 | 0.0074995 | 0.0077713 | 0.00747 | 0.71 |  | 0.78
256 | 0.018787 | 0.0071651 | 0.002729 | 0.97604 | 0.018337 | 0.0013326 | 0.0079251 | 0.0013273 | 0.72 |  | 0.784
257 | 0.013202 | 0.005035 | 0.002691 | 0.96245 | 0.012706 | 0.00093644 | 0.00809 | 0.0009326 | 0.72 |  | 0.791
258 | 0.01076 | 0.0041034 | 0.002653 | 0.94886 | 0.010209 | 0.00076318 | 0.0081185 | 0.00076005 | 0.72 |  | 0.793
259 | 0.0092456 | 0.0035261 | 0.0026145 | 0.93509 | 0.0086454 | 0.0006558 | 0.0083162 | 0.00065304 | 0.735 |  | 0.8
260 | 0.0081671 | 0.0031148 | 0.002576 | 0.92132 | 0.0075245 | 0.0005793 | 0.0082852 | 0.00057688 | 0.725 |  | 0.795
261 | 0.0076305 | 0.0029101 | 0.0025395 | 0.90826 | 0.0069305 | 0.00054124 | 0.0083284 | 0.00053896 | 0.7125 |  | 0.784
262 | 0.0073063 | 0.0027865 | 0.002503 | 0.89521 | 0.0065407 | 0.00051825 | 0.0083959 | 0.00051605 | 0.7 |  | 0.773
263 | 0.0072792 | 0.0027761 | 0.0024625 | 0.88072 | 0.0064109 | 0.00051632 | 0.0083966 | 0.00051413 | 0.6925 |  | 0.762
264 | 0.0076346 | 0.0029117 | 0.002422 | 0.86624 | 0.0066134 | 0.00054153 | 0.0085184 | 0.0005392 | 0.685 |  | 0.751
265 | 0.0074517 | 0.0028419 | 0.002378 | 0.8505 | 0.0063376 | 0.00052855 | 0.0085185 | 0.00052628 | 0.675 |  | 0.74
266 | 0.004898 | 0.001868 | 0.002334 | 0.83476 | 0.0040887 | 0.00034742 | 0.0084257 | 0.00034594 | 0.662 |  | 0.72282
267 | 0.0034172 | 0.0013033 | 0.00229 | 0.81903 | 0.0027988 | 0.00024239 | 0.0084511 | 0.00024135 | 0.649 |  | 0.70563
268 | 0.0031787 | 0.0012123 | 0.002246 | 0.80329 | 0.0025534 | 0.00022547 | 0.0083631 | 0.00022451 | 0.635 |  | 0.68845
269 | 0.0030099 | 0.0011479 | 0.002197 | 0.78577 | 0.0023651 | 0.0002135 | 0.0083403 | 0.0002126 | 0.622 |  | 0.67126
270 | 0.0029572 | 0.0011278 | 0.002148 | 0.76824 | 0.0022718 | 0.00020976 | 0.0082375 | 0.00020888 | 0.6 |  | 0.65408
271 | 0.0022686 | 0.00086518 | 0.002148 | 0.76824 | 0.0017428 | 0.00016091 | 0.0082114 | 0.00016024 | 0.58 |  | 0.62994
272 | 0.0012311 | 0.0004695 | 0.002148 | 0.76824 | 0.00094576 | 8.7321e-05 | 0.0080789 | 8.6964e-05 | 0.56 |  | 0.60579
273 | 0.0010628 | 0.00040534 | 0.002063 | 0.73784 | 0.00078419 | 7.5387e-05 | 0.0079362 | 7.5085e-05 | 0.54 |  | 0.58165
274 | 0.0010294 | 0.00039258 | 0.001978 | 0.70744 | 0.00072823 | 7.3015e-05 | 0.0078569 | 7.2725e-05 | 0.52 |  | 0.5575
275 | 0.0013634 | 0.00051997 | 0.0019365 | 0.6926 | 0.00094428 | 9.6707e-05 | 0.0077657 | 9.6327e-05 | 0.5 |  | 0.53336
276 | 0.0010872 | 0.00041463 | 0.001895 | 0.67775 | 0.00073685 | 7.7116e-05 | 0.0076867 | 7.6816e-05 | 0.48 |  | 0.50718
277 | 0.00080775 | 0.00030806 | 0.00185 | 0.66166 | 0.00053445 | 5.7294e-05 | 0.0074587 | 5.7078e-05 | 0.46 |  | 0.48101
278 | 0.00066459 | 0.00025346 | 0.001805 | 0.64557 | 0.00042904 | 4.714e-05 | 0.007289 | 4.6967e-05 | 0.44 |  | 0.45483
279 | 0.00065469 | 0.00024968 | 0.0017635 | 0.63072 | 0.00041293 | 4.6438e-05 | 0.0071891 | 4.6269e-05 | 0.42 |  | 0.42866
280 | 0.0010812 | 0.00041233 | 0.001722 | 0.61588 | 0.00066587 | 7.6688e-05 | 0.0070557 | 7.6415e-05 | 0.4 |  | 0.40248
281 | 0.00094815 | 0.0003616 | 0.0016815 | 0.60139 | 0.00057021 | 6.7253e-05 | 0.0068292 | 6.7021e-05 | 0.376 |  | 0.37798
282 | 0.00061182 | 0.00023334 | 0.001641 | 0.58691 | 0.00035908 | 4.3397e-05 | 0.0065943 | 4.3252e-05 | 0.352 |  | 0.35349
283 | 0.00046871 | 0.00017876 | 0.001595 | 0.57046 | 0.00026738 | 3.3246e-05 | 0.0062673 | 3.3141e-05 | 0.328 |  | 0.32899
284 | 0.00042416 | 0.00016177 | 0.001549 | 0.55401 | 0.00023499 | 3.0086e-05 | 0.0059948 | 2.9995e-05 | 0.305 |  | 0.3045
285 | 0.00041775 | 0.00015932 | 0.001503 | 0.53755 | 0.00022456 | 2.9631e-05 | 0.0056155 | 2.9547e-05 | 0.28 |  | 0.28
286 | 0.00044183 | 0.0001685 | 0.001457 | 0.5211 | 0.00023024 | 3.1339e-05 | 0.005357 | 3.1254e-05 | 0.255 |  | 0.258
287 | 0.00035598 | 0.00013576 | 0.0014135 | 0.50554 | 0.00017996 | 2.525e-05 | 0.0051804 | 2.5184e-05 | 0.23 |  | 0.236
288 | 0.0007088 | 0.00027032 | 0.00137 | 0.48999 | 0.0003473 | 5.0276e-05 | 0.0049243 | 5.0151e-05 | 0.21 |  | 0.214
289 | 0.0015118 | 0.00057658 | 0.0013245 | 0.47371 | 0.00071617 | 0.00010724 | 0.0046842 | 0.00010698 | 0.19 |  | 0.192
290 | 0.00079336 | 0.00030257 | 0.001279 | 0.45744 | 0.00036291 | 5.6274e-05 | 0.0044139 | 5.6148e-05 | 0.17 |  | 0.17
291 | 0.00034248 | 0.00013061 | 0.0012305 | 0.44009 | 0.00015072 | 2.4292e-05 | 0.0040187 | 2.4243e-05 | 0.15 |  | 0.15
292 | 0.00047913 | 0.00018273 | 0.001182 | 0.42275 | 0.00020255 | 3.3985e-05 | 0.0037011 | 3.3922e-05 | 0.14 |  | 0.13
293 | 0.00042534 | 0.00016222 | 0.001143 | 0.4088 | 0.00017388 | 3.017e-05 | 0.0033761 | 3.0118e-05 | 0.125 |  | 0.11
294 | 0.00030768 | 0.00011734 | 0.001104 | 0.39485 | 0.00012149 | 2.1824e-05 | 0.0030782 | 2.179e-05 | 0.11 |  | 0.095
295 | 0.00079426 | 0.00030291 | 0.001058 | 0.3784 | 0.00030055 | 5.6338e-05 | 0.0029073 | 5.6255e-05 | 0.095 |  | 0.08048
296 | 0.0035535 | 0.0013552 | 0.001012 | 0.36195 | 0.0012862 | 0.00025205 | 0.0027085 | 0.00025171 | 0.08 |  | 0.07
297 | 0.0032764 | 0.0012496 | 0.0009661 | 0.34553 | 0.0011321 | 0.0002324 | 0.0026422 | 0.00023209 | 0.071 |  | 0.06
298 | 0.00035465 | 0.00013525 | 0.0009202 | 0.32911 | 0.00011672 | 2.5155e-05 | 0.0023456 | 2.5125e-05 | 0.062 |  | 0.0534
299 | 0.00017527 | 6.6844e-05 | 0.00088165 | 0.31533 | 5.5267e-05 | 1.2432e-05 | 0.0023566 | 1.2417e-05 | 0.053 |  | 0.0468
300 | 0.00023377 | 8.9155e-05 | 0.0008431 | 0.30154 | 7.0491e-05 | 1.6582e-05 | 0.002246 | 1.6563e-05 | 0.045 |  | 0.04024
301 | 0.0013107 | 0.00049988 | 0.00080045 | 0.28628 | 0.00037524 | 9.2971e-05 | 0.0023133 | 9.2862e-05
302 | 0.002814 | 0.0010732 | 0.0007578 | 0.27103 | 0.00076267 | 0.0001996 | 0.0020654 | 0.00019939
303 | 0.00092139 | 0.0003514 | 0.00071605 | 0.2561 | 0.00023597 | 6.5355e-05 | 0.0020675 | 6.5287e-05
304 | 0.00019589 | 7.471e-05 | 0.0006743 | 0.24117 | 4.7243e-05 | 1.3895e-05 | 0.0019218 | 1.3881e-05
305 | 0.00020657 | 7.878e-05 | 0.00063875 | 0.22845 | 4.7191e-05 | 1.4652e-05 | 0.0019577 | 1.4638e-05
306 | 0.00017702 | 6.751e-05 | 0.0006032 | 0.21574 | 3.8189e-05 | 1.2556e-05 | 0.001946 | 1.2544e-05
307 | 0.00019585 | 7.4692e-05 | 0.00056705 | 0.20281 | 3.9719e-05 | 1.3892e-05 | 0.0018501 | 1.3879e-05
308 | 0.00019418 | 7.4056e-05 | 0.0005309 | 0.18988 | 3.687e-05 | 1.3773e-05 | 0.0017135 | 1.3761e-05
309 | 0.0002132 | 8.131e-05 | 0.0004986 | 0.17833 | 3.8019e-05 | 1.5122e-05 | 0.0018036 | 1.5109e-05
310 | 0.00023537 | 8.9765e-05 | 0.0004663 | 0.16677 | 3.9254e-05 | 1.6695e-05 | 0.0018024 | 1.668e-05
311 | 0.0018814 | 0.00071751 | 0.0004397 | 0.15726 | 0.00029586 | 0.00013345 | 0.0018007 | 0.00013333
312 | 0.01427 | 0.0054421 | 0.0004131 | 0.14775 | 0.0021083 | 0.0010122 | 0.0016782 | 0.0010113
313 | 0.01797 | 0.0068532 | 0.00038805 | 0.13879 | 0.002494 | 0.0012746 | 0.0016608 | 0.0012735
314 | 0.002966 | 0.0011312 | 0.000363 | 0.12983 | 0.00038507 | 0.00021038 | 0.0016318 | 0.00021021
315 | 0.00024039 | 9.1681e-05 | 0.00033755 | 0.12073 | 2.9022e-05 | 1.7051e-05 | 0.0016031 | 1.7038e-05
316 | 0.00017634 | 6.7251e-05 | 0.0003121 | 0.11162 | 1.9684e-05 | 1.2508e-05 | 0.0015152 | 1.2498e-05
317 | 0.00017593 | 6.7097e-05 | 0.00029075 | 0.10399 | 1.8295e-05 | 1.2479e-05 | 0.0015568 | 1.2469e-05
318 | 0.00020544 | 7.835e-05 | 0.0002694 | 0.096352 | 1.9795e-05 | 1.4572e-05 | 0.0015829 | 1.456e-05
319 | 0.00017242 | 6.5758e-05 | 0.0002494 | 0.089199 | 1.538e-05 | 1.223e-05 | 0.0015757 | 1.222e-05
320 | 0.00016338 | 6.231e-05 | 0.0002294 | 0.082046 | 1.3405e-05 | 1.1589e-05 | 0.0014859 | 1.158e-05
321 | 0.00017654 | 6.733e-05 | 0.0002116 | 0.07568 | 1.3361e-05 | 1.2522e-05 | 0.0015423 | 1.2513e-05
322 | 0.00017823 | 6.7972e-05 | 0.0001938 | 0.069313 | 1.2353e-05 | 1.2642e-05 | 0.0015245 | 1.2632e-05
323 | 0.0001692 | 6.4528e-05 | 0.0001783 | 0.06377 | 1.079e-05 | 1.2001e-05 | 0.0014518 | 1.1992e-05
324 | 0.0001029 | 3.9245e-05 | 0.0001628 | 0.058226 | 5.9917e-06 | 7.2991e-06 | 0.0014213 | 7.2938e-06
325 | 0.00011105 | 4.2351e-05 | 0.0001493 | 0.053398 | 5.9297e-06 | 7.8767e-06 | 0.0015273 | 7.8706e-06
326 | 0.00012206 | 4.6551e-05 | 0.0001358 | 0.048569 | 5.9284e-06 | 8.6578e-06 | 0.0014433 | 8.6515e-06
327 | 0.00013667 | 5.2123e-05 | 0.00012465 | 0.044582 | 6.0929e-06 | 9.6941e-06 | 0.0013595 | 9.6874e-06
328 | 0.00013288 | 5.0677e-05 | 0.0001135 | 0.040594 | 5.3941e-06 | 9.4253e-06 | 0.0013031 | 9.419e-06
329 | 6.7644e-05 | 2.5798e-05 | 0.00010425 | 0.037287 | 2.5223e-06 | 4.7981e-06 | 0.0013448 | 4.7948e-06
330 | 8.87e-05 | 3.3828e-05 | 9.501e-05 | 0.033981 | 3.0141e-06 | 6.2916e-06 | 0.0013985 | 6.2871e-06
331 | 0.00010475 | 3.9949e-05 | 8.709e-05 | 0.031148 | 3.2627e-06 | 7.4299e-06 | 0.0013526 | 7.4248e-06
332 | 9.8582e-05 | 3.7597e-05 | 7.917e-05 | 0.028315 | 2.7914e-06 | 6.9925e-06 | 0.0013813 | 6.9876e-06
333 | 0.00063765 | 0.00024319 | 7.2925e-05 | 0.026082 | 1.6631e-05 | 4.5229e-05 | 0.0011794 | 4.5202e-05
334 | 0.0013196 | 0.00050325 | 6.668e-05 | 0.023848 | 3.1469e-05 | 9.3598e-05 | 0.0012495 | 9.3539e-05
335 | 0.00029357 | 0.00011196 | 6.117e-05 | 0.021878 | 6.4225e-06 | 2.0823e-05 | 0.0012781 | 2.0809e-05
336 | 6.381e-05 | 2.4336e-05 | 5.566e-05 | 0.019907 | 1.2703e-06 | 4.5261e-06 | 0.0011952 | 4.5234e-06
337 | 5.743e-05 | 2.1902e-05 | 5.0975e-05 | 0.018231 | 1.047e-06 | 4.0736e-06 | 0.0011471 | 4.0712e-06
338 | 8.6171e-05 | 3.2864e-05 | 4.629e-05 | 0.016556 | 1.4266e-06 | 6.1122e-06 | 0.0011351 | 6.1087e-06
339 | 9.804e-05 | 3.739e-05 | 4.253e-05 | 0.015211 | 1.4913e-06 | 6.9541e-06 | 0.0012439 | 6.9497e-06
340 | 9.0644e-05 | 3.457e-05 | 3.877e-05 | 0.013866 | 1.2569e-06 | 6.4295e-06 | 0.00096916 | 6.4263e-06
341 | 0.00011304 | 4.3111e-05 | 3.5685e-05 | 0.012763 | 1.4427e-06 | 8.018e-06 | 0.0011623 | 8.0133e-06
342 | 0.00011074 | 4.2233e-05 | 3.26e-05 | 0.01166 | 1.2912e-06 | 7.8548e-06 | 0.0010939 | 7.8505e-06
343 | 6.9804e-05 | 2.6622e-05 | 2.9995e-05 | 0.010728 | 7.4885e-07 | 4.9513e-06 | 0.0011048 | 4.9485e-06
344 | 8.3455e-05 | 3.1828e-05 | 2.739e-05 | 0.0097961 | 8.1753e-07 | 5.9195e-06 | 0.00094027 | 5.9167e-06
345 | 6.6646e-05 | 2.5417e-05 | 2.519e-05 | 0.0090093 | 6.0043e-07 | 4.7273e-06 | 0.00099814 | 4.7249e-06
346 | 8.8789e-05 | 3.3862e-05 | 2.299e-05 | 0.0082225 | 7.3007e-07 | 6.2979e-06 | 0.0010038 | 6.2947e-06
347 | 6.703e-05 | 2.5564e-05 | 2.1225e-05 | 0.0075912 | 5.0884e-07 | 4.7545e-06 | 0.0009077 | 4.7523e-06
348 | 5.9403e-05 | 2.2655e-05 | 1.946e-05 | 0.0069599 | 4.1344e-07 | 4.2135e-06 | 0.0011222 | 4.2111e-06
349 | 8.6863e-05 | 3.3128e-05 | 1.8045e-05 | 0.0064539 | 5.606e-07 | 6.1613e-06 | 0.00084142 | 6.1587e-06
350 | 8.0879e-05 | 3.0845e-05 | 1.663e-05 | 0.0059478 | 4.8105e-07 | 5.7368e-06 | 0.0010446 | 5.7338e-06
351 | 6.6617e-05 | 2.5406e-05 | 1.534e-05 | 0.0054864 | 3.6549e-07 | 4.7252e-06 | 0.0010236 | 4.7228e-06
352 | 9.4329e-05 | 3.5975e-05 | 1.405e-05 | 0.005025 | 4.7401e-07 | 6.6909e-06 | 0.00099329 | 6.6875e-06
353 | 6.3042e-05 | 2.4043e-05 | 1.301e-05 | 0.0046531 | 2.9334e-07 | 4.4717e-06 | 0.0010458 | 4.4693e-06
354 | 9.1292e-05 | 3.4817e-05 | 1.197e-05 | 0.0042811 | 3.9083e-07 | 6.4754e-06 | 0.00095677 | 6.4723e-06
355 | 0.0001 | 3.814e-05 | 1.1125e-05 | 0.0039789 | 3.9791e-07 | 7.0934e-06 | 0.0010009 | 7.0898e-06
356 | 0.00011698 | 4.4615e-05 | 1.028e-05 | 0.0036767 | 4.3011e-07 | 8.2977e-06 | 0.00092927 | 8.2938e-06
357 | 0.00013199 | 5.0337e-05 | 9.559e-06 | 0.0034188 | 4.5124e-07 | 9.362e-06 | 0.00088763 | 9.3578e-06
358 | 0.00010328 | 3.939e-05 | 8.838e-06 | 0.0031609 | 3.2647e-07 | 7.326e-06 | 0.00097059 | 7.3224e-06
359 | 0.00013799 | 5.2625e-05 | 8.166e-06 | 0.0029206 | 4.03e-07 | 9.7874e-06 | 0.00080683 | 9.7834e-06
360 | 0.00012523 | 4.7759e-05 | 7.494e-06 | 0.0026803 | 3.3564e-07 | 8.8824e-06 | 0.00091058 | 8.8783e-06
361 | 0.00012576 | 4.7963e-05 | 7.0385e-06 | 0.0025173 | 3.1659e-07 | 8.9205e-06 | 0.00086687 | 8.9166e-06
362 | 0.00017383 | 6.6295e-05 | 6.583e-06 | 0.0023544 | 4.0927e-07 | 1.233e-05 | 0.0008648 | 1.2325e-05
363 | 0.00033959 | 0.00012951 | 6.161e-06 | 0.0022035 | 7.4829e-07 | 2.4087e-05 | 0.00090633 | 2.4076e-05
364 | 0.0075402 | 0.0028757 | 5.739e-06 | 0.0020526 | 1.5477e-05 | 0.00053483 | 0.00083712 | 0.00053461
365 | 0.015614 | 0.0059548 | 5.3745e-06 | 0.0019222 | 3.0013e-05 | 0.0011075 | 0.00080704 | 0.001107
366 | 0.0056879 | 0.0021693 | 5.01e-06 | 0.0017918 | 1.0192e-05 | 0.00040345 | 0.00086933 | 0.00040327
367 | 0.0010433 | 0.00039788 | 4.682e-06 | 0.0016745 | 1.747e-06 | 7.4e-05 | 0.00083938 | 7.3969e-05
368 | 0.00021817 | 8.3205e-05 | 4.354e-06 | 0.0015572 | 3.3974e-07 | 1.5475e-05 | 0.00078049 | 1.5469e-05
369 | 0.00018377 | 7.0085e-05 | 4.1865e-06 | 0.0014973 | 2.7516e-07 | 1.3035e-05 | 0.00078234 | 1.303e-05
370 | 0.00020553 | 7.8384e-05 | 4.019e-06 | 0.0014374 | 2.9543e-07 | 1.4578e-05 | 0.00077894 | 1.4573e-05
371 | 0.00019424 | 7.408e-05 | 3.8535e-06 | 0.0013782 | 2.6771e-07 | 1.3778e-05 | 0.00077278 | 1.3772e-05
372 | 0.00019328 | 7.3713e-05 | 3.688e-06 | 0.001319 | 2.5494e-07 | 1.371e-05 | 0.00076664 | 1.3704e-05
373 | 0.00022584 | 8.6131e-05 | 4.107e-06 | 0.0014689 | 3.3174e-07 | 1.6019e-05 | 0.00075445 | 1.6013e-05
374 | 0.0002082 | 7.9405e-05 | 4.095e-06 | 0.0014646 | 3.0493e-07 | 1.4768e-05 | 0.00072634 | 1.4763e-05
375 | 0.00016218 | 6.1853e-05 | 4.6715e-06 | 0.0016708 | 2.7097e-07 | 1.1504e-05 | 0.00068561 | 1.15e-05
376 | 0.00017473 | 6.6638e-05 | 5.248e-06 | 0.001877 | 3.2796e-07 | 1.2394e-05 | 0.00074424 | 1.2389e-05
377 | 0.00017623 | 6.7212e-05 | 5.043e-06 | 0.0018036 | 3.1787e-07 | 1.2501e-05 | 0.00067934 | 1.2496e-05
378 | 0.00017251 | 6.5791e-05 | 6.483e-06 | 0.0023187 | 3.9999e-07 | 1.2236e-05 | 0.0006805 | 1.2232e-05
379 | 0.00014118 | 5.3843e-05 | 5.763e-06 | 0.0020612 | 2.9099e-07 | 1.0014e-05 | 0.00070633 | 1.001e-05
380 | 0.00015397 | 5.8722e-05 | 5.043e-06 | 0.0018036 | 2.7771e-07 | 1.0922e-05 | 0.00065437 | 1.0918e-05
381 | 0.00017951 | 6.8461e-05 | 5.763e-06 | 0.0020612 | 3.7e-07 | 1.2733e-05 | 0.00067834 | 1.2728e-05
382 | 0.00017294 | 6.5957e-05 | 6.483e-06 | 0.0023187 | 4.01e-07 | 1.2267e-05 | 0.00066288 | 1.2263e-05
383 | 0.00016322 | 6.2247e-05 | 6.2585e-06 | 0.0022384 | 3.6534e-07 | 1.1577e-05 | 0.00060943 | 1.1574e-05
384 | 0.00012315 | 4.6968e-05 | 6.034e-06 | 0.0021581 | 2.6578e-07 | 8.7355e-06 | 0.00063289 | 8.7327e-06
385 | 0.00015031 | 5.7325e-05 | 6.766e-06 | 0.0024199 | 3.6373e-07 | 1.0662e-05 | 0.0006262 | 1.0658e-05
386 | 0.00016587 | 6.3259e-05 | 7.498e-06 | 0.0026817 | 4.4481e-07 | 1.1765e-05 | 0.00060363 | 1.1762e-05
387 | 0.00018133 | 6.9155e-05 | 7.6255e-06 | 0.0027273 | 4.9454e-07 | 1.2862e-05 | 0.00059613 | 1.2858e-05
388 | 0.00016286 | 6.2111e-05 | 7.753e-06 | 0.0027729 | 4.5159e-07 | 1.1552e-05 | 0.0006371 | 1.1548e-05
389 | 0.00018478 | 7.0471e-05 | 8.037e-06 | 0.0028745 | 5.3114e-07 | 1.3107e-05 | 0.00061753 | 1.3103e-05
390 | 0.00026019 | 9.923e-05 | 8.321e-06 | 0.002976 | 7.7433e-07 | 1.8455e-05 | 0.00055253 | 1.845e-05
391 | 0.00023334 | 8.8992e-05 | 8.5965e-06 | 0.0030746 | 7.1743e-07 | 1.6551e-05 | 0.00058807 | 1.6546e-05
392 | 0.00018759 | 7.1543e-05 | 8.872e-06 | 0.0031731 | 5.9525e-07 | 1.3306e-05 | 0.00055449 | 1.3302e-05
393 | 0.00018601 | 7.0939e-05 | 9.0195e-06 | 0.0032259 | 6.0003e-07 | 1.3194e-05 | 0.00056846 | 1.319e-05
394 | 0.0001798 | 6.857e-05 | 9.167e-06 | 0.0032786 | 5.8948e-07 | 1.2753e-05 | 0.00052676 | 1.275e-05
395 | 0.0001887 | 7.1965e-05 | 9.4235e-06 | 0.0033704 | 6.3598e-07 | 1.3384e-05 | 0.00057925 | 1.3381e-05
396 | 0.00016207 | 6.1809e-05 | 9.68e-06 | 0.0034621 | 5.6109e-07 | 1.1496e-05 | 0.00053536 | 1.1492e-05
397 | 0.00018398 | 7.0164e-05 | 9.895e-06 | 0.003539 | 6.5109e-07 | 1.305e-05 | 0.00050619 | 1.3046e-05
398 | 0.00020062 | 7.6514e-05 | 1.011e-05 | 0.0036159 | 7.2543e-07 | 1.423e-05 | 0.00055289 | 1.4226e-05
399 | 0.00020534 | 7.8314e-05 | 1.0295e-05 | 0.003682 | 7.5609e-07 | 1.4565e-05 | 0.00057306 | 1.4561e-05
400 | 0.00019311 | 7.3648e-05 | 1.048e-05 | 0.0037482 | 7.2382e-07 | 1.3697e-05 | 0.0005442 | 1.3694e-05
401 |  | 0 |  | 0 | 0 | 0
402 |  | 0 |  | 0 | 0 | 0
403 |  | 0 |  | 0 | 0 | 0
404 |  | 0 |  | 0 | 0 | 0
405 |  | 0 |  | 0 | 0 | 0
406 |  | 0 |  | 0 | 0 | 0
407 |  | 0 |  | 0 | 0 | 0
408 |  | 0 |  | 0 | 0 | 0
409 |  | 0 |  | 0 | 0 | 0
410 |  | 0 |  | 0 | 0 | 0
411 |  | 0 |  | 0 | 0 | 0
412 |  | 0 |  | 0 | 0 | 0
413 |  | 0 |  | 0 | 0 | 0
414 |  | 0 |  | 0 | 0 | 0
415 |  | 0 |  | 0 | 0 | 0
416 |  | 0 |  | 0 | 0 | 0
417 |  | 0 |  | 0 | 0 | 0
418 |  | 0 |  | 0 | 0 | 0
419 |  | 0 |  | 0 | 0 | 0
420 |  | 0 |  | 0 | 0 | 0
421 |  | 0 |  | 0 | 0 | 0
422 |  | 0 |  | 0 | 0 | 0
423 |  | 0 |  | 0 | 0 | 0
424 |  | 0 |  | 0 | 0 | 0
425 |  | 0 |  | 0 | 0 | 0
426 |  | 0 |  | 0 | 0 | 0
427 |  | 0 |  | 0 | 0 | 0
428 |  | 0 |  | 0 | 0 | 0
429 |  | 0 |  | 0 | 0 | 0
430 |  | 0 |  | 0 | 0 | 0
431 |  | 0 |  | 0 | 0 | 0
432 |  | 0 |  | 0 | 0 | 0
433 |  | 0 |  | 0 | 0 | 0
434 |  | 0 |  | 0 | 0 | 0
435 |  | 0 |  | 0 | 0 | 0
436 |  | 0 |  | 0 | 0 | 0
437 |  | 0 |  | 0 | 0 | 0
438 |  | 0 |  | 0 | 0 | 0
439 |  | 0 |  | 0 | 0 | 0
440 |  | 0 |  | 0 | 0 | 0
441 |  | 0 |  | 0 | 0 | 0
442 |  | 0 |  | 0 | 0 | 0
443 |  | 0 |  | 0 | 0 | 0
444 |  | 0 |  | 0 | 0 | 0
445 |  | 0 |  | 0 | 0 | 0
446 |  | 0 |  | 0 | 0 | 0
447 |  | 0 |  | 0 | 0 | 0
448 |  | 0 |  | 0 | 0 | 0
449 |  | 0 |  | 0 | 0 | 0
450 |  | 0 |  | 0 | 0 | 0
451 |  | 0 |  | 0 | 0 | 0
452 |  | 0 |  | 0 | 0 | 0
453 |  | 0 |  | 0 | 0 | 0
454 |  | 0 |  | 0 | 0 | 0
455 |  | 0 |  | 0 | 0 | 0
456 |  | 0 |  | 0 | 0 | 0
457 |  | 0 |  | 0 | 0 | 0
458 |  | 0 |  | 0 | 0 | 0
459 |  | 0 |  | 0 | 0 | 0
460 |  | 0 |  | 0 | 0 | 0
461 |  | 0 |  | 0 | 0 | 0
462 |  | 0 |  | 0 | 0 | 0
463 |  | 0 |  | 0 | 0 | 0
464 |  | 0 |  | 0 | 0 | 0
465 |  | 0 |  | 0 | 0 | 0
466 |  | 0 |  | 0 | 0 | 0
467 |  | 0 |  | 0 | 0 | 0
468 |  | 0 |  | 0 | 0 | 0
469 |  | 0 |  | 0 | 0 | 0
470 |  | 0 |  | 0 | 0 | 0
471 |  | 0 |  | 0 | 0 | 0
472 |  | 0 |  | 0 | 0 | 0
473 |  | 0 |  | 0 | 0 | 0
474 |  | 0 |  | 0 | 0 | 0
475 |  | 0 |  | 0 | 0 | 0
476 |  | 0 |  | 0 | 0 | 0
477 |  | 0 |  | 0 | 0 | 0
478 |  | 0 |  | 0 | 0 | 0
479 |  | 0 |  | 0 | 0 | 0
480 |  | 0 |  | 0 | 0 | 0
481 |  | 0 |  | 0 | 0 | 0
482 |  | 0 |  | 0 | 0 | 0
483 |  | 0 |  | 0 | 0 | 0
484 |  | 0 |  | 0 | 0 | 0
485 |  | 0 |  | 0 | 0 | 0
486 |  | 0 |  | 0 | 0 | 0
487 |  | 0 |  | 0 | 0 | 0
488 |  | 0 |  | 0 | 0 | 0
489 |  | 0 |  | 0 | 0 | 0
490 |  | 0 |  | 0 | 0 | 0
491 |  | 0 |  | 0 | 0 | 0
492 |  | 0 |  | 0 | 0 | 0
493 |  | 0 |  | 0 | 0 | 0
494 |  | 0 |  | 0 | 0 | 0
495 |  | 0 |  | 0 | 0 | 0
496 |  | 0 |  | 0 | 0 | 0
497 |  | 0 |  | 0 | 0 | 0
498 |  | 0 |  | 0 | 0 | 0
499 |  | 0 |  | 0 | 0 | 0
500 |  | 0 |  | 0 | 0 | 0
501 |  | 0 |  | 0 | 0 | 0
502 |  | 0 |  | 0 | 0 | 0
503 |  | 0 |  | 0 | 0 | 0
504 |  | 0 |  | 0 | 0 | 0
505 |  | 0 |  | 0 | 0 | 0
506 |  | 0 |  | 0 | 0 | 0
507 |  | 0 |  | 0 | 0 | 0
508 |  | 0 |  | 0 | 0 | 0
509 |  | 0 |  | 0 | 0 | 0
510 |  | 0 |  | 0 | 0 | 0
511 |  | 0 |  | 0 | 0 | 0
512 |  | 0 |  | 0 | 0 | 0
513 |  | 0 |  | 0 | 0 | 0
514 |  | 0 |  | 0 | 0 | 0
515 |  | 0 |  | 0 | 0 | 0
516 |  | 0 |  | 0 | 0 | 0
517 |  | 0 |  | 0 | 0 | 0
518 |  | 0 |  | 0 | 0 | 0
519 |  | 0 |  | 0 | 0 | 0
520 |  | 0 |  | 0 | 0 | 0
521 |  | 0 |  | 0 | 0 | 0
522 |  | 0 |  | 0 | 0 | 0
523 |  | 0 |  | 0 | 0 | 0
524 |  | 0 |  | 0 | 0 | 0
525 |  | 0 |  | 0 | 0 | 0
526 |  | 0 |  | 0 | 0 | 0
527 |  | 0 |  | 0 | 0 | 0
528 |  | 0 |  | 0 | 0 | 0
529 |  | 0 |  | 0 | 0 | 0
530 |  | 0 |  | 0 | 0 | 0
531 |  | 0 |  | 0 | 0 | 0
532 |  | 0 |  | 0 | 0 | 0
533 |  | 0 |  | 0 | 0 | 0
534 |  | 0 |  | 0 | 0 | 0
535 |  | 0 |  | 0 | 0 | 0
536 |  | 0 |  | 0 | 0 | 0
537 |  | 0 |  | 0 | 0 | 0
538 |  | 0 |  | 0 | 0 | 0
539 |  | 0 |  | 0 | 0 | 0
540 |  | 0 |  | 0 | 0 | 0
541 |  | 0 |  | 0 | 0 | 0
542 |  | 0 |  | 0 | 0 | 0
543 |  | 0 |  | 0 | 0 | 0
544 |  | 0 |  | 0 | 0 | 0
545 |  | 0 |  | 0 | 0 | 0
546 |  | 0 |  | 0 | 0 | 0
547 |  | 0 |  | 0 | 0 | 0
548 |  | 0 |  | 0 | 0 | 0
549 |  | 0 |  | 0 | 0 | 0
550 |  | 0 |  | 0 | 0 | 0
551 |  | 0 |  | 0 | 0 | 0
552 |  | 0 |  | 0 | 0 | 0
553 |  | 0 |  | 0 | 0 | 0
554 |  | 0 |  | 0 | 0 | 0
555 |  | 0 |  | 0 | 0 | 0
556 |  | 0 |  | 0 | 0 | 0
557 |  | 0 |  | 0 | 0 | 0
558 |  | 0 |  | 0 | 0 | 0
559 |  | 0 |  | 0 | 0 | 0
560 |  | 0 |  | 0 | 0 | 0
561 |  | 0 |  | 0 | 0 | 0
562 |  | 0 |  | 0 | 0 | 0
563 |  | 0 |  | 0 | 0 | 0
564 |  | 0 |  | 0 | 0 | 0
565 |  | 0 |  | 0 | 0 | 0
566 |  | 0 |  | 0 | 0 | 0
567 |  | 0 |  | 0 | 0 | 0
568 |  | 0 |  | 0 | 0 | 0
569 |  | 0 |  | 0 | 0 | 0
570 |  | 0 |  | 0 | 0 | 0
571 |  | 0 |  | 0 | 0 | 0
572 |  | 0 |  | 0 | 0 | 0
573 |  | 0 |  | 0 | 0 | 0
574 |  | 0 |  | 0 | 0 | 0
575 |  | 0 |  | 0 | 0 | 0
576 |  | 0 |  | 0 | 0 | 0
577 |  | 0 |  | 0 | 0 | 0
578 |  | 0 |  | 0 | 0 | 0
579 |  | 0 |  | 0 | 0 | 0
580 |  | 0 |  | 0 | 0 | 0
581 |  | 0 |  | 0 | 0 | 0
582 |  | 0 |  | 0 | 0 | 0
583 |  | 0 |  | 0 | 0 | 0
584 |  | 0 |  | 0 | 0 | 0
585 |  | 0 |  | 0 | 0 | 0
586 |  | 0 |  | 0 | 0 | 0
587 |  | 0 |  | 0 | 0 | 0
588 |  | 0 |  | 0 | 0 | 0
589 |  | 0 |  | 0 | 0 | 0
590 |  | 0 |  | 0 | 0 | 0
591 |  | 0 |  | 0 | 0 | 0
592 |  | 0 |  | 0 | 0 | 0
593 |  | 0 |  | 0 | 0 | 0
594 |  | 0 |  | 0 | 0 | 0
595 |  | 0 |  | 0 | 0 | 0
596 |  | 0 |  | 0 | 0 | 0
597 |  | 0 |  | 0 | 0 | 0
598 |  | 0 |  | 0 | 0 | 0
599 |  | 0 |  | 0 | 0 | 0
600 |  | 0 |  | 0 | 0 | 0
601 |  | 0 |  | 0 | 0 | 0
602 |  | 0 |  | 0 | 0 | 0
603 |  | 0 |  | 0 | 0 | 0
604 |  | 0 |  | 0 | 0 | 0
605 |  | 0 |  | 0 | 0 | 0
606 |  | 0 |  | 0 | 0 | 0
607 |  | 0 |  | 0 | 0 | 0
608 |  | 0 |  | 0 | 0 | 0
609 |  | 0 |  | 0 | 0 | 0
610 |  | 0 |  | 0 | 0 | 0
611 |  | 0 |  | 0 | 0 | 0
612 |  | 0 |  | 0 | 0 | 0
613 |  | 0 |  | 0 | 0 | 0
614 |  | 0 |  | 0 | 0 | 0
615 |  | 0 |  | 0 | 0 | 0
616 |  | 0 |  | 0 | 0 | 0
617 |  | 0 |  | 0 | 0 | 0
618 |  | 0 |  | 0 | 0 | 0
619 |  | 0 |  | 0 | 0 | 0
620 |  | 0 |  | 0 | 0 | 0
621 |  | 0 |  | 0 | 0 | 0
622 |  | 0 |  | 0 | 0 | 0
623 |  | 0 |  | 0 | 0 | 0
624 |  | 0 |  | 0 | 0 | 0
625 |  | 0 |  | 0 | 0 | 0
626 |  | 0 |  | 0 | 0 | 0
627 |  | 0 |  | 0 | 0 | 0
628 |  | 0 |  | 0 | 0 | 0
629 |  | 0 |  | 0 | 0 | 0
630 |  | 0 |  | 0 | 0 | 0
631 |  | 0 |  | 0 | 0 | 0
632 |  | 0 |  | 0 | 0 | 0
633 |  | 0 |  | 0 | 0 | 0
634 |  | 0 |  | 0 | 0 | 0
635 |  | 0 |  | 0 | 0 | 0
636 |  | 0 |  | 0 | 0 | 0
637 |  | 0 |  | 0 | 0 | 0
638 |  | 0 |  | 0 | 0 | 0
639 |  | 0 |  | 0 | 0 | 0
640 |  | 0 |  | 0 | 0 | 0
641 |  | 0 |  | 0 | 0 | 0
642 |  | 0 |  | 0 | 0 | 0
643 |  | 0 |  | 0 | 0 | 0
644 |  | 0 |  | 0 | 0 | 0
645 |  | 0 |  | 0 | 0 | 0
646 |  | 0 |  | 0 | 0 | 0
647 |  | 0 |  | 0 | 0 | 0
648 |  | 0 |  | 0 | 0 | 0
649 |  | 0 |  | 0 | 0 | 0
650 |  | 0 |  | 0 | 0 | 0
651 |  | 0 |  | 0 | 0 | 0
652 |  | 0 |  | 0 | 0 | 0
653 |  | 0 |  | 0 | 0 | 0
654 |  | 0 |  | 0 | 0 | 0
655 |  | 0 |  | 0 | 0 | 0
656 |  | 0 |  | 0 | 0 | 0
657 |  | 0 |  | 0 | 0 | 0
658 |  | 0 |  | 0 | 0 | 0
659 |  | 0 |  | 0 | 0 | 0
660 |  | 0 |  | 0 | 0 | 0
661 |  | 0 |  | 0 | 0 | 0
662 |  | 0 |  | 0 | 0 | 0
663 |  | 0 |  | 0 | 0 | 0
664 |  | 0 |  | 0 | 0 | 0
665 |  | 0 |  | 0 | 0 | 0
666 |  | 0 |  | 0 | 0 | 0
667 |  | 0 |  | 0 | 0 | 0
668 |  | 0 |  | 0 | 0 | 0
669 |  | 0 |  | 0 | 0 | 0
670 |  | 0 |  | 0 | 0 | 0
671 |  | 0 |  | 0 | 0 | 0
672 |  | 0 |  | 0 | 0 | 0
673 |  | 0 |  | 0 | 0 | 0
674 |  | 0 |  | 0 | 0 | 0
675 |  | 0 |  | 0 | 0 | 0
676 |  | 0 |  | 0 | 0 | 0
677 |  | 0 |  | 0 | 0 | 0
678 |  | 0 |  | 0 | 0 | 0
679 |  | 0 |  | 0 | 0 | 0
680 |  | 0 |  | 0 | 0 | 0
681 |  | 0 |  | 0 | 0 | 0
682 |  | 0 |  | 0 | 0 | 0
683 |  | 0 |  | 0 | 0 | 0
684 |  | 0 |  | 0 | 0 | 0
685 |  | 0 |  | 0 | 0 | 0
686 |  | 0 |  | 0 | 0 | 0
687 |  | 0 |  | 0 | 0 | 0
688 |  | 0 |  | 0 | 0 | 0
689 |  | 0 |  | 0 | 0 | 0
690 |  | 0 |  | 0 | 0 | 0
691 |  | 0 |  | 0 | 0 | 0
692 |  | 0 |  | 0 | 0 | 0
693 |  | 0 |  | 0 | 0 | 0
694 |  | 0 |  | 0 | 0 | 0
695 |  | 0 |  | 0 | 0 | 0
696 |  | 0 |  | 0 | 0 | 0
697 |  | 0 |  | 0 | 0 | 0
698 |  | 0 |  | 0 | 0 | 0
699 |  | 0 |  | 0 | 0 | 0
700 |  | 0 |  | 0 | 0 | 0
701 |  | 0 |  | 0 | 0 | 0
702 |  | 0 |  | 0 | 0 | 0
703 |  | 0 |  | 0 | 0 | 0
704 |  | 0 |  | 0 | 0 | 0
705 |  | 0 |  | 0 | 0 | 0
706 |  | 0 |  | 0 | 0 | 0
707 |  | 0 |  | 0 | 0 | 0
708 |  | 0 |  | 0 | 0 | 0
709 |  | 0 |  | 0 | 0 | 0
710 |  | 0 |  | 0 | 0 | 0
711 |  | 0 |  | 0 | 0 | 0
712 |  | 0 |  | 0 | 0 | 0
713 |  | 0 |  | 0 | 0 | 0
714 |  | 0 |  | 0 | 0 | 0
715 |  | 0 |  | 0 | 0 | 0
716 |  | 0 |  | 0 | 0 | 0
717 |  | 0 |  | 0 | 0 | 0
718 |  | 0 |  | 0 | 0 | 0
719 |  | 0 |  | 0 | 0 | 0
720 |  | 0 |  | 0 | 0 | 0
721 |  | 0 |  | 0 | 0 | 0
722 |  | 0 |  | 0 | 0 | 0
723 |  | 0 |  | 0 | 0 | 0
724 |  | 0 |  | 0 | 0 | 0
725 |  | 0 |  | 0 | 0 | 0
726 |  | 0 |  | 0 | 0 | 0
727 |  | 0 |  | 0 | 0 | 0
728 |  | 0 |  | 0 | 0 | 0
729 |  | 0 |  | 0 | 0 | 0
730 |  | 0 |  | 0 | 0 | 0
731 |  | 0 |  | 0 | 0 | 0
732 |  | 0 |  | 0 | 0 | 0
733 |  | 0 |  | 0 | 0 | 0
734 |  | 0 |  | 0 | 0 | 0
735 |  | 0 |  | 0 | 0 | 0
736 |  | 0 |  | 0 | 0 | 0
737 |  | 0 |  | 0 | 0 | 0
738 |  | 0 |  | 0 | 0 | 0
739 |  | 0 |  | 0 | 0 | 0
740 |  | 0 |  | 0 | 0 | 0
741 |  | 0 |  | 0 | 0 | 0
742 |  | 0 |  | 0 | 0 | 0
743 |  | 0 |  | 0 | 0 | 0
744 |  | 0 |  | 0 | 0 | 0
745 |  | 0 |  | 0 | 0 | 0
746 |  | 0 |  | 0 | 0 | 0
747 |  | 0 |  | 0 | 0 | 0
748 |  | 0 |  | 0 | 0 | 0
749 |  | 0 |  | 0 | 0 | 0
750 |  | 0 |  | 0 | 0 | 0
751 |  | 0 |  | 0 | 0 | 0
752 |  | 0 |  | 0 | 0 | 0
753 |  | 0 |  | 0 | 0 | 0
754 |  | 0 |  | 0 | 0 | 0
755 |  | 0 |  | 0 | 0 | 0
756 |  | 0 |  | 0 | 0 | 0
757 |  | 0 |  | 0 | 0 | 0
758 |  | 0 |  | 0 | 0 | 0
759 |  | 0 |  | 0 | 0 | 0
760 |  | 0 |  | 0 | 0 | 0
761 |  | 0 |  | 0 | 0 | 0
762 |  | 0 |  | 0 | 0 | 0
763 |  | 0 |  | 0 | 0 | 0
764 |  | 0 |  | 0 | 0 | 0
765 |  | 0 |  | 0 | 0 | 0
766 |  | 0 |  | 0 | 0 | 0
767 |  | 0 |  | 0 | 0 | 0
768 |  | 0 |  | 0 | 0 | 0
769 |  | 0 |  | 0 | 0 | 0
770 |  | 0 |  | 0 | 0 | 0
771 |  | 0 |  | 0 | 0 | 0
772 |  | 0 |  | 0 | 0 | 0
773 |  | 0 |  | 0 | 0 | 0
774 |  | 0 |  | 0 | 0 | 0
775 |  | 0 |  | 0 | 0 | 0
776 |  | 0 |  | 0 | 0 | 0
777 |  | 0 |  | 0 | 0 | 0
778 |  | 0 |  | 0 | 0 | 0
779 |  | 0 |  | 0 | 0 | 0
780 |  | 0 |  | 0 | 0 | 0
781 |  | 0 |  | 0 | 0 | 0
782 |  | 0 |  | 0 | 0 | 0
783 |  | 0 |  | 0 | 0 | 0
784 |  | 0 |  | 0 | 0 | 0
785 |  | 0 |  | 0 | 0 | 0
786 |  | 0 |  | 0 | 0 | 0
787 |  | 0 |  | 0 | 0 | 0
788 |  | 0 |  | 0 | 0 | 0
789 |  | 0 |  | 0 | 0 | 0
790 |  | 0 |  | 0 | 0 | 0
791 |  | 0 |  | 0 | 0 | 0
792 |  | 0 |  | 0 | 0 | 0
793 |  | 0 |  | 0 | 0 | 0
794 |  | 0 |  | 0 | 0 | 0
795 |  | 0 |  | 0 | 0 | 0
796 |  | 0 |  | 0 | 0 | 0
797 |  | 0 |  | 0 | 0 | 0
798 |  | 0 |  | 0 | 0 | 0
799 |  | 0 |  | 0 | 0 | 0
800 |  | 0 |  | 0 | 0 | 0
199 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
198 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
197 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
196 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
195 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
194 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
193 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
192 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
191 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
190 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
189 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
188 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
187 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
186 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
185 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
184 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
183 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
182 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
181 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
180 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
179 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
178 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!
177 |  | 0 |  | 0 | 0 | 0 | [8 empty] | #VALUE!

## LP_254
Date | June 16 2023 Absolute Irradiance
Spectrometer | Ocean Optics HDX UV-VIS (Last Calibrated 9/6/2022)
Operator(s) | DTM
Lamp(s) | COOSPIDER 6W, 110V LP UV
λavg (200-400) | 258.47
λ peak | 254
Plotted output name | COOSPIDER 6W, 110V LP UV NEW INTERPOLATION
Below are calculations, no additional inputs needed on this tab
Update the details of the spectra measurements above
Calculated
| Sum of NEW INTERPOLATION | SUM OF NEW * Wavelength
| 2.6555 | 686.3668
Wavelength | NEW INTERPOLATION | NEW Interpolation * Wavelength
178 |  | 0
179 |  | 0
180 |  | 0
181 |  | 0
182 |  | 0
183 |  | 0
184 |  | 0
185 |  | 0
186 |  | 0
187 |  | 0
188 |  | 0
189 |  | 0
190 |  | 0
191 |  | 0
192 |  | 0
193 |  | 0
194 |  | 0
195 |  | 0
196 |  | 0
197 |  | 0
198 |  | 0
199 |  | 0
200 | 0 | 0
201 | 0 | 0
202 | 0 | 0
203 | 0 | 0
204 | 0 | 0
205 | 0 | 0
206 | 0 | 0
207 | 0 | 0
208 | 0 | 0
209 | 0 | 0
210 | 0 | 0
211 | 0 | 0
212 | 0 | 0
213 | 0 | 0
214 | 0 | 0
215 | 0 | 0
216 | 0 | 0
217 | 0 | 0
218 | 0 | 0
219 | 0 | 0
220 | 0 | 0
221 | 0 | 0
222 | 0 | 0
223 | 0 | 0
224 | 0 | 0
225 | 0 | 0
226 | 0 | 0
227 | 0 | 0
228 | 0 | 0
229 | 0 | 0
230 | 0.0016841 | 0.38734
231 | 0.0015092 | 0.34862
232 | 0.001238 | 0.28722
233 | 0.0010502 | 0.24469
234 | 0.0013836 | 0.32377
235 | 0.0013284 | 0.31218
236 | 0.0011563 | 0.27289
237 | 0.0011401 | 0.2702
238 | 0.0014046 | 0.3343
239 | 0.0013012 | 0.31099
240 | 0.001468 | 0.35231
241 | 0.0015839 | 0.38172
242 | 0.0015535 | 0.37595
243 | 0.0017261 | 0.41945
244 | 0.0021687 | 0.52917
245 | 0.0027368 | 0.6705
246 | 0.0033488 | 0.8238
247 | 0.0041234 | 1.0185
248 | 0.0057771 | 1.4327
249 | 0.0071633 | 1.7837
250 | 0.010401 | 2.6002
251 | 0.048274 | 12.1167
252 | 0.27607 | 69.5685
253 | 0.90668 | 229.3892
254 | 1 | 254
255 | 0.10573 | 26.9608
256 | 0.018787 | 4.8096
257 | 0.013202 | 3.3929
258 | 0.01076 | 2.776
259 | 0.0092456 | 2.3946
260 | 0.0081671 | 2.1235
261 | 0.0076305 | 1.9916
262 | 0.0073063 | 1.9143
263 | 0.0072792 | 1.9144
264 | 0.0076346 | 2.0155
265 | 0.0074517 | 1.9747
266 | 0.004898 | 1.3029
267 | 0.0034172 | 0.9124
268 | 0.0031787 | 0.85189
269 | 0.0030099 | 0.80968
270 | 0.0029572 | 0.79844
271 | 0.0022686 | 0.61478
272 | 0.0012311 | 0.33485
273 | 0.0010628 | 0.29015
274 | 0.0010294 | 0.28205
275 | 0.0013634 | 0.37493
276 | 0.0010872 | 0.30007
277 | 0.00080775 | 0.22375
278 | 0.00066459 | 0.18476
279 | 0.00065469 | 0.18266
280 | 0.0010812 | 0.30273
281 | 0.00094815 | 0.26643
282 | 0.00061182 | 0.17253
283 | 0.00046871 | 0.13264
284 | 0.00042416 | 0.12046
285 | 0.00041775 | 0.11906
286 | 0.00044183 | 0.12636
287 | 0.00035598 | 0.10217
288 | 0.0007088 | 0.20413
289 | 0.0015118 | 0.43692
290 | 0.00079336 | 0.23007
291 | 0.00034248 | 0.099661
292 | 0.00047913 | 0.13991
293 | 0.00042534 | 0.12462
294 | 0.00030768 | 0.090459
295 | 0.00079426 | 0.23431
296 | 0.0035535 | 1.0518
297 | 0.0032764 | 0.9731
298 | 0.00035465 | 0.10568
299 | 0.00017527 | 0.052406
300 | 0.00023377 | 0.070131
301 | 0.0013107 | 0.39453
302 | 0.002814 | 0.84982
303 | 0.00092139 | 0.27918
304 | 0.00019589 | 0.059552
305 | 0.00020657 | 0.063003
306 | 0.00017702 | 0.054167
307 | 0.00019585 | 0.060125
308 | 0.00019418 | 0.059807
309 | 0.0002132 | 0.065879
310 | 0.00023537 | 0.072965
311 | 0.0018814 | 0.5851
312 | 0.01427 | 4.4521
313 | 0.01797 | 5.6245
314 | 0.002966 | 0.93132
315 | 0.00024039 | 0.075724
316 | 0.00017634 | 0.055723
317 | 0.00017593 | 0.055771
318 | 0.00020544 | 0.06533
319 | 0.00017242 | 0.055003
320 | 0.00016338 | 0.052282
321 | 0.00017654 | 0.05667
322 | 0.00017823 | 0.057389
323 | 0.0001692 | 0.05465
324 | 0.0001029 | 0.033341
325 | 0.00011105 | 0.036091
326 | 0.00012206 | 0.039791
327 | 0.00013667 | 0.044691
328 | 0.00013288 | 0.043584
329 | 6.7644e-05 | 0.022255
330 | 8.87e-05 | 0.029271
331 | 0.00010475 | 0.034672
332 | 9.8582e-05 | 0.032729
333 | 0.00063765 | 0.21234
334 | 0.0013196 | 0.44073
335 | 0.00029357 | 0.098345
336 | 6.381e-05 | 0.02144
337 | 5.743e-05 | 0.019354
338 | 8.6171e-05 | 0.029126
339 | 9.804e-05 | 0.033236
340 | 9.0644e-05 | 0.030819
341 | 0.00011304 | 0.038546
342 | 0.00011074 | 0.037873
343 | 6.9804e-05 | 0.023943
344 | 8.3455e-05 | 0.028708
345 | 6.6646e-05 | 0.022993
346 | 8.8789e-05 | 0.030721
347 | 6.703e-05 | 0.023259
348 | 5.9403e-05 | 0.020672
349 | 8.6863e-05 | 0.030315
350 | 8.0879e-05 | 0.028308
351 | 6.6617e-05 | 0.023383
352 | 9.4329e-05 | 0.033204
353 | 6.3042e-05 | 0.022254
354 | 9.1292e-05 | 0.032317
355 | 0.0001 | 0.035502
356 | 0.00011698 | 0.041646
357 | 0.00013199 | 0.04712
358 | 0.00010328 | 0.036975
359 | 0.00013799 | 0.049537
360 | 0.00012523 | 0.045081
361 | 0.00012576 | 0.0454
362 | 0.00017383 | 0.062927
363 | 0.00033959 | 0.12327
364 | 0.0075402 | 2.7446
365 | 0.015614 | 5.699
366 | 0.0056879 | 2.0818
367 | 0.0010433 | 0.38288
368 | 0.00021817 | 0.080286
369 | 0.00018377 | 0.06781
370 | 0.00020553 | 0.076045
371 | 0.00019424 | 0.072064
372 | 0.00019328 | 0.0719
373 | 0.00022584 | 0.084239
374 | 0.0002082 | 0.077868
375 | 0.00016218 | 0.060819
376 | 0.00017473 | 0.065699
377 | 0.00017623 | 0.066441
378 | 0.00017251 | 0.065208
379 | 0.00014118 | 0.053507
380 | 0.00015397 | 0.05851
381 | 0.00017951 | 0.068393
382 | 0.00017294 | 0.066065
383 | 0.00016322 | 0.062512
384 | 0.00012315 | 0.047291
385 | 0.00015031 | 0.057869
386 | 0.00016587 | 0.064025
387 | 0.00018133 | 0.070175
388 | 0.00016286 | 0.06319
389 | 0.00018478 | 0.071879
390 | 0.00026019 | 0.10147
391 | 0.00023334 | 0.091238
392 | 0.00018759 | 0.073536
393 | 0.00018601 | 0.073101
394 | 0.0001798 | 0.07084
395 | 0.0001887 | 0.074535
396 | 0.00016207 | 0.064179
397 | 0.00018398 | 0.073038
398 | 0.00020062 | 0.079848
399 | 0.00020534 | 0.081933
400 | 0.00019311 | 0.077244
401 | 0.00020245 | 0.08118
402 | 0.00025442 | 0.10228
403 | 0.0016265 | 0.65548
404 | 0.016395 | 6.6234
405 | 0.014952 | 6.0557
406 | 0 | 0
407 | 0 | 0
408 | 0 | 0
409 | 0 | 0
410 | 0 | 0
411 | 0 | 0
412 | 0 | 0
413 | 0 | 0
414 | 0 | 0
415 | 0 | 0
416 | 0 | 0
417 | 0 | 0
418 | 0 | 0
419 | 0 | 0
420 | 0 | 0
421 | 0 | 0
422 | 0 | 0
423 | 0 | 0
424 | 0 | 0
425 | 0 | 0
426 | 0 | 0
427 | 0 | 0
428 | 0 | 0
429 | 0 | 0
430 | 0 | 0
431 | 0 | 0
432 | 0 | 0
433 | 0 | 0
434 | 0 | 0
435 | 0 | 0
436 | 0 | 0
437 | 0 | 0
438 | 0 | 0
439 | 0 | 0
440 | 0 | 0
441 | 0 | 0
442 | 0 | 0
443 | 0 | 0
444 | 0 | 0
445 | 0 | 0
446 | 0 | 0
447 | 0 | 0
448 | 0 | 0
449 | 0 | 0
450 | 0 | 0
451 | 0 | 0
452 | 0 | 0
453 | 0 | 0
454 | 0 | 0
455 | 0 | 0
456 | 0 | 0
457 | 0 | 0
458 | 0 | 0
459 | 0 | 0
460 | 0 | 0
461 | 0 | 0
462 | 0 | 0
463 | 0 | 0
464 | 0 | 0
465 | 0 | 0
466 | 0 | 0
467 | 0 | 0
468 | 0 | 0
469 | 0 | 0
470 | 0 | 0
471 | 0 | 0
472 | 0 | 0
473 | 0 | 0
474 | 0 | 0
475 | 0 | 0
476 | 0 | 0
477 | 0 | 0
478 | 0 | 0
479 | 0 | 0
480 | 0 | 0
481 | 0 | 0
482 | 0 | 0
483 | 0 | 0
484 | 0 | 0
485 | 0 | 0
486 | 0 | 0
487 | 0 | 0
488 | 0 | 0
489 | 0 | 0
490 | 0 | 0
491 | 0 | 0
492 | 0 | 0
493 | 0 | 0
494 | 0 | 0
495 | 0 | 0
496 | 0 | 0
497 | 0 | 0
498 | 0 | 0
499 | 0 | 0
500 | 0 | 0
501 | 0 | 0
502 | 0 | 0
503 | 0 | 0
504 | 0 | 0
505 | 0 | 0
506 | 0 | 0
507 | 0 | 0
508 | 0 | 0
509 | 0 | 0
510 | 0 | 0
511 | 0 | 0
512 | 0 | 0
513 | 0 | 0
514 | 0 | 0
515 | 0 | 0
516 | 0 | 0
517 | 0 | 0
518 | 0 | 0
519 | 0 | 0
520 | 0 | 0
521 | 0 | 0
522 | 0 | 0
523 | 0 | 0
524 | 0 | 0
525 | 0 | 0
526 | 0 | 0
527 | 0 | 0
528 | 0 | 0
529 | 0 | 0
530 | 0 | 0
531 | 0 | 0
532 | 0 | 0
533 | 0 | 0
534 | 0 | 0
535 | 0 | 0
536 | 0 | 0
537 | 0 | 0
538 | 0 | 0
539 | 0 | 0
540 | 0 | 0
541 | 0 | 0
542 | 0 | 0
543 | 0 | 0
544 | 0 | 0
545 | 0 | 0
546 | 0 | 0
547 | 0 | 0
548 | 0 | 0
549 | 0 | 0
550 | 0 | 0
551 | 0 | 0
552 | 0 | 0
553 | 0 | 0
554 | 0 | 0
555 | 0 | 0
556 | 0 | 0
557 | 0 | 0
558 | 0 | 0
559 | 0 | 0
560 | 0 | 0
561 | 0 | 0
562 | 0 | 0
563 | 0 | 0
564 | 0 | 0
565 | 0 | 0
566 | 0 | 0
567 | 0 | 0
568 | 0 | 0
569 | 0 | 0
570 | 0 | 0
571 | 0 | 0
572 | 0 | 0
573 | 0 | 0
574 | 0 | 0
575 | 0 | 0
576 | 0 | 0
577 | 0 | 0
578 | 0 | 0
579 | 0 | 0
580 | 0 | 0
581 | 0 | 0
582 | 0 | 0
583 | 0 | 0
584 | 0 | 0
585 | 0 | 0
586 | 0 | 0
587 | 0 | 0
588 | 0 | 0
589 | 0 | 0
590 | 0 | 0
591 | 0 | 0
592 | 0 | 0
593 | 0 | 0
594 | 0 | 0
595 | 0 | 0
596 | 0 | 0
597 | 0 | 0
598 | 0 | 0
599 | 0 | 0
600 | 0 | 0
601 | 0 | 0
602 | 0 | 0
603 | 0 | 0
604 | 0 | 0
605 | 0 | 0
606 | 0 | 0
607 | 0 | 0
608 | 0 | 0
609 | 0 | 0
610 | 0 | 0
611 | 0 | 0
612 | 0 | 0
613 | 0 | 0
614 | 0 | 0
615 | 0 | 0
616 | 0 | 0
617 | 0 | 0
618 | 0 | 0
619 | 0 | 0
620 | 0 | 0
621 | 0 | 0
622 | 0 | 0
623 | 0 | 0
624 | 0 | 0
625 | 0 | 0
626 | 0 | 0
627 | 0 | 0
628 | 0 | 0
629 | 0 | 0
630 | 0 | 0
631 | 0 | 0
632 | 0 | 0
633 | 0 | 0
634 | 0 | 0
635 | 0 | 0
636 | 0 | 0
637 | 0 | 0
638 | 0 | 0
639 | 0 | 0
640 | 0 | 0
641 | 0 | 0
642 | 0 | 0
643 | 0 | 0
644 | 0 | 0
645 | 0 | 0
646 | 0 | 0
647 | 0 | 0
648 | 0 | 0
649 | 0 | 0
650 | 0 | 0
651 | 0 | 0
652 | 0 | 0
653 | 0 | 0
654 | 0 | 0
655 | 0 | 0
656 | 0 | 0
657 | 0 | 0
658 | 0 | 0
659 | 0 | 0
660 | 0 | 0
661 | 0 | 0
662 | 0 | 0
663 | 0 | 0
664 | 0 | 0
665 | 0 | 0
666 | 0 | 0
667 | 0 | 0
668 | 0 | 0
669 | 0 | 0
670 | 0 | 0
671 | 0 | 0
672 | 0 | 0
673 | 0 | 0
674 | 0 | 0
675 | 0 | 0
676 | 0 | 0
677 | 0 | 0
678 | 0 | 0
679 | 0 | 0
680 | 0 | 0
681 | 0 | 0
682 | 0 | 0
683 | 0 | 0
684 | 0 | 0
685 | 0 | 0
686 | 0 | 0
687 | 0 | 0
688 | 0 | 0
689 | 0 | 0
690 | 0 | 0
691 | 0 | 0
692 | 0 | 0
693 | 0 | 0
694 | 0 | 0
695 | 0 | 0
696 | 0 | 0
697 | 0 | 0
698 | 0 | 0
699 | 0 | 0
700 | 0 | 0
701 | 0 | 0
702 | 0 | 0
703 | 0 | 0
704 | 0 | 0
705 | 0 | 0
706 | 0 | 0
707 | 0 | 0
708 | 0 | 0
709 | 0 | 0
710 | 0 | 0
711 | 0 | 0
712 | 0 | 0
713 | 0 | 0
714 | 0 | 0
715 | 0 | 0
716 | 0 | 0
717 | 0 | 0
718 | 0 | 0
719 | 0 | 0
720 | 0 | 0
721 | 0 | 0
722 | 0 | 0
723 | 0 | 0
724 | 0 | 0
725 | 0 | 0
726 | 0 | 0
727 | 0 | 0
728 | 0 | 0
729 | 0 | 0
730 | 0 | 0
731 | 0 | 0
732 | 0 | 0
733 | 0 | 0
734 | 0 | 0
735 | 0 | 0
736 | 0 | 0
737 | 0 | 0
738 | 0 | 0
739 | 0 | 0
740 | 0 | 0
741 | 0 | 0
742 | 0 | 0
743 | 0 | 0
744 | 0 | 0
745 | 0 | 0
746 | 0 | 0
747 | 0 | 0
748 | 0 | 0
749 | 0 | 0
750 | 0 | 0
751 | 0 | 0
752 | 0 | 0
753 | 0 | 0
754 | 0 | 0
755 | 0 | 0
756 | 0 | 0
757 | 0 | 0
758 | 0 | 0
759 | 0 | 0
760 | 0 | 0
761 | 0 | 0
762 | 0 | 0
763 | 0 | 0
764 | 0 | 0
765 | 0 | 0
766 | 0 | 0
767 | 0 | 0
768 | 0 | 0
769 | 0 | 0
770 | 0 | 0
771 | 0 | 0
772 | 0 | 0
773 | 0 | 0
774 | 0 | 0
775 | 0 | 0
776 | 0 | 0
777 | 0 | 0
778 | 0 | 0
779 | 0 | 0
780 | 0 | 0
781 | 0 | 0
782 | 0 | 0
783 | 0 | 0
784 | 0 | 0
785 | 0 | 0
786 | 0 | 0
787 | 0 | 0
788 | 0 | 0
789 | 0 | 0
790 | 0 | 0
791 | 0 | 0
792 | 0 | 0
793 | 0 | 0
794 | 0 | 0
795 | 0 | 0
796 | 0 | 0
797 | 0 | 0
798 | 0 | 0
799 | 0 | 0
800 | 0 | 0

## UV VIS Meas 500 ms, 10 scan
background
AbsoluteIrradiance:211 |  | Wavelength | background | background 2 | background 3 | 254 | 254 2 | 254 3 | background 4 | background 5 | background 6 | pre-background avg | UV avg | post-background avg | pre- and post-background avg | Wavelength | UV-Background | Relative
195.562 | 0 | 195.562 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 195.562 | 0 | 0
195.934 | 0 | 195.934 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 195.934 | 0 | 0
196.307 | 0 | 196.307 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 196.307 | 0 | 0
196.679 | 0 | 196.679 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 196.679 | 0 | 0
197.051 | 0 | 197.051 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 197.051 | 0 | 0
197.423 | 0 | 197.423 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 197.423 | 0 | 0
197.795 | 0 | 197.795 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 197.795 | 0 | 0
198.167 | 0 | 198.167 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 198.167 | 0 | 0
198.539 | 0 | 198.539 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 198.539 | 0 | 0
198.911 | 0 | 198.911 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 198.911 | 0 | 0
199.283 | 0 | 199.283 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 199.283 | 0 | 0
199.654 | 0 | 199.654 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 199.654 | 0 | 0
200.026 | 0 | 200.026 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 200.026 | 0 | 0
200.397 | 0 | 200.397 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 200.397 | 0 | 0
200.769 | 0 | 200.769 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 200.769 | 0 | 0
201.14 | 0 | 201.14 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 201.14 | 0 | 0
201.511 | 0 | 201.511 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 201.511 | 0 | 0
201.882 | 0 | 201.882 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 201.882 | 0 | 0
202.253 | 0 | 202.253 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 202.253 | 0 | 0
202.624 | 0 | 202.624 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 202.624 | 0 | 0
202.995 | 0 | 202.995 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 202.995 | 0 | 0
203.366 | 0 | 203.366 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 203.366 | 0 | 0
203.736 | 0 | 203.736 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 203.736 | 0 | 0
204.107 | 0 | 204.107 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 204.107 | 0 | 0
204.477 | 0 | 204.477 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 204.477 | 0 | 0
204.848 | 1.642 | 204.848 | 1.642 | 2.213 | 2.213 | 3.052 | 3.052 | 3.275 | 3.686 | 3.677 | 4.275 | 2.0227 | 3.1263 | 3.8793 | 2.951 | 204.848 | 0.17533 | 0.0028542
205.218 | 5.33 | 205.218 | 5.33 | 4.959 | 4.72 | 6.453 | 6.453 | 7.195 | 7.954 | 7.649 | 7.649 | 5.003 | 6.7003 | 7.7507 | 6.3768 | 205.218 | 0.3235 | 0.0052661
205.588 | -0.037 | 205.588 | -0.037 | 0.504 | 0.704 | 2.319 | 2.216 | 2.141 | 1.764 | 2.727 | 2.423 | 0.39033 | 2.2253 | 2.3047 | 1.3475 | 205.588 | 0.87783 | 0.01429
205.959 | 0.228 | 205.959 | 0.228 | 1.38 | 0.918 | 1.759 | 1.98 | 2.532 | 2.663 | 2.498 | 2.76 | 0.842 | 2.0903 | 2.6403 | 1.7412 | 205.959 | 0.34917 | 0.0056839
206.329 | 1.483 | 206.329 | 1.483 | 1.817 | 1.445 | 2.398 | 2.638 | 2.764 | 2.827 | 2.171 | 3.275 | 1.5817 | 2.6 | 2.7577 | 2.1697 | 206.329 | 0.43033 | 0.0070052
206.699 | -0.208 | 206.699 | -0.208 | 0.943 | 0.904 | 1.499 | 2.105 | 1.813 | 2.094 | 1.998 | 2.503 | 0.54633 | 1.8057 | 2.1983 | 1.3723 | 206.699 | 0.43333 | 0.007054
207.068 | 0.937 | 207.068 | 0.937 | 1.535 | 1.499 | 2.342 | 2.154 | 2.919 | 2.368 | 3.08 | 3.408 | 1.3237 | 2.4717 | 2.952 | 2.1378 | 207.068 | 0.33383 | 0.0054343
207.438 | 1.424 | 207.438 | 1.424 | 1.887 | 1.847 | 2.513 | 2.701 | 3.119 | 3.475 | 3.175 | 3.699 | 1.7193 | 2.7777 | 3.4497 | 2.5845 | 207.438 | 0.19317 | 0.0031445
207.808 | 0.169 | 207.808 | 0.169 | 0.52 | 0.776 | 1.534 | 1.378 | 1.62 | 1.935 | 2.132 | 2.068 | 0.48833 | 1.5107 | 2.045 | 1.2667 | 207.808 | 0.244 | 0.003972
208.178 | 1.159 | 208.178 | 1.159 | 1.501 | 1.533 | 2.366 | 2.253 | 2.293 | 2.45 | 2.647 | 2.643 | 1.3977 | 2.304 | 2.58 | 1.9888 | 208.178 | 0.31517 | 0.0051305
208.547 | 0.636 | 208.547 | 0.636 | 1.087 | 1.099 | 1.16 | 1.49 | 1.957 | 1.921 | 1.969 | 1.852 | 0.94067 | 1.5357 | 1.914 | 1.4273 | 208.547 | 0.10833 | 0.0017635
208.917 | 0.561 | 208.917 | 0.561 | 0.776 | 1.082 | 1.364 | 1.575 | 1.579 | 1.822 | 1.957 | 1.762 | 0.80633 | 1.506 | 1.847 | 1.3267 | 208.917 | 0.17933 | 0.0029193
209.286 | 1.727 | 209.286 | 1.727 | 2.026 | 1.799 | 2.651 | 2.826 | 3.308 | 2.977 | 2.961 | 3.236 | 1.8507 | 2.9283 | 3.058 | 2.4543 | 209.286 | 0.474 | 0.007716
209.655 | 1.586 | 209.655 | 1.586 | 1.489 | 1.597 | 2.153 | 2.283 | 2.442 | 2.637 | 2.551 | 2.724 | 1.5573 | 2.2927 | 2.6373 | 2.0973 | 209.655 | 0.19533 | 0.0031797
210.025 | 0.54 | 210.025 | 0.54 | 0.637 | 0.73 | 1.28 | 1.323 | 1.754 | 1.507 | 1.557 | 1.69 | 0.63567 | 1.4523 | 1.5847 | 1.1102 | 210.025 | 0.34217 | 0.00557
210.394 | 0.402 | 210.394 | 0.402 | 0.906 | 0.966 | 1.625 | 1.483 | 1.779 | 1.671 | 1.746 | 1.789 | 0.758 | 1.629 | 1.7353 | 1.2467 | 210.394 | 0.38233 | 0.0062238
210.763 | 1.246 | 210.763 | 1.246 | 1.424 | 1.176 | 1.935 | 1.645 | 2.038 | 2.137 | 1.995 | 2.201 | 1.282 | 1.8727 | 2.111 | 1.6965 | 210.763 | 0.17617 | 0.0028677
211.132 | -0.039 | 211.132 | -0.039 | 0.398 | 0.58 | 0.939 | 1.102 | 1.047 | 0.92 | 1.219 | 1.354 | 0.313 | 1.0293 | 1.1643 | 0.73867 | 211.132 | 0.29067 | 0.0047316
211.501 | 0.35 | 211.501 | 0.35 | 0.374 | 0.168 | 0.903 | 0.652 | 0.914 | 1.173 | 1.109 | 1.189 | 0.29733 | 0.823 | 1.157 | 0.72717 | 211.501 | 0.095833 | 0.00156
211.869 | 0.112 | 211.869 | 0.112 | 0.12 | 0.463 | 0.862 | 0.954 | 1.114 | 1.272 | 1.262 | 1.239 | 0.23167 | 0.97667 | 1.2577 | 0.74467 | 211.869 | 0.232 | 0.0037766
212.238 | 0.865 | 212.238 | 0.865 | 0.737 | 1.037 | 1.381 | 1.628 | 1.206 | 1.657 | 1.602 | 1.558 | 0.87967 | 1.405 | 1.6057 | 1.2427 | 212.238 | 0.16233 | 0.0026426
212.607 | 0.886 | 212.607 | 0.886 | 1.193 | 1.272 | 1.726 | 1.748 | 1.84 | 1.712 | 1.977 | 1.984 | 1.117 | 1.7713 | 1.891 | 1.504 | 212.607 | 0.26733 | 0.0043518
212.975 | 0.15 | 212.975 | 0.15 | 0.322 | 0.416 | 0.712 | 0.705 | 0.893 | 0.912 | 1.175 | 1.255 | 0.296 | 0.77 | 1.114 | 0.705 | 212.975 | 0.065 | 0.0010581
213.344 | 0.323 | 213.344 | 0.323 | 0.231 | 0.404 | 0.605 | 0.807 | 0.694 | 0.893 | 1.133 | 0.932 | 0.31933 | 0.702 | 0.986 | 0.65267 | 213.344 | 0.049333 | 0.00080308
213.712 | -0.067 | 213.712 | -0.067 | 0.245 | 0.438 | 0.61 | 0.673 | 0.718 | 0.823 | 0.953 | 0.872 | 0.20533 | 0.667 | 0.88267 | 0.544 | 213.712 | 0.123 | 0.0020023
214.08 | 0.432 | 214.08 | 0.432 | 0.584 | 0.599 | 0.813 | 0.763 | 1.096 | 1.156 | 1.063 | 1.282 | 0.53833 | 0.89067 | 1.167 | 0.85267 | 214.08 | 0.038 | 0.00061859
214.448 | 0.384 | 214.448 | 0.384 | 0.527 | 0.627 | 0.967 | 0.805 | 1.061 | 1.1 | 0.932 | 1.069 | 0.51267 | 0.94433 | 1.0337 | 0.77317 | 214.448 | 0.17117 | 0.0027863
214.817 | 0.433 | 214.817 | 0.433 | 0.454 | 0.531 | 0.614 | 0.663 | 0.85 | 0.802 | 0.912 | 0.974 | 0.47267 | 0.709 | 0.896 | 0.68433 | 214.817 | 0.024667 | 0.00040154
215.185 | 0.938 | 215.185 | 0.938 | 1.028 | 1.22 | 1.235 | 1.452 | 1.415 | 1.447 | 1.557 | 1.527 | 1.062 | 1.3673 | 1.5103 | 1.2862 | 215.185 | 0.081167 | 0.0013213
215.553 | 0.705 | 215.553 | 0.705 | 0.689 | 0.816 | 0.899 | 0.893 | 1.189 | 1.173 | 1.058 | 1.189 | 0.73667 | 0.99367 | 1.14 | 0.93833 | 215.553 | 0.055333 | 0.00090075
215.92 | 0.511 | 215.92 | 0.511 | 0.729 | 0.782 | 0.94 | 1.028 | 0.979 | 1.32 | 1.102 | 1.229 | 0.674 | 0.98233 | 1.217 | 0.9455 | 215.92 | 0.036833 | 0.00059959
216.288 | 0.459 | 216.288 | 0.459 | 0.585 | 0.687 | 0.83 | 0.864 | 0.991 | 0.962 | 0.974 | 1.198 | 0.577 | 0.895 | 1.0447 | 0.81083 | 216.288 | 0.084167 | 0.0013701
216.656 | 0.479 | 216.656 | 0.479 | 0.7 | 0.747 | 0.97 | 0.995 | 1.142 | 1.076 | 1.08 | 1.363 | 0.642 | 1.0357 | 1.173 | 0.9075 | 216.656 | 0.12817 | 0.0020864
217.023 | 0.311 | 217.023 | 0.311 | 0.35 | 0.412 | 0.677 | 0.733 | 0.766 | 0.717 | 0.722 | 0.811 | 0.35767 | 0.72533 | 0.75 | 0.55383 | 217.023 | 0.1715 | 0.0027918
217.391 | 0.158 | 217.391 | 0.158 | 0.335 | 0.35 | 0.714 | 0.457 | 0.661 | 0.557 | 0.702 | 0.745 | 0.281 | 0.61067 | 0.668 | 0.4745 | 217.391 | 0.13617 | 0.0022166
217.758 | 0.582 | 217.758 | 0.582 | 0.631 | 0.897 | 0.983 | 0.948 | 1.178 | 1.21 | 1.265 | 1.185 | 0.70333 | 1.0363 | 1.22 | 0.96167 | 217.758 | 0.074667 | 0.0012155
218.126 | 0.378 | 218.126 | 0.378 | 0.556 | 0.55 | 0.77 | 0.846 | 0.793 | 0.921 | 0.958 | 0.995 | 0.49467 | 0.803 | 0.958 | 0.72633 | 218.126 | 0.076667 | 0.001248
218.493 | 0.322 | 218.493 | 0.322 | 0.392 | 0.475 | 0.615 | 0.83 | 0.796 | 0.844 | 0.842 | 0.873 | 0.39633 | 0.747 | 0.853 | 0.62467 | 218.493 | 0.12233 | 0.0019914
218.86 | 0.336 | 218.86 | 0.336 | 0.394 | 0.529 | 0.666 | 0.738 | 0.697 | 0.768 | 0.694 | 0.715 | 0.41967 | 0.70033 | 0.72567 | 0.57267 | 218.86 | 0.12767 | 0.0020782
219.227 | 0.071 | 219.227 | 0.071 | 0.22 | 0.228 | 0.406 | 0.437 | 0.577 | 0.543 | 0.618 | 0.755 | 0.173 | 0.47333 | 0.63867 | 0.40583 | 219.227 | 0.0675 | 0.0010988
219.594 | 0.101 | 219.594 | 0.101 | 0.24 | 0.236 | 0.463 | 0.59 | 0.575 | 0.458 | 0.581 | 0.49 | 0.19233 | 0.54267 | 0.50967 | 0.351 | 219.594 | 0.19167 | 0.0031201
219.961 | 0.478 | 219.961 | 0.478 | 0.647 | 0.57 | 0.833 | 0.849 | 0.92 | 0.892 | 0.953 | 0.961 | 0.565 | 0.86733 | 0.93533 | 0.75017 | 219.961 | 0.11717 | 0.0019073
220.328 | 0.198 | 220.328 | 0.198 | 0.198 | 0.334 | 0.429 | 0.47 | 0.584 | 0.479 | 0.529 | 0.579 | 0.24333 | 0.49433 | 0.529 | 0.38617 | 220.328 | 0.10817 | 0.0017608
220.695 | 0.298 | 220.695 | 0.298 | 0.395 | 0.511 | 0.595 | 0.533 | 0.714 | 0.675 | 0.658 | 0.808 | 0.40133 | 0.614 | 0.71367 | 0.5575 | 220.695 | 0.0565 | 0.00091974
221.061 | 0.428 | 221.061 | 0.428 | 0.544 | 0.545 | 0.754 | 0.701 | 0.869 | 0.839 | 0.843 | 0.9 | 0.50567 | 0.77467 | 0.86067 | 0.68317 | 221.061 | 0.0915 | 0.0014895
221.428 | 0.31 | 221.428 | 0.31 | 0.413 | 0.47 | 0.647 | 0.682 | 0.715 | 0.717 | 0.807 | 0.91 | 0.39767 | 0.68133 | 0.81133 | 0.6045 | 221.428 | 0.076833 | 0.0012507
221.795 | 0.469 | 221.795 | 0.469 | 0.532 | 0.615 | 0.73 | 0.72 | 0.789 | 0.735 | 0.897 | 0.83 | 0.53867 | 0.74633 | 0.82067 | 0.67967 | 221.795 | 0.066667 | 0.0010852
222.161 | 0.514 | 222.161 | 0.514 | 0.58 | 0.553 | 0.884 | 0.775 | 0.949 | 0.756 | 1.032 | 0.895 | 0.549 | 0.86933 | 0.89433 | 0.72167 | 222.161 | 0.14767 | 0.0024038
222.527 | 0.131 | 222.527 | 0.131 | 0.271 | 0.295 | 0.324 | 0.4 | 0.486 | 0.547 | 0.561 | 0.612 | 0.23233 | 0.40333 | 0.57333 | 0.40283 | 222.527 | 0.0005 | 8.1393e-06
222.893 | 0.442 | 222.893 | 0.442 | 0.589 | 0.552 | 0.808 | 0.869 | 0.848 | 0.886 | 0.93 | 0.889 | 0.52767 | 0.84167 | 0.90167 | 0.71467 | 222.893 | 0.127 | 0.0020674
223.26 | 0.003 | 223.26 | 0.003 | 0.161 | 0.123 | 0.329 | 0.247 | 0.305 | 0.287 | 0.473 | 0.422 | 0.095667 | 0.29367 | 0.394 | 0.24483 | 223.26 | 0.048833 | 0.00079494
223.626 | 0.499 | 223.626 | 0.499 | 0.646 | 0.707 | 0.755 | 0.789 | 0.88 | 0.859 | 0.872 | 0.91 | 0.61733 | 0.808 | 0.88033 | 0.74883 | 223.626 | 0.059167 | 0.00096315
223.992 | 0.304 | 223.992 | 0.304 | 0.377 | 0.404 | 0.522 | 0.579 | 0.59 | 0.547 | 0.615 | 0.599 | 0.36167 | 0.56367 | 0.587 | 0.47433 | 223.992 | 0.089333 | 0.0014542
224.358 | 0.114 | 224.358 | 0.114 | 0.275 | 0.219 | 0.396 | 0.4 | 0.539 | 0.534 | 0.524 | 0.452 | 0.20267 | 0.445 | 0.50333 | 0.353 | 224.358 | 0.092 | 0.0014976
224.723 | 0.419 | 224.723 | 0.419 | 0.481 | 0.462 | 0.616 | 0.633 | 0.577 | 0.648 | 0.685 | 0.653 | 0.454 | 0.60867 | 0.662 | 0.558 | 224.723 | 0.050667 | 0.00082478
225.089 | -0.032 | 225.089 | -0.032 | 0.014 | 0.025 | 0.197 | 0.216 | 0.27 | 0.219 | 0.256 | 0.205 | 0.0023333 | 0.22767 | 0.22667 | 0.1145 | 225.089 | 0.11317 | 0.0018422
225.455 | 0.294 | 225.455 | 0.294 | 0.329 | 0.353 | 0.469 | 0.522 | 0.548 | 0.485 | 0.568 | 0.568 | 0.32533 | 0.513 | 0.54033 | 0.43283 | 225.455 | 0.080167 | 0.001305
225.82 | 0.112 | 225.82 | 0.112 | 0.207 | 0.213 | 0.354 | 0.428 | 0.375 | 0.359 | 0.423 | 0.443 | 0.17733 | 0.38567 | 0.40833 | 0.29283 | 225.82 | 0.092833 | 0.0015112
226.186 | 0.074 | 226.186 | 0.074 | 0.179 | 0.106 | 0.268 | 0.309 | 0.31 | 0.234 | 0.345 | 0.371 | 0.11967 | 0.29567 | 0.31667 | 0.21817 | 226.186 | 0.0775 | 0.0012616
226.551 | -0.041 | 226.551 | -0.041 | 0.014 | 0.04 | 0.131 | 0.136 | 0.16 | 0.143 | 0.239 | 0.176 | 0.0043333 | 0.14233 | 0.186 | 0.095167 | 226.551 | 0.047167 | 0.00076781
226.917 | 0.074 | 226.917 | 0.074 | 0.134 | 0.179 | 0.248 | 0.289 | 0.339 | 0.327 | 0.338 | 0.413 | 0.129 | 0.292 | 0.35933 | 0.24417 | 226.917 | 0.047833 | 0.00077866
227.282 | 0.147 | 227.282 | 0.147 | 0.145 | 0.246 | 0.352 | 0.334 | 0.33 | 0.343 | 0.301 | 0.392 | 0.17933 | 0.33867 | 0.34533 | 0.26233 | 227.282 | 0.076333 | 0.0012426
227.647 | 0.32 | 227.647 | 0.32 | 0.326 | 0.332 | 0.469 | 0.47 | 0.487 | 0.483 | 0.557 | 0.525 | 0.326 | 0.47533 | 0.52167 | 0.42383 | 227.647 | 0.0515 | 0.00083835
228.012 | 0.204 | 228.012 | 0.204 | 0.229 | 0.249 | 0.356 | 0.339 | 0.366 | 0.385 | 0.425 | 0.391 | 0.22733 | 0.35367 | 0.40033 | 0.31383 | 228.012 | 0.039833 | 0.00064843
228.377 | 0.124 | 228.377 | 0.124 | 0.175 | 0.185 | 0.293 | 0.273 | 0.3 | 0.296 | 0.314 | 0.358 | 0.16133 | 0.28867 | 0.32267 | 0.242 | 228.377 | 0.046667 | 0.00075967
228.742 | 0.163 | 228.742 | 0.163 | 0.205 | 0.234 | 0.318 | 0.306 | 0.338 | 0.364 | 0.348 | 0.319 | 0.20067 | 0.32067 | 0.34367 | 0.27217 | 228.742 | 0.0485 | 0.00078951
229.107 | 0.044 | 229.107 | 0.044 | 0.145 | 0.094 | 0.211 | 0.24 | 0.203 | 0.203 | 0.248 | 0.22 | 0.094333 | 0.218 | 0.22367 | 0.159 | 229.107 | 0.059 | 0.00096043
229.471 | 0.189 | 229.471 | 0.189 | 0.243 | 0.228 | 0.287 | 0.304 | 0.322 | 0.28 | 0.309 | 0.311 | 0.22 | 0.30433 | 0.3 | 0.26 | 229.471 | 0.044333 | 0.00072168
229.836 | 0.066 | 229.836 | 0.066 | 0.11 | 0.107 | 0.198 | 0.217 | 0.235 | 0.174 | 0.217 | 0.215 | 0.094333 | 0.21667 | 0.202 | 0.14817 | 229.836 | 0.0685 | 0.0011151
230.201 | 0.188 | 230.201 | 0.188 | 0.193 | 0.222 | 0.266 | 0.291 | 0.286 | 0.304 | 0.292 | 0.34 | 0.201 | 0.281 | 0.312 | 0.2565 | 230.201 | 0.0245 | 0.00039882
230.565 | 0.149 | 230.565 | 0.149 | 0.2 | 0.199 | 0.305 | 0.325 | 0.278 | 0.302 | 0.291 | 0.319 | 0.18267 | 0.30267 | 0.304 | 0.24333 | 230.565 | 0.059333 | 0.00096586
230.929 | 0.11 | 230.929 | 0.11 | 0.113 | 0.165 | 0.225 | 0.231 | 0.212 | 0.244 | 0.257 | 0.211 | 0.12933 | 0.22267 | 0.23733 | 0.18333 | 230.929 | 0.039333 | 0.00064029
231.294 | 0.04 | 231.294 | 0.04 | 0.044 | 0.066 | 0.159 | 0.173 | 0.144 | 0.139 | 0.181 | 0.186 | 0.05 | 0.15867 | 0.16867 | 0.10933 | 231.294 | 0.049333 | 0.00080308
231.658 | 0.042 | 231.658 | 0.042 | 0.094 | 0.103 | 0.128 | 0.163 | 0.201 | 0.154 | 0.181 | 0.196 | 0.079667 | 0.164 | 0.177 | 0.12833 | 231.658 | 0.035667 | 0.0005806
232.022 | 0.07 | 232.022 | 0.07 | 0.099 | 0.079 | 0.168 | 0.169 | 0.161 | 0.168 | 0.188 | 0.175 | 0.082667 | 0.166 | 0.177 | 0.12983 | 232.022 | 0.036167 | 0.00058874
232.386 | 0.024 | 232.386 | 0.024 | 0.052 | 0.07 | 0.117 | 0.106 | 0.144 | 0.12 | 0.116 | 0.129 | 0.048667 | 0.12233 | 0.12167 | 0.085167 | 232.386 | 0.037167 | 0.00060502
232.75 | -0.11 | 232.75 | -0.11 | -0.099 | -0.069 | -0.053 | -0.04 | -0.005 | -0.028 | -0.016 | -0.018 | -0.092667 | -0.032667 | -0.020667 | -0.056667 | 232.75 | 0.024 | 0.00039069
233.114 | 0.048 | 233.114 | 0.048 | 0.055 | 0.08 | 0.122 | 0.1 | 0.161 | 0.141 | 0.128 | 0.161 | 0.061 | 0.12767 | 0.14333 | 0.10217 | 233.114 | 0.0255 | 0.0004151
233.478 | 0.085 | 233.478 | 0.085 | 0.084 | 0.092 | 0.137 | 0.153 | 0.161 | 0.143 | 0.161 | 0.159 | 0.087 | 0.15033 | 0.15433 | 0.12067 | 233.478 | 0.029667 | 0.00048293
233.842 | 0.088 | 233.842 | 0.088 | 0.101 | 0.118 | 0.179 | 0.191 | 0.183 | 0.145 | 0.175 | 0.186 | 0.10233 | 0.18433 | 0.16867 | 0.1355 | 233.842 | 0.048833 | 0.00079494
234.205 | 0.095 | 234.205 | 0.095 | 0.116 | 0.119 | 0.18 | 0.174 | 0.189 | 0.162 | 0.174 | 0.178 | 0.11 | 0.181 | 0.17133 | 0.14067 | 234.205 | 0.040333 | 0.00065657
234.569 | 0.048 | 234.569 | 0.048 | 0.056 | 0.056 | 0.127 | 0.149 | 0.137 | 0.125 | 0.135 | 0.122 | 0.053333 | 0.13767 | 0.12733 | 0.090333 | 234.569 | 0.047333 | 0.00077052
234.932 | 0.026 | 234.932 | 0.026 | 0.048 | 0.049 | 0.096 | 0.119 | 0.104 | 0.098 | 0.09 | 0.113 | 0.041 | 0.10633 | 0.10033 | 0.070667 | 234.932 | 0.035667 | 0.0005806
235.296 | 0.031 | 235.296 | 0.031 | 0.032 | 0.036 | 0.075 | 0.104 | 0.112 | 0.085 | 0.095 | 0.093 | 0.033 | 0.097 | 0.091 | 0.062 | 235.296 | 0.035 | 0.00056975
235.659 | 0.051 | 235.659 | 0.051 | 0.06 | 0.075 | 0.135 | 0.122 | 0.133 | 0.117 | 0.14 | 0.126 | 0.062 | 0.13 | 0.12767 | 0.094833 | 235.659 | 0.035167 | 0.00057246
236.022 | 0.037 | 236.022 | 0.037 | 0.045 | 0.064 | 0.103 | 0.104 | 0.122 | 0.077 | 0.106 | 0.106 | 0.048667 | 0.10967 | 0.096333 | 0.0725 | 236.022 | 0.037167 | 0.00060502
236.385 | 0.025 | 236.385 | 0.025 | 0.039 | 0.033 | 0.079 | 0.091 | 0.097 | 0.083 | 0.089 | 0.099 | 0.032333 | 0.089 | 0.090333 | 0.061333 | 236.385 | 0.027667 | 0.00045037
236.748 | 0.072 | 236.748 | 0.072 | 0.088 | 0.111 | 0.134 | 0.168 | 0.15 | 0.136 | 0.137 | 0.161 | 0.090333 | 0.15067 | 0.14467 | 0.1175 | 236.748 | 0.033167 | 0.00053991
237.111 | 0.003 | 237.111 | 0.003 | 0.016 | 0.019 | 0.067 | 0.072 | 0.084 | 0.074 | 0.08 | 0.086 | 0.012667 | 0.074333 | 0.08 | 0.046333 | 237.111 | 0.028 | 0.0004558
237.474 | 0.078 | 237.474 | 0.078 | 0.096 | 0.116 | 0.172 | 0.158 | 0.163 | 0.152 | 0.165 | 0.173 | 0.096667 | 0.16433 | 0.16333 | 0.13 | 237.474 | 0.034333 | 0.0005589
237.837 | 0.068 | 237.837 | 0.068 | 0.083 | 0.092 | 0.165 | 0.17 | 0.157 | 0.124 | 0.156 | 0.165 | 0.081 | 0.164 | 0.14833 | 0.11467 | 237.837 | 0.049333 | 0.00080308
238.2 | 0.046 | 238.2 | 0.046 | 0.056 | 0.073 | 0.124 | 0.128 | 0.124 | 0.096 | 0.107 | 0.114 | 0.058333 | 0.12533 | 0.10567 | 0.082 | 238.2 | 0.043333 | 0.0007054
238.562 | 0.039 | 238.562 | 0.039 | 0.048 | 0.05 | 0.105 | 0.105 | 0.121 | 0.103 | 0.101 | 0.103 | 0.045667 | 0.11033 | 0.10233 | 0.074 | 238.562 | 0.036333 | 0.00059145
238.925 | 0.098 | 238.925 | 0.098 | 0.109 | 0.085 | 0.163 | 0.173 | 0.183 | 0.155 | 0.157 | 0.158 | 0.097333 | 0.173 | 0.15667 | 0.127 | 238.925 | 0.046 | 0.00074881
239.287 | 0.076 | 239.287 | 0.076 | 0.077 | 0.098 | 0.133 | 0.155 | 0.139 | 0.131 | 0.15 | 0.156 | 0.083667 | 0.14233 | 0.14567 | 0.11467 | 239.287 | 0.027667 | 0.00045037
239.65 | 0.019 | 239.65 | 0.019 | 0.045 | 0.038 | 0.082 | 0.088 | 0.11 | 0.082 | 0.084 | 0.077 | 0.034 | 0.093333 | 0.081 | 0.0575 | 239.65 | 0.035833 | 0.00058332
240.012 | 0.065 | 240.012 | 0.065 | 0.088 | 0.096 | 0.151 | 0.15 | 0.161 | 0.119 | 0.135 | 0.129 | 0.083 | 0.154 | 0.12767 | 0.10533 | 240.012 | 0.048667 | 0.00079222
240.374 | 0.052 | 240.374 | 0.052 | 0.052 | 0.07 | 0.124 | 0.149 | 0.14 | 0.111 | 0.104 | 0.12 | 0.058 | 0.13767 | 0.11167 | 0.084833 | 240.374 | 0.052833 | 0.00086005
240.736 | 0.076 | 240.736 | 0.076 | 0.097 | 0.083 | 0.15 | 0.153 | 0.149 | 0.12 | 0.134 | 0.139 | 0.085333 | 0.15067 | 0.131 | 0.10817 | 240.736 | 0.0425 | 0.00069184
241.098 | 0.04 | 241.098 | 0.04 | 0.04 | 0.06 | 0.122 | 0.136 | 0.125 | 0.107 | 0.108 | 0.12 | 0.046667 | 0.12767 | 0.11167 | 0.079167 | 241.098 | 0.0485 | 0.00078951
241.46 | 0.052 | 241.46 | 0.052 | 0.06 | 0.103 | 0.125 | 0.139 | 0.153 | 0.115 | 0.125 | 0.119 | 0.071667 | 0.139 | 0.11967 | 0.095667 | 241.46 | 0.043333 | 0.0007054
241.822 | 0.084 | 241.822 | 0.084 | 0.095 | 0.095 | 0.156 | 0.164 | 0.172 | 0.133 | 0.139 | 0.162 | 0.091333 | 0.164 | 0.14467 | 0.118 | 241.822 | 0.046 | 0.00074881
242.184 | 0.002 | 242.184 | 0.002 | 0.013 | 0.036 | 0.071 | 0.082 | 0.083 | 0.04 | 0.045 | 0.068 | 0.017 | 0.078667 | 0.051 | 0.034 | 242.184 | 0.044667 | 0.00072711
242.546 | 0.062 | 242.546 | 0.062 | 0.091 | 0.071 | 0.138 | 0.133 | 0.146 | 0.103 | 0.121 | 0.125 | 0.074667 | 0.139 | 0.11633 | 0.0955 | 242.546 | 0.0435 | 0.00070812
242.907 | 0.033 | 242.907 | 0.033 | 0.038 | 0.053 | 0.101 | 0.128 | 0.122 | 0.083 | 0.084 | 0.09 | 0.041333 | 0.117 | 0.085667 | 0.0635 | 242.907 | 0.0535 | 0.0008709
243.269 | 0.004 | 243.269 | 0.004 | 0.033 | 0.028 | 0.087 | 0.088 | 0.093 | 0.059 | 0.057 | 0.078 | 0.021667 | 0.089333 | 0.064667 | 0.043167 | 243.269 | 0.046167 | 0.00075153
243.63 | 0 | 243.63 | 0 | 0.007 | 0.014 | 0.088 | 0.088 | 0.098 | 0.047 | 0.058 | 0.054 | 0.007 | 0.091333 | 0.053 | 0.03 | 243.63 | 0.061333 | 0.00099842
243.992 | 0.062 | 243.992 | 0.062 | 0.082 | 0.084 | 0.152 | 0.154 | 0.153 | 0.112 | 0.119 | 0.129 | 0.076 | 0.153 | 0.12 | 0.098 | 243.992 | 0.055 | 0.00089532
244.353 | 0.047 | 244.353 | 0.047 | 0.046 | 0.05 | 0.142 | 0.144 | 0.146 | 0.087 | 0.097 | 0.091 | 0.047667 | 0.144 | 0.091667 | 0.069667 | 244.353 | 0.074333 | 0.00121
244.714 | 0.016 | 244.714 | 0.016 | 0.042 | 0.047 | 0.132 | 0.136 | 0.124 | 0.072 | 0.078 | 0.081 | 0.035 | 0.13067 | 0.077 | 0.056 | 244.714 | 0.074667 | 0.0012155
245.075 | 0.05 | 245.075 | 0.05 | 0.042 | 0.067 | 0.156 | 0.141 | 0.164 | 0.087 | 0.084 | 0.114 | 0.053 | 0.15367 | 0.095 | 0.074 | 245.075 | 0.079667 | 0.0012969
245.436 | 0.057 | 245.436 | 0.057 | 0.053 | 0.06 | 0.155 | 0.156 | 0.175 | 0.096 | 0.095 | 0.098 | 0.056667 | 0.162 | 0.096333 | 0.0765 | 245.436 | 0.0855 | 0.0013918
245.797 | 0.01 | 245.797 | 0.01 | 0.028 | 0.032 | 0.138 | 0.138 | 0.147 | 0.055 | 0.076 | 0.079 | 0.023333 | 0.141 | 0.07 | 0.046667 | 245.797 | 0.094333 | 0.0015356
246.158 | -0.009 | 246.158 | -0.009 | -0.009 | -0.007 | 0.113 | 0.105 | 0.108 | 0.011 | 0.03 | 0.023 | -0.0083333 | 0.10867 | 0.021333 | 0.0065 | 246.158 | 0.10217 | 0.0016631
246.519 | 0.037 | 246.519 | 0.037 | 0.066 | 0.058 | 0.178 | 0.182 | 0.188 | 0.09 | 0.105 | 0.095 | 0.053667 | 0.18267 | 0.096667 | 0.075167 | 246.519 | 0.1075 | 0.0017499
246.879 | 0.065 | 246.879 | 0.065 | 0.069 | 0.074 | 0.183 | 0.205 | 0.205 | 0.103 | 0.095 | 0.109 | 0.069333 | 0.19767 | 0.10233 | 0.085833 | 246.879 | 0.11183 | 0.0018205
247.24 | 0.039 | 247.24 | 0.039 | 0.041 | 0.045 | 0.17 | 0.183 | 0.2 | 0.088 | 0.082 | 0.101 | 0.041667 | 0.18433 | 0.090333 | 0.066 | 247.24 | 0.11833 | 0.0019263
247.601 | 0.044 | 247.601 | 0.044 | 0.055 | 0.064 | 0.218 | 0.217 | 0.22 | 0.086 | 0.098 | 0.091 | 0.054333 | 0.21833 | 0.091667 | 0.073 | 247.601 | 0.14533 | 0.0023658
247.961 | 0.054 | 247.961 | 0.054 | 0.063 | 0.062 | 0.246 | 0.259 | 0.252 | 0.092 | 0.082 | 0.106 | 0.059667 | 0.25233 | 0.093333 | 0.0765 | 247.961 | 0.17583 | 0.0028623
248.321 | 0.004 | 248.321 | 0.004 | 0.001 | 0.011 | 0.21 | 0.212 | 0.235 | 0.045 | 0.046 | 0.057 | 0.0053333 | 0.219 | 0.049333 | 0.027333 | 248.321 | 0.19167 | 0.0031201
248.682 | 0.022 | 248.682 | 0.022 | 0.029 | 0.034 | 0.248 | 0.242 | 0.256 | 0.075 | 0.081 | 0.076 | 0.028333 | 0.24867 | 0.077333 | 0.052833 | 248.682 | 0.19583 | 0.0031879
249.042 | 0.015 | 249.042 | 0.015 | 0.018 | 0.028 | 0.224 | 0.242 | 0.234 | 0.05 | 0.062 | 0.063 | 0.020333 | 0.23333 | 0.058333 | 0.039333 | 249.042 | 0.194 | 0.003158
249.402 | 0.019 | 249.402 | 0.019 | 0.035 | 0.02 | 0.266 | 0.266 | 0.277 | 0.063 | 0.061 | 0.079 | 0.024667 | 0.26967 | 0.067667 | 0.046167 | 249.402 | 0.2235 | 0.0036383
249.762 | 0.073 | 249.762 | 0.073 | 0.074 | 0.077 | 0.343 | 0.355 | 0.353 | 0.107 | 0.113 | 0.111 | 0.074667 | 0.35033 | 0.11033 | 0.0925 | 249.762 | 0.25783 | 0.0041972
250.122 | 0.012 | 250.122 | 0.012 | 0.03 | 0.034 | 0.352 | 0.348 | 0.359 | 0.061 | 0.073 | 0.068 | 0.025333 | 0.353 | 0.067333 | 0.046333 | 250.122 | 0.30667 | 0.0049921
250.482 | 0.071 | 250.482 | 0.071 | 0.072 | 0.075 | 0.459 | 0.468 | 0.47 | 0.112 | 0.105 | 0.124 | 0.072667 | 0.46567 | 0.11367 | 0.093167 | 250.482 | 0.3725 | 0.0060638
250.842 | 0.014 | 250.842 | 0.014 | 0.034 | 0.028 | 0.502 | 0.507 | 0.51 | 0.058 | 0.057 | 0.064 | 0.025333 | 0.50633 | 0.059667 | 0.0425 | 250.842 | 0.46383 | 0.0075505
251.201 | 0.026 | 251.201 | 0.026 | 0.043 | 0.033 | 0.824 | 0.847 | 0.84 | 0.071 | 0.082 | 0.082 | 0.034 | 0.837 | 0.078333 | 0.056167 | 251.201 | 0.78083 | 0.012711
251.561 | 0.035 | 251.561 | 0.035 | 0.043 | 0.051 | 2.84 | 2.818 | 2.823 | 0.078 | 0.073 | 0.065 | 0.043 | 2.827 | 0.072 | 0.0575 | 251.561 | 2.7695 | 0.045083
251.921 | 0.036 | 251.921 | 0.036 | 0.042 | 0.05 | 8.632 | 8.547 | 8.569 | 0.079 | 0.081 | 0.091 | 0.042667 | 8.5827 | 0.083667 | 0.063167 | 251.921 | 8.5195 | 0.13869
252.28 | 0.038 | 252.28 | 0.038 | 0.046 | 0.057 | 11.274 | 11.174 | 11.224 | 0.08 | 0.084 | 0.086 | 0.047 | 11.224 | 0.083333 | 0.065167 | 252.28 | 11.1588 | 0.18165
252.639 | 0.039 | 252.639 | 0.039 | 0.045 | 0.039 | 13.195 | 13.064 | 13.037 | 0.082 | 0.078 | 0.072 | 0.041 | 13.0987 | 0.077333 | 0.059167 | 252.639 | 13.0395 | 0.21226
252.999 | 0.041 | 252.999 | 0.041 | 0.051 | 0.054 | 18.092 | 17.959 | 18.001 | 0.076 | 0.087 | 0.08 | 0.048667 | 18.0173 | 0.081 | 0.064833 | 252.999 | 17.9525 | 0.29224
253.358 | 0.031 | 253.358 | 0.031 | 0.035 | 0.038 | 33.694 | 33.301 | 33.415 | 0.068 | 0.077 | 0.077 | 0.034667 | 33.47 | 0.074 | 0.054333 | 253.358 | 33.4157 | 0.54396
253.717 | 0.036 | 253.717 | 0.036 | 0.037 | 0.033 | 61.873 | 61.263 | 61.318 | 0.067 | 0.082 | 0.07 | 0.035333 | 61.4847 | 0.073 | 0.054167 | 253.717 | 61.4305 | 1
254.076 | 0.035 | 254.076 | 0.035 | 0.044 | 0.038 | 28.301 | 28.02 | 28.012 | 0.071 | 0.077 | 0.08 | 0.039 | 28.111 | 0.076 | 0.0575 | 254.076 | 28.0535 | 0.45667
254.435 | 0.009 | 254.435 | 0.009 | 0.027 | 0.026 | 3.729 | 3.696 | 3.684 | 0.045 | 0.052 | 0.06 | 0.020667 | 3.703 | 0.052333 | 0.0365 | 254.435 | 3.6665 | 0.059685
254.794 | 0.002 | 254.794 | 0.002 | 0.005 | 0.003 | 1.215 | 1.235 | 1.216 | 0.045 | 0.031 | 0.034 | 0.0033333 | 1.222 | 0.036667 | 0.02 | 254.794 | 1.202 | 0.019567
255.153 | 0.059 | 255.153 | 0.059 | 0.072 | 0.066 | 0.867 | 0.861 | 0.856 | 0.112 | 0.114 | 0.108 | 0.065667 | 0.86133 | 0.11133 | 0.0885 | 255.153 | 0.77283 | 0.012581
255.511 | 0.032 | 255.511 | 0.032 | 0.039 | 0.047 | 0.723 | 0.715 | 0.713 | 0.072 | 0.087 | 0.086 | 0.039333 | 0.717 | 0.081667 | 0.0605 | 255.511 | 0.6565 | 0.010687
255.87 | -0.001 | 255.87 | -0.001 | 0 | 0.017 | 0.623 | 0.606 | 0.589 | 0.035 | 0.04 | 0.037 | 0.0053333 | 0.606 | 0.037333 | 0.021333 | 255.87 | 0.58467 | 0.0095175
256.228 | 0.033 | 256.228 | 0.033 | 0.037 | 0.044 | 0.538 | 0.543 | 0.541 | 0.073 | 0.078 | 0.075 | 0.038 | 0.54067 | 0.075333 | 0.056667 | 256.228 | 0.484 | 0.0078788
256.587 | 0.026 | 256.587 | 0.026 | 0.03 | 0.034 | 0.474 | 0.465 | 0.485 | 0.075 | 0.069 | 0.078 | 0.03 | 0.47467 | 0.074 | 0.052 | 256.587 | 0.42267 | 0.0068804
256.945 | 0.025 | 256.945 | 0.025 | 0.033 | 0.025 | 0.415 | 0.416 | 0.404 | 0.057 | 0.066 | 0.061 | 0.027667 | 0.41167 | 0.061333 | 0.0445 | 256.945 | 0.36717 | 0.0059769
257.304 | -0.004 | 257.304 | -0.004 | 0 | 0.008 | 0.366 | 0.354 | 0.377 | 0.024 | 0.04 | 0.032 | 0.0013333 | 0.36567 | 0.032 | 0.016667 | 257.304 | 0.349 | 0.0056812
257.662 | 0.011 | 257.662 | 0.011 | 0.014 | 0.032 | 0.374 | 0.387 | 0.389 | 0.054 | 0.057 | 0.067 | 0.019 | 0.38333 | 0.059333 | 0.039167 | 257.662 | 0.34417 | 0.0056025
258.02 | 0.029 | 258.02 | 0.029 | 0.033 | 0.032 | 0.35 | 0.344 | 0.372 | 0.055 | 0.064 | 0.062 | 0.031333 | 0.35533 | 0.060333 | 0.045833 | 258.02 | 0.3095 | 0.0050382
258.378 | 0.038 | 258.378 | 0.038 | 0.049 | 0.048 | 0.35 | 0.342 | 0.355 | 0.072 | 0.078 | 0.08 | 0.045 | 0.349 | 0.076667 | 0.060833 | 258.378 | 0.28817 | 0.0046909
258.736 | 0.005 | 258.736 | 0.005 | 0.002 | 0.015 | 0.288 | 0.304 | 0.301 | 0.036 | 0.041 | 0.048 | 0.0073333 | 0.29767 | 0.041667 | 0.0245 | 258.736 | 0.27317 | 0.0044468
259.094 | 0.071 | 259.094 | 0.071 | 0.072 | 0.068 | 0.36 | 0.364 | 0.369 | 0.103 | 0.106 | 0.112 | 0.070333 | 0.36433 | 0.107 | 0.088667 | 259.094 | 0.27567 | 0.0044875
259.452 | 0.064 | 259.452 | 0.064 | 0.06 | 0.047 | 0.333 | 0.329 | 0.336 | 0.085 | 0.099 | 0.1 | 0.057 | 0.33267 | 0.094667 | 0.075833 | 259.452 | 0.25683 | 0.0041809
259.809 | 0.033 | 259.809 | 0.033 | 0.041 | 0.04 | 0.283 | 0.295 | 0.292 | 0.07 | 0.07 | 0.066 | 0.038 | 0.29 | 0.068667 | 0.053333 | 259.809 | 0.23667 | 0.0038526
260.167 | 0.026 | 260.167 | 0.026 | 0.044 | 0.046 | 0.282 | 0.29 | 0.284 | 0.08 | 0.069 | 0.081 | 0.038667 | 0.28533 | 0.076667 | 0.057667 | 260.167 | 0.22767 | 0.0037061
260.524 | 0.011 | 260.524 | 0.011 | 0.026 | 0.018 | 0.255 | 0.278 | 0.276 | 0.049 | 0.06 | 0.071 | 0.018333 | 0.26967 | 0.06 | 0.039167 | 260.524 | 0.2305 | 0.0037522
260.882 | -0.009 | 260.882 | -0.009 | 0.001 | 0.016 | 0.235 | 0.244 | 0.244 | 0.04 | 0.045 | 0.041 | 0.0026667 | 0.241 | 0.042 | 0.022333 | 260.882 | 0.21867 | 0.0035596
261.239 | 0.015 | 261.239 | 0.015 | 0.023 | 0.029 | 0.279 | 0.265 | 0.266 | 0.052 | 0.065 | 0.057 | 0.022333 | 0.27 | 0.058 | 0.040167 | 261.239 | 0.22983 | 0.0037414
261.597 | 0.013 | 261.597 | 0.013 | 0.029 | 0.034 | 0.248 | 0.25 | 0.251 | 0.048 | 0.055 | 0.061 | 0.025333 | 0.24967 | 0.054667 | 0.04 | 261.597 | 0.20967 | 0.0034131
261.954 | 0.006 | 261.954 | 0.006 | 0.017 | 0.026 | 0.239 | 0.243 | 0.248 | 0.05 | 0.069 | 0.062 | 0.016333 | 0.24333 | 0.060333 | 0.038333 | 261.954 | 0.205 | 0.0033371
262.311 | 0.033 | 262.311 | 0.033 | 0.03 | 0.041 | 0.258 | 0.272 | 0.267 | 0.056 | 0.072 | 0.083 | 0.034667 | 0.26567 | 0.070333 | 0.0525 | 262.311 | 0.21317 | 0.00347
262.668 | 0.042 | 262.668 | 0.042 | 0.035 | 0.038 | 0.27 | 0.271 | 0.277 | 0.061 | 0.062 | 0.08 | 0.038333 | 0.27267 | 0.067667 | 0.053 | 262.668 | 0.21967 | 0.0035759
263.025 | 0.025 | 263.025 | 0.025 | 0.03 | 0.032 | 0.248 | 0.25 | 0.258 | 0.059 | 0.061 | 0.069 | 0.029 | 0.252 | 0.063 | 0.046 | 263.025 | 0.206 | 0.0033534
263.382 | 0.032 | 263.382 | 0.032 | 0.037 | 0.033 | 0.257 | 0.259 | 0.278 | 0.097 | 0.089 | 0.088 | 0.034 | 0.26467 | 0.091333 | 0.062667 | 263.382 | 0.202 | 0.0032883
263.739 | 0.008 | 263.739 | 0.008 | 0.022 | 0.014 | 0.261 | 0.263 | 0.262 | 0.051 | 0.059 | 0.063 | 0.014667 | 0.262 | 0.057667 | 0.036167 | 263.739 | 0.22583 | 0.0036762
264.095 | 0.039 | 264.095 | 0.039 | 0.037 | 0.051 | 0.304 | 0.292 | 0.296 | 0.085 | 0.083 | 0.085 | 0.042333 | 0.29733 | 0.084333 | 0.063333 | 264.095 | 0.234 | 0.0038092
264.452 | 0.012 | 264.452 | 0.012 | 0.014 | 0.002 | 0.241 | 0.26 | 0.254 | 0.033 | 0.038 | 0.039 | 0.0093333 | 0.25167 | 0.036667 | 0.023 | 264.452 | 0.22867 | 0.0037224
264.809 | 0.022 | 264.809 | 0.022 | 0.028 | 0.032 | 0.263 | 0.267 | 0.272 | 0.062 | 0.062 | 0.078 | 0.027333 | 0.26733 | 0.067333 | 0.047333 | 264.809 | 0.22 | 0.0035813
265.165 | 0.014 | 265.165 | 0.014 | 0.024 | 0.026 | 0.269 | 0.275 | 0.291 | 0.046 | 0.062 | 0.068 | 0.021333 | 0.27833 | 0.058667 | 0.04 | 265.165 | 0.23833 | 0.0038797
265.521 | 0.051 | 265.521 | 0.051 | 0.047 | 0.058 | 0.273 | 0.274 | 0.285 | 0.087 | 0.094 | 0.094 | 0.052 | 0.27733 | 0.091667 | 0.071833 | 265.521 | 0.2055 | 0.0033452
265.878 | 0.021 | 265.878 | 0.021 | 0.021 | 0.033 | 0.17 | 0.164 | 0.193 | 0.071 | 0.064 | 0.078 | 0.025 | 0.17567 | 0.071 | 0.048 | 265.878 | 0.12767 | 0.0020782
266.234 | 0.027 | 266.234 | 0.027 | 0.021 | 0.032 | 0.154 | 0.149 | 0.164 | 0.067 | 0.064 | 0.071 | 0.026667 | 0.15567 | 0.067333 | 0.047 | 266.234 | 0.10867 | 0.0017689
266.59 | 0.013 | 266.59 | 0.013 | 0.02 | 0.029 | 0.117 | 0.126 | 0.139 | 0.049 | 0.058 | 0.057 | 0.020667 | 0.12733 | 0.054667 | 0.037667 | 266.59 | 0.089667 | 0.0014596
266.946 | 0.064 | 266.946 | 0.064 | 0.064 | 0.068 | 0.195 | 0.182 | 0.184 | 0.099 | 0.096 | 0.111 | 0.065333 | 0.187 | 0.102 | 0.083667 | 266.946 | 0.10333 | 0.0016821
267.302 | 0.014 | 267.302 | 0.014 | 0.025 | 0.033 | 0.138 | 0.149 | 0.159 | 0.068 | 0.073 | 0.06 | 0.024 | 0.14867 | 0.067 | 0.0455 | 267.302 | 0.10317 | 0.0016794
267.658 | 0.03 | 267.658 | 0.03 | 0.045 | 0.044 | 0.153 | 0.145 | 0.161 | 0.075 | 0.076 | 0.073 | 0.039667 | 0.153 | 0.074667 | 0.057167 | 267.658 | 0.095833 | 0.00156
268.014 | 0.019 | 268.014 | 0.019 | 0.027 | 0.034 | 0.122 | 0.127 | 0.13 | 0.049 | 0.062 | 0.065 | 0.026667 | 0.12633 | 0.058667 | 0.042667 | 268.014 | 0.083667 | 0.001362
268.37 | 0.035 | 268.37 | 0.035 | 0.026 | 0.032 | 0.14 | 0.149 | 0.145 | 0.063 | 0.066 | 0.074 | 0.031 | 0.14467 | 0.067667 | 0.049333 | 268.37 | 0.095333 | 0.0015519
268.725 | 0.018 | 268.725 | 0.018 | 0.03 | 0.029 | 0.12 | 0.139 | 0.131 | 0.056 | 0.066 | 0.063 | 0.025667 | 0.13 | 0.061667 | 0.043667 | 268.725 | 0.086333 | 0.0014054
269.081 | 0.02 | 269.081 | 0.02 | 0.031 | 0.03 | 0.125 | 0.132 | 0.131 | 0.053 | 0.054 | 0.075 | 0.027 | 0.12933 | 0.060667 | 0.043833 | 269.081 | 0.0855 | 0.0013918
269.436 | 0.019 | 269.436 | 0.019 | 0.035 | 0.029 | 0.131 | 0.138 | 0.135 | 0.06 | 0.07 | 0.077 | 0.027667 | 0.13467 | 0.069 | 0.048333 | 269.436 | 0.086333 | 0.0014054
269.792 | 0.003 | 269.792 | 0.003 | 0.001 | 0.006 | 0.104 | 0.107 | 0.115 | 0.042 | 0.028 | 0.039 | 0.0033333 | 0.10867 | 0.036333 | 0.019833 | 269.792 | 0.088833 | 0.0014461
270.147 | 0.029 | 270.147 | 0.029 | 0.022 | 0.025 | 0.122 | 0.133 | 0.137 | 0.068 | 0.072 | 0.076 | 0.025333 | 0.13067 | 0.072 | 0.048667 | 270.147 | 0.082 | 0.0013348
270.502 | 0.041 | 270.502 | 0.041 | 0.048 | 0.053 | 0.151 | 0.154 | 0.148 | 0.087 | 0.084 | 0.076 | 0.047333 | 0.151 | 0.082333 | 0.064833 | 270.502 | 0.086167 | 0.0014027
270.857 | 0.035 | 270.857 | 0.035 | 0.032 | 0.046 | 0.143 | 0.144 | 0.158 | 0.067 | 0.072 | 0.078 | 0.037667 | 0.14833 | 0.072333 | 0.055 | 270.857 | 0.093333 | 0.0015193
271.212 | 0.012 | 271.212 | 0.012 | 0.022 | 0.016 | 0.08 | 0.089 | 0.091 | 0.062 | 0.068 | 0.052 | 0.016667 | 0.086667 | 0.060667 | 0.038667 | 271.212 | 0.048 | 0.00078137
271.567 | -0.012 | 271.567 | -0.012 | -0.007 | 0.002 | 0.04 | 0.05 | 0.058 | 0.027 | 0.023 | 0.03 | -0.0056667 | 0.049333 | 0.026667 | 0.0105 | 271.567 | 0.038833 | 0.00063215
271.922 | 0.06 | 271.922 | 0.06 | 0.061 | 0.065 | 0.101 | 0.113 | 0.116 | 0.096 | 0.093 | 0.109 | 0.062 | 0.11 | 0.099333 | 0.080667 | 271.922 | 0.029333 | 0.0004775
272.277 | 0.022 | 272.277 | 0.022 | 0.021 | 0.018 | 0.071 | 0.07 | 0.088 | 0.058 | 0.062 | 0.073 | 0.020333 | 0.076333 | 0.064333 | 0.042333 | 272.277 | 0.034 | 0.00055347
272.632 | 0.022 | 272.632 | 0.022 | 0.043 | 0.039 | 0.077 | 0.089 | 0.105 | 0.066 | 0.08 | 0.086 | 0.034667 | 0.090333 | 0.077333 | 0.056 | 272.632 | 0.034333 | 0.0005589
272.987 | 0.045 | 272.987 | 0.045 | 0.044 | 0.059 | 0.087 | 0.101 | 0.103 | 0.081 | 0.08 | 0.083 | 0.049333 | 0.097 | 0.081333 | 0.065333 | 272.987 | 0.031667 | 0.00051549
273.341 | 0.039 | 273.341 | 0.039 | 0.05 | 0.035 | 0.077 | 0.093 | 0.092 | 0.078 | 0.073 | 0.084 | 0.041333 | 0.087333 | 0.078333 | 0.059833 | 273.341 | 0.0275 | 0.00044766
273.696 | 0.015 | 273.696 | 0.015 | 0.033 | 0.043 | 0.08 | 0.073 | 0.083 | 0.066 | 0.072 | 0.074 | 0.030333 | 0.078667 | 0.070667 | 0.0505 | 273.696 | 0.028167 | 0.00045851
274.05 | -0.015 | 274.05 | -0.015 | 0.003 | 0.004 | 0.045 | 0.047 | 0.046 | 0.033 | 0.036 | 0.043 | -0.0026667 | 0.046 | 0.037333 | 0.017333 | 274.05 | 0.028667 | 0.00046665
274.404 | 0.007 | 274.404 | 0.007 | 0.022 | 0.021 | 0.071 | 0.079 | 0.064 | 0.053 | 0.062 | 0.067 | 0.016667 | 0.071333 | 0.060667 | 0.038667 | 274.404 | 0.032667 | 0.00053177
274.759 | 0.049 | 274.759 | 0.049 | 0.045 | 0.054 | 0.092 | 0.108 | 0.1 | 0.083 | 0.09 | 0.078 | 0.049333 | 0.1 | 0.083667 | 0.0665 | 274.759 | 0.0335 | 0.00054533
275.113 | 0.046 | 275.113 | 0.046 | 0.057 | 0.047 | 0.128 | 0.117 | 0.137 | 0.083 | 0.108 | 0.101 | 0.05 | 0.12733 | 0.097333 | 0.073667 | 275.113 | 0.053667 | 0.00087362
275.467 | -0.009 | 275.467 | -0.009 | 0.005 | 0.014 | 0.081 | 0.063 | 0.065 | 0.034 | 0.045 | 0.049 | 0.0033333 | 0.069667 | 0.042667 | 0.023 | 275.467 | 0.046667 | 0.00075967
275.821 | 0.002 | 275.821 | 0.002 | 0.014 | 0.013 | 0.046 | 0.057 | 0.068 | 0.049 | 0.046 | 0.045 | 0.0096667 | 0.057 | 0.046667 | 0.028167 | 275.821 | 0.028833 | 0.00046937
276.175 | 0.022 | 276.175 | 0.022 | 0.017 | 0.043 | 0.065 | 0.064 | 0.082 | 0.063 | 0.066 | 0.055 | 0.027333 | 0.070333 | 0.061333 | 0.044333 | 276.175 | 0.026 | 0.00042324
276.529 | 0.035 | 276.529 | 0.035 | 0.046 | 0.053 | 0.09 | 0.079 | 0.098 | 0.084 | 0.093 | 0.082 | 0.044667 | 0.089 | 0.086333 | 0.0655 | 276.529 | 0.0235 | 0.00038255
276.882 | 0.063 | 276.882 | 0.063 | 0.078 | 0.075 | 0.113 | 0.109 | 0.12 | 0.11 | 0.111 | 0.103 | 0.072 | 0.114 | 0.108 | 0.09 | 276.882 | 0.024 | 0.00039069
277.236 | 0.034 | 277.236 | 0.034 | 0.03 | 0.044 | 0.076 | 0.08 | 0.079 | 0.062 | 0.078 | 0.072 | 0.036 | 0.078333 | 0.070667 | 0.053333 | 277.236 | 0.025 | 0.00040696
277.59 | 0.04 | 277.59 | 0.04 | 0.046 | 0.05 | 0.082 | 0.085 | 0.095 | 0.085 | 0.088 | 0.097 | 0.045333 | 0.087333 | 0.09 | 0.067667 | 277.59 | 0.019667 | 0.00032014
277.943 | 0.003 | 277.943 | 0.003 | 0.001 | 0.004 | 0.032 | 0.041 | 0.051 | 0.022 | 0.038 | 0.036 | 0.0026667 | 0.041333 | 0.032 | 0.017333 | 277.943 | 0.024 | 0.00039069
278.297 | 0.077 | 278.297 | 0.077 | 0.082 | 0.081 | 0.107 | 0.116 | 0.115 | 0.106 | 0.122 | 0.127 | 0.08 | 0.11267 | 0.11833 | 0.099167 | 278.297 | 0.0135 | 0.00021976
278.65 | 0.014 | 278.65 | 0.014 | 0.004 | 0.02 | 0.045 | 0.046 | 0.052 | 0.049 | 0.051 | 0.051 | 0.012667 | 0.047667 | 0.050333 | 0.0315 | 278.65 | 0.016167 | 0.00026317
279.003 | -0.003 | 279.003 | -0.003 | 0.018 | 0.008 | 0.045 | 0.046 | 0.038 | 0.037 | 0.049 | 0.045 | 0.0076667 | 0.043 | 0.043667 | 0.025667 | 279.003 | 0.017333 | 0.00028216
279.356 | 0.023 | 279.356 | 0.023 | 0.03 | 0.027 | 0.058 | 0.058 | 0.069 | 0.047 | 0.053 | 0.059 | 0.026667 | 0.061667 | 0.053 | 0.039833 | 279.356 | 0.021833 | 0.00035542
279.709 | 0.023 | 279.709 | 0.023 | 0.037 | 0.034 | 0.077 | 0.067 | 0.082 | 0.061 | 0.073 | 0.072 | 0.031333 | 0.075333 | 0.068667 | 0.05 | 279.709 | 0.025333 | 0.00041239
280.062 | 0.006 | 280.062 | 0.006 | 0.021 | 0.019 | 0.073 | 0.067 | 0.078 | 0.046 | 0.05 | 0.051 | 0.015333 | 0.072667 | 0.049 | 0.032167 | 280.062 | 0.0405 | 0.00065928
280.415 | 0.029 | 280.415 | 0.029 | 0.036 | 0.04 | 0.086 | 0.1 | 0.105 | 0.059 | 0.072 | 0.079 | 0.035 | 0.097 | 0.07 | 0.0525 | 280.415 | 0.0445 | 0.0007244
280.768 | 0.019 | 280.768 | 0.019 | 0.023 | 0.044 | 0.068 | 0.069 | 0.074 | 0.063 | 0.067 | 0.063 | 0.028667 | 0.070333 | 0.064333 | 0.0465 | 280.768 | 0.023833 | 0.00038797
281.121 | 0.05 | 281.121 | 0.05 | 0.057 | 0.067 | 0.088 | 0.089 | 0.108 | 0.084 | 0.083 | 0.094 | 0.058 | 0.095 | 0.087 | 0.0725 | 281.121 | 0.0225 | 0.00036627
281.474 | 0.034 | 281.474 | 0.034 | 0.035 | 0.031 | 0.075 | 0.076 | 0.07 | 0.047 | 0.074 | 0.075 | 0.033333 | 0.073667 | 0.065333 | 0.049333 | 281.474 | 0.024333 | 0.00039611
281.826 | 0.016 | 281.826 | 0.016 | 0.018 | 0.038 | 0.062 | 0.058 | 0.07 | 0.048 | 0.064 | 0.07 | 0.024 | 0.063333 | 0.060667 | 0.042333 | 281.826 | 0.021 | 0.00034185
282.179 | 0.024 | 282.179 | 0.024 | 0.036 | 0.032 | 0.051 | 0.06 | 0.063 | 0.052 | 0.067 | 0.065 | 0.030667 | 0.058 | 0.061333 | 0.046 | 282.179 | 0.012 | 0.00019534
282.531 | 0.005 | 282.531 | 0.005 | 0.011 | 0.017 | 0.043 | 0.043 | 0.043 | 0.035 | 0.043 | 0.051 | 0.011 | 0.043 | 0.043 | 0.027 | 282.531 | 0.016 | 0.00026046
282.884 | 0.037 | 282.884 | 0.037 | 0.038 | 0.04 | 0.063 | 0.071 | 0.063 | 0.07 | 0.077 | 0.078 | 0.038333 | 0.065667 | 0.075 | 0.056667 | 282.884 | 0.009 | 0.00014651
283.236 | 0.024 | 283.236 | 0.024 | 0.026 | 0.032 | 0.057 | 0.061 | 0.073 | 0.067 | 0.061 | 0.07 | 0.027333 | 0.063667 | 0.066 | 0.046667 | 283.236 | 0.017 | 0.00027674
283.588 | 0.068 | 283.588 | 0.068 | 0.064 | 0.066 | 0.083 | 0.097 | 0.1 | 0.092 | 0.096 | 0.093 | 0.066 | 0.093333 | 0.093667 | 0.079833 | 283.588 | 0.0135 | 0.00021976
283.94 | 0.013 | 283.94 | 0.013 | 0.013 | 0.024 | 0.039 | 0.045 | 0.044 | 0.042 | 0.049 | 0.049 | 0.016667 | 0.042667 | 0.046667 | 0.031667 | 283.94 | 0.011 | 0.00017906
284.292 | 0.028 | 284.292 | 0.028 | 0.03 | 0.039 | 0.059 | 0.056 | 0.067 | 0.05 | 0.063 | 0.067 | 0.032333 | 0.060667 | 0.06 | 0.046167 | 284.292 | 0.0145 | 0.00023604
284.644 | 0.009 | 284.644 | 0.009 | 0.016 | 0.016 | 0.042 | 0.038 | 0.037 | 0.043 | 0.049 | 0.052 | 0.013667 | 0.039 | 0.048 | 0.030833 | 284.644 | 0.0081667 | 0.00013294
284.996 | 0.028 | 284.996 | 0.028 | 0.033 | 0.035 | 0.051 | 0.06 | 0.056 | 0.05 | 0.056 | 0.063 | 0.032 | 0.055667 | 0.056333 | 0.044167 | 284.996 | 0.0115 | 0.0001872
285.348 | 0.019 | 285.348 | 0.019 | 0.029 | 0.034 | 0.05 | 0.053 | 0.061 | 0.05 | 0.052 | 0.062 | 0.027333 | 0.054667 | 0.054667 | 0.041 | 285.348 | 0.013667 | 0.00022247
285.7 | 0.04 | 285.7 | 0.04 | 0.043 | 0.056 | 0.078 | 0.069 | 0.084 | 0.067 | 0.083 | 0.082 | 0.046333 | 0.077 | 0.077333 | 0.061833 | 285.7 | 0.015167 | 0.00024689
286.051 | 0.02 | 286.051 | 0.02 | 0.032 | 0.028 | 0.056 | 0.059 | 0.058 | 0.057 | 0.062 | 0.058 | 0.026667 | 0.057667 | 0.059 | 0.042833 | 286.051 | 0.014833 | 0.00024147
286.403 | 0.025 | 286.403 | 0.025 | 0.029 | 0.036 | 0.055 | 0.052 | 0.055 | 0.053 | 0.055 | 0.058 | 0.03 | 0.054 | 0.055333 | 0.042667 | 286.403 | 0.011333 | 0.00018449
286.754 | 0.029 | 286.754 | 0.029 | 0.033 | 0.033 | 0.05 | 0.056 | 0.058 | 0.061 | 0.069 | 0.058 | 0.031667 | 0.054667 | 0.062667 | 0.047167 | 286.754 | 0.0075 | 0.00012209
287.106 | 0.047 | 287.106 | 0.047 | 0.053 | 0.057 | 0.078 | 0.083 | 0.083 | 0.087 | 0.085 | 0.084 | 0.052333 | 0.081333 | 0.085333 | 0.068833 | 287.106 | 0.0125 | 0.00020348
287.457 | 0.037 | 287.457 | 0.037 | 0.05 | 0.051 | 0.069 | 0.069 | 0.072 | 0.074 | 0.072 | 0.082 | 0.046 | 0.07 | 0.076 | 0.061 | 287.457 | 0.009 | 0.00014651
287.808 | 0.037 | 287.808 | 0.037 | 0.041 | 0.036 | 0.056 | 0.065 | 0.07 | 0.066 | 0.069 | 0.065 | 0.038 | 0.063667 | 0.066667 | 0.052333 | 287.808 | 0.011333 | 0.00018449
288.159 | 0.028 | 288.159 | 0.028 | 0.032 | 0.048 | 0.069 | 0.083 | 0.089 | 0.071 | 0.072 | 0.074 | 0.036 | 0.080333 | 0.072333 | 0.054167 | 288.159 | 0.026167 | 0.00042596
288.51 | 0.048 | 288.51 | 0.048 | 0.04 | 0.046 | 0.082 | 0.083 | 0.091 | 0.072 | 0.066 | 0.066 | 0.044667 | 0.085333 | 0.068 | 0.056333 | 288.51 | 0.029 | 0.00047208
288.861 | 0.018 | 288.861 | 0.018 | 0.027 | 0.027 | 0.085 | 0.087 | 0.09 | 0.054 | 0.057 | 0.052 | 0.024 | 0.087333 | 0.054333 | 0.039167 | 288.861 | 0.048167 | 0.00078408
289.212 | 0.003 | 289.212 | 0.003 | -0.007 | 0.009 | 0.083 | 0.087 | 0.088 | 0.025 | 0.026 | 0.035 | 0.0016667 | 0.086 | 0.028667 | 0.015167 | 289.212 | 0.070833 | 0.0011531
289.563 | 0.01 | 289.563 | 0.01 | 0.015 | 0.014 | 0.07 | 0.064 | 0.07 | 0.04 | 0.038 | 0.046 | 0.013 | 0.068 | 0.041333 | 0.027167 | 289.563 | 0.040833 | 0.00066471
289.914 | 0.041 | 289.914 | 0.041 | 0.042 | 0.046 | 0.067 | 0.067 | 0.074 | 0.069 | 0.083 | 0.072 | 0.043 | 0.069333 | 0.074667 | 0.058833 | 289.914 | 0.0105 | 0.00017092
290.264 | -0.002 | 290.264 | -0.002 | 0.011 | 0.019 | 0.037 | 0.03 | 0.037 | 0.037 | 0.039 | 0.039 | 0.0093333 | 0.034667 | 0.038333 | 0.023833 | 290.264 | 0.010833 | 0.00017635
290.615 | 0.027 | 290.615 | 0.027 | 0.03 | 0.039 | 0.053 | 0.054 | 0.052 | 0.059 | 0.056 | 0.056 | 0.032 | 0.053 | 0.057 | 0.0445 | 290.615 | 0.0085 | 0.00013837
290.966 | 0.042 | 290.966 | 0.042 | 0.045 | 0.047 | 0.07 | 0.07 | 0.066 | 0.064 | 0.072 | 0.072 | 0.044667 | 0.068667 | 0.069333 | 0.057 | 290.966 | 0.011667 | 0.00018992
291.316 | 0.011 | 291.316 | 0.011 | 0.027 | 0.019 | 0.033 | 0.034 | 0.036 | 0.032 | 0.036 | 0.037 | 0.019 | 0.034333 | 0.035 | 0.027 | 291.316 | 0.0073333 | 0.00011938
291.666 | 0.036 | 291.666 | 0.036 | 0.041 | 0.048 | 0.064 | 0.068 | 0.071 | 0.066 | 0.072 | 0.064 | 0.041667 | 0.067667 | 0.067333 | 0.0545 | 291.666 | 0.013167 | 0.00021433
292.017 | 0.03 | 292.017 | 0.03 | 0.032 | 0.035 | 0.052 | 0.055 | 0.061 | 0.051 | 0.056 | 0.055 | 0.032333 | 0.056 | 0.054 | 0.043167 | 292.017 | 0.012833 | 0.00020891
292.367 | 0.014 | 292.367 | 0.014 | 0.015 | 0.017 | 0.037 | 0.046 | 0.043 | 0.029 | 0.028 | 0.031 | 0.015333 | 0.042 | 0.029333 | 0.022333 | 292.367 | 0.019667 | 0.00032014
292.717 | 0.003 | 292.717 | 0.003 | 0.015 | 0.013 | 0.035 | 0.046 | 0.039 | 0.038 | 0.042 | 0.037 | 0.010333 | 0.04 | 0.039 | 0.024667 | 292.717 | 0.015333 | 0.0002496
293.067 | 0.022 | 293.067 | 0.022 | 0.027 | 0.025 | 0.041 | 0.042 | 0.042 | 0.043 | 0.043 | 0.051 | 0.024667 | 0.041667 | 0.045667 | 0.035167 | 293.067 | 0.0065 | 0.00010581
293.417 | 0.024 | 293.417 | 0.024 | 0.026 | 0.027 | 0.045 | 0.052 | 0.049 | 0.044 | 0.056 | 0.051 | 0.025667 | 0.048667 | 0.050333 | 0.038 | 293.417 | 0.010667 | 0.00017364
293.767 | 0.013 | 293.767 | 0.013 | 0.021 | 0.024 | 0.038 | 0.04 | 0.046 | 0.043 | 0.045 | 0.046 | 0.019333 | 0.041333 | 0.044667 | 0.032 | 293.767 | 0.0093333 | 0.00015193
294.116 | 0.037 | 294.116 | 0.037 | 0.046 | 0.048 | 0.065 | 0.064 | 0.063 | 0.064 | 0.064 | 0.073 | 0.043667 | 0.064 | 0.067 | 0.055333 | 294.116 | 0.0086667 | 0.00014108
294.466 | 0.018 | 294.466 | 0.018 | 0.025 | 0.026 | 0.037 | 0.045 | 0.046 | 0.04 | 0.052 | 0.05 | 0.023 | 0.042667 | 0.047333 | 0.035167 | 294.466 | 0.0075 | 0.00012209
294.816 | 0.035 | 294.816 | 0.035 | 0.04 | 0.04 | 0.062 | 0.057 | 0.06 | 0.061 | 0.065 | 0.058 | 0.038333 | 0.059667 | 0.061333 | 0.049833 | 294.816 | 0.0098333 | 0.00016007
295.165 | 0.012 | 295.165 | 0.012 | 0.018 | 0.021 | 0.049 | 0.044 | 0.051 | 0.033 | 0.034 | 0.046 | 0.017 | 0.048 | 0.037667 | 0.027333 | 295.165 | 0.020667 | 0.00033642
295.515 | 0.007 | 295.515 | 0.007 | 0.016 | 0.014 | 0.065 | 0.074 | 0.066 | 0.036 | 0.029 | 0.039 | 0.012333 | 0.068333 | 0.034667 | 0.0235 | 295.515 | 0.044833 | 0.00072982
295.864 | 0.024 | 295.864 | 0.024 | 0.034 | 0.032 | 0.106 | 0.109 | 0.109 | 0.048 | 0.055 | 0.048 | 0.03 | 0.108 | 0.050333 | 0.040167 | 295.864 | 0.067833 | 0.0011042
296.213 | 0.012 | 296.213 | 0.012 | 0.015 | 0.018 | 0.154 | 0.152 | 0.15 | 0.03 | 0.036 | 0.035 | 0.015 | 0.152 | 0.033667 | 0.024333 | 296.213 | 0.12767 | 0.0020782
296.562 | 0.006 | 296.562 | 0.006 | 0.01 | 0.011 | 0.236 | 0.239 | 0.248 | 0.029 | 0.028 | 0.032 | 0.009 | 0.241 | 0.029667 | 0.019333 | 296.562 | 0.22167 | 0.0036084
296.911 | 0.029 | 296.911 | 0.029 | 0.03 | 0.039 | 0.165 | 0.155 | 0.162 | 0.054 | 0.059 | 0.06 | 0.032667 | 0.16067 | 0.057667 | 0.045167 | 296.911 | 0.1155 | 0.0018802
297.261 | 0.02 | 297.261 | 0.02 | 0.022 | 0.024 | 0.051 | 0.047 | 0.055 | 0.039 | 0.04 | 0.044 | 0.022 | 0.051 | 0.041 | 0.0315 | 297.261 | 0.0195 | 0.00031743
297.609 | 0.018 | 297.609 | 0.018 | 0.02 | 0.026 | 0.047 | 0.04 | 0.039 | 0.042 | 0.043 | 0.045 | 0.021333 | 0.042 | 0.043333 | 0.032333 | 297.609 | 0.0096667 | 0.00015736
297.958 | -0.012 | 297.958 | -0.012 | -0.001 | -0.005 | 0.014 | 0.015 | 0.015 | 0.014 | 0.016 | 0.021 | -0.006 | 0.014667 | 0.017 | 0.0055 | 297.958 | 0.0091667 | 0.00014922
298.307 | 0.017 | 298.307 | 0.017 | 0.021 | 0.023 | 0.039 | 0.033 | 0.038 | 0.035 | 0.038 | 0.039 | 0.020333 | 0.036667 | 0.037333 | 0.028833 | 298.307 | 0.0078333 | 0.00012752
298.656 | 0.009 | 298.656 | 0.009 | 0.021 | 0.03 | 0.033 | 0.03 | 0.034 | 0.04 | 0.038 | 0.04 | 0.02 | 0.032333 | 0.039333 | 0.029667 | 298.656 | 0.0026667 | 4.3409e-05
299.004 | 0.014 | 299.004 | 0.014 | 0.016 | 0.023 | 0.028 | 0.029 | 0.032 | 0.032 | 0.038 | 0.033 | 0.017667 | 0.029667 | 0.034333 | 0.026 | 299.004 | 0.0036667 | 5.9688e-05
299.353 | -0.013 | 299.353 | -0.013 | -0.013 | -0.005 | 0.004 | 0.003 | 0.003 | 0.003 | 0.003 | 0.007 | -0.010333 | 0.0033333 | 0.0043333 | -0.003 | 299.353 | 0.0063333 | 0.0001031
299.701 | 0.015 | 299.701 | 0.015 | 0.021 | 0.023 | 0.031 | 0.036 | 0.031 | 0.027 | 0.035 | 0.04 | 0.019667 | 0.032667 | 0.034 | 0.026833 | 299.701 | 0.0058333 | 9.4958e-05
300.05 | 0.005 | 300.05 | 0.005 | 0.01 | 0.015 | 0.021 | 0.024 | 0.025 | 0.025 | 0.028 | 0.031 | 0.01 | 0.023333 | 0.028 | 0.019 | 300.05 | 0.0043333 | 7.054e-05
300.398 | 0.024 | 300.398 | 0.024 | 0.026 | 0.027 | 0.039 | 0.037 | 0.04 | 0.04 | 0.045 | 0.048 | 0.025667 | 0.038667 | 0.044333 | 0.035 | 300.398 | 0.0036667 | 5.9688e-05
300.746 | 0.017 | 300.746 | 0.017 | 0.023 | 0.022 | 0.047 | 0.045 | 0.047 | 0.036 | 0.04 | 0.038 | 0.020667 | 0.046333 | 0.038 | 0.029333 | 300.746 | 0.017 | 0.00027674
301.095 | 0.019 | 301.095 | 0.019 | 0.018 | 0.019 | 0.063 | 0.063 | 0.066 | 0.029 | 0.039 | 0.036 | 0.018667 | 0.064 | 0.034667 | 0.026667 | 301.095 | 0.037333 | 0.00060773
301.443 | 0.018 | 301.443 | 0.018 | 0.015 | 0.02 | 0.081 | 0.082 | 0.081 | 0.031 | 0.032 | 0.035 | 0.017667 | 0.081333 | 0.032667 | 0.025167 | 301.443 | 0.056167 | 0.00091431
301.791 | 0.013 | 301.791 | 0.013 | 0.012 | 0.012 | 0.133 | 0.13 | 0.133 | 0.028 | 0.029 | 0.025 | 0.012333 | 0.132 | 0.027333 | 0.019833 | 301.791 | 0.11217 | 0.0018259
302.138 | 0.017 | 302.138 | 0.017 | 0.018 | 0.02 | 0.161 | 0.16 | 0.164 | 0.034 | 0.037 | 0.035 | 0.018333 | 0.16167 | 0.035333 | 0.026833 | 302.138 | 0.13483 | 0.0021949
302.486 | 0.002 | 302.486 | 0.002 | 0.01 | 0.006 | 0.062 | 0.068 | 0.066 | 0.021 | 0.024 | 0.022 | 0.006 | 0.065333 | 0.022333 | 0.014167 | 302.486 | 0.051167 | 0.00083292
302.834 | 0.004 | 302.834 | 0.004 | 0.005 | 0.009 | 0.028 | 0.032 | 0.031 | 0.018 | 0.025 | 0.023 | 0.006 | 0.030333 | 0.022 | 0.014 | 302.834 | 0.016333 | 0.00026588
303.182 | 0.004 | 303.182 | 0.004 | 0.006 | 0.005 | 0.016 | 0.016 | 0.016 | 0.017 | 0.016 | 0.02 | 0.005 | 0.016 | 0.017667 | 0.011333 | 303.182 | 0.0046667 | 7.5967e-05
303.529 | 0.017 | 303.529 | 0.017 | 0.02 | 0.021 | 0.029 | 0.032 | 0.036 | 0.035 | 0.036 | 0.034 | 0.019333 | 0.032333 | 0.035 | 0.027167 | 303.529 | 0.0051667 | 8.4106e-05
303.877 | 0.021 | 303.877 | 0.021 | 0.022 | 0.02 | 0.031 | 0.027 | 0.032 | 0.033 | 0.032 | 0.028 | 0.021 | 0.03 | 0.031 | 0.026 | 303.877 | 0.004 | 6.5114e-05
304.224 | 0.019 | 304.224 | 0.019 | 0.021 | 0.025 | 0.034 | 0.04 | 0.035 | 0.034 | 0.035 | 0.039 | 0.021667 | 0.036333 | 0.036 | 0.028833 | 304.224 | 0.0075 | 0.00012209
304.572 | 0.017 | 304.572 | 0.017 | 0.022 | 0.022 | 0.036 | 0.035 | 0.034 | 0.035 | 0.039 | 0.035 | 0.020333 | 0.035 | 0.036333 | 0.028333 | 304.572 | 0.0066667 | 0.00010852
304.919 | 0 | 304.919 | 0 | 0.004 | 0.005 | 0.011 | 0.023 | 0.014 | 0.018 | 0.019 | 0.017 | 0.003 | 0.016 | 0.018 | 0.0105 | 304.919 | 0.0055 | 8.9532e-05
305.266 | 0.018 | 305.266 | 0.018 | 0.021 | 0.019 | 0.039 | 0.038 | 0.036 | 0.038 | 0.041 | 0.043 | 0.019333 | 0.037667 | 0.040667 | 0.03 | 305.266 | 0.0076667 | 0.0001248
305.613 | 0.009 | 305.613 | 0.009 | 0.007 | 0.012 | 0.023 | 0.02 | 0.024 | 0.028 | 0.029 | 0.026 | 0.0093333 | 0.022333 | 0.027667 | 0.0185 | 305.613 | 0.0038333 | 6.2401e-05
305.96 | 0.027 | 305.96 | 0.027 | 0.03 | 0.029 | 0.04 | 0.043 | 0.043 | 0.047 | 0.045 | 0.049 | 0.028667 | 0.042 | 0.047 | 0.037833 | 305.96 | 0.0041667 | 6.7827e-05
306.307 | 0.01 | 306.307 | 0.01 | 0.005 | 0.015 | 0.02 | 0.028 | 0.026 | 0.024 | 0.03 | 0.031 | 0.01 | 0.024667 | 0.028333 | 0.019167 | 306.307 | 0.0055 | 8.9532e-05
306.654 | 0.025 | 306.654 | 0.025 | 0.031 | 0.03 | 0.042 | 0.041 | 0.049 | 0.046 | 0.05 | 0.045 | 0.028667 | 0.044 | 0.047 | 0.037833 | 306.654 | 0.0061667 | 0.00010038
307.001 | 0.027 | 307.001 | 0.027 | 0.025 | 0.028 | 0.039 | 0.036 | 0.043 | 0.039 | 0.042 | 0.042 | 0.026667 | 0.039333 | 0.041 | 0.033833 | 307.001 | 0.0055 | 8.9532e-05
307.347 | 0.014 | 307.347 | 0.014 | 0.02 | 0.016 | 0.027 | 0.03 | 0.033 | 0.031 | 0.034 | 0.031 | 0.016667 | 0.03 | 0.032 | 0.024333 | 307.347 | 0.0056667 | 9.2245e-05
307.694 | 0.004 | 307.694 | 0.004 | 0.003 | 0.011 | 0.018 | 0.02 | 0.017 | 0.02 | 0.018 | 0.019 | 0.006 | 0.018333 | 0.019 | 0.0125 | 307.694 | 0.0058333 | 9.4958e-05
308.04 | 0.007 | 308.04 | 0.007 | 0.012 | 0.014 | 0.025 | 0.023 | 0.022 | 0.024 | 0.024 | 0.023 | 0.011 | 0.023333 | 0.023667 | 0.017333 | 308.04 | 0.006 | 9.7671e-05
308.387 | 0.023 | 308.387 | 0.023 | 0.024 | 0.027 | 0.034 | 0.034 | 0.04 | 0.036 | 0.036 | 0.039 | 0.024667 | 0.036 | 0.037 | 0.030833 | 308.387 | 0.0051667 | 8.4106e-05
308.733 | 0.011 | 308.733 | 0.011 | 0.016 | 0.018 | 0.027 | 0.026 | 0.031 | 0.032 | 0.029 | 0.028 | 0.015 | 0.028 | 0.029667 | 0.022333 | 308.733 | 0.0056667 | 9.2245e-05
309.08 | 0.015 | 309.08 | 0.015 | 0.016 | 0.014 | 0.025 | 0.028 | 0.029 | 0.029 | 0.03 | 0.025 | 0.015 | 0.027333 | 0.028 | 0.0215 | 309.08 | 0.0058333 | 9.4958e-05
309.426 | 0.021 | 309.426 | 0.021 | 0.026 | 0.031 | 0.045 | 0.041 | 0.046 | 0.041 | 0.046 | 0.048 | 0.026 | 0.044 | 0.045 | 0.0355 | 309.426 | 0.0085 | 0.00013837
309.772 | 0.016 | 309.772 | 0.016 | 0.019 | 0.022 | 0.032 | 0.033 | 0.033 | 0.033 | 0.035 | 0.038 | 0.019 | 0.032667 | 0.035333 | 0.027167 | 309.772 | 0.0055 | 8.9532e-05
310.118 | 0.024 | 310.118 | 0.024 | 0.027 | 0.028 | 0.036 | 0.041 | 0.037 | 0.042 | 0.046 | 0.038 | 0.026333 | 0.038 | 0.042 | 0.034167 | 310.118 | 0.0038333 | 6.2401e-05
310.464 | 0.009 | 310.464 | 0.009 | 0.008 | 0.013 | 0.024 | 0.021 | 0.027 | 0.023 | 0.022 | 0.025 | 0.01 | 0.024 | 0.023333 | 0.016667 | 310.464 | 0.0073333 | 0.00011938
310.81 | 0.016 | 310.81 | 0.016 | 0.017 | 0.022 | 0.038 | 0.039 | 0.04 | 0.031 | 0.036 | 0.032 | 0.018333 | 0.039 | 0.033 | 0.025667 | 310.81 | 0.013333 | 0.00021705
311.156 | 0.011 | 311.156 | 0.011 | 0.016 | 0.014 | 0.052 | 0.055 | 0.054 | 0.031 | 0.031 | 0.025 | 0.013667 | 0.053667 | 0.029 | 0.021333 | 311.156 | 0.032333 | 0.00052634
311.501 | 0.021 | 311.501 | 0.021 | 0.021 | 0.023 | 0.137 | 0.142 | 0.142 | 0.034 | 0.034 | 0.034 | 0.021667 | 0.14033 | 0.034 | 0.027833 | 311.501 | 0.1125 | 0.0018313
311.847 | 0.015 | 311.847 | 0.015 | 0.017 | 0.018 | 0.281 | 0.281 | 0.288 | 0.032 | 0.03 | 0.031 | 0.016667 | 0.28333 | 0.031 | 0.023833 | 311.847 | 0.2595 | 0.0042243
312.193 | 0.006 | 312.193 | 0.006 | 0.011 | 0.014 | 0.598 | 0.594 | 0.601 | 0.024 | 0.024 | 0.026 | 0.010333 | 0.59767 | 0.024667 | 0.0175 | 312.193 | 0.58017 | 0.0094443
312.538 | 0.014 | 312.538 | 0.014 | 0.021 | 0.016 | 0.799 | 0.8 | 0.797 | 0.024 | 0.029 | 0.038 | 0.017 | 0.79867 | 0.030333 | 0.023667 | 312.538 | 0.775 | 0.012616
312.883 | 0.015 | 312.883 | 0.015 | 0.015 | 0.019 | 0.796 | 0.786 | 0.784 | 0.024 | 0.029 | 0.03 | 0.016333 | 0.78867 | 0.027667 | 0.022 | 312.883 | 0.76667 | 0.01248
313.229 | 0.013 | 313.229 | 0.013 | 0.015 | 0.018 | 0.521 | 0.524 | 0.525 | 0.027 | 0.025 | 0.026 | 0.015333 | 0.52333 | 0.026 | 0.020667 | 313.229 | 0.50267 | 0.0081827
313.574 | 0.005 | 313.574 | 0.005 | 0.012 | 0.013 | 0.107 | 0.108 | 0.107 | 0.024 | 0.024 | 0.024 | 0.01 | 0.10733 | 0.024 | 0.017 | 313.574 | 0.090333 | 0.0014705
313.919 | 0.016 | 313.919 | 0.016 | 0.016 | 0.018 | 0.04 | 0.036 | 0.041 | 0.028 | 0.029 | 0.03 | 0.016667 | 0.039 | 0.029 | 0.022833 | 313.919 | 0.016167 | 0.00026317
314.264 | 0.022 | 314.264 | 0.022 | 0.024 | 0.027 | 0.034 | 0.036 | 0.039 | 0.038 | 0.036 | 0.032 | 0.024333 | 0.036333 | 0.035333 | 0.029833 | 314.264 | 0.0065 | 0.00010581
314.609 | 0.021 | 314.609 | 0.021 | 0.024 | 0.023 | 0.038 | 0.035 | 0.038 | 0.036 | 0.038 | 0.036 | 0.022667 | 0.037 | 0.036667 | 0.029667 | 314.609 | 0.0073333 | 0.00011938
314.954 | 0.01 | 314.954 | 0.01 | 0.012 | 0.012 | 0.026 | 0.025 | 0.029 | 0.024 | 0.027 | 0.025 | 0.011333 | 0.026667 | 0.025333 | 0.018333 | 314.954 | 0.0083333 | 0.00013565
315.299 | 0.008 | 315.299 | 0.008 | 0.014 | 0.015 | 0.024 | 0.02 | 0.028 | 0.019 | 0.025 | 0.028 | 0.012333 | 0.024 | 0.024 | 0.018167 | 315.299 | 0.0058333 | 9.4958e-05
315.644 | 0.007 | 315.644 | 0.007 | 0.007 | 0.009 | 0.017 | 0.023 | 0.021 | 0.019 | 0.018 | 0.022 | 0.0076667 | 0.020333 | 0.019667 | 0.013667 | 315.644 | 0.0066667 | 0.00010852
315.988 | 0.019 | 315.988 | 0.019 | 0.023 | 0.026 | 0.03 | 0.029 | 0.034 | 0.033 | 0.034 | 0.036 | 0.022667 | 0.031 | 0.034333 | 0.0285 | 315.988 | 0.0025 | 4.0696e-05
316.333 | 0.021 | 316.333 | 0.021 | 0.02 | 0.028 | 0.034 | 0.035 | 0.036 | 0.036 | 0.036 | 0.038 | 0.023 | 0.035 | 0.036667 | 0.029833 | 316.333 | 0.0051667 | 8.4106e-05
316.678 | 0.014 | 316.678 | 0.014 | 0.016 | 0.016 | 0.028 | 0.023 | 0.027 | 0.021 | 0.028 | 0.028 | 0.015333 | 0.026 | 0.025667 | 0.0205 | 316.678 | 0.0055 | 8.9532e-05
317.022 | 0.005 | 317.022 | 0.005 | 0.016 | 0.017 | 0.023 | 0.023 | 0.026 | 0.025 | 0.028 | 0.027 | 0.012667 | 0.024 | 0.026667 | 0.019667 | 317.022 | 0.0043333 | 7.054e-05
317.366 | 0.009 | 317.366 | 0.009 | 0.013 | 0.014 | 0.023 | 0.024 | 0.026 | 0.023 | 0.025 | 0.023 | 0.012 | 0.024333 | 0.023667 | 0.017833 | 317.366 | 0.0065 | 0.00010581
317.711 | 0.017 | 317.711 | 0.017 | 0.015 | 0.021 | 0.029 | 0.026 | 0.027 | 0.03 | 0.026 | 0.028 | 0.017667 | 0.027333 | 0.028 | 0.022833 | 317.711 | 0.0045 | 7.3254e-05
318.055 | 0.007 | 318.055 | 0.007 | 0.006 | 0.009 | 0.02 | 0.024 | 0.021 | 0.016 | 0.02 | 0.02 | 0.0073333 | 0.021667 | 0.018667 | 0.013 | 318.055 | 0.0086667 | 0.00014108
318.399 | 0.007 | 318.399 | 0.007 | 0.013 | 0.014 | 0.018 | 0.021 | 0.028 | 0.02 | 0.027 | 0.023 | 0.011333 | 0.022333 | 0.023333 | 0.017333 | 318.399 | 0.005 | 8.1393e-05
318.743 | 0.012 | 318.743 | 0.012 | 0.016 | 0.019 | 0.03 | 0.027 | 0.029 | 0.028 | 0.028 | 0.03 | 0.015667 | 0.028667 | 0.028667 | 0.022167 | 318.743 | 0.0065 | 0.00010581
319.087 | 0.012 | 319.087 | 0.012 | 0.016 | 0.019 | 0.024 | 0.023 | 0.025 | 0.025 | 0.029 | 0.03 | 0.015667 | 0.024 | 0.028 | 0.021833 | 319.087 | 0.0021667 | 3.527e-05
319.431 | 0.016 | 319.431 | 0.016 | 0.023 | 0.021 | 0.033 | 0.032 | 0.033 | 0.031 | 0.036 | 0.033 | 0.02 | 0.032667 | 0.033333 | 0.026667 | 319.431 | 0.006 | 9.7671e-05
319.775 | 0.012 | 319.775 | 0.012 | 0.017 | 0.017 | 0.02 | 0.025 | 0.025 | 0.023 | 0.023 | 0.027 | 0.015333 | 0.023333 | 0.024333 | 0.019833 | 319.775 | 0.0035 | 5.6975e-05
320.118 | 0.022 | 320.118 | 0.022 | 0.023 | 0.024 | 0.038 | 0.036 | 0.04 | 0.039 | 0.041 | 0.04 | 0.023 | 0.038 | 0.04 | 0.0315 | 320.118 | 0.0065 | 0.00010581
320.462 | 0.012 | 320.462 | 0.012 | 0.015 | 0.016 | 0.024 | 0.022 | 0.026 | 0.027 | 0.026 | 0.027 | 0.014333 | 0.024 | 0.026667 | 0.0205 | 320.462 | 0.0035 | 5.6975e-05
320.806 | 0.013 | 320.806 | 0.013 | 0.014 | 0.013 | 0.021 | 0.023 | 0.026 | 0.019 | 0.024 | 0.025 | 0.013333 | 0.023333 | 0.022667 | 0.018 | 320.806 | 0.0053333 | 8.6819e-05
321.149 | 0.017 | 321.149 | 0.017 | 0.022 | 0.02 | 0.032 | 0.034 | 0.031 | 0.031 | 0.033 | 0.036 | 0.019667 | 0.032333 | 0.033333 | 0.0265 | 321.149 | 0.0058333 | 9.4958e-05
321.493 | 0.011 | 321.493 | 0.011 | 0.013 | 0.021 | 0.027 | 0.025 | 0.027 | 0.026 | 0.026 | 0.027 | 0.015 | 0.026333 | 0.026333 | 0.020667 | 321.493 | 0.0056667 | 9.2245e-05
321.836 | 0.01 | 321.836 | 0.01 | 0.007 | 0.013 | 0.018 | 0.025 | 0.021 | 0.022 | 0.025 | 0.025 | 0.01 | 0.021333 | 0.024 | 0.017 | 321.836 | 0.0043333 | 7.054e-05
322.179 | -0.001 | 322.179 | -0.001 | 0 | 0.004 | 0.009 | 0.012 | 0.011 | 0.009 | 0.01 | 0.012 | 0.001 | 0.010667 | 0.010333 | 0.0056667 | 322.179 | 0.005 | 8.1393e-05
322.522 | 0.011 | 322.522 | 0.011 | 0.014 | 0.016 | 0.025 | 0.027 | 0.026 | 0.031 | 0.024 | 0.028 | 0.013667 | 0.026 | 0.027667 | 0.020667 | 322.522 | 0.0053333 | 8.6819e-05
322.865 | 0.013 | 322.865 | 0.013 | 0.014 | 0.019 | 0.025 | 0.027 | 0.031 | 0.026 | 0.025 | 0.03 | 0.015333 | 0.027667 | 0.027 | 0.021167 | 322.865 | 0.0065 | 0.00010581
323.208 | 0.011 | 323.208 | 0.011 | 0.015 | 0.018 | 0.024 | 0.024 | 0.024 | 0.026 | 0.025 | 0.027 | 0.014667 | 0.024 | 0.026 | 0.020333 | 323.208 | 0.0036667 | 5.9688e-05
323.551 | 0.02 | 323.551 | 0.02 | 0.025 | 0.023 | 0.033 | 0.032 | 0.037 | 0.035 | 0.034 | 0.036 | 0.022667 | 0.034 | 0.035 | 0.028833 | 323.551 | 0.0051667 | 8.4106e-05
323.894 | 0.01 | 323.894 | 0.01 | 0.011 | 0.013 | 0.018 | 0.02 | 0.021 | 0.023 | 0.023 | 0.023 | 0.011333 | 0.019667 | 0.023 | 0.017167 | 323.894 | 0.0025 | 4.0696e-05
324.237 | 0.005 | 324.237 | 0.005 | 0.006 | 0.01 | 0.013 | 0.014 | 0.019 | 0.019 | 0.019 | 0.022 | 0.007 | 0.015333 | 0.02 | 0.0135 | 324.237 | 0.0018333 | 2.9844e-05
324.58 | 0.011 | 324.58 | 0.011 | 0.017 | 0.015 | 0.017 | 0.023 | 0.023 | 0.022 | 0.026 | 0.026 | 0.014333 | 0.021 | 0.024667 | 0.0195 | 324.58 | 0.0015 | 2.4418e-05
324.922 | 0.015 | 324.922 | 0.015 | 0.014 | 0.017 | 0.025 | 0.025 | 0.026 | 0.025 | 0.025 | 0.03 | 0.015333 | 0.025333 | 0.026667 | 0.021 | 324.922 | 0.0043333 | 7.054e-05
325.265 | 0.018 | 325.265 | 0.018 | 0.02 | 0.021 | 0.029 | 0.033 | 0.032 | 0.03 | 0.03 | 0.037 | 0.019667 | 0.031333 | 0.032333 | 0.026 | 325.265 | 0.0053333 | 8.6819e-05
325.607 | 0.021 | 325.607 | 0.021 | 0.021 | 0.024 | 0.029 | 0.03 | 0.027 | 0.03 | 0.031 | 0.03 | 0.022 | 0.028667 | 0.030333 | 0.026167 | 325.607 | 0.0025 | 4.0696e-05
325.95 | 0.018 | 325.95 | 0.018 | 0.023 | 0.024 | 0.03 | 0.033 | 0.03 | 0.036 | 0.035 | 0.033 | 0.021667 | 0.031 | 0.034667 | 0.028167 | 325.95 | 0.0028333 | 4.6123e-05
326.292 | -0.003 | 326.292 | -0.003 | 0.003 | 0.001 | 0.009 | 0.016 | 0.013 | 0.01 | 0.015 | 0.017 | 0.00033333 | 0.012667 | 0.014 | 0.0071667 | 326.292 | 0.0055 | 8.9532e-05
326.634 | 0.007 | 326.634 | 0.007 | 0.017 | 0.014 | 0.022 | 0.02 | 0.02 | 0.026 | 0.023 | 0.023 | 0.012667 | 0.020667 | 0.024 | 0.018333 | 326.634 | 0.0023333 | 3.7983e-05
326.976 | 0.008 | 326.976 | 0.008 | 0.006 | 0.01 | 0.016 | 0.019 | 0.019 | 0.017 | 0.016 | 0.021 | 0.008 | 0.018 | 0.018 | 0.013 | 326.976 | 0.005 | 8.1393e-05
327.318 | 0.002 | 327.318 | 0.002 | 0.011 | 0.012 | 0.019 | 0.015 | 0.016 | 0.015 | 0.02 | 0.022 | 0.0083333 | 0.016667 | 0.019 | 0.013667 | 327.318 | 0.003 | 4.8836e-05
327.66 | 0.016 | 327.66 | 0.016 | 0.016 | 0.014 | 0.028 | 0.028 | 0.028 | 0.026 | 0.032 | 0.03 | 0.015333 | 0.028 | 0.029333 | 0.022333 | 327.66 | 0.0056667 | 9.2245e-05
328.002 | -0.006 | 328.002 | -0.006 | 0.002 | -0.002 | 0.008 | 0.004 | 0.013 | 0.01 | 0.008 | 0.012 | -0.002 | 0.0083333 | 0.01 | 0.004 | 328.002 | 0.0043333 | 7.054e-05
328.344 | 0.011 | 328.344 | 0.011 | 0.017 | 0.016 | 0.024 | 0.022 | 0.026 | 0.025 | 0.024 | 0.03 | 0.014667 | 0.024 | 0.026333 | 0.0205 | 328.344 | 0.0035 | 5.6975e-05
328.686 | 0.017 | 328.686 | 0.017 | 0.015 | 0.019 | 0.024 | 0.023 | 0.028 | 0.028 | 0.027 | 0.031 | 0.017 | 0.025 | 0.028667 | 0.022833 | 328.686 | 0.0021667 | 3.527e-05
329.027 | 0.002 | 329.027 | 0.002 | 0.007 | 0.006 | 0.007 | 0.01 | 0.013 | 0.015 | 0.015 | 0.016 | 0.005 | 0.01 | 0.015333 | 0.010167 | 329.027 | -0.00016667 | -2.7131e-06
329.369 | 0.003 | 329.369 | 0.003 | 0.006 | 0.008 | 0.012 | 0.015 | 0.013 | 0.016 | 0.013 | 0.016 | 0.0056667 | 0.013333 | 0.015 | 0.010333 | 329.369 | 0.003 | 4.8836e-05
329.711 | 0.014 | 329.711 | 0.014 | 0.017 | 0.018 | 0.022 | 0.022 | 0.026 | 0.025 | 0.03 | 0.028 | 0.016333 | 0.023333 | 0.027667 | 0.022 | 329.711 | 0.0013333 | 2.1705e-05
330.052 | 0.012 | 330.052 | 0.012 | 0.012 | 0.016 | 0.021 | 0.022 | 0.025 | 0.023 | 0.022 | 0.029 | 0.013333 | 0.022667 | 0.024667 | 0.019 | 330.052 | 0.0036667 | 5.9688e-05
330.393 | 0.002 | 330.393 | 0.002 | 0.005 | 0.008 | 0.012 | 0.015 | 0.015 | 0.014 | 0.015 | 0.015 | 0.005 | 0.014 | 0.014667 | 0.0098333 | 330.393 | 0.0041667 | 6.7827e-05
330.735 | 0.004 | 330.735 | 0.004 | 0.005 | 0.009 | 0.011 | 0.014 | 0.018 | 0.018 | 0.021 | 0.02 | 0.006 | 0.014333 | 0.019667 | 0.012833 | 330.735 | 0.0015 | 2.4418e-05
331.076 | 0.01 | 331.076 | 0.01 | 0.015 | 0.016 | 0.023 | 0.023 | 0.023 | 0.024 | 0.023 | 0.023 | 0.013667 | 0.023 | 0.023333 | 0.0185 | 331.076 | 0.0045 | 7.3254e-05
331.417 | 0.005 | 331.417 | 0.005 | 0.007 | 0.009 | 0.014 | 0.013 | 0.017 | 0.014 | 0.016 | 0.019 | 0.007 | 0.014667 | 0.016333 | 0.011667 | 331.417 | 0.003 | 4.8836e-05
331.758 | 0.013 | 331.758 | 0.013 | 0.016 | 0.015 | 0.023 | 0.022 | 0.023 | 0.027 | 0.024 | 0.025 | 0.014667 | 0.022667 | 0.025333 | 0.02 | 331.758 | 0.0026667 | 4.3409e-05
332.099 | -0.003 | 332.099 | -0.003 | 0.002 | 0.001 | 0.007 | 0.008 | 0.009 | 0.011 | 0.011 | 0.014 | 0 | 0.008 | 0.012 | 0.006 | 332.099 | 0.002 | 3.2557e-05
332.44 | 0.015 | 332.44 | 0.015 | 0.017 | 0.018 | 0.025 | 0.025 | 0.025 | 0.028 | 0.027 | 0.031 | 0.016667 | 0.025 | 0.028667 | 0.022667 | 332.44 | 0.0023333 | 3.7983e-05
332.78 | -0.002 | 332.78 | -0.002 | 0.006 | 0.006 | 0.013 | 0.012 | 0.017 | 0.015 | 0.014 | 0.017 | 0.0033333 | 0.014 | 0.015333 | 0.0093333 | 332.78 | 0.0046667 | 7.5967e-05
333.121 | 0.011 | 333.121 | 0.011 | 0.017 | 0.014 | 0.032 | 0.031 | 0.033 | 0.022 | 0.023 | 0.029 | 0.014 | 0.032 | 0.024667 | 0.019333 | 333.121 | 0.012667 | 0.0002062
333.462 | 0.008 | 333.462 | 0.008 | 0.007 | 0.013 | 0.048 | 0.051 | 0.051 | 0.018 | 0.019 | 0.024 | 0.0093333 | 0.05 | 0.020333 | 0.014833 | 333.462 | 0.035167 | 0.00057246
333.802 | 0.013 | 333.802 | 0.013 | 0.015 | 0.013 | 0.086 | 0.086 | 0.086 | 0.023 | 0.025 | 0.031 | 0.013667 | 0.086 | 0.026333 | 0.02 | 333.802 | 0.066 | 0.0010744
334.143 | 0.01 | 334.143 | 0.01 | 0.009 | 0.011 | 0.07 | 0.068 | 0.071 | 0.019 | 0.019 | 0.026 | 0.01 | 0.069667 | 0.021333 | 0.015667 | 334.143 | 0.054 | 0.00087904
334.483 | 0.005 | 334.483 | 0.005 | 0.011 | 0.012 | 0.028 | 0.028 | 0.03 | 0.021 | 0.019 | 0.018 | 0.0093333 | 0.028667 | 0.019333 | 0.014333 | 334.483 | 0.014333 | 0.00023333
334.823 | 0 | 334.823 | 0 | 0.003 | 0.007 | 0.01 | 0.011 | 0.016 | 0.012 | 0.016 | 0.019 | 0.0033333 | 0.012333 | 0.015667 | 0.0095 | 334.823 | 0.0028333 | 4.6123e-05
335.164 | 0.014 | 335.164 | 0.014 | 0.014 | 0.015 | 0.022 | 0.021 | 0.025 | 0.024 | 0.026 | 0.031 | 0.014333 | 0.022667 | 0.027 | 0.020667 | 335.164 | 0.002 | 3.2557e-05
335.504 | 0.004 | 335.504 | 0.004 | 0.004 | 0.008 | 0.011 | 0.012 | 0.013 | 0.019 | 0.017 | 0.018 | 0.0053333 | 0.012 | 0.018 | 0.011667 | 335.504 | 0.00033333 | 5.4262e-06
335.844 | 0.014 | 335.844 | 0.014 | 0.017 | 0.017 | 0.024 | 0.026 | 0.027 | 0.029 | 0.028 | 0.026 | 0.016 | 0.025667 | 0.027667 | 0.021833 | 335.844 | 0.0038333 | 6.2401e-05
336.184 | 0.002 | 336.184 | 0.002 | 0.005 | 0.006 | 0.01 | 0.013 | 0.015 | 0.015 | 0.016 | 0.014 | 0.0043333 | 0.012667 | 0.015 | 0.0096667 | 336.184 | 0.003 | 4.8836e-05
336.524 | 0.007 | 336.524 | 0.007 | 0.006 | 0.011 | 0.014 | 0.012 | 0.014 | 0.014 | 0.014 | 0.023 | 0.008 | 0.013333 | 0.017 | 0.0125 | 336.524 | 0.00083333 | 1.3565e-05
336.864 | 0.01 | 336.864 | 0.01 | 0.009 | 0.014 | 0.013 | 0.017 | 0.017 | 0.02 | 0.018 | 0.02 | 0.011 | 0.015667 | 0.019333 | 0.015167 | 336.864 | 0.0005 | 8.1393e-06
337.203 | 0.007 | 337.203 | 0.007 | 0.011 | 0.011 | 0.015 | 0.018 | 0.018 | 0.02 | 0.018 | 0.021 | 0.0096667 | 0.017 | 0.019667 | 0.014667 | 337.203 | 0.0023333 | 3.7983e-05
337.543 | 0.017 | 337.543 | 0.017 | 0.022 | 0.022 | 0.025 | 0.028 | 0.033 | 0.032 | 0.032 | 0.033 | 0.020333 | 0.028667 | 0.032333 | 0.026333 | 337.543 | 0.0023333 | 3.7983e-05
337.883 | 0.011 | 337.883 | 0.011 | 0.016 | 0.014 | 0.018 | 0.022 | 0.022 | 0.02 | 0.025 | 0.026 | 0.013667 | 0.020667 | 0.023667 | 0.018667 | 337.883 | 0.002 | 3.2557e-05
338.222 | 0.022 | 338.222 | 0.022 | 0.022 | 0.024 | 0.029 | 0.035 | 0.034 | 0.034 | 0.038 | 0.034 | 0.022667 | 0.032667 | 0.035333 | 0.029 | 338.222 | 0.0036667 | 5.9688e-05
338.562 | 0.005 | 338.562 | 0.005 | 0.009 | 0.005 | 0.012 | 0.011 | 0.016 | 0.016 | 0.014 | 0.019 | 0.0063333 | 0.013 | 0.016333 | 0.011333 | 338.562 | 0.0016667 | 2.7131e-05
338.901 | 0.015 | 338.901 | 0.015 | 0.016 | 0.018 | 0.02 | 0.027 | 0.027 | 0.027 | 0.026 | 0.026 | 0.016333 | 0.024667 | 0.026333 | 0.021333 | 338.901 | 0.0033333 | 5.4262e-05
339.24 | -0.003 | 339.24 | -0.003 | -0.002 | 0.001 | 0.009 | 0.009 | 0.01 | 0.007 | 0.011 | 0.012 | -0.0013333 | 0.0093333 | 0.01 | 0.0043333 | 339.24 | 0.005 | 8.1393e-05
339.58 | 0.007 | 339.58 | 0.007 | 0.009 | 0.013 | 0.018 | 0.014 | 0.019 | 0.019 | 0.024 | 0.028 | 0.0096667 | 0.017 | 0.023667 | 0.016667 | 339.58 | 0.00033333 | 5.4262e-06
339.919 | 0.005 | 339.919 | 0.005 | 0.007 | 0.005 | 0.014 | 0.015 | 0.015 | 0.015 | 0.014 | 0.018 | 0.0056667 | 0.014667 | 0.015667 | 0.010667 | 339.919 | 0.004 | 6.5114e-05
340.258 | 0.01 | 340.258 | 0.01 | 0.016 | 0.017 | 0.023 | 0.022 | 0.025 | 0.027 | 0.027 | 0.032 | 0.014333 | 0.023333 | 0.028667 | 0.0215 | 340.258 | 0.0018333 | 2.9844e-05
340.597 | 0.008 | 340.597 | 0.008 | 0.01 | 0.013 | 0.019 | 0.016 | 0.018 | 0.019 | 0.019 | 0.02 | 0.010333 | 0.017667 | 0.019333 | 0.014833 | 340.597 | 0.0028333 | 4.6123e-05
340.936 | 0.006 | 340.936 | 0.006 | 0.003 | 0.005 | 0.012 | 0.012 | 0.014 | 0.011 | 0.014 | 0.014 | 0.0046667 | 0.012667 | 0.013 | 0.0088333 | 340.936 | 0.0038333 | 6.2401e-05
341.274 | 0.007 | 341.274 | 0.007 | 0.012 | 0.011 | 0.017 | 0.021 | 0.022 | 0.023 | 0.022 | 0.024 | 0.01 | 0.02 | 0.023 | 0.0165 | 341.274 | 0.0035 | 5.6975e-05
341.613 | 0.013 | 341.613 | 0.013 | 0.016 | 0.021 | 0.025 | 0.026 | 0.026 | 0.024 | 0.028 | 0.03 | 0.016667 | 0.025667 | 0.027333 | 0.022 | 341.613 | 0.0036667 | 5.9688e-05
341.952 | 0.01 | 341.952 | 0.01 | 0.012 | 0.017 | 0.018 | 0.022 | 0.023 | 0.023 | 0.021 | 0.022 | 0.013 | 0.021 | 0.022 | 0.0175 | 341.952 | 0.0035 | 5.6975e-05
342.29 | 0.011 | 342.29 | 0.011 | 0.011 | 0.012 | 0.018 | 0.02 | 0.021 | 0.021 | 0.019 | 0.022 | 0.011333 | 0.019667 | 0.020667 | 0.016 | 342.29 | 0.0036667 | 5.9688e-05
342.629 | 0.021 | 342.629 | 0.021 | 0.021 | 0.024 | 0.026 | 0.031 | 0.027 | 0.031 | 0.028 | 0.029 | 0.022 | 0.028 | 0.029333 | 0.025667 | 342.629 | 0.0023333 | 3.7983e-05
342.967 | 0.007 | 342.967 | 0.007 | 0.011 | 0.009 | 0.01 | 0.01 | 0.018 | 0.017 | 0.018 | 0.021 | 0.009 | 0.012667 | 0.018667 | 0.013833 | 342.967 | -0.0011667 | -1.8992e-05
343.306 | 0.008 | 343.306 | 0.008 | 0.007 | 0.014 | 0.018 | 0.017 | 0.02 | 0.021 | 0.022 | 0.025 | 0.0096667 | 0.018333 | 0.022667 | 0.016167 | 343.306 | 0.0021667 | 3.527e-05
343.644 | 0.012 | 343.644 | 0.012 | 0.011 | 0.015 | 0.019 | 0.026 | 0.021 | 0.024 | 0.02 | 0.024 | 0.012667 | 0.022 | 0.022667 | 0.017667 | 343.644 | 0.0043333 | 7.054e-05
343.982 | 0.003 | 343.982 | 0.003 | 0.005 | 0.011 | 0.014 | 0.01 | 0.019 | 0.018 | 0.019 | 0.018 | 0.0063333 | 0.014333 | 0.018333 | 0.012333 | 343.982 | 0.002 | 3.2557e-05
344.32 | 0.01 | 344.32 | 0.01 | 0.014 | 0.015 | 0.023 | 0.023 | 0.021 | 0.026 | 0.026 | 0.028 | 0.013 | 0.022333 | 0.026667 | 0.019833 | 344.32 | 0.0025 | 4.0696e-05
344.658 | 0.004 | 344.658 | 0.004 | 0.006 | 0.007 | 0.012 | 0.013 | 0.013 | 0.017 | 0.017 | 0.018 | 0.0056667 | 0.012667 | 0.017333 | 0.0115 | 344.658 | 0.0011667 | 1.8992e-05
344.996 | 0.007 | 344.996 | 0.007 | 0.012 | 0.013 | 0.016 | 0.018 | 0.019 | 0.019 | 0.019 | 0.024 | 0.010667 | 0.017667 | 0.020667 | 0.015667 | 344.996 | 0.002 | 3.2557e-05
345.334 | 0.007 | 345.334 | 0.007 | 0.007 | 0.008 | 0.015 | 0.014 | 0.017 | 0.017 | 0.022 | 0.022 | 0.0073333 | 0.015333 | 0.020333 | 0.013833 | 345.334 | 0.0015 | 2.4418e-05
345.672 | 0.014 | 345.672 | 0.014 | 0.013 | 0.02 | 0.027 | 0.023 | 0.027 | 0.027 | 0.028 | 0.033 | 0.015667 | 0.025667 | 0.029333 | 0.0225 | 345.672 | 0.0031667 | 5.1549e-05
346.01 | 0.003 | 346.01 | 0.003 | 0.007 | 0.006 | 0.007 | 0.012 | 0.018 | 0.014 | 0.014 | 0.015 | 0.0053333 | 0.012333 | 0.014333 | 0.0098333 | 346.01 | 0.0025 | 4.0696e-05
346.347 | 0.021 | 346.347 | 0.021 | 0.024 | 0.03 | 0.032 | 0.032 | 0.037 | 0.036 | 0.037 | 0.038 | 0.025 | 0.033667 | 0.037 | 0.031 | 346.347 | 0.0026667 | 4.3409e-05
346.685 | 0.01 | 346.685 | 0.01 | 0.01 | 0.014 | 0.019 | 0.023 | 0.022 | 0.024 | 0.026 | 0.024 | 0.011333 | 0.021333 | 0.024667 | 0.018 | 346.685 | 0.0033333 | 5.4262e-05
347.022 | 0.015 | 347.022 | 0.015 | 0.023 | 0.022 | 0.025 | 0.027 | 0.027 | 0.028 | 0.036 | 0.032 | 0.02 | 0.026333 | 0.032 | 0.026 | 347.022 | 0.00033333 | 5.4262e-06
347.36 | 0.018 | 347.36 | 0.018 | 0.015 | 0.023 | 0.025 | 0.027 | 0.028 | 0.028 | 0.032 | 0.031 | 0.018667 | 0.026667 | 0.030333 | 0.0245 | 347.36 | 0.0021667 | 3.527e-05
347.697 | 0.009 | 347.697 | 0.009 | 0.015 | 0.015 | 0.017 | 0.017 | 0.024 | 0.024 | 0.02 | 0.025 | 0.013 | 0.019333 | 0.023 | 0.018 | 347.697 | 0.0013333 | 2.1705e-05
348.034 | 0.011 | 348.034 | 0.011 | 0.012 | 0.012 | 0.019 | 0.021 | 0.022 | 0.026 | 0.029 | 0.027 | 0.011667 | 0.020667 | 0.027333 | 0.0195 | 348.034 | 0.0011667 | 1.8992e-05
348.371 | 0.012 | 348.371 | 0.012 | 0.01 | 0.016 | 0.017 | 0.021 | 0.023 | 0.019 | 0.022 | 0.025 | 0.012667 | 0.020333 | 0.022 | 0.017333 | 348.371 | 0.003 | 4.8836e-05
348.708 | 0.01 | 348.708 | 0.01 | 0.011 | 0.011 | 0.015 | 0.019 | 0.018 | 0.021 | 0.021 | 0.022 | 0.010667 | 0.017333 | 0.021333 | 0.016 | 348.708 | 0.0013333 | 2.1705e-05
349.045 | 0.017 | 349.045 | 0.017 | 0.016 | 0.02 | 0.022 | 0.026 | 0.03 | 0.026 | 0.029 | 0.029 | 0.017667 | 0.026 | 0.028 | 0.022833 | 349.045 | 0.0031667 | 5.1549e-05
349.382 | 0.009 | 349.382 | 0.009 | 0.017 | 0.018 | 0.023 | 0.021 | 0.024 | 0.026 | 0.026 | 0.028 | 0.014667 | 0.022667 | 0.026667 | 0.020667 | 349.382 | 0.002 | 3.2557e-05
349.719 | -0.001 | 349.719 | -0.001 | 0.004 | 0.008 | 0.011 | 0.014 | 0.016 | 0.015 | 0.011 | 0.019 | 0.0036667 | 0.013667 | 0.015 | 0.0093333 | 349.719 | 0.0043333 | 7.054e-05
350.056 | 0.01 | 350.056 | 0.01 | 0.009 | 0.011 | 0.018 | 0.014 | 0.02 | 0.018 | 0.023 | 0.024 | 0.01 | 0.017333 | 0.021667 | 0.015833 | 350.056 | 0.0015 | 2.4418e-05
350.393 | 0.009 | 350.393 | 0.009 | 0.011 | 0.009 | 0.013 | 0.017 | 0.016 | 0.017 | 0.017 | 0.02 | 0.0096667 | 0.015333 | 0.018 | 0.013833 | 350.393 | 0.0015 | 2.4418e-05
350.729 | 0.015 | 350.729 | 0.015 | 0.019 | 0.021 | 0.025 | 0.025 | 0.031 | 0.031 | 0.034 | 0.034 | 0.018333 | 0.027 | 0.033 | 0.025667 | 350.729 | 0.0013333 | 2.1705e-05
351.066 | 0.005 | 351.066 | 0.005 | 0.007 | 0.013 | 0.015 | 0.015 | 0.02 | 0.016 | 0.02 | 0.024 | 0.0083333 | 0.016667 | 0.02 | 0.014167 | 351.066 | 0.0025 | 4.0696e-05
351.402 | 0.01 | 351.402 | 0.01 | 0.011 | 0.015 | 0.018 | 0.021 | 0.018 | 0.023 | 0.024 | 0.026 | 0.012 | 0.019 | 0.024333 | 0.018167 | 351.402 | 0.00083333 | 1.3565e-05
351.739 | 0.019 | 351.739 | 0.019 | 0.021 | 0.023 | 0.028 | 0.032 | 0.033 | 0.029 | 0.032 | 0.034 | 0.021 | 0.031 | 0.031667 | 0.026333 | 351.739 | 0.0046667 | 7.5967e-05
352.075 | 0.014 | 352.075 | 0.014 | 0.016 | 0.019 | 0.024 | 0.024 | 0.026 | 0.025 | 0.028 | 0.029 | 0.016333 | 0.024667 | 0.027333 | 0.021833 | 352.075 | 0.0028333 | 4.6123e-05
352.411 | 0.008 | 352.411 | 0.008 | 0.01 | 0.012 | 0.019 | 0.021 | 0.018 | 0.019 | 0.022 | 0.022 | 0.01 | 0.019333 | 0.021 | 0.0155 | 352.411 | 0.0038333 | 6.2401e-05
352.747 | 0.007 | 352.747 | 0.007 | 0.009 | 0.011 | 0.013 | 0.016 | 0.015 | 0.019 | 0.02 | 0.022 | 0.009 | 0.014667 | 0.020333 | 0.014667 | 352.747 | 0 | 0
353.083 | 0.007 | 353.083 | 0.007 | 0.013 | 0.014 | 0.018 | 0.016 | 0.02 | 0.022 | 0.024 | 0.024 | 0.011333 | 0.018 | 0.023333 | 0.017333 | 353.083 | 0.00066667 | 1.0852e-05
353.419 | 0.003 | 353.419 | 0.003 | 0.011 | 0.01 | 0.017 | 0.015 | 0.019 | 0.022 | 0.023 | 0.023 | 0.008 | 0.017 | 0.022667 | 0.015333 | 353.419 | 0.0016667 | 2.7131e-05
353.755 | 0.009 | 353.755 | 0.009 | 0.008 | 0.011 | 0.019 | 0.022 | 0.018 | 0.019 | 0.021 | 0.025 | 0.0093333 | 0.019667 | 0.021667 | 0.0155 | 353.755 | 0.0041667 | 6.7827e-05
354.091 | 0.014 | 354.091 | 0.014 | 0.017 | 0.013 | 0.02 | 0.024 | 0.027 | 0.023 | 0.027 | 0.029 | 0.014667 | 0.023667 | 0.026333 | 0.0205 | 354.091 | 0.0031667 | 5.1549e-05
354.427 | 0.013 | 354.427 | 0.013 | 0.011 | 0.012 | 0.016 | 0.021 | 0.024 | 0.026 | 0.023 | 0.025 | 0.012 | 0.020333 | 0.024667 | 0.018333 | 354.427 | 0.002 | 3.2557e-05
354.763 | 0.013 | 354.763 | 0.013 | 0.015 | 0.015 | 0.02 | 0.022 | 0.023 | 0.024 | 0.024 | 0.025 | 0.014333 | 0.021667 | 0.024333 | 0.019333 | 354.763 | 0.0023333 | 3.7983e-05
355.098 | 0.015 | 355.098 | 0.015 | 0.018 | 0.019 | 0.021 | 0.026 | 0.029 | 0.025 | 0.026 | 0.028 | 0.017333 | 0.025333 | 0.026333 | 0.021833 | 355.098 | 0.0035 | 5.6975e-05
355.434 | 0.002 | 355.434 | 0.002 | 0.004 | 0.004 | 0.009 | 0.012 | 0.017 | 0.014 | 0.017 | 0.017 | 0.0033333 | 0.012667 | 0.016 | 0.0096667 | 355.434 | 0.003 | 4.8836e-05
355.769 | 0.019 | 355.769 | 0.019 | 0.025 | 0.025 | 0.032 | 0.034 | 0.035 | 0.032 | 0.037 | 0.039 | 0.023 | 0.033667 | 0.036 | 0.0295 | 355.769 | 0.0041667 | 6.7827e-05
356.104 | 0.001 | 356.104 | 0.001 | 0 | 0.004 | 0.008 | 0.008 | 0.011 | 0.012 | 0.012 | 0.015 | 0.0016667 | 0.009 | 0.013 | 0.0073333 | 356.104 | 0.0016667 | 2.7131e-05
356.44 | 0.002 | 356.44 | 0.002 | 0.001 | 0.007 | 0.012 | 0.014 | 0.016 | 0.019 | 0.016 | 0.02 | 0.0033333 | 0.014 | 0.018333 | 0.010833 | 356.44 | 0.0031667 | 5.1549e-05
356.775 | 0.01 | 356.775 | 0.01 | 0.015 | 0.015 | 0.025 | 0.024 | 0.026 | 0.022 | 0.026 | 0.026 | 0.013333 | 0.025 | 0.024667 | 0.019 | 356.775 | 0.006 | 9.7671e-05
357.11 | 0.012 | 357.11 | 0.012 | 0.013 | 0.014 | 0.021 | 0.021 | 0.023 | 0.021 | 0.02 | 0.024 | 0.013 | 0.021667 | 0.021667 | 0.017333 | 357.11 | 0.0043333 | 7.054e-05
357.445 | 0.007 | 357.445 | 0.007 | 0.009 | 0.011 | 0.017 | 0.02 | 0.02 | 0.019 | 0.02 | 0.025 | 0.009 | 0.019 | 0.021333 | 0.015167 | 357.445 | 0.0038333 | 6.2401e-05
357.78 | 0.003 | 357.78 | 0.003 | 0.006 | 0.007 | 0.014 | 0.01 | 0.015 | 0.02 | 0.019 | 0.018 | 0.0053333 | 0.013 | 0.019 | 0.012167 | 357.78 | 0.00083333 | 1.3565e-05
358.115 | 0.013 | 358.115 | 0.013 | 0.016 | 0.018 | 0.019 | 0.025 | 0.027 | 0.026 | 0.024 | 0.027 | 0.015667 | 0.023667 | 0.025667 | 0.020667 | 358.115 | 0.003 | 4.8836e-05
358.45 | 0.015 | 358.45 | 0.015 | 0.015 | 0.016 | 0.022 | 0.024 | 0.023 | 0.021 | 0.026 | 0.028 | 0.015333 | 0.023 | 0.025 | 0.020167 | 358.45 | 0.0028333 | 4.6123e-05
358.785 | 0.013 | 358.785 | 0.013 | 0.014 | 0.015 | 0.026 | 0.025 | 0.026 | 0.023 | 0.028 | 0.027 | 0.014 | 0.025667 | 0.026 | 0.02 | 358.785 | 0.0056667 | 9.2245e-05
359.119 | 0.003 | 359.119 | 0.003 | 0.005 | 0.007 | 0.013 | 0.015 | 0.016 | 0.014 | 0.018 | 0.016 | 0.005 | 0.014667 | 0.016 | 0.0105 | 359.119 | 0.0041667 | 6.7827e-05
359.454 | 0.003 | 359.454 | 0.003 | 0.009 | 0.014 | 0.016 | 0.018 | 0.019 | 0.015 | 0.02 | 0.023 | 0.0086667 | 0.017667 | 0.019333 | 0.014 | 359.454 | 0.0036667 | 5.9688e-05
359.788 | 0.009 | 359.788 | 0.009 | 0.013 | 0.013 | 0.021 | 0.017 | 0.023 | 0.019 | 0.025 | 0.023 | 0.011667 | 0.020333 | 0.022333 | 0.017 | 359.788 | 0.0033333 | 5.4262e-05
360.123 | 0.011 | 360.123 | 0.011 | 0.015 | 0.011 | 0.021 | 0.021 | 0.023 | 0.021 | 0.024 | 0.024 | 0.012333 | 0.021667 | 0.023 | 0.017667 | 360.123 | 0.004 | 6.5114e-05
360.457 | 0.001 | 360.457 | 0.001 | 0.003 | 0.004 | 0.013 | 0.011 | 0.014 | 0.011 | 0.014 | 0.016 | 0.0026667 | 0.012667 | 0.013667 | 0.0081667 | 360.457 | 0.0045 | 7.3254e-05
360.791 | 0.016 | 360.791 | 0.016 | 0.021 | 0.019 | 0.028 | 0.03 | 0.024 | 0.03 | 0.031 | 0.035 | 0.018667 | 0.027333 | 0.032 | 0.025333 | 360.791 | 0.002 | 3.2557e-05
361.126 | 0.001 | 361.126 | 0.001 | 0.005 | 0.007 | 0.011 | 0.016 | 0.012 | 0.015 | 0.017 | 0.017 | 0.0043333 | 0.013 | 0.016333 | 0.010333 | 361.126 | 0.0026667 | 4.3409e-05
361.46 | 0.008 | 361.46 | 0.008 | 0.005 | 0.008 | 0.016 | 0.017 | 0.017 | 0.015 | 0.014 | 0.016 | 0.007 | 0.016667 | 0.015 | 0.011 | 361.46 | 0.0056667 | 9.2245e-05
361.794 | 0.011 | 361.794 | 0.011 | 0.012 | 0.014 | 0.025 | 0.026 | 0.023 | 0.026 | 0.028 | 0.029 | 0.012333 | 0.024667 | 0.027667 | 0.02 | 361.794 | 0.0046667 | 7.5967e-05
362.128 | 0.016 | 362.128 | 0.016 | 0.017 | 0.018 | 0.024 | 0.031 | 0.034 | 0.029 | 0.029 | 0.031 | 0.017 | 0.029667 | 0.029667 | 0.023333 | 362.128 | 0.0063333 | 0.0001031
362.462 | 0.018 | 362.462 | 0.018 | 0.025 | 0.024 | 0.037 | 0.032 | 0.035 | 0.036 | 0.036 | 0.037 | 0.022333 | 0.034667 | 0.036333 | 0.029333 | 362.462 | 0.0053333 | 8.6819e-05
362.795 | 0.017 | 362.795 | 0.017 | 0.023 | 0.021 | 0.03 | 0.032 | 0.029 | 0.032 | 0.031 | 0.035 | 0.020333 | 0.030333 | 0.032667 | 0.0265 | 362.795 | 0.0038333 | 6.2401e-05
363.129 | 0.01 | 363.129 | 0.01 | 0.014 | 0.015 | 0.025 | 0.03 | 0.027 | 0.023 | 0.027 | 0.025 | 0.013 | 0.027333 | 0.025 | 0.019 | 363.129 | 0.0083333 | 0.00013565
363.463 | -0.001 | 363.463 | -0.001 | 0.001 | 0.004 | 0.016 | 0.019 | 0.021 | 0.012 | 0.013 | 0.014 | 0.0013333 | 0.018667 | 0.013 | 0.0071667 | 363.463 | 0.0115 | 0.0001872
363.796 | 0.013 | 363.796 | 0.013 | 0.016 | 0.019 | 0.053 | 0.051 | 0.054 | 0.024 | 0.028 | 0.028 | 0.016 | 0.052667 | 0.026667 | 0.021333 | 363.796 | 0.031333 | 0.00051006
364.13 | 0.021 | 364.13 | 0.021 | 0.022 | 0.024 | 0.187 | 0.191 | 0.191 | 0.034 | 0.033 | 0.035 | 0.022333 | 0.18967 | 0.034 | 0.028167 | 364.13 | 0.1615 | 0.002629
364.463 | 0.015 | 364.463 | 0.015 | 0.016 | 0.018 | 0.524 | 0.526 | 0.527 | 0.026 | 0.029 | 0.033 | 0.016333 | 0.52567 | 0.029333 | 0.022833 | 364.463 | 0.50283 | 0.0081854
364.797 | 0.014 | 364.797 | 0.014 | 0.017 | 0.016 | 0.746 | 0.737 | 0.741 | 0.028 | 0.028 | 0.028 | 0.015667 | 0.74133 | 0.028 | 0.021833 | 364.797 | 0.7195 | 0.011712
365.13 | 0.015 | 365.13 | 0.015 | 0.02 | 0.02 | 0.597 | 0.59 | 0.595 | 0.029 | 0.025 | 0.03 | 0.018333 | 0.594 | 0.028 | 0.023167 | 365.13 | 0.57083 | 0.0092923
365.463 | 0.001 | 365.463 | 0.001 | 0.005 | 0.006 | 0.256 | 0.254 | 0.253 | 0.017 | 0.015 | 0.016 | 0.004 | 0.25433 | 0.016 | 0.01 | 365.463 | 0.24433 | 0.0039774
365.796 | 0.009 | 365.796 | 0.009 | 0.009 | 0.007 | 0.165 | 0.166 | 0.164 | 0.018 | 0.022 | 0.019 | 0.0083333 | 0.165 | 0.019667 | 0.014 | 365.796 | 0.151 | 0.0024581
366.129 | 0.012 | 366.129 | 0.012 | 0.015 | 0.015 | 0.147 | 0.148 | 0.148 | 0.022 | 0.023 | 0.026 | 0.014 | 0.14767 | 0.023667 | 0.018833 | 366.129 | 0.12883 | 0.0020972
366.462 | 0.004 | 366.462 | 0.004 | 0.008 | 0.006 | 0.087 | 0.083 | 0.09 | 0.012 | 0.017 | 0.015 | 0.006 | 0.086667 | 0.014667 | 0.010333 | 366.462 | 0.076333 | 0.0012426
366.795 | 0.006 | 366.795 | 0.006 | 0.008 | 0.011 | 0.03 | 0.037 | 0.036 | 0.023 | 0.022 | 0.024 | 0.0083333 | 0.034333 | 0.023 | 0.015667 | 366.795 | 0.018667 | 0.00030387
367.128 | 0.016 | 367.128 | 0.016 | 0.021 | 0.021 | 0.031 | 0.031 | 0.037 | 0.03 | 0.031 | 0.029 | 0.019333 | 0.033 | 0.03 | 0.024667 | 367.128 | 0.0083333 | 0.00013565
367.461 | 0.004 | 367.461 | 0.004 | 0.008 | 0.004 | 0.018 | 0.017 | 0.02 | 0.019 | 0.017 | 0.023 | 0.0053333 | 0.018333 | 0.019667 | 0.0125 | 367.461 | 0.0058333 | 9.4958e-05
367.793 | 0.021 | 367.793 | 0.021 | 0.018 | 0.025 | 0.029 | 0.033 | 0.032 | 0.026 | 0.032 | 0.032 | 0.021333 | 0.031333 | 0.03 | 0.025667 | 367.793 | 0.0056667 | 9.2245e-05
368.126 | 0.007 | 368.126 | 0.007 | 0.011 | 0.01 | 0.023 | 0.02 | 0.029 | 0.021 | 0.022 | 0.024 | 0.0093333 | 0.024 | 0.022333 | 0.015833 | 368.126 | 0.0081667 | 0.00013294
368.458 | 0.017 | 368.458 | 0.017 | 0.016 | 0.021 | 0.027 | 0.033 | 0.026 | 0.03 | 0.027 | 0.032 | 0.018 | 0.028667 | 0.029667 | 0.023833 | 368.458 | 0.0048333 | 7.868e-05
368.791 | 0.012 | 368.791 | 0.012 | 0.014 | 0.014 | 0.025 | 0.025 | 0.028 | 0.026 | 0.027 | 0.026 | 0.013333 | 0.026 | 0.026333 | 0.019833 | 368.791 | 0.0061667 | 0.00010038
369.123 | 0.012 | 369.123 | 0.012 | 0.013 | 0.015 | 0.024 | 0.025 | 0.025 | 0.026 | 0.029 | 0.027 | 0.013333 | 0.024667 | 0.027333 | 0.020333 | 369.123 | 0.0043333 | 7.054e-05
369.455 | 0.008 | 369.455 | 0.008 | 0.007 | 0.015 | 0.023 | 0.023 | 0.02 | 0.019 | 0.025 | 0.026 | 0.01 | 0.022 | 0.023333 | 0.016667 | 369.455 | 0.0053333 | 8.6819e-05
369.787 | 0.008 | 369.787 | 0.008 | 0.007 | 0.009 | 0.019 | 0.02 | 0.023 | 0.023 | 0.025 | 0.023 | 0.008 | 0.020667 | 0.023667 | 0.015833 | 369.787 | 0.0048333 | 7.868e-05
370.12 | 0.003 | 370.12 | 0.003 | 0.005 | 0.006 | 0.019 | 0.021 | 0.022 | 0.017 | 0.021 | 0.017 | 0.0046667 | 0.020667 | 0.018333 | 0.0115 | 370.12 | 0.0091667 | 0.00014922
370.452 | 0.002 | 370.452 | 0.002 | 0.003 | 0.003 | 0.013 | 0.014 | 0.013 | 0.012 | 0.013 | 0.016 | 0.0026667 | 0.013333 | 0.013667 | 0.0081667 | 370.452 | 0.0051667 | 8.4106e-05
370.784 | 0.005 | 370.784 | 0.005 | 0.008 | 0.011 | 0.015 | 0.021 | 0.019 | 0.014 | 0.02 | 0.02 | 0.008 | 0.018333 | 0.018 | 0.013 | 370.784 | 0.0053333 | 8.6819e-05
371.115 | 0.006 | 371.115 | 0.006 | 0.007 | 0.007 | 0.021 | 0.022 | 0.023 | 0.02 | 0.023 | 0.024 | 0.0066667 | 0.022 | 0.022333 | 0.0145 | 371.115 | 0.0075 | 0.00012209
371.447 | 0.004 | 371.447 | 0.004 | 0 | 0.003 | 0.011 | 0.011 | 0.013 | 0.011 | 0.014 | 0.015 | 0.0023333 | 0.011667 | 0.013333 | 0.0078333 | 371.447 | 0.0038333 | 6.2401e-05
371.779 | 0.007 | 371.779 | 0.007 | 0.011 | 0.009 | 0.021 | 0.021 | 0.019 | 0.021 | 0.024 | 0.024 | 0.009 | 0.020333 | 0.023 | 0.016 | 371.779 | 0.0043333 | 7.054e-05
372.111 | 0.01 | 372.111 | 0.01 | 0.011 | 0.017 | 0.02 | 0.026 | 0.025 | 0.022 | 0.024 | 0.023 | 0.012667 | 0.023667 | 0.023 | 0.017833 | 372.111 | 0.0058333 | 9.4958e-05
372.442 | 0.005 | 372.442 | 0.005 | 0.007 | 0.011 | 0.019 | 0.021 | 0.02 | 0.015 | 0.019 | 0.019 | 0.0076667 | 0.02 | 0.017667 | 0.012667 | 372.442 | 0.0073333 | 0.00011938
372.774 | 0.017 | 372.774 | 0.017 | 0.018 | 0.021 | 0.034 | 0.032 | 0.03 | 0.031 | 0.034 | 0.031 | 0.018667 | 0.032 | 0.032 | 0.025333 | 372.774 | 0.0066667 | 0.00010852
373.105 | 0.011 | 373.105 | 0.011 | 0.012 | 0.016 | 0.023 | 0.026 | 0.023 | 0.022 | 0.026 | 0.026 | 0.013 | 0.024 | 0.024667 | 0.018833 | 373.105 | 0.0051667 | 8.4106e-05
373.436 | 0.006 | 373.436 | 0.006 | 0.009 | 0.011 | 0.02 | 0.02 | 0.023 | 0.017 | 0.017 | 0.021 | 0.0086667 | 0.021 | 0.018333 | 0.0135 | 373.436 | 0.0075 | 0.00012209
373.768 | 0.006 | 373.768 | 0.006 | 0.005 | 0.009 | 0.019 | 0.019 | 0.022 | 0.018 | 0.019 | 0.02 | 0.0066667 | 0.02 | 0.019 | 0.012833 | 373.768 | 0.0071667 | 0.00011666
374.099 | -0.004 | 374.099 | -0.004 | 0.001 | 0.003 | 0.007 | 0.014 | 0.015 | 0.012 | 0.012 | 0.012 | 0 | 0.012 | 0.012 | 0.006 | 374.099 | 0.006 | 9.7671e-05
374.43 | 0.017 | 374.43 | 0.017 | 0.017 | 0.019 | 0.028 | 0.027 | 0.029 | 0.026 | 0.028 | 0.028 | 0.017667 | 0.028 | 0.027333 | 0.0225 | 374.43 | 0.0055 | 8.9532e-05
374.761 | 0.008 | 374.761 | 0.008 | 0.006 | 0.01 | 0.016 | 0.016 | 0.019 | 0.021 | 0.016 | 0.019 | 0.008 | 0.017 | 0.018667 | 0.013333 | 374.761 | 0.0036667 | 5.9688e-05
375.092 | 0.013 | 375.092 | 0.013 | 0.01 | 0.014 | 0.023 | 0.023 | 0.021 | 0.024 | 0.027 | 0.024 | 0.012333 | 0.022333 | 0.025 | 0.018667 | 375.092 | 0.0036667 | 5.9688e-05
375.423 | 0.003 | 375.423 | 0.003 | 0.005 | 0.012 | 0.019 | 0.015 | 0.02 | 0.019 | 0.02 | 0.021 | 0.0066667 | 0.018 | 0.02 | 0.013333 | 375.423 | 0.0046667 | 7.5967e-05
375.754 | 0.004 | 375.754 | 0.004 | 0.006 | 0.01 | 0.019 | 0.018 | 0.021 | 0.019 | 0.017 | 0.02 | 0.0066667 | 0.019333 | 0.018667 | 0.012667 | 375.754 | 0.0066667 | 0.00010852
376.084 | 0.004 | 376.084 | 0.004 | 0.01 | 0.009 | 0.015 | 0.019 | 0.019 | 0.02 | 0.018 | 0.019 | 0.0076667 | 0.017667 | 0.019 | 0.013333 | 376.084 | 0.0043333 | 7.054e-05
376.415 | 0.007 | 376.415 | 0.007 | 0.007 | 0.01 | 0.016 | 0.02 | 0.018 | 0.016 | 0.017 | 0.018 | 0.008 | 0.018 | 0.017 | 0.0125 | 376.415 | 0.0055 | 8.9532e-05
376.745 | 0.008 | 376.745 | 0.008 | 0.008 | 0.013 | 0.015 | 0.02 | 0.02 | 0.017 | 0.019 | 0.022 | 0.0096667 | 0.018333 | 0.019333 | 0.0145 | 376.745 | 0.0038333 | 6.2401e-05
377.076 | 0.007 | 377.076 | 0.007 | 0.015 | 0.014 | 0.022 | 0.022 | 0.027 | 0.021 | 0.024 | 0.027 | 0.012 | 0.023667 | 0.024 | 0.018 | 377.076 | 0.0056667 | 9.2245e-05
377.406 | 0.012 | 377.406 | 0.012 | 0.013 | 0.012 | 0.021 | 0.021 | 0.024 | 0.019 | 0.023 | 0.022 | 0.012333 | 0.022 | 0.021333 | 0.016833 | 377.406 | 0.0051667 | 8.4106e-05
377.737 | 0.021 | 377.737 | 0.021 | 0.025 | 0.024 | 0.037 | 0.034 | 0.037 | 0.036 | 0.035 | 0.036 | 0.023333 | 0.036 | 0.035667 | 0.0295 | 377.737 | 0.0065 | 0.00010581
378.067 | 0.012 | 378.067 | 0.012 | 0.013 | 0.018 | 0.024 | 0.026 | 0.026 | 0.026 | 0.029 | 0.025 | 0.014333 | 0.025333 | 0.026667 | 0.0205 | 378.067 | 0.0048333 | 7.868e-05
378.397 | 0.005 | 378.397 | 0.005 | 0.008 | 0.008 | 0.015 | 0.017 | 0.019 | 0.019 | 0.019 | 0.02 | 0.007 | 0.017 | 0.019333 | 0.013167 | 378.397 | 0.0038333 | 6.2401e-05
378.727 | 0.008 | 378.727 | 0.008 | 0.007 | 0.011 | 0.017 | 0.017 | 0.019 | 0.017 | 0.018 | 0.02 | 0.0086667 | 0.017667 | 0.018333 | 0.0135 | 378.727 | 0.0041667 | 6.7827e-05
379.057 | 0.001 | 379.057 | 0.001 | 0.004 | 0.009 | 0.015 | 0.013 | 0.018 | 0.015 | 0.017 | 0.016 | 0.0046667 | 0.015333 | 0.016 | 0.010333 | 379.057 | 0.005 | 8.1393e-05
379.387 | 0.005 | 379.387 | 0.005 | 0.006 | 0.012 | 0.014 | 0.016 | 0.018 | 0.017 | 0.018 | 0.019 | 0.0076667 | 0.016 | 0.018 | 0.012833 | 379.387 | 0.0031667 | 5.1549e-05
379.717 | 0.007 | 379.717 | 0.007 | 0.011 | 0.011 | 0.018 | 0.02 | 0.02 | 0.017 | 0.024 | 0.02 | 0.0096667 | 0.019333 | 0.020333 | 0.015 | 379.717 | 0.0043333 | 7.054e-05
380.047 | 0.012 | 380.047 | 0.012 | 0.013 | 0.012 | 0.02 | 0.024 | 0.024 | 0.023 | 0.024 | 0.024 | 0.012333 | 0.022667 | 0.023667 | 0.018 | 380.047 | 0.0046667 | 7.5967e-05
380.377 | 0.01 | 380.377 | 0.01 | 0.009 | 0.012 | 0.021 | 0.022 | 0.024 | 0.025 | 0.023 | 0.021 | 0.010333 | 0.022333 | 0.023 | 0.016667 | 380.377 | 0.0056667 | 9.2245e-05
380.706 | 0.015 | 380.706 | 0.015 | 0.01 | 0.019 | 0.023 | 0.027 | 0.024 | 0.024 | 0.028 | 0.027 | 0.014667 | 0.024667 | 0.026333 | 0.0205 | 380.706 | 0.0041667 | 6.7827e-05
381.036 | 0.009 | 381.036 | 0.009 | 0.011 | 0.013 | 0.017 | 0.018 | 0.025 | 0.018 | 0.019 | 0.018 | 0.011 | 0.02 | 0.018333 | 0.014667 | 381.036 | 0.0053333 | 8.6819e-05
381.365 | 0.023 | 381.365 | 0.023 | 0.023 | 0.023 | 0.032 | 0.036 | 0.033 | 0.032 | 0.033 | 0.032 | 0.023 | 0.033667 | 0.032333 | 0.027667 | 381.365 | 0.006 | 9.7671e-05
381.695 | 0.006 | 381.695 | 0.006 | 0.006 | 0.011 | 0.016 | 0.02 | 0.019 | 0.017 | 0.017 | 0.018 | 0.0076667 | 0.018333 | 0.017333 | 0.0125 | 381.695 | 0.0058333 | 9.4958e-05
382.024 | 0.005 | 382.024 | 0.005 | 0.006 | 0.008 | 0.012 | 0.016 | 0.016 | 0.011 | 0.016 | 0.018 | 0.0063333 | 0.014667 | 0.015 | 0.010667 | 382.024 | 0.004 | 6.5114e-05
382.353 | -0.002 | 382.353 | -0.002 | -0.002 | -0.001 | 0.007 | 0.007 | 0.009 | 0.006 | 0.007 | 0.008 | -0.0016667 | 0.0076667 | 0.007 | 0.0026667 | 382.353 | 0.005 | 8.1393e-05
382.682 | 0.009 | 382.682 | 0.009 | 0.008 | 0.01 | 0.019 | 0.02 | 0.02 | 0.021 | 0.023 | 0.021 | 0.009 | 0.019667 | 0.021667 | 0.015333 | 382.682 | 0.0043333 | 7.054e-05
383.012 | 0.004 | 383.012 | 0.004 | 0.002 | 0.005 | 0.016 | 0.016 | 0.017 | 0.015 | 0.019 | 0.018 | 0.0036667 | 0.016333 | 0.017333 | 0.0105 | 383.012 | 0.0058333 | 9.4958e-05
383.341 | 0.013 | 383.341 | 0.013 | 0.014 | 0.015 | 0.021 | 0.026 | 0.029 | 0.026 | 0.028 | 0.026 | 0.014 | 0.025333 | 0.026667 | 0.020333 | 383.341 | 0.005 | 8.1393e-05
383.67 | 0.013 | 383.67 | 0.013 | 0.013 | 0.015 | 0.021 | 0.023 | 0.023 | 0.023 | 0.022 | 0.025 | 0.013667 | 0.022333 | 0.023333 | 0.0185 | 383.67 | 0.0038333 | 6.2401e-05
383.998 | 0.003 | 383.998 | 0.003 | 0.004 | 0.005 | 0.012 | 0.011 | 0.011 | 0.012 | 0.017 | 0.015 | 0.004 | 0.011333 | 0.014667 | 0.0093333 | 383.998 | 0.002 | 3.2557e-05
384.327 | 0.019 | 384.327 | 0.019 | 0.017 | 0.017 | 0.025 | 0.025 | 0.023 | 0.024 | 0.025 | 0.028 | 0.017667 | 0.024333 | 0.025667 | 0.021667 | 384.327 | 0.0026667 | 4.3409e-05
384.656 | 0.004 | 384.656 | 0.004 | 0.005 | 0.006 | 0.012 | 0.017 | 0.014 | 0.013 | 0.014 | 0.016 | 0.005 | 0.014333 | 0.014333 | 0.0096667 | 384.656 | 0.0046667 | 7.5967e-05
384.985 | 0 | 384.985 | 0 | 0 | 0.002 | 0.01 | 0.009 | 0.014 | 0.011 | 0.008 | 0.013 | 0.00066667 | 0.011 | 0.010667 | 0.0056667 | 384.985 | 0.0053333 | 8.6819e-05
385.313 | -0.001 | 385.313 | -0.001 | 0 | 0.002 | 0.01 | 0.012 | 0.01 | 0.01 | 0.014 | 0.011 | 0.00033333 | 0.010667 | 0.011667 | 0.006 | 385.313 | 0.0046667 | 7.5967e-05
385.642 | 0.007 | 385.642 | 0.007 | 0.004 | 0.006 | 0.014 | 0.015 | 0.017 | 0.015 | 0.017 | 0.016 | 0.0056667 | 0.015333 | 0.016 | 0.010833 | 385.642 | 0.0045 | 7.3254e-05
385.97 | 0.008 | 385.97 | 0.008 | 0.008 | 0.009 | 0.014 | 0.016 | 0.017 | 0.014 | 0.016 | 0.016 | 0.0083333 | 0.015667 | 0.015333 | 0.011833 | 385.97 | 0.0038333 | 6.2401e-05
386.298 | 0.004 | 386.298 | 0.004 | 0.005 | 0.005 | 0.014 | 0.016 | 0.016 | 0.016 | 0.017 | 0.016 | 0.0046667 | 0.015333 | 0.016333 | 0.0105 | 386.298 | 0.0048333 | 7.868e-05
386.627 | 0.002 | 386.627 | 0.002 | 0.004 | 0.005 | 0.016 | 0.015 | 0.015 | 0.012 | 0.016 | 0.015 | 0.0036667 | 0.015333 | 0.014333 | 0.009 | 386.627 | 0.0063333 | 0.0001031
386.955 | 0.014 | 386.955 | 0.014 | 0.014 | 0.02 | 0.024 | 0.027 | 0.027 | 0.025 | 0.022 | 0.029 | 0.016 | 0.026 | 0.025333 | 0.020667 | 386.955 | 0.0053333 | 8.6819e-05
387.283 | 0.003 | 387.283 | 0.003 | 0.003 | 0.006 | 0.014 | 0.014 | 0.014 | 0.012 | 0.016 | 0.015 | 0.004 | 0.014 | 0.014333 | 0.0091667 | 387.283 | 0.0048333 | 7.868e-05
387.611 | 0.001 | 387.611 | 0.001 | 0.004 | 0.006 | 0.012 | 0.014 | 0.014 | 0.013 | 0.013 | 0.013 | 0.0036667 | 0.013333 | 0.013 | 0.0083333 | 387.611 | 0.005 | 8.1393e-05
387.939 | 0.007 | 387.939 | 0.007 | 0.008 | 0.009 | 0.014 | 0.018 | 0.019 | 0.016 | 0.018 | 0.017 | 0.008 | 0.017 | 0.017 | 0.0125 | 387.939 | 0.0045 | 7.3254e-05
388.267 | 0.003 | 388.267 | 0.003 | 0.006 | 0.008 | 0.017 | 0.018 | 0.015 | 0.017 | 0.019 | 0.019 | 0.0056667 | 0.016667 | 0.018333 | 0.012 | 388.267 | 0.0046667 | 7.5967e-05
388.595 | 0.015 | 388.595 | 0.015 | 0.018 | 0.02 | 0.026 | 0.028 | 0.026 | 0.027 | 0.026 | 0.026 | 0.017667 | 0.026667 | 0.026333 | 0.022 | 388.595 | 0.0046667 | 7.5967e-05
388.922 | -0.006 | 388.922 | -0.006 | -0.003 | -0.002 | 0.004 | 0.006 | 0.008 | 0.003 | 0.006 | 0.009 | -0.0036667 | 0.006 | 0.006 | 0.0011667 | 388.922 | 0.0048333 | 7.868e-05
389.25 | 0.021 | 389.25 | 0.021 | 0.019 | 0.021 | 0.029 | 0.031 | 0.033 | 0.029 | 0.031 | 0.028 | 0.020333 | 0.031 | 0.029333 | 0.024833 | 389.25 | 0.0061667 | 0.00010038
389.578 | 0.006 | 389.578 | 0.006 | 0.007 | 0.011 | 0.017 | 0.018 | 0.019 | 0.015 | 0.014 | 0.019 | 0.008 | 0.018 | 0.016 | 0.012 | 389.578 | 0.006 | 9.7671e-05
389.905 | 0.01 | 389.905 | 0.01 | 0.011 | 0.011 | 0.021 | 0.024 | 0.024 | 0.021 | 0.02 | 0.021 | 0.010667 | 0.023 | 0.020667 | 0.015667 | 389.905 | 0.0073333 | 0.00011938
390.233 | 0.004 | 390.233 | 0.004 | 0.005 | 0.01 | 0.017 | 0.019 | 0.022 | 0.013 | 0.015 | 0.015 | 0.0063333 | 0.019333 | 0.014333 | 0.010333 | 390.233 | 0.009 | 0.00014651
390.56 | 0.009 | 390.56 | 0.009 | 0.009 | 0.01 | 0.025 | 0.024 | 0.027 | 0.02 | 0.021 | 0.022 | 0.0093333 | 0.025333 | 0.021 | 0.015167 | 390.56 | 0.010167 | 0.0001655
390.887 | 0.017 | 390.887 | 0.017 | 0.015 | 0.022 | 0.029 | 0.029 | 0.027 | 0.029 | 0.026 | 0.032 | 0.018 | 0.028333 | 0.029 | 0.0235 | 390.887 | 0.0048333 | 7.868e-05
391.214 | 0.009 | 391.214 | 0.009 | 0.014 | 0.014 | 0.021 | 0.022 | 0.023 | 0.021 | 0.021 | 0.021 | 0.012333 | 0.022 | 0.021 | 0.016667 | 391.214 | 0.0053333 | 8.6819e-05
391.542 | 0.003 | 391.542 | 0.003 | 0.002 | 0.001 | 0.013 | 0.011 | 0.014 | 0.01 | 0.01 | 0.012 | 0.002 | 0.012667 | 0.010667 | 0.0063333 | 391.542 | 0.0063333 | 0.0001031
391.869 | 0.005 | 391.869 | 0.005 | 0.007 | 0.009 | 0.017 | 0.017 | 0.016 | 0.017 | 0.014 | 0.018 | 0.007 | 0.016667 | 0.016333 | 0.011667 | 391.869 | 0.005 | 8.1393e-05
392.196 | 0 | 392.196 | 0 | 0.002 | 0.004 | 0.009 | 0.011 | 0.015 | 0.01 | 0.011 | 0.012 | 0.002 | 0.011667 | 0.011 | 0.0065 | 392.196 | 0.0051667 | 8.4106e-05
392.522 | -0.001 | 392.522 | -0.001 | 0.001 | 0.004 | 0.013 | 0.013 | 0.01 | 0.011 | 0.012 | 0.011 | 0.0013333 | 0.012 | 0.011333 | 0.0063333 | 392.522 | 0.0056667 | 9.2245e-05
392.849 | 0.005 | 392.849 | 0.005 | 0.006 | 0.006 | 0.014 | 0.016 | 0.017 | 0.016 | 0.017 | 0.015 | 0.0056667 | 0.015667 | 0.016 | 0.010833 | 392.849 | 0.0048333 | 7.868e-05
393.176 | -0.002 | 393.176 | -0.002 | 0.001 | 0.003 | 0.011 | 0.015 | 0.015 | 0.011 | 0.011 | 0.012 | 0.00066667 | 0.013667 | 0.011333 | 0.006 | 393.176 | 0.0076667 | 0.0001248
393.503 | 0.004 | 393.503 | 0.004 | 0.008 | 0.011 | 0.016 | 0.016 | 0.018 | 0.017 | 0.02 | 0.017 | 0.0076667 | 0.016667 | 0.018 | 0.012833 | 393.503 | 0.0038333 | 6.2401e-05
393.829 | 0.003 | 393.829 | 0.003 | 0.007 | 0.009 | 0.014 | 0.017 | 0.016 | 0.015 | 0.016 | 0.016 | 0.0063333 | 0.015667 | 0.015667 | 0.011 | 393.829 | 0.0046667 | 7.5967e-05
394.156 | 0.004 | 394.156 | 0.004 | 0.005 | 0.007 | 0.014 | 0.017 | 0.018 | 0.013 | 0.015 | 0.014 | 0.0053333 | 0.016333 | 0.014 | 0.0096667 | 394.156 | 0.0066667 | 0.00010852
394.482 | 0.009 | 394.482 | 0.009 | 0.009 | 0.014 | 0.018 | 0.016 | 0.021 | 0.019 | 0.019 | 0.021 | 0.010667 | 0.018333 | 0.019667 | 0.015167 | 394.482 | 0.0031667 | 5.1549e-05
394.809 | 0.007 | 394.809 | 0.007 | 0.01 | 0.009 | 0.018 | 0.022 | 0.024 | 0.017 | 0.018 | 0.021 | 0.0086667 | 0.021333 | 0.018667 | 0.013667 | 394.809 | 0.0076667 | 0.0001248
395.135 | 0.012 | 395.135 | 0.012 | 0.011 | 0.014 | 0.021 | 0.022 | 0.023 | 0.021 | 0.021 | 0.022 | 0.012333 | 0.022 | 0.021333 | 0.016833 | 395.135 | 0.0051667 | 8.4106e-05
395.461 | 0.001 | 395.461 | 0.001 | 0.003 | 0.005 | 0.013 | 0.011 | 0.013 | 0.011 | 0.01 | 0.013 | 0.003 | 0.012333 | 0.011333 | 0.0071667 | 395.461 | 0.0051667 | 8.4106e-05
395.787 | 0.012 | 395.787 | 0.012 | 0.014 | 0.016 | 0.02 | 0.025 | 0.024 | 0.023 | 0.022 | 0.023 | 0.014 | 0.023 | 0.022667 | 0.018333 | 395.787 | 0.0046667 | 7.5967e-05
396.113 | -0.001 | 396.113 | -0.001 | 0 | 0.002 | 0.008 | 0.01 | 0.011 | 0.01 | 0.009 | 0.012 | 0.00033333 | 0.0096667 | 0.010333 | 0.0053333 | 396.113 | 0.0043333 | 7.054e-05
396.439 | 0.005 | 396.439 | 0.005 | 0.009 | 0.009 | 0.017 | 0.015 | 0.014 | 0.014 | 0.016 | 0.016 | 0.0076667 | 0.015333 | 0.015333 | 0.0115 | 396.439 | 0.0038333 | 6.2401e-05
396.765 | 0.005 | 396.765 | 0.005 | 0.005 | 0.004 | 0.014 | 0.016 | 0.015 | 0.012 | 0.014 | 0.015 | 0.0046667 | 0.015 | 0.013667 | 0.0091667 | 396.765 | 0.0058333 | 9.4958e-05
397.091 | 0.003 | 397.091 | 0.003 | 0.008 | 0.007 | 0.016 | 0.017 | 0.014 | 0.014 | 0.01 | 0.016 | 0.006 | 0.015667 | 0.013333 | 0.0096667 | 397.091 | 0.006 | 9.7671e-05
397.417 | 0.005 | 397.417 | 0.005 | 0.002 | 0.005 | 0.013 | 0.014 | 0.013 | 0.01 | 0.012 | 0.014 | 0.004 | 0.013333 | 0.012 | 0.008 | 397.417 | 0.0053333 | 8.6819e-05
397.742 | 0.004 | 397.742 | 0.004 | 0.007 | 0.01 | 0.016 | 0.016 | 0.02 | 0.015 | 0.014 | 0.017 | 0.007 | 0.017333 | 0.015333 | 0.011167 | 397.742 | 0.0061667 | 0.00010038
398.068 | 0.014 | 398.068 | 0.014 | 0.014 | 0.013 | 0.021 | 0.027 | 0.025 | 0.024 | 0.021 | 0.022 | 0.013667 | 0.024333 | 0.022333 | 0.018 | 398.068 | 0.0063333 | 0.0001031
398.394 | 0.008 | 398.394 | 0.008 | 0.012 | 0.009 | 0.018 | 0.017 | 0.021 | 0.019 | 0.016 | 0.019 | 0.0096667 | 0.018667 | 0.018 | 0.013833 | 398.394 | 0.0048333 | 7.868e-05
398.719 | 0.003 | 398.719 | 0.003 | 0.007 | 0.005 | 0.017 | 0.015 | 0.018 | 0.014 | 0.013 | 0.018 | 0.005 | 0.016667 | 0.015 | 0.01 | 398.719 | 0.0066667 | 0.00010852
399.044 | 0.005 | 399.044 | 0.005 | 0.005 | 0.007 | 0.013 | 0.017 | 0.017 | 0.016 | 0.013 | 0.016 | 0.0056667 | 0.015667 | 0.015 | 0.010333 | 399.044 | 0.0053333 | 8.6819e-05
399.37 | 0.014 | 399.37 | 0.014 | 0.013 | 0.015 | 0.023 | 0.024 | 0.027 | 0.019 | 0.023 | 0.023 | 0.014 | 0.024667 | 0.021667 | 0.017833 | 399.37 | 0.0068333 | 0.00011124
399.695 | 0.006 | 399.695 | 0.006 | 0.007 | 0.01 | 0.016 | 0.017 | 0.018 | 0.016 | 0.013 | 0.017 | 0.0076667 | 0.017 | 0.015333 | 0.0115 | 399.695 | 0.0055 | 8.9532e-05
400.02 | 0.008 | 400.02 | 0.008 | 0.013 | 0.015 | 0.021 | 0.023 | 0.023 | 0.019 | 0.022 | 0.022 | 0.012 | 0.022333 | 0.021 | 0.0165 | 400.02 | 0.0058333 | 9.4958e-05
400.345 | 0.014 | 400.345 | 0.014 | 0.016 | 0.016 | 0.021 | 0.026 | 0.027 | 0.023 | 0.027 | 0.024 | 0.015333 | 0.024667 | 0.024667 | 0.02 | 400.345 | 0.0046667 | 7.5967e-05
400.67 | 0.011 | 400.67 | 0.011 | 0.012 | 0.014 | 0.023 | 0.021 | 0.022 | 0.019 | 0.021 | 0.022 | 0.012333 | 0.022 | 0.020667 | 0.0165 | 400.67 | 0.0055 | 8.9532e-05
400.995 | 0.014 | 400.995 | 0.014 | 0.015 | 0.021 | 0.03 | 0.029 | 0.029 | 0.025 | 0.027 | 0.028 | 0.016667 | 0.029333 | 0.026667 | 0.021667 | 400.995 | 0.0076667 | 0.0001248
401.32 | 0.009 | 401.32 | 0.009 | 0.01 | 0.013 | 0.019 | 0.019 | 0.02 | 0.017 | 0.017 | 0.021 | 0.010667 | 0.019333 | 0.018333 | 0.0145 | 401.32 | 0.0048333 | 7.868e-05
401.645 | 0.007 | 401.645 | 0.007 | 0.013 | 0.008 | 0.018 | 0.021 | 0.022 | 0.016 | 0.017 | 0.019 | 0.0093333 | 0.020333 | 0.017333 | 0.013333 | 401.645 | 0.007 | 0.00011395
401.969 | 0.002 | 401.969 | 0.002 | 0.006 | 0.006 | 0.016 | 0.015 | 0.014 | 0.011 | 0.014 | 0.015 | 0.0046667 | 0.015 | 0.013333 | 0.009 | 401.969 | 0.006 | 9.7671e-05
402.294 | 0.005 | 402.294 | 0.005 | 0.006 | 0.007 | 0.015 | 0.018 | 0.019 | 0.012 | 0.014 | 0.016 | 0.006 | 0.017333 | 0.014 | 0.01 | 402.294 | 0.0073333 | 0.00011938
402.618 | 0.015 | 402.618 | 0.015 | 0.02 | 0.02 | 0.031 | 0.033 | 0.038 | 0.029 | 0.029 | 0.029 | 0.018333 | 0.034 | 0.029 | 0.023667 | 402.618 | 0.010333 | 0.00016821
402.943 | -0.003 | 402.943 | -0.003 | 0.002 | 0.002 | 0.016 | 0.019 | 0.021 | 0.01 | 0.009 | 0.008 | 0.00033333 | 0.018667 | 0.009 | 0.0046667 | 402.943 | 0.014 | 0.0002279
403.267 | 0.002 | 403.267 | 0.002 | 0.002 | 0.004 | 0.029 | 0.032 | 0.031 | 0.012 | 0.011 | 0.012 | 0.0026667 | 0.030667 | 0.011667 | 0.0071667 | 403.267 | 0.0235 | 0.00038255
403.592 | 0.006 | 403.592 | 0.006 | 0.005 | 0.008 | 0.111 | 0.112 | 0.114 | 0.014 | 0.015 | 0.016 | 0.0063333 | 0.11233 | 0.015 | 0.010667 | 403.592 | 0.10167 | 0.001655
403.916 | 0.016 | 403.916 | 0.016 | 0.02 | 0.018 | 0.38 | 0.377 | 0.379 | 0.026 | 0.027 | 0.027 | 0.018 | 0.37867 | 0.026667 | 0.022333 | 403.916 | 0.35633 | 0.0058006
404.24 | 0.009 | 404.24 | 0.009 | 0.009 | 0.009 | 0.791 | 0.783 | 0.787 | 0.014 | 0.016 | 0.02 | 0.009 | 0.787 | 0.016667 | 0.012833 | 404.24 | 0.77417 | 0.012602
404.564 | 0.011 | 404.564 | 0.011 | 0.005 | 0.006 | 1.038 | 1.031 | 1.04 | 0.016 | 0.018 | 0.016 | 0.0073333 | 1.0363 | 0.016667 | 0.012 | 404.564 | 1.0243 | 0.016675
404.888 | 0.021 | 404.888 | 0.021 | 0.02 | 0.023 | 0.615 | 0.606 | 0.608 | 0.03 | 0.031 | 0.031 | 0.021333 | 0.60967 | 0.030667 | 0.026 | 404.888 | 0.58367 | 0.0095013
405.212 | 0.011 | 405.212 | 0.011 | 0.014 | 0.013 | 0.116 | 0.114 | 0.117 | 0.023 | 0.022 | 0.027 | 0.012667 | 0.11567 | 0.024 | 0.018333 | 405.212 | 0.097333 | 0.0015844
405.536 | 0.002 | 405.536 | 0.002 | 0.002 | 0.002 | 0.019 | 0.024 | 0.023 | 0.011 | 0.011 | 0.011 | 0.002 | 0.022 | 0.011 | 0.0065 | 405.536 | 0.0155 | 0.00025232
405.86 | 0.01 | 405.86 | 0.01 | 0.013 | 0.014 | 0.028 | 0.028 | 0.026 | 0.02 | 0.023 | 0.02 | 0.012333 | 0.027333 | 0.021 | 0.016667 | 405.86 | 0.010667 | 0.00017364
406.183 | 0.007 | 406.183 | 0.007 | 0.007 | 0.01 | 0.018 | 0.021 | 0.023 | 0.016 | 0.017 | 0.017 | 0.008 | 0.020667 | 0.016667 | 0.012333 | 406.183 | 0.0083333 | 0.00013565
406.507 | 0.008 | 406.507 | 0.008 | 0.01 | 0.011 | 0.025 | 0.027 | 0.027 | 0.017 | 0.019 | 0.022 | 0.0096667 | 0.026333 | 0.019333 | 0.0145 | 406.507 | 0.011833 | 0.00019263
406.831 | 0.014 | 406.831 | 0.014 | 0.013 | 0.015 | 0.035 | 0.038 | 0.038 | 0.022 | 0.023 | 0.023 | 0.014 | 0.037 | 0.022667 | 0.018333 | 406.831 | 0.018667 | 0.00030387
407.154 | 0.008 | 407.154 | 0.008 | 0.006 | 0.01 | 0.056 | 0.056 | 0.058 | 0.014 | 0.017 | 0.019 | 0.008 | 0.056667 | 0.016667 | 0.012333 | 407.154 | 0.044333 | 0.00072168
407.478 | 0.008 | 407.478 | 0.008 | 0.009 | 0.009 | 0.091 | 0.094 | 0.092 | 0.015 | 0.015 | 0.017 | 0.0086667 | 0.092333 | 0.015667 | 0.012167 | 407.478 | 0.080167 | 0.001305
407.801 | 0.005 | 407.801 | 0.005 | 0.007 | 0.009 | 0.094 | 0.098 | 0.094 | 0.017 | 0.018 | 0.018 | 0.007 | 0.095333 | 0.017667 | 0.012333 | 407.801 | 0.083 | 0.0013511
408.124 | 0 | 408.124 | 0 | 0.001 | 0.001 | 0.042 | 0.045 | 0.042 | 0.011 | 0.01 | 0.012 | 0.00066667 | 0.043 | 0.011 | 0.0058333 | 408.124 | 0.037167 | 0.00060502
408.447 | 0.007 | 408.447 | 0.007 | 0.007 | 0.009 | 0.018 | 0.022 | 0.022 | 0.013 | 0.016 | 0.01 | 0.0076667 | 0.020667 | 0.013 | 0.010333 | 408.447 | 0.010333 | 0.00016821
408.77 | 0.012 | 408.77 | 0.012 | 0.014 | 0.013 | 0.022 | 0.023 | 0.024 | 0.019 | 0.021 | 0.021 | 0.013 | 0.023 | 0.020333 | 0.016667 | 408.77 | 0.0063333 | 0.0001031
409.093 | 0.01 | 409.093 | 0.01 | 0.011 | 0.014 | 0.021 | 0.024 | 0.024 | 0.019 | 0.019 | 0.021 | 0.011667 | 0.023 | 0.019667 | 0.015667 | 409.093 | 0.0073333 | 0.00011938
409.416 | 0.003 | 409.416 | 0.003 | 0.005 | 0.004 | 0.015 | 0.011 | 0.017 | 0.013 | 0.013 | 0.015 | 0.004 | 0.014333 | 0.013667 | 0.0088333 | 409.416 | 0.0055 | 8.9532e-05
409.739 | 0.013 | 409.739 | 0.013 | 0.011 | 0.015 | 0.022 | 0.025 | 0.026 | 0.02 | 0.024 | 0.022 | 0.013 | 0.024333 | 0.022 | 0.0175 | 409.739 | 0.0068333 | 0.00011124
410.062 | 0.001 | 410.062 | 0.001 | 0.001 | 0.001 | 0.012 | 0.013 | 0.011 | 0.007 | 0.007 | 0.01 | 0.001 | 0.012 | 0.008 | 0.0045 | 410.062 | 0.0075 | 0.00012209
410.385 | 0.011 | 410.385 | 0.011 | 0.013 | 0.013 | 0.023 | 0.023 | 0.028 | 0.02 | 0.021 | 0.02 | 0.012333 | 0.024667 | 0.020333 | 0.016333 | 410.385 | 0.0083333 | 0.00013565
410.708 | 0.01 | 410.708 | 0.01 | 0.014 | 0.014 | 0.021 | 0.023 | 0.023 | 0.021 | 0.021 | 0.022 | 0.012667 | 0.022333 | 0.021333 | 0.017 | 410.708 | 0.0053333 | 8.6819e-05
411.03 | 0.01 | 411.03 | 0.01 | 0.012 | 0.016 | 0.023 | 0.025 | 0.026 | 0.023 | 0.022 | 0.024 | 0.012667 | 0.024667 | 0.023 | 0.017833 | 411.03 | 0.0068333 | 0.00011124
411.353 | 0.003 | 411.353 | 0.003 | 0.004 | 0.002 | 0.013 | 0.011 | 0.013 | 0.012 | 0.013 | 0.011 | 0.003 | 0.012333 | 0.012 | 0.0075 | 411.353 | 0.0048333 | 7.868e-05
411.675 | 0.012 | 411.675 | 0.012 | 0.01 | 0.012 | 0.021 | 0.025 | 0.025 | 0.022 | 0.022 | 0.021 | 0.011333 | 0.023667 | 0.021667 | 0.0165 | 411.675 | 0.0071667 | 0.00011666
411.997 | 0.008 | 411.997 | 0.008 | 0.008 | 0.008 | 0.018 | 0.016 | 0.022 | 0.016 | 0.019 | 0.018 | 0.008 | 0.018667 | 0.017667 | 0.012833 | 411.997 | 0.0058333 | 9.4958e-05
412.32 | 0.001 | 412.32 | 0.001 | 0.004 | 0.006 | 0.014 | 0.015 | 0.014 | 0.012 | 0.013 | 0.011 | 0.0036667 | 0.014333 | 0.012 | 0.0078333 | 412.32 | 0.0065 | 0.00010581
412.642 | 0.012 | 412.642 | 0.012 | 0.01 | 0.01 | 0.022 | 0.024 | 0.024 | 0.02 | 0.019 | 0.02 | 0.010667 | 0.023333 | 0.019667 | 0.015167 | 412.642 | 0.0081667 | 0.00013294
412.964 | 0.002 | 412.964 | 0.002 | 0.007 | 0.009 | 0.018 | 0.017 | 0.019 | 0.015 | 0.015 | 0.016 | 0.006 | 0.018 | 0.015333 | 0.010667 | 412.964 | 0.0073333 | 0.00011938
413.286 | 0.006 | 413.286 | 0.006 | 0.006 | 0.009 | 0.017 | 0.018 | 0.017 | 0.016 | 0.015 | 0.016 | 0.007 | 0.017333 | 0.015667 | 0.011333 | 413.286 | 0.006 | 9.7671e-05
413.608 | 0.01 | 413.608 | 0.01 | 0.011 | 0.011 | 0.02 | 0.02 | 0.02 | 0.016 | 0.021 | 0.02 | 0.010667 | 0.02 | 0.019 | 0.014833 | 413.608 | 0.0051667 | 8.4106e-05
413.93 | 0.007 | 413.93 | 0.007 | 0.012 | 0.01 | 0.019 | 0.023 | 0.021 | 0.019 | 0.022 | 0.021 | 0.0096667 | 0.021 | 0.020667 | 0.015167 | 413.93 | 0.0058333 | 9.4958e-05
414.252 | -0.001 | 414.252 | -0.001 | 0.001 | 0.001 | 0.012 | 0.01 | 0.011 | 0.009 | 0.009 | 0.01 | 0.00033333 | 0.011 | 0.0093333 | 0.0048333 | 414.252 | 0.0061667 | 0.00010038
414.574 | -0.001 | 414.574 | -0.001 | 0.002 | 0.004 | 0.011 | 0.013 | 0.015 | 0.009 | 0.011 | 0.01 | 0.0016667 | 0.013 | 0.01 | 0.0058333 | 414.574 | 0.0071667 | 0.00011666
414.895 | 0.004 | 414.895 | 0.004 | 0.004 | 0.008 | 0.013 | 0.016 | 0.015 | 0.012 | 0.013 | 0.015 | 0.0053333 | 0.014667 | 0.013333 | 0.0093333 | 414.895 | 0.0053333 | 8.6819e-05
415.217 | 0.011 | 415.217 | 0.011 | 0.014 | 0.014 | 0.023 | 0.024 | 0.025 | 0.023 | 0.023 | 0.023 | 0.013 | 0.024 | 0.023 | 0.018 | 415.217 | 0.006 | 9.7671e-05
415.538 | 0.004 | 415.538 | 0.004 | 0.005 | 0.008 | 0.013 | 0.012 | 0.015 | 0.011 | 0.013 | 0.013 | 0.0056667 | 0.013333 | 0.012333 | 0.009 | 415.538 | 0.0043333 | 7.054e-05
415.86 | 0.003 | 415.86 | 0.003 | 0.006 | 0.008 | 0.017 | 0.018 | 0.016 | 0.016 | 0.017 | 0.017 | 0.0056667 | 0.017 | 0.016667 | 0.011167 | 415.86 | 0.0058333 | 9.4958e-05
416.181 | 0.012 | 416.181 | 0.012 | 0.011 | 0.014 | 0.024 | 0.025 | 0.027 | 0.022 | 0.022 | 0.024 | 0.012333 | 0.025333 | 0.022667 | 0.0175 | 416.181 | 0.0078333 | 0.00012752
416.503 | 0.007 | 416.503 | 0.007 | 0.011 | 0.009 | 0.015 | 0.016 | 0.017 | 0.013 | 0.018 | 0.016 | 0.009 | 0.016 | 0.015667 | 0.012333 | 416.503 | 0.0036667 | 5.9688e-05
416.824 | 0.007 | 416.824 | 0.007 | 0.006 | 0.01 | 0.018 | 0.014 | 0.017 | 0.015 | 0.017 | 0.017 | 0.0076667 | 0.016333 | 0.016333 | 0.012 | 416.824 | 0.0043333 | 7.054e-05
417.145 | 0.004 | 417.145 | 0.004 | 0.004 | 0.007 | 0.011 | 0.017 | 0.015 | 0.012 | 0.014 | 0.014 | 0.005 | 0.014333 | 0.013333 | 0.0091667 | 417.145 | 0.0051667 | 8.4106e-05
417.466 | 0.007 | 417.466 | 0.007 | 0.011 | 0.008 | 0.02 | 0.017 | 0.019 | 0.014 | 0.016 | 0.017 | 0.0086667 | 0.018667 | 0.015667 | 0.012167 | 417.466 | 0.0065 | 0.00010581
417.787 | 0.013 | 417.787 | 0.013 | 0.017 | 0.016 | 0.024 | 0.027 | 0.023 | 0.024 | 0.024 | 0.028 | 0.015333 | 0.024667 | 0.025333 | 0.020333 | 417.787 | 0.0043333 | 7.054e-05
418.108 | 0.003 | 418.108 | 0.003 | 0.004 | 0.004 | 0.008 | 0.01 | 0.012 | 0.012 | 0.009 | 0.012 | 0.0036667 | 0.01 | 0.011 | 0.0073333 | 418.108 | 0.0026667 | 4.3409e-05
418.429 | 0.004 | 418.429 | 0.004 | 0.006 | 0.006 | 0.016 | 0.015 | 0.015 | 0.014 | 0.015 | 0.016 | 0.0053333 | 0.015333 | 0.015 | 0.010167 | 418.429 | 0.0051667 | 8.4106e-05
418.75 | 0.005 | 418.75 | 0.005 | 0.007 | 0.01 | 0.014 | 0.017 | 0.019 | 0.015 | 0.016 | 0.015 | 0.0073333 | 0.016667 | 0.015333 | 0.011333 | 418.75 | 0.0053333 | 8.6819e-05
419.07 | 0.004 | 419.07 | 0.004 | 0.006 | 0.006 | 0.014 | 0.016 | 0.018 | 0.013 | 0.015 | 0.013 | 0.0053333 | 0.016 | 0.013667 | 0.0095 | 419.07 | 0.0065 | 0.00010581
419.391 | 0.01 | 419.391 | 0.01 | 0.013 | 0.013 | 0.018 | 0.02 | 0.019 | 0.018 | 0.018 | 0.018 | 0.012 | 0.019 | 0.018 | 0.015 | 419.391 | 0.004 | 6.5114e-05
419.712 | 0.009 | 419.712 | 0.009 | 0.011 | 0.01 | 0.018 | 0.02 | 0.018 | 0.018 | 0.019 | 0.018 | 0.01 | 0.018667 | 0.018333 | 0.014167 | 419.712 | 0.0045 | 7.3254e-05
420.032 | 0.005 | 420.032 | 0.005 | 0.007 | 0.007 | 0.015 | 0.016 | 0.017 | 0.014 | 0.017 | 0.016 | 0.0063333 | 0.016 | 0.015667 | 0.011 | 420.032 | 0.005 | 8.1393e-05
420.353 | 0.005 | 420.353 | 0.005 | 0.005 | 0.005 | 0.016 | 0.016 | 0.018 | 0.015 | 0.015 | 0.017 | 0.005 | 0.016667 | 0.015667 | 0.010333 | 420.353 | 0.0063333 | 0.0001031
420.673 | 0.006 | 420.673 | 0.006 | 0.005 | 0.009 | 0.014 | 0.013 | 0.014 | 0.013 | 0.015 | 0.015 | 0.0066667 | 0.013667 | 0.014333 | 0.0105 | 420.673 | 0.0031667 | 5.1549e-05
420.993 | 0.01 | 420.993 | 0.01 | 0.01 | 0.012 | 0.016 | 0.019 | 0.021 | 0.018 | 0.016 | 0.018 | 0.010667 | 0.018667 | 0.017333 | 0.014 | 420.993 | 0.0046667 | 7.5967e-05
421.313 | 0.011 | 421.313 | 0.011 | 0.015 | 0.012 | 0.022 | 0.02 | 0.022 | 0.021 | 0.02 | 0.021 | 0.012667 | 0.021333 | 0.020667 | 0.016667 | 421.313 | 0.0046667 | 7.5967e-05
421.633 | 0.007 | 421.633 | 0.007 | 0.012 | 0.01 | 0.017 | 0.021 | 0.02 | 0.02 | 0.021 | 0.021 | 0.0096667 | 0.019333 | 0.020667 | 0.015167 | 421.633 | 0.0041667 | 6.7827e-05
421.953 | 0.002 | 421.953 | 0.002 | 0.002 | 0.007 | 0.009 | 0.012 | 0.013 | 0.007 | 0.012 | 0.015 | 0.0036667 | 0.011333 | 0.011333 | 0.0075 | 421.953 | 0.0038333 | 6.2401e-05
422.273 | 0.009 | 422.273 | 0.009 | 0.007 | 0.009 | 0.016 | 0.014 | 0.017 | 0.016 | 0.017 | 0.02 | 0.0083333 | 0.015667 | 0.017667 | 0.013 | 422.273 | 0.0026667 | 4.3409e-05
422.593 | 0.008 | 422.593 | 0.008 | 0.01 | 0.011 | 0.018 | 0.017 | 0.021 | 0.015 | 0.018 | 0.019 | 0.0096667 | 0.018667 | 0.017333 | 0.0135 | 422.593 | 0.0051667 | 8.4106e-05
422.913 | 0.006 | 422.913 | 0.006 | 0.006 | 0.006 | 0.015 | 0.016 | 0.015 | 0.013 | 0.017 | 0.016 | 0.006 | 0.015333 | 0.015333 | 0.010667 | 422.913 | 0.0046667 | 7.5967e-05
423.233 | 0.004 | 423.233 | 0.004 | 0.01 | 0.009 | 0.012 | 0.014 | 0.014 | 0.014 | 0.014 | 0.016 | 0.0076667 | 0.013333 | 0.014667 | 0.011167 | 423.233 | 0.0021667 | 3.527e-05
423.553 | 0.007 | 423.553 | 0.007 | 0.005 | 0.007 | 0.012 | 0.015 | 0.013 | 0.01 | 0.016 | 0.015 | 0.0063333 | 0.013333 | 0.013667 | 0.01 | 423.553 | 0.0033333 | 5.4262e-05
423.872 | 0.004 | 423.872 | 0.004 | 0.004 | 0.006 | 0.01 | 0.013 | 0.014 | 0.011 | 0.013 | 0.012 | 0.0046667 | 0.012333 | 0.012 | 0.0083333 | 423.872 | 0.004 | 6.5114e-05
424.192 | 0.008 | 424.192 | 0.008 | 0.007 | 0.008 | 0.013 | 0.014 | 0.015 | 0.013 | 0.015 | 0.016 | 0.0076667 | 0.014 | 0.014667 | 0.011167 | 424.192 | 0.0028333 | 4.6123e-05
424.511 | 0.004 | 424.511 | 0.004 | 0.005 | 0.006 | 0.007 | 0.012 | 0.012 | 0.009 | 0.011 | 0.012 | 0.005 | 0.010333 | 0.010667 | 0.0078333 | 424.511 | 0.0025 | 4.0696e-05
424.83 | 0.008 | 424.83 | 0.008 | 0.012 | 0.012 | 0.015 | 0.022 | 0.021 | 0.02 | 0.021 | 0.02 | 0.010667 | 0.019333 | 0.020333 | 0.0155 | 424.83 | 0.0038333 | 6.2401e-05
425.15 | 0.014 | 425.15 | 0.014 | 0.014 | 0.014 | 0.02 | 0.02 | 0.021 | 0.019 | 0.023 | 0.028 | 0.014 | 0.020333 | 0.023333 | 0.018667 | 425.15 | 0.0016667 | 2.7131e-05
425.469 | 0.012 | 425.469 | 0.012 | 0.01 | 0.014 | 0.016 | 0.021 | 0.019 | 0.018 | 0.019 | 0.021 | 0.012 | 0.018667 | 0.019333 | 0.015667 | 425.469 | 0.003 | 4.8836e-05
425.788 | 0.012 | 425.788 | 0.012 | 0.015 | 0.016 | 0.02 | 0.022 | 0.024 | 0.023 | 0.021 | 0.024 | 0.014333 | 0.022 | 0.022667 | 0.0185 | 425.788 | 0.0035 | 5.6975e-05
426.107 | 0.007 | 426.107 | 0.007 | 0.009 | 0.011 | 0.017 | 0.021 | 0.018 | 0.016 | 0.02 | 0.019 | 0.009 | 0.018667 | 0.018333 | 0.013667 | 426.107 | 0.005 | 8.1393e-05
426.426 | 0.013 | 426.426 | 0.013 | 0.018 | 0.015 | 0.021 | 0.024 | 0.024 | 0.022 | 0.026 | 0.025 | 0.015333 | 0.023 | 0.024333 | 0.019833 | 426.426 | 0.0031667 | 5.1549e-05
426.745 | 0.017 | 426.745 | 0.017 | 0.019 | 0.018 | 0.024 | 0.027 | 0.027 | 0.023 | 0.027 | 0.029 | 0.018 | 0.026 | 0.026333 | 0.022167 | 426.745 | 0.0038333 | 6.2401e-05
427.064 | 0.009 | 427.064 | 0.009 | 0.012 | 0.012 | 0.02 | 0.02 | 0.019 | 0.018 | 0.019 | 0.022 | 0.011 | 0.019667 | 0.019667 | 0.015333 | 427.064 | 0.0043333 | 7.054e-05
427.383 | 0.012 | 427.383 | 0.012 | 0.012 | 0.013 | 0.022 | 0.023 | 0.023 | 0.02 | 0.021 | 0.022 | 0.012333 | 0.022667 | 0.021 | 0.016667 | 427.383 | 0.006 | 9.7671e-05
427.702 | -0.001 | 427.702 | -0.001 | 0.001 | 0.001 | 0.007 | 0.009 | 0.009 | 0.01 | 0.008 | 0.014 | 0.00033333 | 0.0083333 | 0.010667 | 0.0055 | 427.702 | 0.0028333 | 4.6123e-05
428.02 | 0.009 | 428.02 | 0.009 | 0.013 | 0.012 | 0.021 | 0.022 | 0.024 | 0.019 | 0.021 | 0.024 | 0.011333 | 0.022333 | 0.021333 | 0.016333 | 428.02 | 0.006 | 9.7671e-05
428.339 | 0.007 | 428.339 | 0.007 | 0.009 | 0.012 | 0.015 | 0.018 | 0.018 | 0.015 | 0.016 | 0.016 | 0.0093333 | 0.017 | 0.015667 | 0.0125 | 428.339 | 0.0045 | 7.3254e-05
428.657 | 0.008 | 428.657 | 0.008 | 0.01 | 0.01 | 0.017 | 0.021 | 0.019 | 0.015 | 0.018 | 0.018 | 0.0093333 | 0.019 | 0.017 | 0.013167 | 428.657 | 0.0058333 | 9.4958e-05
428.976 | 0.013 | 428.976 | 0.013 | 0.012 | 0.013 | 0.019 | 0.021 | 0.022 | 0.02 | 0.022 | 0.022 | 0.012667 | 0.020667 | 0.021333 | 0.017 | 428.976 | 0.0036667 | 5.9688e-05
429.294 | 0.008 | 429.294 | 0.008 | 0.007 | 0.01 | 0.011 | 0.016 | 0.017 | 0.014 | 0.018 | 0.017 | 0.0083333 | 0.014667 | 0.016333 | 0.012333 | 429.294 | 0.0023333 | 3.7983e-05
429.612 | 0.003 | 429.612 | 0.003 | 0.004 | 0.005 | 0.011 | 0.012 | 0.013 | 0.011 | 0.011 | 0.011 | 0.004 | 0.012 | 0.011 | 0.0075 | 429.612 | 0.0045 | 7.3254e-05
429.93 | 0.009 | 429.93 | 0.009 | 0.012 | 0.012 | 0.016 | 0.02 | 0.022 | 0.019 | 0.021 | 0.02 | 0.011 | 0.019333 | 0.02 | 0.0155 | 429.93 | 0.0038333 | 6.2401e-05
430.248 | 0.016 | 430.248 | 0.016 | 0.017 | 0.018 | 0.022 | 0.025 | 0.027 | 0.023 | 0.021 | 0.024 | 0.017 | 0.024667 | 0.022667 | 0.019833 | 430.248 | 0.0048333 | 7.868e-05
430.567 | 0.008 | 430.567 | 0.008 | 0.007 | 0.011 | 0.016 | 0.018 | 0.02 | 0.016 | 0.016 | 0.014 | 0.0086667 | 0.018 | 0.015333 | 0.012 | 430.567 | 0.006 | 9.7671e-05
430.884 | 0.007 | 430.884 | 0.007 | 0.009 | 0.009 | 0.016 | 0.017 | 0.02 | 0.015 | 0.016 | 0.016 | 0.0083333 | 0.017667 | 0.015667 | 0.012 | 430.884 | 0.0056667 | 9.2245e-05
431.202 | 0.007 | 431.202 | 0.007 | 0.006 | 0.006 | 0.011 | 0.013 | 0.014 | 0.011 | 0.013 | 0.015 | 0.0063333 | 0.012667 | 0.013 | 0.0096667 | 431.202 | 0.003 | 4.8836e-05
431.52 | 0.005 | 431.52 | 0.005 | 0.007 | 0.005 | 0.012 | 0.014 | 0.015 | 0.01 | 0.01 | 0.016 | 0.0056667 | 0.013667 | 0.012 | 0.0088333 | 431.52 | 0.0048333 | 7.868e-05
431.838 | 0.003 | 431.838 | 0.003 | 0.005 | 0.006 | 0.013 | 0.015 | 0.015 | 0.012 | 0.013 | 0.013 | 0.0046667 | 0.014333 | 0.012667 | 0.0086667 | 431.838 | 0.0056667 | 9.2245e-05
432.156 | 0.007 | 432.156 | 0.007 | 0.009 | 0.007 | 0.018 | 0.019 | 0.018 | 0.016 | 0.019 | 0.018 | 0.0076667 | 0.018333 | 0.017667 | 0.012667 | 432.156 | 0.0056667 | 9.2245e-05
432.473 | 0.01 | 432.473 | 0.01 | 0.009 | 0.01 | 0.018 | 0.019 | 0.022 | 0.016 | 0.019 | 0.019 | 0.0096667 | 0.019667 | 0.018 | 0.013833 | 432.473 | 0.0058333 | 9.4958e-05
432.791 | 0.012 | 432.791 | 0.012 | 0.011 | 0.013 | 0.024 | 0.025 | 0.027 | 0.019 | 0.022 | 0.024 | 0.012 | 0.025333 | 0.021667 | 0.016833 | 432.791 | 0.0085 | 0.00013837
433.108 | 0.006 | 433.108 | 0.006 | 0.005 | 0.007 | 0.016 | 0.019 | 0.021 | 0.015 | 0.016 | 0.017 | 0.006 | 0.018667 | 0.016 | 0.011 | 433.108 | 0.0076667 | 0.0001248
433.426 | 0.017 | 433.426 | 0.017 | 0.018 | 0.017 | 0.033 | 0.035 | 0.038 | 0.022 | 0.024 | 0.028 | 0.017333 | 0.035333 | 0.024667 | 0.021 | 433.426 | 0.014333 | 0.00023333
433.743 | 0.013 | 433.743 | 0.013 | 0.014 | 0.013 | 0.035 | 0.036 | 0.04 | 0.016 | 0.019 | 0.02 | 0.013333 | 0.037 | 0.018333 | 0.015833 | 433.743 | 0.021167 | 0.00034456
434.06 | 0.009 | 434.06 | 0.009 | 0.009 | 0.009 | 0.045 | 0.049 | 0.045 | 0.015 | 0.016 | 0.016 | 0.009 | 0.046333 | 0.015667 | 0.012333 | 434.06 | 0.034 | 0.00055347
434.377 | 0.012 | 434.377 | 0.012 | 0.013 | 0.013 | 0.073 | 0.074 | 0.074 | 0.019 | 0.024 | 0.023 | 0.012667 | 0.073667 | 0.022 | 0.017333 | 434.377 | 0.056333 | 0.00091703
434.694 | 0.012 | 434.694 | 0.012 | 0.015 | 0.015 | 0.199 | 0.197 | 0.196 | 0.02 | 0.023 | 0.026 | 0.014 | 0.19733 | 0.023 | 0.0185 | 434.694 | 0.17883 | 0.0029111
435.011 | 0.006 | 435.011 | 0.006 | 0.006 | 0.005 | 0.646 | 0.637 | 0.636 | 0.013 | 0.012 | 0.016 | 0.0056667 | 0.63967 | 0.013667 | 0.0096667 | 435.011 | 0.63 | 0.010255
435.328 | 0.008 | 435.328 | 0.008 | 0.008 | 0.01 | 1.259 | 1.25 | 1.247 | 0.016 | 0.017 | 0.017 | 0.0086667 | 1.252 | 0.016667 | 0.012667 | 435.328 | 1.2393 | 0.020175
435.645 | 0.006 | 435.645 | 0.006 | 0.008 | 0.011 | 2.167 | 2.162 | 2.16 | 0.014 | 0.014 | 0.016 | 0.0083333 | 2.163 | 0.014667 | 0.0115 | 435.645 | 2.1515 | 0.035023
435.962 | 0.007 | 435.962 | 0.007 | 0.008 | 0.007 | 2.374 | 2.342 | 2.355 | 0.016 | 0.02 | 0.017 | 0.0073333 | 2.357 | 0.017667 | 0.0125 | 435.962 | 2.3445 | 0.038165
436.279 | 0.007 | 436.279 | 0.007 | 0.007 | 0.01 | 0.76 | 0.747 | 0.753 | 0.016 | 0.017 | 0.016 | 0.008 | 0.75333 | 0.016333 | 0.012167 | 436.279 | 0.74117 | 0.012065
436.595 | 0.003 | 436.595 | 0.003 | 0.004 | 0.003 | 0.085 | 0.09 | 0.092 | 0.013 | 0.013 | 0.013 | 0.0033333 | 0.089 | 0.013 | 0.0081667 | 436.595 | 0.080833 | 0.0013159
436.912 | 0.005 | 436.912 | 0.005 | 0.007 | 0.005 | 0.029 | 0.03 | 0.032 | 0.015 | 0.012 | 0.016 | 0.0056667 | 0.030333 | 0.014333 | 0.01 | 436.912 | 0.020333 | 0.000331
437.228 | 0.007 | 437.228 | 0.007 | 0.007 | 0.009 | 0.024 | 0.023 | 0.025 | 0.015 | 0.017 | 0.016 | 0.0076667 | 0.024 | 0.016 | 0.011833 | 437.228 | 0.012167 | 0.00019806
437.545 | 0 | 437.545 | 0 | 0.009 | 0.003 | 0.017 | 0.02 | 0.02 | 0.011 | 0.014 | 0.013 | 0.004 | 0.019 | 0.012667 | 0.0083333 | 437.545 | 0.010667 | 0.00017364
437.861 | 0.005 | 437.861 | 0.005 | 0.005 | 0.007 | 0.018 | 0.02 | 0.02 | 0.013 | 0.014 | 0.011 | 0.0056667 | 0.019333 | 0.012667 | 0.0091667 | 437.861 | 0.010167 | 0.0001655
438.177 | 0.003 | 438.177 | 0.003 | 0.002 | 0.004 | 0.019 | 0.018 | 0.018 | 0.011 | 0.012 | 0.013 | 0.003 | 0.018333 | 0.012 | 0.0075 | 438.177 | 0.010833 | 0.00017635
438.493 | 0.002 | 438.493 | 0.002 | 0.003 | 0.003 | 0.017 | 0.015 | 0.017 | 0.009 | 0.011 | 0.009 | 0.0026667 | 0.016333 | 0.0096667 | 0.0061667 | 438.493 | 0.010167 | 0.0001655
438.81 | 0.009 | 438.81 | 0.009 | 0.011 | 0.013 | 0.025 | 0.024 | 0.027 | 0.016 | 0.019 | 0.02 | 0.011 | 0.025333 | 0.018333 | 0.014667 | 438.81 | 0.010667 | 0.00017364
439.126 | 0.006 | 439.126 | 0.006 | 0.005 | 0.007 | 0.02 | 0.025 | 0.025 | 0.014 | 0.015 | 0.016 | 0.006 | 0.023333 | 0.015 | 0.0105 | 439.126 | 0.012833 | 0.00020891
439.442 | 0.011 | 439.442 | 0.011 | 0.013 | 0.012 | 0.027 | 0.03 | 0.029 | 0.018 | 0.02 | 0.021 | 0.012 | 0.028667 | 0.019667 | 0.015833 | 439.442 | 0.012833 | 0.00020891
439.757 | -0.006 | 439.757 | -0.006 | -0.001 | -0.002 | 0.013 | 0.014 | 0.014 | 0.003 | 0.007 | 0.004 | -0.003 | 0.013667 | 0.0046667 | 0.00083333 | 439.757 | 0.012833 | 0.00020891
440.073 | 0.006 | 440.073 | 0.006 | 0.009 | 0.005 | 0.02 | 0.024 | 0.022 | 0.011 | 0.012 | 0.015 | 0.0066667 | 0.022 | 0.012667 | 0.0096667 | 440.073 | 0.012333 | 0.00020077
440.389 | 0.006 | 440.389 | 0.006 | 0.005 | 0.006 | 0.02 | 0.023 | 0.022 | 0.014 | 0.014 | 0.014 | 0.0056667 | 0.021667 | 0.014 | 0.0098333 | 440.389 | 0.011833 | 0.00019263
440.705 | 0.016 | 440.705 | 0.016 | 0.015 | 0.018 | 0.033 | 0.034 | 0.032 | 0.023 | 0.024 | 0.026 | 0.016333 | 0.033 | 0.024333 | 0.020333 | 440.705 | 0.012667 | 0.0002062
441.02 | 0.005 | 441.02 | 0.005 | 0.008 | 0.008 | 0.02 | 0.022 | 0.025 | 0.013 | 0.016 | 0.015 | 0.007 | 0.022333 | 0.014667 | 0.010833 | 441.02 | 0.0115 | 0.0001872
441.336 | 0.004 | 441.336 | 0.004 | 0.004 | 0.005 | 0.017 | 0.02 | 0.022 | 0.013 | 0.013 | 0.013 | 0.0043333 | 0.019667 | 0.013 | 0.0086667 | 441.336 | 0.011 | 0.00017906
441.651 | 0.01 | 441.651 | 0.01 | 0.013 | 0.012 | 0.024 | 0.024 | 0.024 | 0.017 | 0.018 | 0.019 | 0.011667 | 0.024 | 0.018 | 0.014833 | 441.651 | 0.0091667 | 0.00014922
441.967 | 0.004 | 441.967 | 0.004 | 0.003 | 0.002 | 0.017 | 0.017 | 0.017 | 0.011 | 0.011 | 0.013 | 0.003 | 0.017 | 0.011667 | 0.0073333 | 441.967 | 0.0096667 | 0.00015736
442.282 | -0.002 | 442.282 | -0.002 | -0.001 | 0.001 | 0.011 | 0.015 | 0.014 | 0.007 | 0.008 | 0.006 | -0.00066667 | 0.013333 | 0.007 | 0.0031667 | 442.282 | 0.010167 | 0.0001655
442.597 | 0.007 | 442.597 | 0.007 | 0.01 | 0.009 | 0.02 | 0.019 | 0.023 | 0.018 | 0.018 | 0.015 | 0.0086667 | 0.020667 | 0.017 | 0.012833 | 442.597 | 0.0078333 | 0.00012752
442.912 | 0.014 | 442.912 | 0.014 | 0.012 | 0.015 | 0.028 | 0.03 | 0.03 | 0.02 | 0.022 | 0.022 | 0.013667 | 0.029333 | 0.021333 | 0.0175 | 442.912 | 0.011833 | 0.00019263
443.227 | 0.008 | 443.227 | 0.008 | 0.01 | 0.011 | 0.022 | 0.026 | 0.026 | 0.017 | 0.017 | 0.019 | 0.0096667 | 0.024667 | 0.017667 | 0.013667 | 443.227 | 0.011 | 0.00017906
443.543 | 0.012 | 443.543 | 0.012 | 0.012 | 0.014 | 0.024 | 0.026 | 0.026 | 0.016 | 0.018 | 0.019 | 0.012667 | 0.025333 | 0.017667 | 0.015167 | 443.543 | 0.010167 | 0.0001655
443.857 | 0.008 | 443.857 | 0.008 | 0.012 | 0.013 | 0.028 | 0.025 | 0.03 | 0.019 | 0.019 | 0.021 | 0.011 | 0.027667 | 0.019667 | 0.015333 | 443.857 | 0.012333 | 0.00020077
444.172 | 0.004 | 444.172 | 0.004 | 0.005 | 0.006 | 0.021 | 0.021 | 0.024 | 0.012 | 0.015 | 0.014 | 0.005 | 0.022 | 0.013667 | 0.0093333 | 444.172 | 0.012667 | 0.0002062
444.487 | 0.009 | 444.487 | 0.009 | 0.012 | 0.01 | 0.024 | 0.024 | 0.026 | 0.017 | 0.018 | 0.016 | 0.010333 | 0.024667 | 0.017 | 0.013667 | 444.487 | 0.011 | 0.00017906
444.802 | 0.015 | 444.802 | 0.015 | 0.015 | 0.014 | 0.032 | 0.027 | 0.032 | 0.025 | 0.026 | 0.023 | 0.014667 | 0.030333 | 0.024667 | 0.019667 | 444.802 | 0.010667 | 0.00017364
445.116 | 0.012 | 445.116 | 0.012 | 0.015 | 0.015 | 0.028 | 0.03 | 0.028 | 0.021 | 0.024 | 0.024 | 0.014 | 0.028667 | 0.023 | 0.0185 | 445.116 | 0.010167 | 0.0001655
445.431 | -0.001 | 445.431 | -0.001 | 0 | 0.003 | 0.013 | 0.015 | 0.015 | 0.01 | 0.011 | 0.009 | 0.00066667 | 0.014333 | 0.01 | 0.0053333 | 445.431 | 0.009 | 0.00014651
445.746 | 0.003 | 445.746 | 0.003 | 0.002 | 0.002 | 0.014 | 0.013 | 0.015 | 0.01 | 0.011 | 0.01 | 0.0023333 | 0.014 | 0.010333 | 0.0063333 | 445.746 | 0.0076667 | 0.0001248
446.06 | 0.007 | 446.06 | 0.007 | 0.008 | 0.008 | 0.017 | 0.018 | 0.018 | 0.017 | 0.018 | 0.015 | 0.0076667 | 0.017667 | 0.016667 | 0.012167 | 446.06 | 0.0055 | 8.9532e-05
446.374 | 0.007 | 446.374 | 0.007 | 0.007 | 0.007 | 0.017 | 0.017 | 0.016 | 0.013 | 0.014 | 0.014 | 0.007 | 0.016667 | 0.013667 | 0.010333 | 446.374 | 0.0063333 | 0.0001031
446.689 | 0.009 | 446.689 | 0.009 | 0.014 | 0.015 | 0.021 | 0.021 | 0.023 | 0.022 | 0.023 | 0.024 | 0.012667 | 0.021667 | 0.023 | 0.017833 | 446.689 | 0.0038333 | 6.2401e-05
447.003 | 0.008 | 447.003 | 0.008 | 0.007 | 0.009 | 0.016 | 0.017 | 0.018 | 0.013 | 0.018 | 0.015 | 0.008 | 0.017 | 0.015333 | 0.011667 | 447.003 | 0.0053333 | 8.6819e-05
447.317 | 0.009 | 447.317 | 0.009 | 0.011 | 0.011 | 0.015 | 0.017 | 0.015 | 0.016 | 0.015 | 0.017 | 0.010333 | 0.015667 | 0.016 | 0.013167 | 447.317 | 0.0025 | 4.0696e-05
447.631 | -0.001 | 447.631 | -0.001 | 0.001 | -0.002 | 0.008 | 0.008 | 0.008 | 0.009 | 0.008 | 0.006 | -0.00066667 | 0.008 | 0.0076667 | 0.0035 | 447.631 | 0.0045 | 7.3254e-05
447.945 | 0.011 | 447.945 | 0.011 | 0.01 | 0.011 | 0.016 | 0.018 | 0.019 | 0.018 | 0.017 | 0.019 | 0.010667 | 0.017667 | 0.018 | 0.014333 | 447.945 | 0.0033333 | 5.4262e-05
448.259 | 0.01 | 448.259 | 0.01 | 0.013 | 0.014 | 0.021 | 0.019 | 0.021 | 0.018 | 0.021 | 0.022 | 0.012333 | 0.020333 | 0.020333 | 0.016333 | 448.259 | 0.004 | 6.5114e-05
448.573 | 0.014 | 448.573 | 0.014 | 0.012 | 0.016 | 0.018 | 0.022 | 0.022 | 0.025 | 0.025 | 0.026 | 0.014 | 0.020667 | 0.025333 | 0.019667 | 448.573 | 0.001 | 1.6279e-05
448.887 | 0.005 | 448.887 | 0.005 | 0.004 | 0.003 | 0.011 | 0.012 | 0.01 | 0.012 | 0.013 | 0.014 | 0.004 | 0.011 | 0.013 | 0.0085 | 448.887 | 0.0025 | 4.0696e-05
449.2 | 0.014 | 449.2 | 0.014 | 0.018 | 0.017 | 0.022 | 0.024 | 0.027 | 0.023 | 0.028 | 0.024 | 0.016333 | 0.024333 | 0.025 | 0.020667 | 449.2 | 0.0036667 | 5.9688e-05
449.514 | 0.011 | 449.514 | 0.011 | 0.01 | 0.014 | 0.017 | 0.021 | 0.021 | 0.021 | 0.022 | 0.021 | 0.011667 | 0.019667 | 0.021333 | 0.0165 | 449.514 | 0.0031667 | 5.1549e-05
449.827 | 0.008 | 449.827 | 0.008 | 0.01 | 0.012 | 0.015 | 0.018 | 0.018 | 0.015 | 0.016 | 0.018 | 0.01 | 0.017 | 0.016333 | 0.013167 | 449.827 | 0.0038333 | 6.2401e-05
450.141 | 0.008 | 450.141 | 0.008 | 0.007 | 0.01 | 0.015 | 0.019 | 0.018 | 0.016 | 0.018 | 0.018 | 0.0083333 | 0.017333 | 0.017333 | 0.012833 | 450.141 | 0.0045 | 7.3254e-05
450.454 | 0.005 | 450.454 | 0.005 | 0.006 | 0.009 | 0.016 | 0.012 | 0.013 | 0.014 | 0.013 | 0.016 | 0.0066667 | 0.013667 | 0.014333 | 0.0105 | 450.454 | 0.0031667 | 5.1549e-05
450.768 | 0.009 | 450.768 | 0.009 | 0.007 | 0.008 | 0.014 | 0.015 | 0.019 | 0.015 | 0.017 | 0.017 | 0.008 | 0.016 | 0.016333 | 0.012167 | 450.768 | 0.0038333 | 6.2401e-05
451.081 | 0.013 | 451.081 | 0.013 | 0.012 | 0.013 | 0.023 | 0.021 | 0.019 | 0.021 | 0.023 | 0.024 | 0.012667 | 0.021 | 0.022667 | 0.017667 | 451.081 | 0.0033333 | 5.4262e-05
451.394 | 0.005 | 451.394 | 0.005 | 0.007 | 0.006 | 0.012 | 0.013 | 0.014 | 0.013 | 0.016 | 0.015 | 0.006 | 0.013 | 0.014667 | 0.010333 | 451.394 | 0.0026667 | 4.3409e-05
451.707 | 0.003 | 451.707 | 0.003 | 0.008 | 0.006 | 0.011 | 0.012 | 0.014 | 0.011 | 0.011 | 0.014 | 0.0056667 | 0.012333 | 0.012 | 0.0088333 | 451.707 | 0.0035 | 5.6975e-05
452.02 | 0.015 | 452.02 | 0.015 | 0.015 | 0.015 | 0.02 | 0.024 | 0.022 | 0.021 | 0.023 | 0.024 | 0.015 | 0.022 | 0.022667 | 0.018833 | 452.02 | 0.0031667 | 5.1549e-05
452.333 | 0.006 | 452.333 | 0.006 | 0.01 | 0.009 | 0.016 | 0.015 | 0.015 | 0.016 | 0.014 | 0.014 | 0.0083333 | 0.015333 | 0.014667 | 0.0115 | 452.333 | 0.0038333 | 6.2401e-05
452.646 | 0.004 | 452.646 | 0.004 | 0.007 | 0.006 | 0.012 | 0.012 | 0.013 | 0.012 | 0.013 | 0.015 | 0.0056667 | 0.012333 | 0.013333 | 0.0095 | 452.646 | 0.0028333 | 4.6123e-05
452.959 | 0.003 | 452.959 | 0.003 | 0.005 | 0.005 | 0.011 | 0.013 | 0.012 | 0.015 | 0.016 | 0.015 | 0.0043333 | 0.012 | 0.015333 | 0.0098333 | 452.959 | 0.0021667 | 3.527e-05
453.272 | 0.003 | 453.272 | 0.003 | 0.007 | 0.008 | 0.011 | 0.015 | 0.015 | 0.016 | 0.017 | 0.014 | 0.006 | 0.013667 | 0.015667 | 0.010833 | 453.272 | 0.0028333 | 4.6123e-05
453.584 | 0.005 | 453.584 | 0.005 | 0.007 | 0.007 | 0.013 | 0.013 | 0.013 | 0.013 | 0.015 | 0.014 | 0.0063333 | 0.013 | 0.014 | 0.010167 | 453.584 | 0.0028333 | 4.6123e-05
453.897 | 0.003 | 453.897 | 0.003 | 0.007 | 0.007 | 0.014 | 0.016 | 0.013 | 0.016 | 0.017 | 0.015 | 0.0056667 | 0.014333 | 0.016 | 0.010833 | 453.897 | 0.0035 | 5.6975e-05
454.21 | 0.01 | 454.21 | 0.01 | 0.011 | 0.013 | 0.016 | 0.017 | 0.019 | 0.017 | 0.019 | 0.021 | 0.011333 | 0.017333 | 0.019 | 0.015167 | 454.21 | 0.0021667 | 3.527e-05
454.522 | 0.003 | 454.522 | 0.003 | 0.006 | 0.007 | 0.011 | 0.013 | 0.013 | 0.012 | 0.014 | 0.016 | 0.0053333 | 0.012333 | 0.014 | 0.0096667 | 454.522 | 0.0026667 | 4.3409e-05
454.834 | 0.001 | 454.834 | 0.001 | 0.005 | 0.007 | 0.008 | 0.01 | 0.011 | 0.012 | 0.011 | 0.016 | 0.0043333 | 0.0096667 | 0.013 | 0.0086667 | 454.834 | 0.001 | 1.6279e-05
455.147 | 0.007 | 455.147 | 0.007 | 0.007 | 0.009 | 0.012 | 0.011 | 0.014 | 0.016 | 0.015 | 0.016 | 0.0076667 | 0.012333 | 0.015667 | 0.011667 | 455.147 | 0.00066667 | 1.0852e-05
455.459 | -0.002 | 455.459 | -0.002 | 0.001 | -0.001 | 0.003 | 0.006 | 0.005 | 0.005 | 0.008 | 0.009 | -0.00066667 | 0.0046667 | 0.0073333 | 0.0033333 | 455.459 | 0.0013333 | 2.1705e-05
455.771 | -0.006 | 455.771 | -0.006 | -0.004 | -0.002 | 0.001 | 0.004 | 0.006 | 0.003 | 0.004 | 0.006 | -0.004 | 0.0036667 | 0.0043333 | 0.00016667 | 455.771 | 0.0035 | 5.6975e-05
456.083 | 0.004 | 456.083 | 0.004 | 0.002 | 0.004 | 0.008 | 0.008 | 0.011 | 0.012 | 0.012 | 0.012 | 0.0033333 | 0.009 | 0.012 | 0.0076667 | 456.083 | 0.0013333 | 2.1705e-05
456.395 | 0.002 | 456.395 | 0.002 | 0.006 | 0.008 | 0.013 | 0.013 | 0.012 | 0.014 | 0.017 | 0.013 | 0.0053333 | 0.012667 | 0.014667 | 0.01 | 456.395 | 0.0026667 | 4.3409e-05
456.707 | 0.008 | 456.707 | 0.008 | 0.007 | 0.01 | 0.013 | 0.013 | 0.015 | 0.015 | 0.013 | 0.017 | 0.0083333 | 0.013667 | 0.015 | 0.011667 | 456.707 | 0.002 | 3.2557e-05
457.019 | 0.011 | 457.019 | 0.011 | 0.011 | 0.013 | 0.016 | 0.017 | 0.016 | 0.017 | 0.018 | 0.021 | 0.011667 | 0.016333 | 0.018667 | 0.015167 | 457.019 | 0.0011667 | 1.8992e-05
457.331 | 0.007 | 457.331 | 0.007 | 0.01 | 0.009 | 0.014 | 0.013 | 0.017 | 0.016 | 0.018 | 0.018 | 0.0086667 | 0.014667 | 0.017333 | 0.013 | 457.331 | 0.0016667 | 2.7131e-05
457.643 | 0.002 | 457.643 | 0.002 | 0.007 | 0.004 | 0.01 | 0.011 | 0.012 | 0.011 | 0.015 | 0.014 | 0.0043333 | 0.011 | 0.013333 | 0.0088333 | 457.643 | 0.0021667 | 3.527e-05
457.954 | 0.008 | 457.954 | 0.008 | 0.009 | 0.011 | 0.014 | 0.018 | 0.017 | 0.017 | 0.017 | 0.018 | 0.0093333 | 0.016333 | 0.017333 | 0.013333 | 457.954 | 0.003 | 4.8836e-05
458.266 | 0.004 | 458.266 | 0.004 | 0.004 | 0.008 | 0.01 | 0.012 | 0.014 | 0.012 | 0.012 | 0.016 | 0.0053333 | 0.012 | 0.013333 | 0.0093333 | 458.266 | 0.0026667 | 4.3409e-05
458.577 | 0.007 | 458.577 | 0.007 | 0.006 | 0.005 | 0.013 | 0.012 | 0.014 | 0.012 | 0.015 | 0.015 | 0.006 | 0.013 | 0.014 | 0.01 | 458.577 | 0.003 | 4.8836e-05
458.889 | 0.004 | 458.889 | 0.004 | 0.003 | 0.005 | 0.01 | 0.009 | 0.011 | 0.008 | 0.013 | 0.013 | 0.004 | 0.01 | 0.011333 | 0.0076667 | 458.889 | 0.0023333 | 3.7983e-05
459.2 | 0.006 | 459.2 | 0.006 | 0.009 | 0.009 | 0.013 | 0.015 | 0.016 | 0.018 | 0.017 | 0.018 | 0.008 | 0.014667 | 0.017667 | 0.012833 | 459.2 | 0.0018333 | 2.9844e-05
459.512 | 0.002 | 459.512 | 0.002 | 0.003 | 0.005 | 0.01 | 0.01 | 0.009 | 0.011 | 0.012 | 0.014 | 0.0033333 | 0.0096667 | 0.012333 | 0.0078333 | 459.512 | 0.0018333 | 2.9844e-05
459.823 | 0.007 | 459.823 | 0.007 | 0.006 | 0.007 | 0.016 | 0.015 | 0.013 | 0.016 | 0.016 | 0.017 | 0.0066667 | 0.014667 | 0.016333 | 0.0115 | 459.823 | 0.0031667 | 5.1549e-05
460.134 | 0.008 | 460.134 | 0.008 | 0.009 | 0.01 | 0.016 | 0.013 | 0.015 | 0.016 | 0.016 | 0.018 | 0.009 | 0.014667 | 0.016667 | 0.012833 | 460.134 | 0.0018333 | 2.9844e-05
460.445 | 0.012 | 460.445 | 0.012 | 0.013 | 0.012 | 0.018 | 0.018 | 0.02 | 0.017 | 0.019 | 0.02 | 0.012333 | 0.018667 | 0.018667 | 0.0155 | 460.445 | 0.0031667 | 5.1549e-05
460.756 | 0.006 | 460.756 | 0.006 | 0.008 | 0.007 | 0.014 | 0.016 | 0.015 | 0.013 | 0.016 | 0.014 | 0.007 | 0.015 | 0.014333 | 0.010667 | 460.756 | 0.0043333 | 7.054e-05
461.067 | 0.007 | 461.067 | 0.007 | 0.009 | 0.011 | 0.014 | 0.014 | 0.016 | 0.016 | 0.017 | 0.014 | 0.009 | 0.014667 | 0.015667 | 0.012333 | 461.067 | 0.0023333 | 3.7983e-05
461.378 | 0 | 461.378 | 0 | 0.005 | 0.007 | 0.009 | 0.01 | 0.011 | 0.011 | 0.014 | 0.01 | 0.004 | 0.01 | 0.011667 | 0.0078333 | 461.378 | 0.0021667 | 3.527e-05
461.689 | 0.004 | 461.689 | 0.004 | 0.005 | 0.005 | 0.008 | 0.01 | 0.011 | 0.011 | 0.014 | 0.014 | 0.0046667 | 0.0096667 | 0.013 | 0.0088333 | 461.689 | 0.00083333 | 1.3565e-05
461.999 | 0.008 | 461.999 | 0.008 | 0.016 | 0.017 | 0.022 | 0.025 | 0.025 | 0.028 | 0.029 | 0.031 | 0.013667 | 0.024 | 0.029333 | 0.0215 | 461.999 | 0.0025 | 4.0696e-05
462.31 | 0.01 | 462.31 | 0.01 | 0.016 | 0.011 | 0.018 | 0.019 | 0.021 | 0.021 | 0.022 | 0.019 | 0.012333 | 0.019333 | 0.020667 | 0.0165 | 462.31 | 0.0028333 | 4.6123e-05
462.621 | 0.004 | 462.621 | 0.004 | 0.003 | 0.003 | 0.011 | 0.011 | 0.008 | 0.011 | 0.012 | 0.013 | 0.0033333 | 0.01 | 0.012 | 0.0076667 | 462.621 | 0.0023333 | 3.7983e-05
462.931 | 0.005 | 462.931 | 0.005 | 0.008 | 0.009 | 0.013 | 0.013 | 0.015 | 0.018 | 0.017 | 0.018 | 0.0073333 | 0.013667 | 0.017667 | 0.0125 | 462.931 | 0.0011667 | 1.8992e-05
463.242 | 0.009 | 463.242 | 0.009 | 0.012 | 0.011 | 0.015 | 0.018 | 0.016 | 0.017 | 0.02 | 0.018 | 0.010667 | 0.016333 | 0.018333 | 0.0145 | 463.242 | 0.0018333 | 2.9844e-05
463.552 | 0.01 | 463.552 | 0.01 | 0.013 | 0.015 | 0.02 | 0.021 | 0.019 | 0.019 | 0.021 | 0.021 | 0.012667 | 0.02 | 0.020333 | 0.0165 | 463.552 | 0.0035 | 5.6975e-05
463.862 | 0 | 463.862 | 0 | 0.002 | 0.005 | 0.008 | 0.01 | 0.01 | 0.01 | 0.012 | 0.012 | 0.0023333 | 0.0093333 | 0.011333 | 0.0068333 | 463.862 | 0.0025 | 4.0696e-05
464.172 | 0.009 | 464.172 | 0.009 | 0.008 | 0.01 | 0.013 | 0.014 | 0.013 | 0.015 | 0.02 | 0.02 | 0.009 | 0.013333 | 0.018333 | 0.013667 | 464.172 | -0.00033333 | -5.4262e-06
464.483 | 0.014 | 464.483 | 0.014 | 0.018 | 0.017 | 0.021 | 0.023 | 0.024 | 0.025 | 0.025 | 0.023 | 0.016333 | 0.022667 | 0.024333 | 0.020333 | 464.483 | 0.0023333 | 3.7983e-05
464.793 | 0.004 | 464.793 | 0.004 | 0.007 | 0.007 | 0.011 | 0.009 | 0.012 | 0.013 | 0.012 | 0.015 | 0.006 | 0.010667 | 0.013333 | 0.0096667 | 464.793 | 0.001 | 1.6279e-05
465.103 | 0.008 | 465.103 | 0.008 | 0.008 | 0.012 | 0.014 | 0.016 | 0.018 | 0.018 | 0.017 | 0.017 | 0.0093333 | 0.016 | 0.017333 | 0.013333 | 465.103 | 0.0026667 | 4.3409e-05
465.413 | 0.011 | 465.413 | 0.011 | 0.011 | 0.016 | 0.017 | 0.016 | 0.02 | 0.018 | 0.018 | 0.024 | 0.012667 | 0.017667 | 0.02 | 0.016333 | 465.413 | 0.0013333 | 2.1705e-05
465.722 | 0.008 | 465.722 | 0.008 | 0.009 | 0.008 | 0.012 | 0.014 | 0.014 | 0.015 | 0.015 | 0.014 | 0.0083333 | 0.013333 | 0.014667 | 0.0115 | 465.722 | 0.0018333 | 2.9844e-05
466.032 | 0.008 | 466.032 | 0.008 | 0.009 | 0.011 | 0.019 | 0.018 | 0.018 | 0.019 | 0.022 | 0.022 | 0.0093333 | 0.018333 | 0.021 | 0.015167 | 466.032 | 0.0031667 | 5.1549e-05
466.342 | 0.006 | 466.342 | 0.006 | 0.007 | 0.008 | 0.012 | 0.014 | 0.013 | 0.013 | 0.016 | 0.014 | 0.007 | 0.013 | 0.014333 | 0.010667 | 466.342 | 0.0023333 | 3.7983e-05
466.652 | 0.005 | 466.652 | 0.005 | 0.008 | 0.012 | 0.014 | 0.015 | 0.016 | 0.015 | 0.015 | 0.019 | 0.0083333 | 0.015 | 0.016333 | 0.012333 | 466.652 | 0.0026667 | 4.3409e-05
466.961 | 0.01 | 466.961 | 0.01 | 0.012 | 0.012 | 0.015 | 0.015 | 0.018 | 0.016 | 0.017 | 0.02 | 0.011333 | 0.016 | 0.017667 | 0.0145 | 466.961 | 0.0015 | 2.4418e-05
467.271 | 0.012 | 467.271 | 0.012 | 0.012 | 0.013 | 0.015 | 0.019 | 0.02 | 0.019 | 0.021 | 0.022 | 0.012333 | 0.018 | 0.020667 | 0.0165 | 467.271 | 0.0015 | 2.4418e-05
467.58 | 0.008 | 467.58 | 0.008 | 0.009 | 0.012 | 0.015 | 0.016 | 0.015 | 0.016 | 0.017 | 0.018 | 0.0096667 | 0.015333 | 0.017 | 0.013333 | 467.58 | 0.002 | 3.2557e-05
467.889 | 0.006 | 467.889 | 0.006 | 0.01 | 0.013 | 0.015 | 0.017 | 0.019 | 0.019 | 0.019 | 0.019 | 0.0096667 | 0.017 | 0.019 | 0.014333 | 467.889 | 0.0026667 | 4.3409e-05
468.199 | 0.01 | 468.199 | 0.01 | 0.011 | 0.012 | 0.016 | 0.017 | 0.017 | 0.018 | 0.02 | 0.02 | 0.011 | 0.016667 | 0.019333 | 0.015167 | 468.199 | 0.0015 | 2.4418e-05
468.508 | 0.005 | 468.508 | 0.005 | 0.007 | 0.008 | 0.012 | 0.012 | 0.015 | 0.014 | 0.015 | 0.017 | 0.0066667 | 0.013 | 0.015333 | 0.011 | 468.508 | 0.002 | 3.2557e-05
468.817 | 0.01 | 468.817 | 0.01 | 0.012 | 0.011 | 0.018 | 0.016 | 0.019 | 0.02 | 0.022 | 0.022 | 0.011 | 0.017667 | 0.021333 | 0.016167 | 468.817 | 0.0015 | 2.4418e-05
469.126 | 0.005 | 469.126 | 0.005 | 0.01 | 0.01 | 0.016 | 0.015 | 0.013 | 0.017 | 0.016 | 0.019 | 0.0083333 | 0.014667 | 0.017333 | 0.012833 | 469.126 | 0.0018333 | 2.9844e-05
469.435 | 0.014 | 469.435 | 0.014 | 0.017 | 0.016 | 0.021 | 0.021 | 0.023 | 0.021 | 0.023 | 0.025 | 0.015667 | 0.021667 | 0.023 | 0.019333 | 469.435 | 0.0023333 | 3.7983e-05
469.744 | 0.005 | 469.744 | 0.005 | 0.009 | 0.011 | 0.014 | 0.013 | 0.015 | 0.016 | 0.016 | 0.018 | 0.0083333 | 0.014 | 0.016667 | 0.0125 | 469.744 | 0.0015 | 2.4418e-05
470.053 | 0.008 | 470.053 | 0.008 | 0.009 | 0.01 | 0.012 | 0.012 | 0.014 | 0.013 | 0.015 | 0.017 | 0.009 | 0.012667 | 0.015 | 0.012 | 470.053 | 0.00066667 | 1.0852e-05
470.362 | 0.005 | 470.362 | 0.005 | 0.006 | 0.008 | 0.009 | 0.016 | 0.013 | 0.015 | 0.015 | 0.016 | 0.0063333 | 0.012667 | 0.015333 | 0.010833 | 470.362 | 0.0018333 | 2.9844e-05
470.67 | 0.003 | 470.67 | 0.003 | 0.006 | 0.002 | 0.01 | 0.011 | 0.011 | 0.012 | 0.015 | 0.013 | 0.0036667 | 0.010667 | 0.013333 | 0.0085 | 470.67 | 0.0021667 | 3.527e-05
470.979 | 0.01 | 470.979 | 0.01 | 0.01 | 0.011 | 0.015 | 0.015 | 0.019 | 0.019 | 0.019 | 0.021 | 0.010333 | 0.016333 | 0.019667 | 0.015 | 470.979 | 0.0013333 | 2.1705e-05
471.287 | 0.003 | 471.287 | 0.003 | 0.003 | 0.006 | 0.008 | 0.009 | 0.009 | 0.011 | 0.014 | 0.011 | 0.004 | 0.0086667 | 0.012 | 0.008 | 471.287 | 0.00066667 | 1.0852e-05
471.596 | 0.007 | 471.596 | 0.007 | 0.007 | 0.007 | 0.012 | 0.011 | 0.013 | 0.013 | 0.015 | 0.015 | 0.007 | 0.012 | 0.014333 | 0.010667 | 471.596 | 0.0013333 | 2.1705e-05
471.904 | 0.014 | 471.904 | 0.014 | 0.015 | 0.017 | 0.02 | 0.021 | 0.021 | 0.024 | 0.026 | 0.026 | 0.015333 | 0.020667 | 0.025333 | 0.020333 | 471.904 | 0.00033333 | 5.4262e-06
472.213 | 0.002 | 472.213 | 0.002 | 0.002 | 0.004 | 0.01 | 0.01 | 0.008 | 0.009 | 0.01 | 0.013 | 0.0026667 | 0.0093333 | 0.010667 | 0.0066667 | 472.213 | 0.0026667 | 4.3409e-05
472.521 | 0.011 | 472.521 | 0.011 | 0.015 | 0.012 | 0.018 | 0.019 | 0.02 | 0.02 | 0.022 | 0.023 | 0.012667 | 0.019 | 0.021667 | 0.017167 | 472.521 | 0.0018333 | 2.9844e-05
472.829 | 0.007 | 472.829 | 0.007 | 0.01 | 0.01 | 0.015 | 0.018 | 0.017 | 0.017 | 0.018 | 0.02 | 0.009 | 0.016667 | 0.018333 | 0.013667 | 472.829 | 0.003 | 4.8836e-05
473.137 | 0.003 | 473.137 | 0.003 | 0.004 | 0.008 | 0.014 | 0.014 | 0.017 | 0.016 | 0.02 | 0.02 | 0.005 | 0.015 | 0.018667 | 0.011833 | 473.137 | 0.0031667 | 5.1549e-05
473.445 | 0 | 473.445 | 0 | 0.003 | 0.003 | 0.01 | 0.01 | 0.011 | 0.01 | 0.011 | 0.013 | 0.002 | 0.010333 | 0.011333 | 0.0066667 | 473.445 | 0.0036667 | 5.9688e-05
473.753 | 0.018 | 473.753 | 0.018 | 0.021 | 0.023 | 0.023 | 0.025 | 0.026 | 0.026 | 0.026 | 0.026 | 0.020667 | 0.024667 | 0.026 | 0.023333 | 473.753 | 0.0013333 | 2.1705e-05
474.061 | 0.005 | 474.061 | 0.005 | 0.005 | 0.006 | 0.007 | 0.011 | 0.01 | 0.012 | 0.012 | 0.013 | 0.0053333 | 0.0093333 | 0.012333 | 0.0088333 | 474.061 | 0.0005 | 8.1393e-06
474.369 | 0.009 | 474.369 | 0.009 | 0.007 | 0.011 | 0.016 | 0.014 | 0.016 | 0.016 | 0.017 | 0.021 | 0.009 | 0.015333 | 0.018 | 0.0135 | 474.369 | 0.0018333 | 2.9844e-05
474.677 | 0.009 | 474.677 | 0.009 | 0.015 | 0.014 | 0.021 | 0.014 | 0.02 | 0.019 | 0.02 | 0.02 | 0.012667 | 0.018333 | 0.019667 | 0.016167 | 474.677 | 0.0021667 | 3.527e-05
474.985 | 0.001 | 474.985 | 0.001 | 0.003 | 0.005 | 0.007 | 0.009 | 0.01 | 0.01 | 0.012 | 0.012 | 0.003 | 0.0086667 | 0.011333 | 0.0071667 | 474.985 | 0.0015 | 2.4418e-05
475.292 | 0.001 | 475.292 | 0.001 | 0.005 | 0.003 | 0.007 | 0.007 | 0.007 | 0.009 | 0.008 | 0.012 | 0.003 | 0.007 | 0.0096667 | 0.0063333 | 475.292 | 0.00066667 | 1.0852e-05
475.6 | 0.012 | 475.6 | 0.012 | 0.012 | 0.015 | 0.019 | 0.019 | 0.021 | 0.02 | 0.022 | 0.021 | 0.013 | 0.019667 | 0.021 | 0.017 | 475.6 | 0.0026667 | 4.3409e-05
475.907 | 0.009 | 475.907 | 0.009 | 0.007 | 0.011 | 0.015 | 0.016 | 0.017 | 0.019 | 0.019 | 0.02 | 0.009 | 0.016 | 0.019333 | 0.014167 | 475.907 | 0.0018333 | 2.9844e-05
476.215 | 0.005 | 476.215 | 0.005 | 0.009 | 0.01 | 0.01 | 0.011 | 0.014 | 0.013 | 0.016 | 0.017 | 0.008 | 0.011667 | 0.015333 | 0.011667 | 476.215 | 0 | 0
476.522 | 0.015 | 476.522 | 0.015 | 0.018 | 0.016 | 0.022 | 0.022 | 0.024 | 0.023 | 0.024 | 0.026 | 0.016333 | 0.022667 | 0.024333 | 0.020333 | 476.522 | 0.0023333 | 3.7983e-05
476.829 | 0.01 | 476.829 | 0.01 | 0.008 | 0.011 | 0.014 | 0.016 | 0.019 | 0.017 | 0.016 | 0.018 | 0.0096667 | 0.016333 | 0.017 | 0.013333 | 476.829 | 0.003 | 4.8836e-05
477.137 | 0.014 | 477.137 | 0.014 | 0.016 | 0.014 | 0.017 | 0.02 | 0.021 | 0.019 | 0.021 | 0.023 | 0.014667 | 0.019333 | 0.021 | 0.017833 | 477.137 | 0.0015 | 2.4418e-05
477.444 | 0.002 | 477.444 | 0.002 | 0.004 | 0.005 | 0.008 | 0.009 | 0.012 | 0.014 | 0.013 | 0.013 | 0.0036667 | 0.0096667 | 0.013333 | 0.0085 | 477.444 | 0.0011667 | 1.8992e-05
477.751 | 0.013 | 477.751 | 0.013 | 0.016 | 0.014 | 0.019 | 0.02 | 0.019 | 0.025 | 0.026 | 0.025 | 0.014333 | 0.019333 | 0.025333 | 0.019833 | 477.751 | -0.0005 | -8.1393e-06
478.058 | 0.002 | 478.058 | 0.002 | 0.003 | 0.005 | 0.006 | 0.012 | 0.01 | 0.011 | 0.008 | 0.013 | 0.0033333 | 0.0093333 | 0.010667 | 0.007 | 478.058 | 0.0023333 | 3.7983e-05
478.365 | 0.008 | 478.365 | 0.008 | 0.01 | 0.009 | 0.013 | 0.015 | 0.015 | 0.018 | 0.017 | 0.018 | 0.009 | 0.014333 | 0.017667 | 0.013333 | 478.365 | 0.001 | 1.6279e-05
478.671 | 0.011 | 478.671 | 0.011 | 0.014 | 0.016 | 0.019 | 0.021 | 0.023 | 0.02 | 0.024 | 0.026 | 0.013667 | 0.021 | 0.023333 | 0.0185 | 478.671 | 0.0025 | 4.0696e-05
478.978 | 0.01 | 478.978 | 0.01 | 0.014 | 0.013 | 0.016 | 0.019 | 0.02 | 0.018 | 0.019 | 0.021 | 0.012333 | 0.018333 | 0.019333 | 0.015833 | 478.978 | 0.0025 | 4.0696e-05
479.285 | 0.003 | 479.285 | 0.003 | 0.005 | 0.004 | 0.01 | 0.011 | 0.012 | 0.01 | 0.012 | 0.013 | 0.004 | 0.011 | 0.011667 | 0.0078333 | 479.285 | 0.0031667 | 5.1549e-05
479.592 | 0.004 | 479.592 | 0.004 | 0.006 | 0.007 | 0.011 | 0.012 | 0.011 | 0.014 | 0.016 | 0.014 | 0.0056667 | 0.011333 | 0.014667 | 0.010167 | 479.592 | 0.0011667 | 1.8992e-05
479.898 | 0 | 479.898 | 0 | 0.005 | 0.004 | 0.008 | 0.007 | 0.01 | 0.011 | 0.014 | 0.011 | 0.003 | 0.0083333 | 0.012 | 0.0075 | 479.898 | 0.00083333 | 1.3565e-05
480.205 | 0.007 | 480.205 | 0.007 | 0.01 | 0.012 | 0.014 | 0.017 | 0.015 | 0.018 | 0.017 | 0.018 | 0.0096667 | 0.015333 | 0.017667 | 0.013667 | 480.205 | 0.0016667 | 2.7131e-05
480.511 | -0.003 | 480.511 | -0.003 | 0.001 | 0 | 0.002 | 0.004 | 0.004 | 0.005 | 0.005 | 0.005 | -0.00066667 | 0.0033333 | 0.005 | 0.0021667 | 480.511 | 0.0011667 | 1.8992e-05
480.817 | 0.012 | 480.817 | 0.012 | 0.012 | 0.016 | 0.019 | 0.019 | 0.022 | 0.019 | 0.025 | 0.025 | 0.013333 | 0.02 | 0.023 | 0.018167 | 480.817 | 0.0018333 | 2.9844e-05
481.124 | 0.005 | 481.124 | 0.005 | 0.009 | 0.008 | 0.011 | 0.015 | 0.015 | 0.016 | 0.017 | 0.015 | 0.0073333 | 0.013667 | 0.016 | 0.011667 | 481.124 | 0.002 | 3.2557e-05
481.43 | 0.008 | 481.43 | 0.008 | 0.012 | 0.01 | 0.017 | 0.018 | 0.017 | 0.021 | 0.018 | 0.023 | 0.01 | 0.017333 | 0.020667 | 0.015333 | 481.43 | 0.002 | 3.2557e-05
481.736 | 0.018 | 481.736 | 0.018 | 0.02 | 0.021 | 0.023 | 0.027 | 0.025 | 0.028 | 0.029 | 0.03 | 0.019667 | 0.025 | 0.029 | 0.024333 | 481.736 | 0.00066667 | 1.0852e-05
482.042 | 0.006 | 482.042 | 0.006 | 0.008 | 0.01 | 0.014 | 0.012 | 0.015 | 0.014 | 0.017 | 0.017 | 0.008 | 0.013667 | 0.016 | 0.012 | 482.042 | 0.0016667 | 2.7131e-05
482.348 | 0.005 | 482.348 | 0.005 | 0.006 | 0.008 | 0.011 | 0.013 | 0.012 | 0.013 | 0.013 | 0.017 | 0.0063333 | 0.012 | 0.014333 | 0.010333 | 482.348 | 0.0016667 | 2.7131e-05
482.654 | 0.012 | 482.654 | 0.012 | 0.017 | 0.016 | 0.021 | 0.02 | 0.022 | 0.022 | 0.022 | 0.022 | 0.015 | 0.021 | 0.022 | 0.0185 | 482.654 | 0.0025 | 4.0696e-05
482.96 | 0.012 | 482.96 | 0.012 | 0.015 | 0.013 | 0.019 | 0.019 | 0.019 | 0.02 | 0.022 | 0.02 | 0.013333 | 0.019 | 0.020667 | 0.017 | 482.96 | 0.002 | 3.2557e-05
483.266 | 0.006 | 483.266 | 0.006 | 0.009 | 0.012 | 0.015 | 0.019 | 0.016 | 0.022 | 0.019 | 0.023 | 0.009 | 0.016667 | 0.021333 | 0.015167 | 483.266 | 0.0015 | 2.4418e-05
483.571 | 0.012 | 483.571 | 0.012 | 0.012 | 0.015 | 0.019 | 0.022 | 0.025 | 0.024 | 0.02 | 0.023 | 0.013 | 0.022 | 0.022333 | 0.017667 | 483.571 | 0.0043333 | 7.054e-05
483.877 | -0.002 | 483.877 | -0.002 | 0.001 | 0 | 0.007 | 0.007 | 0.007 | 0.006 | 0.008 | 0.007 | -0.00033333 | 0.007 | 0.007 | 0.0033333 | 483.877 | 0.0036667 | 5.9688e-05
484.182 | 0.001 | 484.182 | 0.001 | 0.005 | 0.006 | 0.006 | 0.008 | 0.008 | 0.009 | 0.009 | 0.013 | 0.004 | 0.0073333 | 0.010333 | 0.0071667 | 484.182 | 0.00016667 | 2.7131e-06
484.488 | 0.007 | 484.488 | 0.007 | 0.01 | 0.009 | 0.016 | 0.015 | 0.017 | 0.015 | 0.016 | 0.018 | 0.0086667 | 0.016 | 0.016333 | 0.0125 | 484.488 | 0.0035 | 5.6975e-05
484.793 | -0.003 | 484.793 | -0.003 | 0.003 | 0.004 | 0.005 | 0.01 | 0.006 | 0.012 | 0.011 | 0.013 | 0.0013333 | 0.007 | 0.012 | 0.0066667 | 484.793 | 0.00033333 | 5.4262e-06
485.099 | 0.006 | 485.099 | 0.006 | 0.008 | 0.005 | 0.013 | 0.012 | 0.015 | 0.017 | 0.02 | 0.017 | 0.0063333 | 0.013333 | 0.018 | 0.012167 | 485.099 | 0.0011667 | 1.8992e-05
485.404 | 0.008 | 485.404 | 0.008 | 0.011 | 0.012 | 0.014 | 0.014 | 0.016 | 0.019 | 0.016 | 0.02 | 0.010333 | 0.014667 | 0.018333 | 0.014333 | 485.404 | 0.00033333 | 5.4262e-06
485.709 | 0.011 | 485.709 | 0.011 | 0.015 | 0.014 | 0.017 | 0.019 | 0.019 | 0.021 | 0.022 | 0.023 | 0.013333 | 0.018333 | 0.022 | 0.017667 | 485.709 | 0.00066667 | 1.0852e-05
486.014 | 0.006 | 486.014 | 0.006 | 0.008 | 0.011 | 0.012 | 0.015 | 0.017 | 0.017 | 0.017 | 0.017 | 0.0083333 | 0.014667 | 0.017 | 0.012667 | 486.014 | 0.002 | 3.2557e-05
486.319 | 0.014 | 486.319 | 0.014 | 0.017 | 0.018 | 0.022 | 0.024 | 0.02 | 0.024 | 0.023 | 0.027 | 0.016333 | 0.022 | 0.024667 | 0.0205 | 486.319 | 0.0015 | 2.4418e-05
486.624 | 0.001 | 486.624 | 0.001 | 0.006 | 0.008 | 0.01 | 0.012 | 0.011 | 0.015 | 0.015 | 0.015 | 0.005 | 0.011 | 0.015 | 0.01 | 486.624 | 0.001 | 1.6279e-05
486.929 | 0.001 | 486.929 | 0.001 | 0.004 | 0.006 | 0.008 | 0.007 | 0.01 | 0.009 | 0.012 | 0.014 | 0.0036667 | 0.0083333 | 0.011667 | 0.0076667 | 486.929 | 0.00066667 | 1.0852e-05
487.234 | -0.006 | 487.234 | -0.006 | -0.001 | -0.003 | 0.002 | 0.004 | 0.004 | 0.006 | 0.008 | 0.01 | -0.0033333 | 0.0033333 | 0.008 | 0.0023333 | 487.234 | 0.001 | 1.6279e-05
487.539 | 0.007 | 487.539 | 0.007 | 0.01 | 0.01 | 0.016 | 0.018 | 0.019 | 0.019 | 0.022 | 0.021 | 0.009 | 0.017667 | 0.020667 | 0.014833 | 487.539 | 0.0028333 | 4.6123e-05
487.844 | 0.013 | 487.844 | 0.013 | 0.012 | 0.016 | 0.018 | 0.018 | 0.02 | 0.019 | 0.02 | 0.023 | 0.013667 | 0.018667 | 0.020667 | 0.017167 | 487.844 | 0.0015 | 2.4418e-05
488.148 | 0.006 | 488.148 | 0.006 | 0.007 | 0.009 | 0.011 | 0.014 | 0.017 | 0.015 | 0.016 | 0.019 | 0.0073333 | 0.014 | 0.016667 | 0.012 | 488.148 | 0.002 | 3.2557e-05
488.453 | 0.015 | 488.453 | 0.015 | 0.017 | 0.018 | 0.021 | 0.024 | 0.023 | 0.024 | 0.029 | 0.029 | 0.016667 | 0.022667 | 0.027333 | 0.022 | 488.453 | 0.00066667 | 1.0852e-05
488.757 | 0.008 | 488.757 | 0.008 | 0.01 | 0.012 | 0.013 | 0.018 | 0.018 | 0.021 | 0.019 | 0.021 | 0.01 | 0.016333 | 0.020333 | 0.015167 | 488.757 | 0.0011667 | 1.8992e-05
489.062 | 0.002 | 489.062 | 0.002 | 0.005 | 0.006 | 0.011 | 0.013 | 0.012 | 0.015 | 0.018 | 0.016 | 0.0043333 | 0.012 | 0.016333 | 0.010333 | 489.062 | 0.0016667 | 2.7131e-05
489.366 | 0.011 | 489.366 | 0.011 | 0.011 | 0.014 | 0.015 | 0.017 | 0.019 | 0.021 | 0.023 | 0.022 | 0.012 | 0.017 | 0.022 | 0.017 | 489.366 | 0 | 0
489.67 | 0.004 | 489.67 | 0.004 | 0.004 | 0.009 | 0.014 | 0.012 | 0.014 | 0.013 | 0.016 | 0.015 | 0.0056667 | 0.013333 | 0.014667 | 0.010167 | 489.67 | 0.0031667 | 5.1549e-05
489.975 | 0.002 | 489.975 | 0.002 | 0.002 | 0.003 | 0.011 | 0.009 | 0.014 | 0.012 | 0.013 | 0.013 | 0.0023333 | 0.011333 | 0.012667 | 0.0075 | 489.975 | 0.0038333 | 6.2401e-05
490.279 | 0.007 | 490.279 | 0.007 | 0.008 | 0.008 | 0.012 | 0.014 | 0.014 | 0.016 | 0.014 | 0.02 | 0.0076667 | 0.013333 | 0.016667 | 0.012167 | 490.279 | 0.0011667 | 1.8992e-05
490.583 | 0.006 | 490.583 | 0.006 | 0.011 | 0.011 | 0.015 | 0.02 | 0.017 | 0.017 | 0.018 | 0.021 | 0.0093333 | 0.017333 | 0.018667 | 0.014 | 490.583 | 0.0033333 | 5.4262e-05
490.887 | 0.005 | 490.887 | 0.005 | 0.005 | 0.008 | 0.014 | 0.015 | 0.016 | 0.014 | 0.013 | 0.015 | 0.006 | 0.015 | 0.014 | 0.01 | 490.887 | 0.005 | 8.1393e-05
491.191 | 0.007 | 491.191 | 0.007 | 0.011 | 0.008 | 0.015 | 0.016 | 0.02 | 0.017 | 0.016 | 0.021 | 0.0086667 | 0.017 | 0.018 | 0.013333 | 491.191 | 0.0036667 | 5.9688e-05
491.495 | 0.01 | 491.495 | 0.01 | 0.012 | 0.014 | 0.023 | 0.028 | 0.026 | 0.02 | 0.02 | 0.021 | 0.012 | 0.025667 | 0.020333 | 0.016167 | 491.495 | 0.0095 | 0.00015465
491.798 | 0.013 | 491.798 | 0.013 | 0.016 | 0.014 | 0.027 | 0.029 | 0.032 | 0.022 | 0.018 | 0.026 | 0.014333 | 0.029333 | 0.022 | 0.018167 | 491.798 | 0.011167 | 0.00018178
492.102 | 0.009 | 492.102 | 0.009 | 0.013 | 0.013 | 0.023 | 0.021 | 0.022 | 0.019 | 0.021 | 0.023 | 0.011667 | 0.022 | 0.021 | 0.016333 | 492.102 | 0.0056667 | 9.2245e-05
492.406 | 0.002 | 492.406 | 0.002 | 0.007 | 0.007 | 0.011 | 0.01 | 0.013 | 0.011 | 0.017 | 0.013 | 0.0053333 | 0.011333 | 0.013667 | 0.0095 | 492.406 | 0.0018333 | 2.9844e-05
492.709 | 0.01 | 492.709 | 0.01 | 0.012 | 0.013 | 0.015 | 0.016 | 0.018 | 0.018 | 0.019 | 0.02 | 0.011667 | 0.016333 | 0.019 | 0.015333 | 492.709 | 0.001 | 1.6279e-05
493.013 | 0.017 | 493.013 | 0.017 | 0.017 | 0.018 | 0.02 | 0.02 | 0.023 | 0.022 | 0.024 | 0.026 | 0.017333 | 0.021 | 0.024 | 0.020667 | 493.013 | 0.00033333 | 5.4262e-06
493.316 | 0.009 | 493.316 | 0.009 | 0.01 | 0.012 | 0.014 | 0.016 | 0.017 | 0.02 | 0.016 | 0.019 | 0.010333 | 0.015667 | 0.018333 | 0.014333 | 493.316 | 0.0013333 | 2.1705e-05
493.62 | 0.011 | 493.62 | 0.011 | 0.016 | 0.019 | 0.02 | 0.022 | 0.023 | 0.023 | 0.024 | 0.026 | 0.015333 | 0.021667 | 0.024333 | 0.019833 | 493.62 | 0.0018333 | 2.9844e-05
493.923 | 0.006 | 493.923 | 0.006 | 0.011 | 0.011 | 0.013 | 0.013 | 0.014 | 0.016 | 0.017 | 0.017 | 0.0093333 | 0.013333 | 0.016667 | 0.013 | 493.923 | 0.00033333 | 5.4262e-06
494.226 | 0.015 | 494.226 | 0.015 | 0.018 | 0.015 | 0.022 | 0.021 | 0.022 | 0.02 | 0.023 | 0.023 | 0.016 | 0.021667 | 0.022 | 0.019 | 494.226 | 0.0026667 | 4.3409e-05
494.529 | 0.009 | 494.529 | 0.009 | 0.01 | 0.012 | 0.012 | 0.015 | 0.016 | 0.016 | 0.018 | 0.019 | 0.010333 | 0.014333 | 0.017667 | 0.014 | 494.529 | 0.00033333 | 5.4262e-06
494.832 | 0.007 | 494.832 | 0.007 | 0.007 | 0.01 | 0.01 | 0.016 | 0.018 | 0.019 | 0.018 | 0.02 | 0.008 | 0.014667 | 0.019 | 0.0135 | 494.832 | 0.0011667 | 1.8992e-05
495.135 | 0.004 | 495.135 | 0.004 | 0.007 | 0.01 | 0.012 | 0.013 | 0.015 | 0.015 | 0.016 | 0.018 | 0.007 | 0.013333 | 0.016333 | 0.011667 | 495.135 | 0.0016667 | 2.7131e-05
495.438 | 0.012 | 495.438 | 0.012 | 0.014 | 0.015 | 0.021 | 0.02 | 0.022 | 0.022 | 0.022 | 0.024 | 0.013667 | 0.021 | 0.022667 | 0.018167 | 495.438 | 0.0028333 | 4.6123e-05
495.741 | 0.018 | 495.741 | 0.018 | 0.022 | 0.023 | 0.027 | 0.028 | 0.028 | 0.028 | 0.033 | 0.03 | 0.021 | 0.027667 | 0.030333 | 0.025667 | 495.741 | 0.002 | 3.2557e-05
496.044 | 0.006 | 496.044 | 0.006 | 0.007 | 0.008 | 0.012 | 0.011 | 0.011 | 0.014 | 0.014 | 0.018 | 0.007 | 0.011333 | 0.015333 | 0.011167 | 496.044 | 0.00016667 | 2.7131e-06
496.347 | -0.001 | 496.347 | -0.001 | 0.004 | 0.004 | 0.01 | 0.008 | 0.011 | 0.012 | 0.013 | 0.015 | 0.0023333 | 0.0096667 | 0.013333 | 0.0078333 | 496.347 | 0.0018333 | 2.9844e-05
496.649 | 0.016 | 496.649 | 0.016 | 0.014 | 0.018 | 0.019 | 0.025 | 0.023 | 0.026 | 0.026 | 0.026 | 0.016 | 0.022333 | 0.026 | 0.021 | 496.649 | 0.0013333 | 2.1705e-05
496.952 | 0.006 | 496.952 | 0.006 | 0.009 | 0.009 | 0.012 | 0.014 | 0.015 | 0.013 | 0.017 | 0.016 | 0.008 | 0.013667 | 0.015333 | 0.011667 | 496.952 | 0.002 | 3.2557e-05
497.255 | 0.007 | 497.255 | 0.007 | 0.008 | 0.008 | 0.012 | 0.013 | 0.016 | 0.013 | 0.012 | 0.016 | 0.0076667 | 0.013667 | 0.013667 | 0.010667 | 497.255 | 0.003 | 4.8836e-05
497.557 | 0.006 | 497.557 | 0.006 | 0.007 | 0.009 | 0.012 | 0.014 | 0.017 | 0.015 | 0.016 | 0.019 | 0.0073333 | 0.014333 | 0.016667 | 0.012 | 497.557 | 0.0023333 | 3.7983e-05
497.859 | 0.006 | 497.859 | 0.006 | 0.009 | 0.011 | 0.015 | 0.012 | 0.016 | 0.014 | 0.014 | 0.021 | 0.0086667 | 0.014333 | 0.016333 | 0.0125 | 497.859 | 0.0018333 | 2.9844e-05
498.162 | -0.001 | 498.162 | -0.001 | 0 | 0.002 | 0.003 | 0.006 | 0.005 | 0.006 | 0.006 | 0.007 | 0.00033333 | 0.0046667 | 0.0063333 | 0.0033333 | 498.162 | 0.0013333 | 2.1705e-05
498.464 | 0.013 | 498.464 | 0.013 | 0.016 | 0.017 | 0.024 | 0.022 | 0.023 | 0.024 | 0.027 | 0.033 | 0.015333 | 0.023 | 0.028 | 0.021667 | 498.464 | 0.0013333 | 2.1705e-05
498.766 | 0.008 | 498.766 | 0.008 | 0.013 | 0.012 | 0.014 | 0.019 | 0.016 | 0.016 | 0.02 | 0.023 | 0.011 | 0.016333 | 0.019667 | 0.015333 | 498.766 | 0.001 | 1.6279e-05
499.068 | 0.005 | 499.068 | 0.005 | 0.006 | 0.008 | 0.011 | 0.013 | 0.015 | 0.015 | 0.018 | 0.016 | 0.0063333 | 0.013 | 0.016333 | 0.011333 | 499.068 | 0.0016667 | 2.7131e-05
499.37 | 0.007 | 499.37 | 0.007 | 0.009 | 0.011 | 0.013 | 0.012 | 0.016 | 0.017 | 0.017 | 0.016 | 0.009 | 0.013667 | 0.016667 | 0.012833 | 499.37 | 0.00083333 | 1.3565e-05
499.672 | 0.015 | 499.672 | 0.015 | 0.017 | 0.016 | 0.019 | 0.02 | 0.02 | 0.024 | 0.021 | 0.022 | 0.016 | 0.019667 | 0.022333 | 0.019167 | 499.672 | 0.0005 | 8.1393e-06
499.974 | 0.007 | 499.974 | 0.007 | 0.009 | 0.009 | 0.012 | 0.013 | 0.014 | 0.017 | 0.015 | 0.017 | 0.0083333 | 0.013 | 0.016333 | 0.012333 | 499.974 | 0.00066667 | 1.0852e-05
500.276 | 0.01 | 500.276 | 0.01 | 0.01 | 0.012 | 0.015 | 0.017 | 0.017 | 0.022 | 0.02 | 0.022 | 0.010667 | 0.016333 | 0.021333 | 0.016 | 500.276 | 0.00033333 | 5.4262e-06
500.578 | 0.01 | 500.578 | 0.01 | 0.012 | 0.016 | 0.018 | 0.018 | 0.017 | 0.02 | 0.021 | 0.023 | 0.012667 | 0.017667 | 0.021333 | 0.017 | 500.578 | 0.00066667 | 1.0852e-05
500.879 | -0.001 | 500.879 | -0.001 | 0.003 | 0.002 | 0.006 | 0.006 | 0.008 | 0.008 | 0.01 | 0.011 | 0.0013333 | 0.0066667 | 0.0096667 | 0.0055 | 500.879 | 0.0011667 | 1.8992e-05
501.181 | 0.015 | 501.181 | 0.015 | 0.014 | 0.019 | 0.022 | 0.024 | 0.022 | 0.025 | 0.023 | 0.022 | 0.016 | 0.022667 | 0.023333 | 0.019667 | 501.181 | 0.003 | 4.8836e-05
501.483 | 0.005 | 501.483 | 0.005 | 0.01 | 0.008 | 0.01 | 0.013 | 0.015 | 0.013 | 0.016 | 0.017 | 0.0076667 | 0.012667 | 0.015333 | 0.0115 | 501.483 | 0.0011667 | 1.8992e-05
501.784 | 0.012 | 501.784 | 0.012 | 0.013 | 0.015 | 0.018 | 0.021 | 0.022 | 0.022 | 0.024 | 0.024 | 0.013333 | 0.020333 | 0.023333 | 0.018333 | 501.784 | 0.002 | 3.2557e-05
502.085 | 0.001 | 502.085 | 0.001 | 0.007 | 0.006 | 0.006 | 0.009 | 0.012 | 0.011 | 0.013 | 0.013 | 0.0046667 | 0.009 | 0.012333 | 0.0085 | 502.085 | 0.0005 | 8.1393e-06
502.387 | 0.003 | 502.387 | 0.003 | 0.003 | 0.005 | 0.007 | 0.009 | 0.01 | 0.01 | 0.012 | 0.014 | 0.0036667 | 0.0086667 | 0.012 | 0.0078333 | 502.387 | 0.00083333 | 1.3565e-05
502.688 | -0.001 | 502.688 | -0.001 | 0.003 | 0.004 | 0.01 | 0.009 | 0.01 | 0.012 | 0.015 | 0.014 | 0.002 | 0.0096667 | 0.013667 | 0.0078333 | 502.688 | 0.0018333 | 2.9844e-05
502.989 | 0.009 | 502.989 | 0.009 | 0.01 | 0.014 | 0.02 | 0.018 | 0.02 | 0.018 | 0.022 | 0.025 | 0.011 | 0.019333 | 0.021667 | 0.016333 | 502.989 | 0.003 | 4.8836e-05
503.29 | 0.001 | 503.29 | 0.001 | 0.001 | 0.002 | 0.007 | 0.009 | 0.008 | 0.009 | 0.012 | 0.013 | 0.0013333 | 0.008 | 0.011333 | 0.0063333 | 503.29 | 0.0016667 | 2.7131e-05
503.591 | 0 | 503.591 | 0 | 0.004 | 0.003 | 0.009 | 0.011 | 0.009 | 0.01 | 0.012 | 0.016 | 0.0023333 | 0.0096667 | 0.012667 | 0.0075 | 503.591 | 0.0021667 | 3.527e-05
503.892 | 0.001 | 503.892 | 0.001 | 0 | 0.006 | 0.008 | 0.008 | 0.007 | 0.01 | 0.013 | 0.011 | 0.0023333 | 0.0076667 | 0.011333 | 0.0068333 | 503.892 | 0.00083333 | 1.3565e-05
504.193 | 0.016 | 504.193 | 0.016 | 0.017 | 0.017 | 0.02 | 0.019 | 0.024 | 0.024 | 0.027 | 0.026 | 0.016667 | 0.021 | 0.025667 | 0.021167 | 504.193 | -0.00016667 | -2.7131e-06
504.494 | 0.007 | 504.494 | 0.007 | 0.007 | 0.008 | 0.013 | 0.014 | 0.016 | 0.015 | 0.018 | 0.016 | 0.0073333 | 0.014333 | 0.016333 | 0.011833 | 504.494 | 0.0025 | 4.0696e-05
504.795 | 0.003 | 504.795 | 0.003 | 0.001 | 0.003 | 0.009 | 0.006 | 0.008 | 0.009 | 0.007 | 0.012 | 0.0023333 | 0.0076667 | 0.0093333 | 0.0058333 | 504.795 | 0.0018333 | 2.9844e-05
505.096 | 0.015 | 505.096 | 0.015 | 0.016 | 0.022 | 0.02 | 0.023 | 0.024 | 0.025 | 0.026 | 0.031 | 0.017667 | 0.022333 | 0.027333 | 0.0225 | 505.096 | -0.00016667 | -2.7131e-06
505.396 | 0.002 | 505.396 | 0.002 | 0.001 | 0.004 | 0.01 | 0.008 | 0.012 | 0.011 | 0.012 | 0.014 | 0.0023333 | 0.01 | 0.012333 | 0.0073333 | 505.396 | 0.0026667 | 4.3409e-05
505.697 | -0.005 | 505.697 | -0.005 | -0.003 | -0.001 | 0.002 | 0.005 | 0.004 | 0.007 | 0.003 | 0.007 | -0.003 | 0.0036667 | 0.0056667 | 0.0013333 | 505.697 | 0.0023333 | 3.7983e-05
505.997 | 0.011 | 505.997 | 0.011 | 0.006 | 0.009 | 0.011 | 0.015 | 0.017 | 0.018 | 0.017 | 0.016 | 0.0086667 | 0.014333 | 0.017 | 0.012833 | 505.997 | 0.0015 | 2.4418e-05
506.298 | 0.01 | 506.298 | 0.01 | 0.01 | 0.012 | 0.017 | 0.016 | 0.017 | 0.016 | 0.019 | 0.021 | 0.010667 | 0.016667 | 0.018667 | 0.014667 | 506.298 | 0.002 | 3.2557e-05
506.598 | 0.002 | 506.598 | 0.002 | 0 | 0.008 | 0.012 | 0.009 | 0.015 | 0.012 | 0.012 | 0.017 | 0.0033333 | 0.012 | 0.013667 | 0.0085 | 506.598 | 0.0035 | 5.6975e-05
506.898 | 0.012 | 506.898 | 0.012 | 0.009 | 0.011 | 0.02 | 0.02 | 0.023 | 0.021 | 0.024 | 0.021 | 0.010667 | 0.021 | 0.022 | 0.016333 | 506.898 | 0.0046667 | 7.5967e-05
507.199 | 0.004 | 507.199 | 0.004 | 0.008 | 0.006 | 0.012 | 0.012 | 0.012 | 0.015 | 0.015 | 0.016 | 0.006 | 0.012 | 0.015333 | 0.010667 | 507.199 | 0.0013333 | 2.1705e-05
507.499 | 0.009 | 507.499 | 0.009 | 0.009 | 0.01 | 0.013 | 0.017 | 0.017 | 0.019 | 0.018 | 0.021 | 0.0093333 | 0.015667 | 0.019333 | 0.014333 | 507.499 | 0.0013333 | 2.1705e-05
507.799 | 0.007 | 507.799 | 0.007 | 0.004 | 0.009 | 0.01 | 0.01 | 0.012 | 0.013 | 0.014 | 0.016 | 0.0066667 | 0.010667 | 0.014333 | 0.0105 | 507.799 | 0.00016667 | 2.7131e-06
508.099 | 0.004 | 508.099 | 0.004 | 0.005 | 0.008 | 0.011 | 0.014 | 0.012 | 0.015 | 0.016 | 0.015 | 0.0056667 | 0.012333 | 0.015333 | 0.0105 | 508.099 | 0.0018333 | 2.9844e-05
508.399 | 0.011 | 508.399 | 0.011 | 0.011 | 0.012 | 0.017 | 0.016 | 0.016 | 0.015 | 0.02 | 0.019 | 0.011333 | 0.016333 | 0.018 | 0.014667 | 508.399 | 0.0016667 | 2.7131e-05
508.699 | 0.01 | 508.699 | 0.01 | 0.011 | 0.012 | 0.014 | 0.018 | 0.014 | 0.016 | 0.018 | 0.02 | 0.011 | 0.015333 | 0.018 | 0.0145 | 508.699 | 0.00083333 | 1.3565e-05
508.998 | 0.004 | 508.998 | 0.004 | 0.008 | 0.011 | 0.01 | 0.014 | 0.014 | 0.012 | 0.015 | 0.019 | 0.0076667 | 0.012667 | 0.015333 | 0.0115 | 508.998 | 0.0011667 | 1.8992e-05
509.298 | 0.011 | 509.298 | 0.011 | 0.014 | 0.017 | 0.022 | 0.021 | 0.02 | 0.026 | 0.024 | 0.029 | 0.014 | 0.021 | 0.026333 | 0.020167 | 509.298 | 0.00083333 | 1.3565e-05
509.598 | 0.003 | 509.598 | 0.003 | 0.008 | 0.004 | 0.011 | 0.009 | 0.012 | 0.012 | 0.013 | 0.015 | 0.005 | 0.010667 | 0.013333 | 0.0091667 | 509.598 | 0.0015 | 2.4418e-05
509.897 | 0.01 | 509.897 | 0.01 | 0.011 | 0.013 | 0.015 | 0.017 | 0.019 | 0.018 | 0.021 | 0.024 | 0.011333 | 0.017 | 0.021 | 0.016167 | 509.897 | 0.00083333 | 1.3565e-05
510.197 | 0.001 | 510.197 | 0.001 | 0.004 | 0.005 | 0.008 | 0.01 | 0.011 | 0.014 | 0.013 | 0.014 | 0.0033333 | 0.0096667 | 0.013667 | 0.0085 | 510.197 | 0.0011667 | 1.8992e-05
510.496 | 0.006 | 510.496 | 0.006 | 0.007 | 0.009 | 0.012 | 0.013 | 0.015 | 0.015 | 0.017 | 0.017 | 0.0073333 | 0.013333 | 0.016333 | 0.011833 | 510.496 | 0.0015 | 2.4418e-05
510.796 | -0.001 | 510.796 | -0.001 | 0 | 0.005 | 0.007 | 0.01 | 0.012 | 0.008 | 0.009 | 0.012 | 0.0013333 | 0.0096667 | 0.0096667 | 0.0055 | 510.796 | 0.0041667 | 6.7827e-05
511.095 | 0.008 | 511.095 | 0.008 | 0.009 | 0.008 | 0.013 | 0.017 | 0.013 | 0.016 | 0.016 | 0.016 | 0.0083333 | 0.014333 | 0.016 | 0.012167 | 511.095 | 0.0021667 | 3.527e-05
511.394 | 0.005 | 511.394 | 0.005 | 0.008 | 0.012 | 0.011 | 0.015 | 0.015 | 0.017 | 0.02 | 0.019 | 0.0083333 | 0.013667 | 0.018667 | 0.0135 | 511.394 | 0.00016667 | 2.7131e-06
511.693 | 0.013 | 511.693 | 0.013 | 0.015 | 0.016 | 0.019 | 0.024 | 0.025 | 0.026 | 0.031 | 0.031 | 0.014667 | 0.022667 | 0.029333 | 0.022 | 511.693 | 0.00066667 | 1.0852e-05
511.992 | 0.009 | 511.992 | 0.009 | 0.012 | 0.011 | 0.016 | 0.015 | 0.016 | 0.021 | 0.018 | 0.022 | 0.010667 | 0.015667 | 0.020333 | 0.0155 | 511.992 | 0.00016667 | 2.7131e-06
512.291 | 0.01 | 512.291 | 0.01 | 0.011 | 0.011 | 0.016 | 0.019 | 0.018 | 0.017 | 0.02 | 0.022 | 0.010667 | 0.017667 | 0.019667 | 0.015167 | 512.291 | 0.0025 | 4.0696e-05
512.59 | 0.009 | 512.59 | 0.009 | 0.014 | 0.014 | 0.017 | 0.02 | 0.019 | 0.021 | 0.02 | 0.021 | 0.012333 | 0.018667 | 0.020667 | 0.0165 | 512.59 | 0.0021667 | 3.527e-05
512.889 | 0.008 | 512.889 | 0.008 | 0.014 | 0.015 | 0.019 | 0.018 | 0.018 | 0.021 | 0.023 | 0.022 | 0.012333 | 0.018333 | 0.022 | 0.017167 | 512.889 | 0.0011667 | 1.8992e-05
513.188 | 0.009 | 513.188 | 0.009 | 0.011 | 0.014 | 0.018 | 0.02 | 0.021 | 0.02 | 0.022 | 0.025 | 0.011333 | 0.019667 | 0.022333 | 0.016833 | 513.188 | 0.0028333 | 4.6123e-05
513.487 | 0.007 | 513.487 | 0.007 | 0.008 | 0.011 | 0.014 | 0.019 | 0.016 | 0.02 | 0.019 | 0.02 | 0.0086667 | 0.016333 | 0.019667 | 0.014167 | 513.487 | 0.0021667 | 3.527e-05
513.786 | 0.011 | 513.786 | 0.011 | 0.011 | 0.014 | 0.016 | 0.016 | 0.022 | 0.022 | 0.023 | 0.026 | 0.012 | 0.018 | 0.023667 | 0.017833 | 513.786 | 0.00016667 | 2.7131e-06
514.084 | 0.01 | 514.084 | 0.01 | 0.012 | 0.011 | 0.015 | 0.017 | 0.019 | 0.02 | 0.022 | 0.023 | 0.011 | 0.017 | 0.021667 | 0.016333 | 514.084 | 0.00066667 | 1.0852e-05
514.383 | 0.014 | 514.383 | 0.014 | 0.014 | 0.015 | 0.018 | 0.021 | 0.022 | 0.022 | 0.023 | 0.025 | 0.014333 | 0.020333 | 0.023333 | 0.018833 | 514.383 | 0.0015 | 2.4418e-05
514.681 | 0.018 | 514.681 | 0.018 | 0.018 | 0.02 | 0.021 | 0.025 | 0.026 | 0.024 | 0.027 | 0.03 | 0.018667 | 0.024 | 0.027 | 0.022833 | 514.681 | 0.0011667 | 1.8992e-05
514.98 | 0.008 | 514.98 | 0.008 | 0.009 | 0.011 | 0.016 | 0.014 | 0.015 | 0.019 | 0.017 | 0.022 | 0.0093333 | 0.015 | 0.019333 | 0.014333 | 514.98 | 0.00066667 | 1.0852e-05
515.278 | 0.011 | 515.278 | 0.011 | 0.013 | 0.017 | 0.021 | 0.02 | 0.024 | 0.023 | 0.024 | 0.024 | 0.013667 | 0.021667 | 0.023667 | 0.018667 | 515.278 | 0.003 | 4.8836e-05
515.576 | 0.018 | 515.576 | 0.018 | 0.016 | 0.019 | 0.02 | 0.023 | 0.024 | 0.027 | 0.029 | 0.026 | 0.017667 | 0.022333 | 0.027333 | 0.0225 | 515.576 | -0.00016667 | -2.7131e-06
515.874 | 0.01 | 515.874 | 0.01 | 0.011 | 0.009 | 0.014 | 0.018 | 0.02 | 0.021 | 0.022 | 0.023 | 0.01 | 0.017333 | 0.022 | 0.016 | 515.874 | 0.0013333 | 2.1705e-05
516.173 | 0.001 | 516.173 | 0.001 | 0.004 | 0.005 | 0.012 | 0.013 | 0.012 | 0.012 | 0.013 | 0.014 | 0.0033333 | 0.012333 | 0.013 | 0.0081667 | 516.173 | 0.0041667 | 6.7827e-05
516.471 | 0.015 | 516.471 | 0.015 | 0.017 | 0.018 | 0.016 | 0.023 | 0.023 | 0.026 | 0.025 | 0.026 | 0.016667 | 0.020667 | 0.025667 | 0.021167 | 516.471 | -0.0005 | -8.1393e-06
516.769 | 0.01 | 516.769 | 0.01 | 0.011 | 0.012 | 0.013 | 0.017 | 0.019 | 0.022 | 0.019 | 0.022 | 0.011 | 0.016333 | 0.021 | 0.016 | 516.769 | 0.00033333 | 5.4262e-06
517.067 | 0.006 | 517.067 | 0.006 | 0.009 | 0.011 | 0.012 | 0.016 | 0.016 | 0.016 | 0.015 | 0.018 | 0.0086667 | 0.014667 | 0.016333 | 0.0125 | 517.067 | 0.0021667 | 3.527e-05
517.364 | -0.001 | 517.364 | -0.001 | 0.003 | 0.006 | 0.008 | 0.011 | 0.012 | 0.013 | 0.014 | 0.019 | 0.0026667 | 0.010333 | 0.015333 | 0.009 | 517.364 | 0.0013333 | 2.1705e-05
517.662 | 0.002 | 517.662 | 0.002 | 0.003 | 0.001 | 0.007 | 0.011 | 0.006 | 0.01 | 0.012 | 0.01 | 0.002 | 0.008 | 0.010667 | 0.0063333 | 517.662 | 0.0016667 | 2.7131e-05
517.96 | 0 | 517.96 | 0 | 0.007 | 0.005 | 0.007 | 0.013 | 0.016 | 0.014 | 0.014 | 0.015 | 0.004 | 0.012 | 0.014333 | 0.0091667 | 517.96 | 0.0028333 | 4.6123e-05
518.257 | 0.003 | 518.257 | 0.003 | 0.006 | 0.007 | 0.011 | 0.013 | 0.015 | 0.014 | 0.016 | 0.017 | 0.0053333 | 0.013 | 0.015667 | 0.0105 | 518.257 | 0.0025 | 4.0696e-05
518.555 | 0.006 | 518.555 | 0.006 | 0.002 | 0.007 | 0.014 | 0.015 | 0.013 | 0.017 | 0.016 | 0.018 | 0.005 | 0.014 | 0.017 | 0.011 | 518.555 | 0.003 | 4.8836e-05
518.853 | 0.005 | 518.853 | 0.005 | 0.005 | 0.007 | 0.007 | 0.011 | 0.011 | 0.013 | 0.013 | 0.015 | 0.0056667 | 0.0096667 | 0.013667 | 0.0096667 | 518.853 | 0 | 0
519.15 | 0.004 | 519.15 | 0.004 | 0.005 | 0.005 | 0.009 | 0.013 | 0.011 | 0.011 | 0.015 | 0.016 | 0.0046667 | 0.011 | 0.014 | 0.0093333 | 519.15 | 0.0016667 | 2.7131e-05
519.447 | 0.008 | 519.447 | 0.008 | 0.008 | 0.008 | 0.01 | 0.017 | 0.015 | 0.019 | 0.017 | 0.019 | 0.008 | 0.014 | 0.018333 | 0.013167 | 519.447 | 0.00083333 | 1.3565e-05
519.745 | 0.009 | 519.745 | 0.009 | 0.009 | 0.01 | 0.014 | 0.013 | 0.016 | 0.017 | 0.014 | 0.017 | 0.0093333 | 0.014333 | 0.016 | 0.012667 | 519.745 | 0.0016667 | 2.7131e-05
520.042 | 0.011 | 520.042 | 0.011 | 0.011 | 0.017 | 0.018 | 0.02 | 0.024 | 0.023 | 0.023 | 0.023 | 0.013 | 0.020667 | 0.023 | 0.018 | 520.042 | 0.0026667 | 4.3409e-05
520.339 | 0.003 | 520.339 | 0.003 | 0.003 | 0.004 | 0.009 | 0.009 | 0.012 | 0.009 | 0.011 | 0.013 | 0.0033333 | 0.01 | 0.011 | 0.0071667 | 520.339 | 0.0028333 | 4.6123e-05
520.636 | 0.01 | 520.636 | 0.01 | 0.013 | 0.014 | 0.015 | 0.018 | 0.021 | 0.021 | 0.02 | 0.023 | 0.012333 | 0.018 | 0.021333 | 0.016833 | 520.636 | 0.0011667 | 1.8992e-05
520.933 | 0.012 | 520.933 | 0.012 | 0.02 | 0.02 | 0.022 | 0.023 | 0.023 | 0.027 | 0.028 | 0.034 | 0.017333 | 0.022667 | 0.029667 | 0.0235 | 520.933 | -0.00083333 | -1.3565e-05
521.23 | 0.005 | 521.23 | 0.005 | 0.005 | 0.006 | 0.011 | 0.012 | 0.009 | 0.01 | 0.018 | 0.018 | 0.0053333 | 0.010667 | 0.015333 | 0.010333 | 521.23 | 0.00033333 | 5.4262e-06
521.527 | 0.003 | 521.527 | 0.003 | 0.003 | 0.005 | 0.014 | 0.013 | 0.013 | 0.014 | 0.018 | 0.014 | 0.0036667 | 0.013333 | 0.015333 | 0.0095 | 521.527 | 0.0038333 | 6.2401e-05
521.824 | 0.009 | 521.824 | 0.009 | 0.014 | 0.012 | 0.017 | 0.02 | 0.019 | 0.021 | 0.021 | 0.025 | 0.011667 | 0.018667 | 0.022333 | 0.017 | 521.824 | 0.0016667 | 2.7131e-05
522.121 | 0.005 | 522.121 | 0.005 | 0.005 | 0.007 | 0.008 | 0.011 | 0.014 | 0.011 | 0.012 | 0.015 | 0.0056667 | 0.011 | 0.012667 | 0.0091667 | 522.121 | 0.0018333 | 2.9844e-05
522.417 | 0.001 | 522.417 | 0.001 | 0.004 | 0.008 | 0.011 | 0.008 | 0.01 | 0.014 | 0.015 | 0.017 | 0.0043333 | 0.0096667 | 0.015333 | 0.0098333 | 522.417 | -0.00016667 | -2.7131e-06
522.714 | -0.001 | 522.714 | -0.001 | 0.001 | 0.003 | 0.005 | 0.006 | 0.008 | 0.011 | 0.009 | 0.014 | 0.001 | 0.0063333 | 0.011333 | 0.0061667 | 522.714 | 0.00016667 | 2.7131e-06
523.01 | 0.006 | 523.01 | 0.006 | 0.009 | 0.01 | 0.013 | 0.015 | 0.019 | 0.017 | 0.017 | 0.022 | 0.0083333 | 0.015667 | 0.018667 | 0.0135 | 523.01 | 0.0021667 | 3.527e-05
523.307 | 0.003 | 523.307 | 0.003 | 0.008 | 0.011 | 0.009 | 0.012 | 0.013 | 0.016 | 0.016 | 0.018 | 0.0073333 | 0.011333 | 0.016667 | 0.012 | 523.307 | -0.00066667 | -1.0852e-05
523.603 | 0.01 | 523.603 | 0.01 | 0.011 | 0.012 | 0.018 | 0.015 | 0.021 | 0.021 | 0.021 | 0.026 | 0.011 | 0.018 | 0.022667 | 0.016833 | 523.603 | 0.0011667 | 1.8992e-05
523.9 | 0.008 | 523.9 | 0.008 | 0.011 | 0.012 | 0.017 | 0.013 | 0.018 | 0.02 | 0.018 | 0.023 | 0.010333 | 0.016 | 0.020333 | 0.015333 | 523.9 | 0.00066667 | 1.0852e-05
524.196 | 0.009 | 524.196 | 0.009 | 0.012 | 0.013 | 0.019 | 0.019 | 0.02 | 0.024 | 0.022 | 0.023 | 0.011333 | 0.019333 | 0.023 | 0.017167 | 524.196 | 0.0021667 | 3.527e-05
524.492 | 0 | 524.492 | 0 | 0.002 | 0 | 0.008 | 0.008 | 0.006 | 0.01 | 0.009 | 0.011 | 0.00066667 | 0.0073333 | 0.01 | 0.0053333 | 524.492 | 0.002 | 3.2557e-05
524.788 | 0.016 | 524.788 | 0.016 | 0.018 | 0.02 | 0.022 | 0.024 | 0.023 | 0.027 | 0.028 | 0.029 | 0.018 | 0.023 | 0.028 | 0.023 | 524.788 | 0 | 0
525.084 | 0.009 | 525.084 | 0.009 | 0.012 | 0.01 | 0.015 | 0.018 | 0.017 | 0.02 | 0.02 | 0.025 | 0.010333 | 0.016667 | 0.021667 | 0.016 | 525.084 | 0.00066667 | 1.0852e-05
525.38 | 0.012 | 525.38 | 0.012 | 0.016 | 0.016 | 0.02 | 0.023 | 0.024 | 0.028 | 0.027 | 0.026 | 0.014667 | 0.022333 | 0.027 | 0.020833 | 525.38 | 0.0015 | 2.4418e-05
525.676 | 0.003 | 525.676 | 0.003 | 0.002 | 0.003 | 0.009 | 0.011 | 0.01 | 0.012 | 0.012 | 0.015 | 0.0026667 | 0.01 | 0.013 | 0.0078333 | 525.676 | 0.0021667 | 3.527e-05
525.972 | 0.009 | 525.972 | 0.009 | 0.007 | 0.01 | 0.012 | 0.017 | 0.017 | 0.018 | 0.019 | 0.019 | 0.0086667 | 0.015333 | 0.018667 | 0.013667 | 525.972 | 0.0016667 | 2.7131e-05
526.268 | 0.007 | 526.268 | 0.007 | 0.011 | 0.01 | 0.014 | 0.016 | 0.018 | 0.018 | 0.021 | 0.021 | 0.0093333 | 0.016 | 0.02 | 0.014667 | 526.268 | 0.0013333 | 2.1705e-05
526.564 | 0.007 | 526.564 | 0.007 | 0.009 | 0.01 | 0.018 | 0.017 | 0.019 | 0.016 | 0.024 | 0.024 | 0.0086667 | 0.018 | 0.021333 | 0.015 | 526.564 | 0.003 | 4.8836e-05
526.859 | 0.005 | 526.859 | 0.005 | 0.009 | 0.015 | 0.014 | 0.018 | 0.015 | 0.02 | 0.022 | 0.024 | 0.0096667 | 0.015667 | 0.022 | 0.015833 | 526.859 | -0.00016667 | -2.7131e-06
527.155 | 0.007 | 527.155 | 0.007 | 0.005 | 0.005 | 0.013 | 0.01 | 0.014 | 0.014 | 0.018 | 0.018 | 0.0056667 | 0.012333 | 0.016667 | 0.011167 | 527.155 | 0.0011667 | 1.8992e-05
527.45 | 0.013 | 527.45 | 0.013 | 0.012 | 0.015 | 0.021 | 0.017 | 0.019 | 0.023 | 0.02 | 0.024 | 0.013333 | 0.019 | 0.022333 | 0.017833 | 527.45 | 0.0011667 | 1.8992e-05
527.746 | 0.01 | 527.746 | 0.01 | 0.013 | 0.01 | 0.015 | 0.018 | 0.02 | 0.019 | 0.021 | 0.024 | 0.011 | 0.017667 | 0.021333 | 0.016167 | 527.746 | 0.0015 | 2.4418e-05
528.041 | 0.015 | 528.041 | 0.015 | 0.015 | 0.015 | 0.023 | 0.02 | 0.023 | 0.026 | 0.025 | 0.029 | 0.015 | 0.022 | 0.026667 | 0.020833 | 528.041 | 0.0011667 | 1.8992e-05
528.336 | 0.01 | 528.336 | 0.01 | 0.013 | 0.015 | 0.018 | 0.016 | 0.022 | 0.019 | 0.021 | 0.021 | 0.012667 | 0.018667 | 0.020333 | 0.0165 | 528.336 | 0.0021667 | 3.527e-05
528.632 | 0.003 | 528.632 | 0.003 | 0.005 | 0.007 | 0.009 | 0.01 | 0.009 | 0.009 | 0.013 | 0.011 | 0.005 | 0.0093333 | 0.011 | 0.008 | 528.632 | 0.0013333 | 2.1705e-05
528.927 | 0.005 | 528.927 | 0.005 | 0.004 | 0.007 | 0.008 | 0.012 | 0.012 | 0.01 | 0.014 | 0.014 | 0.0053333 | 0.010667 | 0.012667 | 0.009 | 528.927 | 0.0016667 | 2.7131e-05
529.222 | 0.006 | 529.222 | 0.006 | 0.008 | 0.011 | 0.014 | 0.017 | 0.02 | 0.02 | 0.016 | 0.02 | 0.0083333 | 0.017 | 0.018667 | 0.0135 | 529.222 | 0.0035 | 5.6975e-05
529.517 | 0.012 | 529.517 | 0.012 | 0.015 | 0.015 | 0.017 | 0.019 | 0.023 | 0.019 | 0.025 | 0.026 | 0.014 | 0.019667 | 0.023333 | 0.018667 | 529.517 | 0.001 | 1.6279e-05
529.812 | 0.006 | 529.812 | 0.006 | 0.007 | 0.009 | 0.013 | 0.014 | 0.015 | 0.016 | 0.018 | 0.017 | 0.0073333 | 0.014 | 0.017 | 0.012167 | 529.812 | 0.0018333 | 2.9844e-05
530.107 | 0.005 | 530.107 | 0.005 | 0.005 | 0.007 | 0.01 | 0.014 | 0.011 | 0.013 | 0.013 | 0.018 | 0.0056667 | 0.011667 | 0.014667 | 0.010167 | 530.107 | 0.0015 | 2.4418e-05
530.402 | 0.008 | 530.402 | 0.008 | 0.01 | 0.009 | 0.015 | 0.016 | 0.018 | 0.02 | 0.018 | 0.018 | 0.009 | 0.016333 | 0.018667 | 0.013833 | 530.402 | 0.0025 | 4.0696e-05
530.696 | 0.013 | 530.696 | 0.013 | 0.014 | 0.014 | 0.018 | 0.021 | 0.025 | 0.023 | 0.025 | 0.027 | 0.013667 | 0.021333 | 0.025 | 0.019333 | 530.696 | 0.002 | 3.2557e-05
530.991 | 0.018 | 530.991 | 0.018 | 0.016 | 0.018 | 0.022 | 0.025 | 0.024 | 0.023 | 0.026 | 0.026 | 0.017333 | 0.023667 | 0.025 | 0.021167 | 530.991 | 0.0025 | 4.0696e-05
531.286 | 0.007 | 531.286 | 0.007 | 0.012 | 0.013 | 0.014 | 0.018 | 0.017 | 0.022 | 0.024 | 0.021 | 0.010667 | 0.016333 | 0.022333 | 0.0165 | 531.286 | -0.00016667 | -2.7131e-06
531.58 | 0.012 | 531.58 | 0.012 | 0.013 | 0.015 | 0.015 | 0.017 | 0.019 | 0.021 | 0.023 | 0.023 | 0.013333 | 0.017 | 0.022333 | 0.017833 | 531.58 | -0.00083333 | -1.3565e-05
531.875 | 0.004 | 531.875 | 0.004 | 0.009 | 0.007 | 0.01 | 0.014 | 0.016 | 0.016 | 0.016 | 0.015 | 0.0066667 | 0.013333 | 0.015667 | 0.011167 | 531.875 | 0.0021667 | 3.527e-05
532.169 | 0.008 | 532.169 | 0.008 | 0.011 | 0.012 | 0.014 | 0.017 | 0.019 | 0.02 | 0.023 | 0.02 | 0.010333 | 0.016667 | 0.021 | 0.015667 | 532.169 | 0.001 | 1.6279e-05
532.463 | 0.004 | 532.463 | 0.004 | 0.007 | 0.006 | 0.011 | 0.012 | 0.013 | 0.016 | 0.017 | 0.019 | 0.0056667 | 0.012 | 0.017333 | 0.0115 | 532.463 | 0.0005 | 8.1393e-06
532.758 | -0.001 | 532.758 | -0.001 | 0.001 | 0.002 | 0.01 | 0.009 | 0.009 | 0.007 | 0.012 | 0.011 | 0.00066667 | 0.0093333 | 0.01 | 0.0053333 | 532.758 | 0.004 | 6.5114e-05
533.052 | 0.011 | 533.052 | 0.011 | 0.01 | 0.009 | 0.015 | 0.018 | 0.015 | 0.019 | 0.021 | 0.022 | 0.01 | 0.016 | 0.020667 | 0.015333 | 533.052 | 0.00066667 | 1.0852e-05
533.346 | -0.002 | 533.346 | -0.002 | 0.004 | 0.005 | 0.008 | 0.01 | 0.01 | 0.012 | 0.014 | 0.013 | 0.0023333 | 0.0093333 | 0.013 | 0.0076667 | 533.346 | 0.0016667 | 2.7131e-05
533.64 | 0.005 | 533.64 | 0.005 | 0.013 | 0.011 | 0.014 | 0.017 | 0.018 | 0.021 | 0.018 | 0.019 | 0.0096667 | 0.016333 | 0.019333 | 0.0145 | 533.64 | 0.0018333 | 2.9844e-05
533.934 | 0.009 | 533.934 | 0.009 | 0.009 | 0.011 | 0.018 | 0.019 | 0.021 | 0.02 | 0.026 | 0.022 | 0.0096667 | 0.019333 | 0.022667 | 0.016167 | 533.934 | 0.0031667 | 5.1549e-05
534.228 | 0.013 | 534.228 | 0.013 | 0.016 | 0.017 | 0.018 | 0.021 | 0.023 | 0.023 | 0.024 | 0.024 | 0.015333 | 0.020667 | 0.023667 | 0.0195 | 534.228 | 0.0011667 | 1.8992e-05
534.522 | 0.01 | 534.522 | 0.01 | 0.014 | 0.016 | 0.019 | 0.018 | 0.021 | 0.023 | 0.024 | 0.022 | 0.013333 | 0.019333 | 0.023 | 0.018167 | 534.522 | 0.0011667 | 1.8992e-05
534.816 | 0.005 | 534.816 | 0.005 | 0.013 | 0.01 | 0.019 | 0.017 | 0.019 | 0.022 | 0.023 | 0.025 | 0.0093333 | 0.018333 | 0.023333 | 0.016333 | 534.816 | 0.002 | 3.2557e-05
535.11 | 0.002 | 535.11 | 0.002 | 0.007 | 0.004 | 0.008 | 0.008 | 0.012 | 0.012 | 0.012 | 0.012 | 0.0043333 | 0.0093333 | 0.012 | 0.0081667 | 535.11 | 0.0011667 | 1.8992e-05
535.403 | 0.017 | 535.403 | 0.017 | 0.021 | 0.019 | 0.022 | 0.019 | 0.024 | 0.025 | 0.028 | 0.028 | 0.019 | 0.021667 | 0.027 | 0.023 | 535.403 | -0.0013333 | -2.1705e-05
535.697 | 0.013 | 535.697 | 0.013 | 0.014 | 0.015 | 0.023 | 0.023 | 0.021 | 0.024 | 0.028 | 0.025 | 0.014 | 0.022333 | 0.025667 | 0.019833 | 535.697 | 0.0025 | 4.0696e-05
535.99 | 0.01 | 535.99 | 0.01 | 0.014 | 0.014 | 0.017 | 0.018 | 0.02 | 0.018 | 0.025 | 0.025 | 0.012667 | 0.018333 | 0.022667 | 0.017667 | 535.99 | 0.00066667 | 1.0852e-05
536.284 | 0.006 | 536.284 | 0.006 | 0.005 | 0.008 | 0.014 | 0.015 | 0.015 | 0.018 | 0.02 | 0.021 | 0.0063333 | 0.014667 | 0.019667 | 0.013 | 536.284 | 0.0016667 | 2.7131e-05
536.577 | 0.017 | 536.577 | 0.017 | 0.021 | 0.025 | 0.027 | 0.027 | 0.026 | 0.028 | 0.029 | 0.031 | 0.021 | 0.026667 | 0.029333 | 0.025167 | 536.577 | 0.0015 | 2.4418e-05
536.871 | 0.007 | 536.871 | 0.007 | 0.006 | 0.009 | 0.014 | 0.015 | 0.018 | 0.019 | 0.02 | 0.023 | 0.0073333 | 0.015667 | 0.020667 | 0.014 | 536.871 | 0.0016667 | 2.7131e-05
537.164 | 0.008 | 537.164 | 0.008 | 0.009 | 0.009 | 0.014 | 0.015 | 0.015 | 0.017 | 0.021 | 0.021 | 0.0086667 | 0.014667 | 0.019667 | 0.014167 | 537.164 | 0.0005 | 8.1393e-06
537.457 | 0.016 | 537.457 | 0.016 | 0.016 | 0.015 | 0.021 | 0.018 | 0.018 | 0.024 | 0.024 | 0.022 | 0.015667 | 0.019 | 0.023333 | 0.0195 | 537.457 | -0.0005 | -8.1393e-06
537.75 | 0.005 | 537.75 | 0.005 | 0.007 | 0.011 | 0.014 | 0.015 | 0.013 | 0.019 | 0.018 | 0.019 | 0.0076667 | 0.014 | 0.018667 | 0.013167 | 537.75 | 0.00083333 | 1.3565e-05
538.043 | 0.005 | 538.043 | 0.005 | 0.008 | 0.011 | 0.017 | 0.017 | 0.018 | 0.019 | 0.023 | 0.022 | 0.008 | 0.017333 | 0.021333 | 0.014667 | 538.043 | 0.0026667 | 4.3409e-05
538.336 | 0.012 | 538.336 | 0.012 | 0.013 | 0.013 | 0.016 | 0.018 | 0.021 | 0.021 | 0.02 | 0.021 | 0.012667 | 0.018333 | 0.020667 | 0.016667 | 538.336 | 0.0016667 | 2.7131e-05
538.629 | 0.007 | 538.629 | 0.007 | 0.008 | 0.01 | 0.015 | 0.016 | 0.015 | 0.019 | 0.014 | 0.022 | 0.0083333 | 0.015333 | 0.018333 | 0.013333 | 538.629 | 0.002 | 3.2557e-05
538.922 | 0.01 | 538.922 | 0.01 | 0.011 | 0.013 | 0.017 | 0.016 | 0.018 | 0.018 | 0.021 | 0.02 | 0.011333 | 0.017 | 0.019667 | 0.0155 | 538.922 | 0.0015 | 2.4418e-05
539.215 | -0.002 | 539.215 | -0.002 | 0.001 | 0.004 | 0.003 | 0.003 | 0.008 | 0.009 | 0.008 | 0.011 | 0.001 | 0.0046667 | 0.0093333 | 0.0051667 | 539.215 | -0.0005 | -8.1393e-06
539.507 | 0.008 | 539.507 | 0.008 | 0.007 | 0.009 | 0.015 | 0.011 | 0.019 | 0.012 | 0.017 | 0.016 | 0.008 | 0.015 | 0.015 | 0.0115 | 539.507 | 0.0035 | 5.6975e-05
539.8 | -0.002 | 539.8 | -0.002 | 0.005 | 0.001 | 0.006 | 0.007 | 0.007 | 0.006 | 0.009 | 0.008 | 0.0013333 | 0.0066667 | 0.0076667 | 0.0045 | 539.8 | 0.0021667 | 3.527e-05
540.093 | 0.016 | 540.093 | 0.016 | 0.017 | 0.016 | 0.025 | 0.025 | 0.023 | 0.026 | 0.028 | 0.03 | 0.016333 | 0.024333 | 0.028 | 0.022167 | 540.093 | 0.0021667 | 3.527e-05
540.385 | 0.002 | 540.385 | 0.002 | 0.006 | 0.005 | 0.01 | 0.011 | 0.014 | 0.013 | 0.014 | 0.012 | 0.0043333 | 0.011667 | 0.013 | 0.0086667 | 540.385 | 0.003 | 4.8836e-05
540.678 | 0.004 | 540.678 | 0.004 | 0.004 | 0.006 | 0.015 | 0.015 | 0.014 | 0.015 | 0.016 | 0.016 | 0.0046667 | 0.014667 | 0.015667 | 0.010167 | 540.678 | 0.0045 | 7.3254e-05
540.97 | -0.001 | 540.97 | -0.001 | -0.001 | -0.001 | 0.004 | 0.006 | 0.006 | 0.005 | 0.006 | 0.009 | -0.001 | 0.0053333 | 0.0066667 | 0.0028333 | 540.97 | 0.0025 | 4.0696e-05
541.262 | 0.008 | 541.262 | 0.008 | 0.009 | 0.011 | 0.016 | 0.017 | 0.024 | 0.017 | 0.019 | 0.023 | 0.0093333 | 0.019 | 0.019667 | 0.0145 | 541.262 | 0.0045 | 7.3254e-05
541.554 | 0.009 | 541.554 | 0.009 | 0.008 | 0.01 | 0.013 | 0.014 | 0.015 | 0.015 | 0.017 | 0.018 | 0.009 | 0.014 | 0.016667 | 0.012833 | 541.554 | 0.0011667 | 1.8992e-05
541.847 | 0.01 | 541.847 | 0.01 | 0.011 | 0.011 | 0.018 | 0.018 | 0.016 | 0.022 | 0.022 | 0.022 | 0.010667 | 0.017333 | 0.022 | 0.016333 | 541.847 | 0.001 | 1.6279e-05
542.139 | 0.011 | 542.139 | 0.011 | 0.01 | 0.012 | 0.02 | 0.021 | 0.022 | 0.021 | 0.022 | 0.022 | 0.011 | 0.021 | 0.021667 | 0.016333 | 542.139 | 0.0046667 | 7.5967e-05
542.431 | 0.015 | 542.431 | 0.015 | 0.014 | 0.013 | 0.02 | 0.021 | 0.021 | 0.024 | 0.027 | 0.024 | 0.014 | 0.020667 | 0.025 | 0.0195 | 542.431 | 0.0011667 | 1.8992e-05
542.723 | 0.008 | 542.723 | 0.008 | 0.009 | 0.012 | 0.014 | 0.015 | 0.015 | 0.018 | 0.022 | 0.021 | 0.0096667 | 0.014667 | 0.020333 | 0.015 | 542.723 | -0.00033333 | -5.4262e-06
543.015 | 0.011 | 543.015 | 0.011 | 0.014 | 0.013 | 0.019 | 0.02 | 0.024 | 0.022 | 0.023 | 0.025 | 0.012667 | 0.021 | 0.023333 | 0.018 | 543.015 | 0.003 | 4.8836e-05
543.307 | 0.012 | 543.307 | 0.012 | 0.016 | 0.016 | 0.023 | 0.023 | 0.029 | 0.021 | 0.026 | 0.022 | 0.014667 | 0.025 | 0.023 | 0.018833 | 543.307 | 0.0061667 | 0.00010038
543.598 | 0.01 | 543.598 | 0.01 | 0.014 | 0.015 | 0.023 | 0.026 | 0.025 | 0.02 | 0.022 | 0.025 | 0.013 | 0.024667 | 0.022333 | 0.017667 | 543.598 | 0.007 | 0.00011395
543.89 | 0.013 | 543.89 | 0.013 | 0.017 | 0.016 | 0.022 | 0.024 | 0.025 | 0.022 | 0.027 | 0.026 | 0.015333 | 0.023667 | 0.025 | 0.020167 | 543.89 | 0.0035 | 5.6975e-05
544.182 | 0.005 | 544.182 | 0.005 | 0.01 | 0.011 | 0.018 | 0.018 | 0.02 | 0.017 | 0.016 | 0.02 | 0.0086667 | 0.018667 | 0.017667 | 0.013167 | 544.182 | 0.0055 | 8.9532e-05
544.473 | 0.014 | 544.473 | 0.014 | 0.013 | 0.017 | 0.026 | 0.028 | 0.031 | 0.024 | 0.023 | 0.025 | 0.014667 | 0.028333 | 0.024 | 0.019333 | 544.473 | 0.009 | 0.00014651
544.765 | 0.017 | 544.765 | 0.017 | 0.015 | 0.021 | 0.042 | 0.044 | 0.046 | 0.028 | 0.029 | 0.03 | 0.017667 | 0.044 | 0.029 | 0.023333 | 544.765 | 0.020667 | 0.00033642
545.056 | -0.001 | 545.056 | -0.001 | -0.001 | 0.003 | 0.082 | 0.079 | 0.081 | 0.008 | 0.011 | 0.008 | 0.00033333 | 0.080667 | 0.009 | 0.0046667 | 545.056 | 0.076 | 0.0012372
545.347 | 0.006 | 545.347 | 0.006 | 0.009 | 0.007 | 0.336 | 0.34 | 0.344 | 0.015 | 0.017 | 0.017 | 0.0073333 | 0.34 | 0.016333 | 0.011833 | 545.347 | 0.32817 | 0.0053421
545.639 | -0.002 | 545.639 | -0.002 | 0.003 | 0.007 | 0.708 | 0.696 | 0.704 | 0.013 | 0.012 | 0.015 | 0.0026667 | 0.70267 | 0.013333 | 0.008 | 545.639 | 0.69467 | 0.011308
545.93 | 0.004 | 545.93 | 0.004 | 0.003 | 0.005 | 0.959 | 0.952 | 0.959 | 0.01 | 0.016 | 0.014 | 0.004 | 0.95667 | 0.013333 | 0.0086667 | 545.93 | 0.948 | 0.015432
546.221 | 0.011 | 546.221 | 0.011 | 0.01 | 0.007 | 1.487 | 1.464 | 1.471 | 0.018 | 0.019 | 0.023 | 0.0093333 | 1.474 | 0.02 | 0.014667 | 546.221 | 1.4593 | 0.023756
546.512 | 0.012 | 546.512 | 0.012 | 0.011 | 0.011 | 0.773 | 0.748 | 0.756 | 0.018 | 0.021 | 0.023 | 0.011333 | 0.759 | 0.020667 | 0.016 | 546.512 | 0.743 | 0.012095
546.803 | 0.002 | 546.803 | 0.002 | 0.003 | 0.003 | 0.122 | 0.121 | 0.119 | 0.011 | 0.014 | 0.016 | 0.0026667 | 0.12067 | 0.013667 | 0.0081667 | 546.803 | 0.1125 | 0.0018313
547.094 | 0.001 | 547.094 | 0.001 | 0.003 | 0.005 | 0.034 | 0.039 | 0.041 | 0.015 | 0.018 | 0.015 | 0.003 | 0.038 | 0.016 | 0.0095 | 547.094 | 0.0285 | 0.00046394
547.385 | 0.003 | 547.385 | 0.003 | 0.007 | 0.007 | 0.027 | 0.019 | 0.029 | 0.017 | 0.021 | 0.022 | 0.0056667 | 0.025 | 0.02 | 0.012833 | 547.385 | 0.012167 | 0.00019806
547.676 | 0.012 | 547.676 | 0.012 | 0.014 | 0.013 | 0.026 | 0.026 | 0.026 | 0.02 | 0.021 | 0.02 | 0.013 | 0.026 | 0.020333 | 0.016667 | 547.676 | 0.0093333 | 0.00015193
547.967 | 0.022 | 547.967 | 0.022 | 0.021 | 0.026 | 0.034 | 0.035 | 0.036 | 0.034 | 0.031 | 0.032 | 0.023 | 0.035 | 0.032333 | 0.027667 | 547.967 | 0.0073333 | 0.00011938
548.257 | 0.007 | 548.257 | 0.007 | 0.006 | 0.004 | 0.018 | 0.015 | 0.021 | 0.017 | 0.017 | 0.018 | 0.0056667 | 0.018 | 0.017333 | 0.0115 | 548.257 | 0.0065 | 0.00010581
548.548 | 0.013 | 548.548 | 0.013 | 0.013 | 0.017 | 0.024 | 0.023 | 0.023 | 0.021 | 0.026 | 0.026 | 0.014333 | 0.023333 | 0.024333 | 0.019333 | 548.548 | 0.004 | 6.5114e-05
548.839 | -0.001 | 548.839 | -0.001 | 0.003 | 0.004 | 0.007 | 0.008 | 0.011 | 0.01 | 0.013 | 0.007 | 0.002 | 0.0086667 | 0.01 | 0.006 | 548.839 | 0.0026667 | 4.3409e-05
549.129 | 0.005 | 549.129 | 0.005 | 0.005 | 0.009 | 0.014 | 0.015 | 0.018 | 0.016 | 0.015 | 0.021 | 0.0063333 | 0.015667 | 0.017333 | 0.011833 | 549.129 | 0.0038333 | 6.2401e-05
549.419 | 0.015 | 549.419 | 0.015 | 0.015 | 0.016 | 0.023 | 0.028 | 0.025 | 0.026 | 0.026 | 0.025 | 0.015333 | 0.025333 | 0.025667 | 0.0205 | 549.419 | 0.0048333 | 7.868e-05
549.71 | 0.008 | 549.71 | 0.008 | 0.009 | 0.012 | 0.019 | 0.022 | 0.021 | 0.02 | 0.02 | 0.022 | 0.0096667 | 0.020667 | 0.020667 | 0.015167 | 549.71 | 0.0055 | 8.9532e-05
550 | 0.015 | 550 | 0.015 | 0.013 | 0.017 | 0.026 | 0.031 | 0.026 | 0.024 | 0.026 | 0.027 | 0.015 | 0.027667 | 0.025667 | 0.020333 | 550 | 0.0073333 | 0.00011938
550.29 | 0.006 | 550.29 | 0.006 | 0.011 | 0.011 | 0.018 | 0.022 | 0.021 | 0.015 | 0.015 | 0.017 | 0.0093333 | 0.020333 | 0.015667 | 0.0125 | 550.29 | 0.0078333 | 0.00012752
550.58 | -0.001 | 550.58 | -0.001 | 0.001 | 0.002 | 0.014 | 0.009 | 0.013 | 0.011 | 0.01 | 0.009 | 0.00066667 | 0.012 | 0.01 | 0.0053333 | 550.58 | 0.0066667 | 0.00010852
550.871 | 0.005 | 550.871 | 0.005 | 0.003 | 0.004 | 0.016 | 0.02 | 0.017 | 0.017 | 0.013 | 0.019 | 0.004 | 0.017667 | 0.016333 | 0.010167 | 550.871 | 0.0075 | 0.00012209
551.161 | 0.006 | 551.161 | 0.006 | 0.003 | 0.004 | 0.014 | 0.016 | 0.018 | 0.013 | 0.014 | 0.014 | 0.0043333 | 0.016 | 0.013667 | 0.009 | 551.161 | 0.007 | 0.00011395
551.45 | 0.002 | 551.45 | 0.002 | 0 | 0 | 0.008 | 0.012 | 0.015 | 0.01 | 0.013 | 0.013 | 0.00066667 | 0.011667 | 0.012 | 0.0063333 | 551.45 | 0.0053333 | 8.6819e-05
551.74 | 0.005 | 551.74 | 0.005 | 0.007 | 0.009 | 0.014 | 0.018 | 0.018 | 0.015 | 0.017 | 0.017 | 0.007 | 0.016667 | 0.016333 | 0.011667 | 551.74 | 0.005 | 8.1393e-05
552.03 | 0 | 552.03 | 0 | 0.005 | 0.004 | 0.013 | 0.011 | 0.015 | 0.012 | 0.011 | 0.009 | 0.003 | 0.013 | 0.010667 | 0.0068333 | 552.03 | 0.0061667 | 0.00010038
552.32 | 0.005 | 552.32 | 0.005 | 0.004 | 0.01 | 0.019 | 0.019 | 0.022 | 0.017 | 0.019 | 0.02 | 0.0063333 | 0.02 | 0.018667 | 0.0125 | 552.32 | 0.0075 | 0.00012209
552.61 | 0.009 | 552.61 | 0.009 | 0.007 | 0.01 | 0.016 | 0.022 | 0.024 | 0.02 | 0.024 | 0.021 | 0.0086667 | 0.020667 | 0.021667 | 0.015167 | 552.61 | 0.0055 | 8.9532e-05
552.899 | 0.019 | 552.899 | 0.019 | 0.018 | 0.021 | 0.027 | 0.029 | 0.029 | 0.027 | 0.029 | 0.028 | 0.019333 | 0.028333 | 0.028 | 0.023667 | 552.899 | 0.0046667 | 7.5967e-05
553.189 | 0.005 | 553.189 | 0.005 | 0.008 | 0.006 | 0.015 | 0.018 | 0.016 | 0.014 | 0.017 | 0.013 | 0.0063333 | 0.016333 | 0.014667 | 0.0105 | 553.189 | 0.0058333 | 9.4958e-05
553.478 | 0.005 | 553.478 | 0.005 | 0.01 | 0.007 | 0.017 | 0.019 | 0.018 | 0.014 | 0.018 | 0.017 | 0.0073333 | 0.018 | 0.016333 | 0.011833 | 553.478 | 0.0061667 | 0.00010038
553.768 | 0.001 | 553.768 | 0.001 | 0.004 | 0.006 | 0.012 | 0.02 | 0.016 | 0.012 | 0.011 | 0.013 | 0.0036667 | 0.016 | 0.012 | 0.0078333 | 553.768 | 0.0081667 | 0.00013294
554.057 | 0.014 | 554.057 | 0.014 | 0.018 | 0.02 | 0.027 | 0.031 | 0.032 | 0.03 | 0.029 | 0.031 | 0.017333 | 0.03 | 0.03 | 0.023667 | 554.057 | 0.0063333 | 0.0001031
554.346 | 0.012 | 554.346 | 0.012 | 0.016 | 0.016 | 0.025 | 0.026 | 0.027 | 0.024 | 0.025 | 0.027 | 0.014667 | 0.026 | 0.025333 | 0.02 | 554.346 | 0.006 | 9.7671e-05
554.635 | 0.009 | 554.635 | 0.009 | 0.006 | 0.011 | 0.019 | 0.024 | 0.023 | 0.02 | 0.022 | 0.022 | 0.0086667 | 0.022 | 0.021333 | 0.015 | 554.635 | 0.007 | 0.00011395
554.925 | 0.004 | 554.925 | 0.004 | 0.009 | 0.01 | 0.019 | 0.024 | 0.022 | 0.017 | 0.019 | 0.021 | 0.0076667 | 0.021667 | 0.019 | 0.013333 | 554.925 | 0.0083333 | 0.00013565
555.214 | 0.013 | 555.214 | 0.013 | 0.015 | 0.016 | 0.023 | 0.024 | 0.026 | 0.023 | 0.028 | 0.024 | 0.014667 | 0.024333 | 0.025 | 0.019833 | 555.214 | 0.0045 | 7.3254e-05
555.503 | 0 | 555.503 | 0 | 0.003 | 0.004 | 0.014 | 0.013 | 0.014 | 0.01 | 0.014 | 0.013 | 0.0023333 | 0.013667 | 0.012333 | 0.0073333 | 555.503 | 0.0063333 | 0.0001031
555.792 | 0.011 | 555.792 | 0.011 | 0.011 | 0.017 | 0.021 | 0.024 | 0.021 | 0.02 | 0.026 | 0.026 | 0.013 | 0.022 | 0.024 | 0.0185 | 555.792 | 0.0035 | 5.6975e-05
556.081 | -0.002 | 556.081 | -0.002 | 0 | 0.001 | 0.011 | 0.009 | 0.011 | 0.012 | 0.01 | 0.01 | -0.00033333 | 0.010333 | 0.010667 | 0.0051667 | 556.081 | 0.0051667 | 8.4106e-05
556.369 | 0.007 | 556.369 | 0.007 | 0.005 | 0.01 | 0.01 | 0.017 | 0.018 | 0.017 | 0.02 | 0.021 | 0.0073333 | 0.015 | 0.019333 | 0.013333 | 556.369 | 0.0016667 | 2.7131e-05
556.658 | 0.006 | 556.658 | 0.006 | 0.005 | 0.006 | 0.012 | 0.013 | 0.015 | 0.019 | 0.018 | 0.017 | 0.0056667 | 0.013333 | 0.018 | 0.011833 | 556.658 | 0.0015 | 2.4418e-05
556.947 | 0.013 | 556.947 | 0.013 | 0.015 | 0.019 | 0.019 | 0.023 | 0.022 | 0.021 | 0.029 | 0.023 | 0.015667 | 0.021333 | 0.024333 | 0.02 | 556.947 | 0.0013333 | 2.1705e-05
557.235 | 0 | 557.235 | 0 | 0.003 | 0.004 | 0.01 | 0.017 | 0.012 | 0.01 | 0.017 | 0.017 | 0.0023333 | 0.013 | 0.014667 | 0.0085 | 557.235 | 0.0045 | 7.3254e-05
557.524 | 0.019 | 557.524 | 0.019 | 0.021 | 0.026 | 0.03 | 0.027 | 0.032 | 0.03 | 0.032 | 0.029 | 0.022 | 0.029667 | 0.030333 | 0.026167 | 557.524 | 0.0035 | 5.6975e-05
557.812 | 0.005 | 557.812 | 0.005 | 0.007 | 0.011 | 0.016 | 0.019 | 0.017 | 0.017 | 0.019 | 0.019 | 0.0076667 | 0.017333 | 0.018333 | 0.013 | 557.812 | 0.0043333 | 7.054e-05
558.101 | 0.008 | 558.101 | 0.008 | 0.009 | 0.009 | 0.015 | 0.016 | 0.022 | 0.021 | 0.019 | 0.02 | 0.0086667 | 0.017667 | 0.02 | 0.014333 | 558.101 | 0.0033333 | 5.4262e-05
558.389 | 0.012 | 558.389 | 0.012 | 0.015 | 0.019 | 0.021 | 0.022 | 0.024 | 0.023 | 0.028 | 0.026 | 0.015333 | 0.022333 | 0.025667 | 0.0205 | 558.389 | 0.0018333 | 2.9844e-05
558.677 | 0.01 | 558.677 | 0.01 | 0.011 | 0.015 | 0.013 | 0.018 | 0.018 | 0.019 | 0.024 | 0.02 | 0.012 | 0.016333 | 0.021 | 0.0165 | 558.677 | -0.00016667 | -2.7131e-06
558.966 | 0.015 | 558.966 | 0.015 | 0.014 | 0.014 | 0.022 | 0.023 | 0.027 | 0.024 | 0.026 | 0.032 | 0.014333 | 0.024 | 0.027333 | 0.020833 | 558.966 | 0.0031667 | 5.1549e-05
559.254 | 0.018 | 559.254 | 0.018 | 0.019 | 0.02 | 0.026 | 0.028 | 0.03 | 0.03 | 0.03 | 0.029 | 0.019 | 0.028 | 0.029667 | 0.024333 | 559.254 | 0.0036667 | 5.9688e-05
559.542 | 0.008 | 559.542 | 0.008 | 0.01 | 0.014 | 0.019 | 0.021 | 0.019 | 0.018 | 0.021 | 0.019 | 0.010667 | 0.019667 | 0.019333 | 0.015 | 559.542 | 0.0046667 | 7.5967e-05
559.83 | 0.013 | 559.83 | 0.013 | 0.014 | 0.013 | 0.016 | 0.024 | 0.018 | 0.022 | 0.029 | 0.025 | 0.013333 | 0.019333 | 0.025333 | 0.019333 | 559.83 | 0 | 0
560.118 | 0.011 | 560.118 | 0.011 | 0.011 | 0.017 | 0.017 | 0.021 | 0.024 | 0.024 | 0.022 | 0.024 | 0.013 | 0.020667 | 0.023333 | 0.018167 | 560.118 | 0.0025 | 4.0696e-05
560.406 | 0 | 560.406 | 0 | 0.005 | 0.005 | 0.01 | 0.013 | 0.017 | 0.017 | 0.018 | 0.015 | 0.0033333 | 0.013333 | 0.016667 | 0.01 | 560.406 | 0.0033333 | 5.4262e-05
560.693 | 0.024 | 560.693 | 0.024 | 0.025 | 0.027 | 0.031 | 0.032 | 0.032 | 0.033 | 0.036 | 0.038 | 0.025333 | 0.031667 | 0.035667 | 0.0305 | 560.693 | 0.0011667 | 1.8992e-05
560.981 | 0.015 | 560.981 | 0.015 | 0.018 | 0.017 | 0.024 | 0.032 | 0.028 | 0.027 | 0.029 | 0.029 | 0.016667 | 0.028 | 0.028333 | 0.0225 | 560.981 | 0.0055 | 8.9532e-05
561.269 | 0.006 | 561.269 | 0.006 | 0.006 | 0.011 | 0.015 | 0.015 | 0.018 | 0.017 | 0.02 | 0.023 | 0.0076667 | 0.016 | 0.02 | 0.013833 | 561.269 | 0.0021667 | 3.527e-05
561.556 | 0.006 | 561.556 | 0.006 | 0.01 | 0.016 | 0.016 | 0.017 | 0.019 | 0.018 | 0.018 | 0.02 | 0.010667 | 0.017333 | 0.018667 | 0.014667 | 561.556 | 0.0026667 | 4.3409e-05
561.844 | 0.015 | 561.844 | 0.015 | 0.014 | 0.015 | 0.021 | 0.022 | 0.021 | 0.023 | 0.026 | 0.024 | 0.014667 | 0.021333 | 0.024333 | 0.0195 | 561.844 | 0.0018333 | 2.9844e-05
562.131 | 0.008 | 562.131 | 0.008 | 0.013 | 0.014 | 0.022 | 0.022 | 0.02 | 0.022 | 0.023 | 0.025 | 0.011667 | 0.021333 | 0.023333 | 0.0175 | 562.131 | 0.0038333 | 6.2401e-05
562.419 | 0.018 | 562.419 | 0.018 | 0.015 | 0.021 | 0.023 | 0.027 | 0.027 | 0.025 | 0.026 | 0.025 | 0.018 | 0.025667 | 0.025333 | 0.021667 | 562.419 | 0.004 | 6.5114e-05
562.706 | 0.009 | 562.706 | 0.009 | 0.012 | 0.011 | 0.017 | 0.017 | 0.021 | 0.023 | 0.024 | 0.024 | 0.010667 | 0.018333 | 0.023667 | 0.017167 | 562.706 | 0.0011667 | 1.8992e-05
562.994 | 0.003 | 562.994 | 0.003 | 0.006 | 0.01 | 0.012 | 0.015 | 0.017 | 0.017 | 0.019 | 0.017 | 0.0063333 | 0.014667 | 0.017667 | 0.012 | 562.994 | 0.0026667 | 4.3409e-05
563.281 | 0.017 | 563.281 | 0.017 | 0.017 | 0.021 | 0.025 | 0.027 | 0.028 | 0.029 | 0.03 | 0.031 | 0.018333 | 0.026667 | 0.03 | 0.024167 | 563.281 | 0.0025 | 4.0696e-05
563.568 | 0.011 | 563.568 | 0.011 | 0.012 | 0.016 | 0.021 | 0.025 | 0.022 | 0.021 | 0.025 | 0.023 | 0.013 | 0.022667 | 0.023 | 0.018 | 563.568 | 0.0046667 | 7.5967e-05
563.855 | 0.02 | 563.855 | 0.02 | 0.021 | 0.024 | 0.026 | 0.029 | 0.03 | 0.037 | 0.032 | 0.032 | 0.021667 | 0.028333 | 0.033667 | 0.027667 | 563.855 | 0.00066667 | 1.0852e-05
564.142 | 0.01 | 564.142 | 0.01 | 0.011 | 0.014 | 0.011 | 0.015 | 0.015 | 0.019 | 0.023 | 0.018 | 0.011667 | 0.013667 | 0.02 | 0.015833 | 564.142 | -0.0021667 | -3.527e-05
564.429 | 0.013 | 564.429 | 0.013 | 0.016 | 0.019 | 0.023 | 0.026 | 0.024 | 0.027 | 0.029 | 0.029 | 0.016 | 0.024333 | 0.028333 | 0.022167 | 564.429 | 0.0021667 | 3.527e-05
564.716 | 0.011 | 564.716 | 0.011 | 0.016 | 0.016 | 0.019 | 0.025 | 0.024 | 0.025 | 0.027 | 0.027 | 0.014333 | 0.022667 | 0.026333 | 0.020333 | 564.716 | 0.0023333 | 3.7983e-05
565.003 | 0.018 | 565.003 | 0.018 | 0.021 | 0.023 | 0.025 | 0.03 | 0.027 | 0.03 | 0.032 | 0.03 | 0.020667 | 0.027333 | 0.030667 | 0.025667 | 565.003 | 0.0016667 | 2.7131e-05
565.289 | 0.015 | 565.289 | 0.015 | 0.013 | 0.016 | 0.023 | 0.022 | 0.023 | 0.024 | 0.024 | 0.025 | 0.014667 | 0.022667 | 0.024333 | 0.0195 | 565.289 | 0.0031667 | 5.1549e-05
565.576 | 0.012 | 565.576 | 0.012 | 0.019 | 0.021 | 0.019 | 0.022 | 0.027 | 0.026 | 0.028 | 0.029 | 0.017333 | 0.022667 | 0.027667 | 0.0225 | 565.576 | 0.00016667 | 2.7131e-06
565.863 | 0.012 | 565.863 | 0.012 | 0.014 | 0.019 | 0.019 | 0.021 | 0.026 | 0.023 | 0.022 | 0.025 | 0.015 | 0.022 | 0.023333 | 0.019167 | 565.863 | 0.0028333 | 4.6123e-05
566.149 | 0.015 | 566.149 | 0.015 | 0.016 | 0.017 | 0.024 | 0.023 | 0.025 | 0.029 | 0.027 | 0.026 | 0.016 | 0.024 | 0.027333 | 0.021667 | 566.149 | 0.0023333 | 3.7983e-05
566.436 | 0.017 | 566.436 | 0.017 | 0.018 | 0.021 | 0.022 | 0.025 | 0.029 | 0.026 | 0.027 | 0.029 | 0.018667 | 0.025333 | 0.027333 | 0.023 | 566.436 | 0.0023333 | 3.7983e-05
566.722 | 0.007 | 566.722 | 0.007 | 0.01 | 0.009 | 0.011 | 0.017 | 0.015 | 0.022 | 0.02 | 0.02 | 0.0086667 | 0.014333 | 0.020667 | 0.014667 | 566.722 | -0.00033333 | -5.4262e-06
567.008 | 0.007 | 567.008 | 0.007 | 0.009 | 0.012 | 0.017 | 0.017 | 0.018 | 0.021 | 0.021 | 0.024 | 0.0093333 | 0.017333 | 0.022 | 0.015667 | 567.008 | 0.0016667 | 2.7131e-05
567.295 | 0.019 | 567.295 | 0.019 | 0.018 | 0.021 | 0.024 | 0.026 | 0.027 | 0.031 | 0.031 | 0.029 | 0.019333 | 0.025667 | 0.030333 | 0.024833 | 567.295 | 0.00083333 | 1.3565e-05
567.581 | 0.007 | 567.581 | 0.007 | 0.011 | 0.012 | 0.015 | 0.016 | 0.016 | 0.017 | 0.022 | 0.02 | 0.01 | 0.015667 | 0.019667 | 0.014833 | 567.581 | 0.00083333 | 1.3565e-05
567.867 | 0.011 | 567.867 | 0.011 | 0.009 | 0.012 | 0.015 | 0.019 | 0.018 | 0.021 | 0.022 | 0.025 | 0.010667 | 0.017333 | 0.022667 | 0.016667 | 567.867 | 0.00066667 | 1.0852e-05
568.153 | 0.004 | 568.153 | 0.004 | 0.007 | 0.006 | 0.01 | 0.014 | 0.012 | 0.016 | 0.018 | 0.022 | 0.0056667 | 0.012 | 0.018667 | 0.012167 | 568.153 | -0.00016667 | -2.7131e-06
568.439 | 0.013 | 568.439 | 0.013 | 0.016 | 0.016 | 0.017 | 0.021 | 0.023 | 0.024 | 0.027 | 0.022 | 0.015 | 0.020333 | 0.024333 | 0.019667 | 568.439 | 0.00066667 | 1.0852e-05
568.725 | 0.009 | 568.725 | 0.009 | 0.009 | 0.011 | 0.017 | 0.017 | 0.018 | 0.019 | 0.02 | 0.023 | 0.0096667 | 0.017333 | 0.020667 | 0.015167 | 568.725 | 0.0021667 | 3.527e-05
569.011 | 0.018 | 569.011 | 0.018 | 0.016 | 0.015 | 0.021 | 0.024 | 0.025 | 0.024 | 0.029 | 0.028 | 0.016333 | 0.023333 | 0.027 | 0.021667 | 569.011 | 0.0016667 | 2.7131e-05
569.297 | 0.012 | 569.297 | 0.012 | 0.013 | 0.012 | 0.02 | 0.019 | 0.021 | 0.022 | 0.026 | 0.025 | 0.012333 | 0.02 | 0.024333 | 0.018333 | 569.297 | 0.0016667 | 2.7131e-05
569.583 | 0.01 | 569.583 | 0.01 | 0.013 | 0.016 | 0.019 | 0.021 | 0.022 | 0.024 | 0.024 | 0.024 | 0.013 | 0.020667 | 0.024 | 0.0185 | 569.583 | 0.0021667 | 3.527e-05
569.868 | 0.011 | 569.868 | 0.011 | 0.01 | 0.015 | 0.014 | 0.017 | 0.021 | 0.021 | 0.023 | 0.025 | 0.012 | 0.017333 | 0.023 | 0.0175 | 569.868 | -0.00016667 | -2.7131e-06
570.154 | 0.008 | 570.154 | 0.008 | 0.011 | 0.006 | 0.02 | 0.02 | 0.019 | 0.02 | 0.02 | 0.023 | 0.0083333 | 0.019667 | 0.021 | 0.014667 | 570.154 | 0.005 | 8.1393e-05
570.439 | 0.005 | 570.439 | 0.005 | 0.007 | 0.008 | 0.011 | 0.012 | 0.014 | 0.015 | 0.015 | 0.018 | 0.0066667 | 0.012333 | 0.016 | 0.011333 | 570.439 | 0.001 | 1.6279e-05
570.725 | 0.012 | 570.725 | 0.012 | 0.014 | 0.013 | 0.021 | 0.02 | 0.021 | 0.025 | 0.024 | 0.024 | 0.013 | 0.020667 | 0.024333 | 0.018667 | 570.725 | 0.002 | 3.2557e-05
571.01 | 0.009 | 571.01 | 0.009 | 0.013 | 0.015 | 0.018 | 0.019 | 0.019 | 0.02 | 0.023 | 0.023 | 0.012333 | 0.018667 | 0.022 | 0.017167 | 571.01 | 0.0015 | 2.4418e-05
571.296 | 0.016 | 571.296 | 0.016 | 0.02 | 0.018 | 0.027 | 0.024 | 0.029 | 0.027 | 0.027 | 0.031 | 0.018 | 0.026667 | 0.028333 | 0.023167 | 571.296 | 0.0035 | 5.6975e-05
571.581 | 0.012 | 571.581 | 0.012 | 0.012 | 0.012 | 0.017 | 0.015 | 0.018 | 0.017 | 0.018 | 0.022 | 0.012 | 0.016667 | 0.019 | 0.0155 | 571.581 | 0.0011667 | 1.8992e-05
571.866 | 0.018 | 571.866 | 0.018 | 0.019 | 0.023 | 0.021 | 0.027 | 0.029 | 0.029 | 0.029 | 0.033 | 0.02 | 0.025667 | 0.030333 | 0.025167 | 571.866 | 0.0005 | 8.1393e-06
572.151 | 0.017 | 572.151 | 0.017 | 0.017 | 0.022 | 0.023 | 0.026 | 0.024 | 0.027 | 0.027 | 0.032 | 0.018667 | 0.024333 | 0.028667 | 0.023667 | 572.151 | 0.00066667 | 1.0852e-05
572.436 | 0.007 | 572.436 | 0.007 | 0.008 | 0.007 | 0.015 | 0.016 | 0.016 | 0.02 | 0.018 | 0.022 | 0.0073333 | 0.015667 | 0.02 | 0.013667 | 572.436 | 0.002 | 3.2557e-05
572.721 | 0.008 | 572.721 | 0.008 | 0.017 | 0.012 | 0.018 | 0.017 | 0.022 | 0.021 | 0.02 | 0.025 | 0.012333 | 0.019 | 0.022 | 0.017167 | 572.721 | 0.0018333 | 2.9844e-05
573.006 | 0.008 | 573.006 | 0.008 | 0.013 | 0.013 | 0.015 | 0.017 | 0.021 | 0.02 | 0.021 | 0.026 | 0.011333 | 0.017667 | 0.022333 | 0.016833 | 573.006 | 0.00083333 | 1.3565e-05
573.291 | 0.015 | 573.291 | 0.015 | 0.02 | 0.02 | 0.025 | 0.026 | 0.029 | 0.032 | 0.033 | 0.031 | 0.018333 | 0.026667 | 0.032 | 0.025167 | 573.291 | 0.0015 | 2.4418e-05
573.576 | 0.013 | 573.576 | 0.013 | 0.014 | 0.015 | 0.018 | 0.023 | 0.024 | 0.02 | 0.025 | 0.023 | 0.014 | 0.021667 | 0.022667 | 0.018333 | 573.576 | 0.0033333 | 5.4262e-05
573.861 | 0.016 | 573.861 | 0.016 | 0.018 | 0.018 | 0.024 | 0.025 | 0.027 | 0.028 | 0.029 | 0.03 | 0.017333 | 0.025333 | 0.029 | 0.023167 | 573.861 | 0.0021667 | 3.527e-05
574.146 | 0.006 | 574.146 | 0.006 | 0.007 | 0.011 | 0.012 | 0.012 | 0.016 | 0.018 | 0.021 | 0.022 | 0.008 | 0.013333 | 0.020333 | 0.014167 | 574.146 | -0.00083333 | -1.3565e-05
574.43 | 0.016 | 574.43 | 0.016 | 0.015 | 0.017 | 0.027 | 0.023 | 0.025 | 0.025 | 0.026 | 0.031 | 0.016 | 0.025 | 0.027333 | 0.021667 | 574.43 | 0.0033333 | 5.4262e-05
574.715 | 0.018 | 574.715 | 0.018 | 0.018 | 0.024 | 0.029 | 0.028 | 0.03 | 0.03 | 0.029 | 0.036 | 0.02 | 0.029 | 0.031667 | 0.025833 | 574.715 | 0.0031667 | 5.1549e-05
574.999 | 0.016 | 574.999 | 0.016 | 0.019 | 0.022 | 0.022 | 0.029 | 0.031 | 0.032 | 0.03 | 0.031 | 0.019 | 0.027333 | 0.031 | 0.025 | 574.999 | 0.0023333 | 3.7983e-05
575.284 | 0.016 | 575.284 | 0.016 | 0.025 | 0.021 | 0.029 | 0.03 | 0.032 | 0.03 | 0.028 | 0.033 | 0.020667 | 0.030333 | 0.030333 | 0.0255 | 575.284 | 0.0048333 | 7.868e-05
575.568 | 0.009 | 575.568 | 0.009 | 0.009 | 0.008 | 0.015 | 0.017 | 0.022 | 0.021 | 0.021 | 0.022 | 0.0086667 | 0.018 | 0.021333 | 0.015 | 575.568 | 0.003 | 4.8836e-05
575.852 | 0.011 | 575.852 | 0.011 | 0.013 | 0.012 | 0.025 | 0.023 | 0.024 | 0.021 | 0.025 | 0.025 | 0.012 | 0.024 | 0.023667 | 0.017833 | 575.852 | 0.0061667 | 0.00010038
576.137 | 0.019 | 576.137 | 0.019 | 0.021 | 0.024 | 0.052 | 0.054 | 0.054 | 0.031 | 0.028 | 0.029 | 0.021333 | 0.053333 | 0.029333 | 0.025333 | 576.137 | 0.028 | 0.0004558
576.421 | 0.023 | 576.421 | 0.023 | 0.026 | 0.032 | 0.13 | 0.127 | 0.13 | 0.033 | 0.035 | 0.038 | 0.027 | 0.129 | 0.035333 | 0.031167 | 576.421 | 0.097833 | 0.0015926
576.705 | 0.005 | 576.705 | 0.005 | 0.01 | 0.01 | 0.159 | 0.162 | 0.168 | 0.021 | 0.017 | 0.022 | 0.0083333 | 0.163 | 0.02 | 0.014167 | 576.705 | 0.14883 | 0.0024228
576.989 | 0.012 | 576.989 | 0.012 | 0.016 | 0.016 | 0.249 | 0.247 | 0.249 | 0.025 | 0.026 | 0.028 | 0.014667 | 0.24833 | 0.026333 | 0.0205 | 576.989 | 0.22783 | 0.0037088
577.273 | 0.001 | 577.273 | 0.001 | 0.003 | 0.003 | 0.208 | 0.205 | 0.207 | 0.012 | 0.014 | 0.015 | 0.0023333 | 0.20667 | 0.013667 | 0.008 | 577.273 | 0.19867 | 0.003234
577.557 | 0.008 | 577.557 | 0.008 | 0.011 | 0.015 | 0.06 | 0.063 | 0.063 | 0.026 | 0.024 | 0.027 | 0.011333 | 0.062 | 0.025667 | 0.0185 | 577.557 | 0.0435 | 0.00070812
577.841 | 0.008 | 577.841 | 0.008 | 0.011 | 0.013 | 0.027 | 0.03 | 0.032 | 0.022 | 0.022 | 0.026 | 0.010667 | 0.029667 | 0.023333 | 0.017 | 577.841 | 0.012667 | 0.0002062
578.124 | 0.018 | 578.124 | 0.018 | 0.02 | 0.023 | 0.041 | 0.039 | 0.046 | 0.031 | 0.032 | 0.032 | 0.020333 | 0.042 | 0.031667 | 0.026 | 578.124 | 0.016 | 0.00026046
578.408 | 0.015 | 578.408 | 0.015 | 0.02 | 0.02 | 0.089 | 0.09 | 0.09 | 0.026 | 0.033 | 0.034 | 0.018333 | 0.089667 | 0.031 | 0.024667 | 578.408 | 0.065 | 0.0010581
578.692 | -0.001 | 578.692 | -0.001 | 0.002 | 0.002 | 0.136 | 0.138 | 0.135 | 0.018 | 0.016 | 0.018 | 0.001 | 0.13633 | 0.017333 | 0.0091667 | 578.692 | 0.12717 | 0.0020701
578.975 | 0.018 | 578.975 | 0.018 | 0.015 | 0.018 | 0.205 | 0.197 | 0.199 | 0.025 | 0.026 | 0.03 | 0.017 | 0.20033 | 0.027 | 0.022 | 578.975 | 0.17833 | 0.002903
579.259 | 0.004 | 579.259 | 0.004 | 0.01 | 0.011 | 0.246 | 0.244 | 0.245 | 0.018 | 0.019 | 0.022 | 0.0083333 | 0.245 | 0.019667 | 0.014 | 579.259 | 0.231 | 0.0037603
579.542 | 0.004 | 579.542 | 0.004 | 0.007 | 0.012 | 0.098 | 0.093 | 0.1 | 0.019 | 0.022 | 0.026 | 0.0076667 | 0.097 | 0.022333 | 0.015 | 579.542 | 0.082 | 0.0013348
579.826 | 0.004 | 579.826 | 0.004 | 0.007 | 0.008 | 0.026 | 0.03 | 0.028 | 0.014 | 0.023 | 0.02 | 0.0063333 | 0.028 | 0.019 | 0.012667 | 579.826 | 0.015333 | 0.0002496
580.109 | 0.011 | 580.109 | 0.011 | 0.015 | 0.014 | 0.027 | 0.025 | 0.023 | 0.025 | 0.027 | 0.026 | 0.013333 | 0.025 | 0.026 | 0.019667 | 580.109 | 0.0053333 | 8.6819e-05
580.392 | 0.017 | 580.392 | 0.017 | 0.019 | 0.02 | 0.024 | 0.029 | 0.028 | 0.029 | 0.031 | 0.03 | 0.018667 | 0.027 | 0.03 | 0.024333 | 580.392 | 0.0026667 | 4.3409e-05
580.676 | 0.009 | 580.676 | 0.009 | 0.007 | 0.007 | 0.018 | 0.019 | 0.019 | 0.016 | 0.017 | 0.021 | 0.0076667 | 0.018667 | 0.018 | 0.012833 | 580.676 | 0.0058333 | 9.4958e-05
580.959 | 0.009 | 580.959 | 0.009 | 0.014 | 0.015 | 0.02 | 0.025 | 0.023 | 0.021 | 0.028 | 0.027 | 0.012667 | 0.022667 | 0.025333 | 0.019 | 580.959 | 0.0036667 | 5.9688e-05
581.242 | 0.01 | 581.242 | 0.01 | 0.013 | 0.015 | 0.02 | 0.028 | 0.028 | 0.019 | 0.025 | 0.027 | 0.012667 | 0.025333 | 0.023667 | 0.018167 | 581.242 | 0.0071667 | 0.00011666
581.525 | 0.013 | 581.525 | 0.013 | 0.012 | 0.01 | 0.017 | 0.021 | 0.024 | 0.021 | 0.026 | 0.027 | 0.011667 | 0.020667 | 0.024667 | 0.018167 | 581.525 | 0.0025 | 4.0696e-05
581.808 | 0.009 | 581.808 | 0.009 | 0.014 | 0.013 | 0.02 | 0.021 | 0.023 | 0.027 | 0.026 | 0.022 | 0.012 | 0.021333 | 0.025 | 0.0185 | 581.808 | 0.0028333 | 4.6123e-05
582.091 | 0.005 | 582.091 | 0.005 | 0.01 | 0.011 | 0.017 | 0.018 | 0.018 | 0.022 | 0.02 | 0.025 | 0.0086667 | 0.017667 | 0.022333 | 0.0155 | 582.091 | 0.0021667 | 3.527e-05
582.373 | 0.003 | 582.373 | 0.003 | 0.007 | 0.008 | 0.012 | 0.016 | 0.013 | 0.016 | 0.019 | 0.021 | 0.006 | 0.013667 | 0.018667 | 0.012333 | 582.373 | 0.0013333 | 2.1705e-05
582.656 | 0.012 | 582.656 | 0.012 | 0.017 | 0.015 | 0.021 | 0.026 | 0.024 | 0.025 | 0.031 | 0.028 | 0.014667 | 0.023667 | 0.028 | 0.021333 | 582.656 | 0.0023333 | 3.7983e-05
582.939 | 0.008 | 582.939 | 0.008 | 0.014 | 0.011 | 0.019 | 0.021 | 0.02 | 0.022 | 0.026 | 0.025 | 0.011 | 0.02 | 0.024333 | 0.017667 | 582.939 | 0.0023333 | 3.7983e-05
583.221 | -0.001 | 583.221 | -0.001 | 0.004 | 0.005 | 0.01 | 0.01 | 0.015 | 0.015 | 0.01 | 0.011 | 0.0026667 | 0.011667 | 0.012 | 0.0073333 | 583.221 | 0.0043333 | 7.054e-05
583.504 | 0.008 | 583.504 | 0.008 | 0.015 | 0.012 | 0.016 | 0.021 | 0.024 | 0.026 | 0.028 | 0.026 | 0.011667 | 0.020333 | 0.026667 | 0.019167 | 583.504 | 0.0011667 | 1.8992e-05
583.786 | 0.009 | 583.786 | 0.009 | 0.012 | 0.016 | 0.019 | 0.02 | 0.023 | 0.024 | 0.023 | 0.024 | 0.012333 | 0.020667 | 0.023667 | 0.018 | 583.786 | 0.0026667 | 4.3409e-05
584.069 | 0.006 | 584.069 | 0.006 | 0.005 | 0.008 | 0.013 | 0.014 | 0.017 | 0.018 | 0.02 | 0.024 | 0.0063333 | 0.014667 | 0.020667 | 0.0135 | 584.069 | 0.0011667 | 1.8992e-05
584.351 | 0.006 | 584.351 | 0.006 | 0.01 | 0.013 | 0.015 | 0.02 | 0.021 | 0.017 | 0.024 | 0.022 | 0.0096667 | 0.018667 | 0.021 | 0.015333 | 584.351 | 0.0033333 | 5.4262e-05
584.633 | 0.007 | 584.633 | 0.007 | 0.004 | 0.007 | 0.015 | 0.013 | 0.016 | 0.017 | 0.018 | 0.016 | 0.006 | 0.014667 | 0.017 | 0.0115 | 584.633 | 0.0031667 | 5.1549e-05
584.916 | -0.002 | 584.916 | -0.002 | 0.001 | 0 | 0.008 | 0.012 | 0.014 | 0.007 | 0.013 | 0.013 | -0.00033333 | 0.011333 | 0.011 | 0.0053333 | 584.916 | 0.006 | 9.7671e-05
585.198 | 0.007 | 585.198 | 0.007 | 0.014 | 0.013 | 0.018 | 0.021 | 0.02 | 0.022 | 0.023 | 0.027 | 0.011333 | 0.019667 | 0.024 | 0.017667 | 585.198 | 0.002 | 3.2557e-05
585.48 | 0.01 | 585.48 | 0.01 | 0.013 | 0.009 | 0.019 | 0.021 | 0.024 | 0.024 | 0.022 | 0.021 | 0.010667 | 0.021333 | 0.022333 | 0.0165 | 585.48 | 0.0048333 | 7.868e-05
585.762 | -0.004 | 585.762 | -0.004 | -0.003 | 0 | 0.006 | 0.005 | 0.012 | 0.006 | 0.006 | 0.012 | -0.0023333 | 0.0076667 | 0.008 | 0.0028333 | 585.762 | 0.0048333 | 7.868e-05
586.044 | 0.002 | 586.044 | 0.002 | 0.004 | 0.004 | 0.014 | 0.014 | 0.017 | 0.015 | 0.017 | 0.019 | 0.0033333 | 0.015 | 0.017 | 0.010167 | 586.044 | 0.0048333 | 7.868e-05
586.326 | 0.015 | 586.326 | 0.015 | 0.015 | 0.017 | 0.018 | 0.022 | 0.025 | 0.027 | 0.027 | 0.027 | 0.015667 | 0.021667 | 0.027 | 0.021333 | 586.326 | 0.00033333 | 5.4262e-06
586.607 | 0.01 | 586.607 | 0.01 | 0.008 | 0.01 | 0.016 | 0.021 | 0.019 | 0.022 | 0.021 | 0.024 | 0.0093333 | 0.018667 | 0.022333 | 0.015833 | 586.607 | 0.0028333 | 4.6123e-05
586.889 | 0.024 | 586.889 | 0.024 | 0.023 | 0.026 | 0.035 | 0.036 | 0.036 | 0.037 | 0.036 | 0.041 | 0.024333 | 0.035667 | 0.038 | 0.031167 | 586.889 | 0.0045 | 7.3254e-05
587.171 | 0.017 | 587.171 | 0.017 | 0.016 | 0.019 | 0.021 | 0.025 | 0.028 | 0.028 | 0.035 | 0.03 | 0.017333 | 0.024667 | 0.031 | 0.024167 | 587.171 | 0.0005 | 8.1393e-06
587.453 | 0.003 | 587.453 | 0.003 | 0.007 | 0.005 | 0.013 | 0.015 | 0.018 | 0.018 | 0.022 | 0.017 | 0.005 | 0.015333 | 0.019 | 0.012 | 587.453 | 0.0033333 | 5.4262e-05
587.734 | 0.011 | 587.734 | 0.011 | 0.014 | 0.013 | 0.021 | 0.022 | 0.025 | 0.02 | 0.028 | 0.026 | 0.012667 | 0.022667 | 0.024667 | 0.018667 | 587.734 | 0.004 | 6.5114e-05
588.016 | 0.014 | 588.016 | 0.014 | 0.016 | 0.014 | 0.023 | 0.028 | 0.027 | 0.024 | 0.028 | 0.028 | 0.014667 | 0.026 | 0.026667 | 0.020667 | 588.016 | 0.0053333 | 8.6819e-05
588.297 | 0 | 588.297 | 0 | 0.003 | 0.005 | 0.011 | 0.011 | 0.013 | 0.009 | 0.014 | 0.014 | 0.0026667 | 0.011667 | 0.012333 | 0.0075 | 588.297 | 0.0041667 | 6.7827e-05
588.578 | 0.015 | 588.578 | 0.015 | 0.012 | 0.02 | 0.024 | 0.023 | 0.021 | 0.03 | 0.029 | 0.027 | 0.015667 | 0.022667 | 0.028667 | 0.022167 | 588.578 | 0.0005 | 8.1393e-06
588.86 | 0.013 | 588.86 | 0.013 | 0.014 | 0.017 | 0.024 | 0.024 | 0.025 | 0.027 | 0.027 | 0.029 | 0.014667 | 0.024333 | 0.027667 | 0.021167 | 588.86 | 0.0031667 | 5.1549e-05
589.141 | 0.005 | 589.141 | 0.005 | 0.01 | 0.009 | 0.014 | 0.018 | 0.016 | 0.023 | 0.019 | 0.024 | 0.008 | 0.016 | 0.022 | 0.015 | 589.141 | 0.001 | 1.6279e-05
589.422 | 0.011 | 589.422 | 0.011 | 0.013 | 0.017 | 0.017 | 0.022 | 0.024 | 0.021 | 0.025 | 0.024 | 0.013667 | 0.021 | 0.023333 | 0.0185 | 589.422 | 0.0025 | 4.0696e-05
589.703 | 0.02 | 589.703 | 0.02 | 0.02 | 0.019 | 0.029 | 0.026 | 0.029 | 0.033 | 0.04 | 0.033 | 0.019667 | 0.028 | 0.035333 | 0.0275 | 589.703 | 0.0005 | 8.1393e-06
589.984 | 0.002 | 589.984 | 0.002 | 0.005 | 0.004 | 0.01 | 0.011 | 0.013 | 0.015 | 0.018 | 0.013 | 0.0036667 | 0.011333 | 0.015333 | 0.0095 | 589.984 | 0.0018333 | 2.9844e-05
590.265 | 0.01 | 590.265 | 0.01 | 0.009 | 0.01 | 0.02 | 0.02 | 0.015 | 0.02 | 0.019 | 0.022 | 0.0096667 | 0.018333 | 0.020333 | 0.015 | 590.265 | 0.0033333 | 5.4262e-05
590.546 | 0.01 | 590.546 | 0.01 | 0.012 | 0.017 | 0.02 | 0.021 | 0.021 | 0.023 | 0.025 | 0.028 | 0.013 | 0.020667 | 0.025333 | 0.019167 | 590.546 | 0.0015 | 2.4418e-05
590.827 | 0.007 | 590.827 | 0.007 | 0.004 | 0.006 | 0.012 | 0.012 | 0.012 | 0.016 | 0.017 | 0.017 | 0.0056667 | 0.012 | 0.016667 | 0.011167 | 590.827 | 0.00083333 | 1.3565e-05
591.108 | 0.004 | 591.108 | 0.004 | 0.007 | 0.008 | 0.012 | 0.011 | 0.013 | 0.019 | 0.015 | 0.02 | 0.0063333 | 0.012 | 0.018 | 0.012167 | 591.108 | -0.00016667 | -2.7131e-06
591.389 | 0.014 | 591.389 | 0.014 | 0.018 | 0.021 | 0.026 | 0.024 | 0.026 | 0.028 | 0.028 | 0.026 | 0.017667 | 0.025333 | 0.027333 | 0.0225 | 591.389 | 0.0028333 | 4.6123e-05
591.669 | 0.004 | 591.669 | 0.004 | 0.005 | 0.005 | 0.012 | 0.013 | 0.016 | 0.017 | 0.021 | 0.02 | 0.0046667 | 0.013667 | 0.019333 | 0.012 | 591.669 | 0.0016667 | 2.7131e-05
591.95 | 0.011 | 591.95 | 0.011 | 0.014 | 0.009 | 0.018 | 0.025 | 0.024 | 0.024 | 0.025 | 0.027 | 0.011333 | 0.022333 | 0.025333 | 0.018333 | 591.95 | 0.004 | 6.5114e-05
592.23 | 0.012 | 592.23 | 0.012 | 0.01 | 0.016 | 0.018 | 0.023 | 0.027 | 0.025 | 0.027 | 0.028 | 0.012667 | 0.022667 | 0.026667 | 0.019667 | 592.23 | 0.003 | 4.8836e-05
592.511 | 0.019 | 592.511 | 0.019 | 0.02 | 0.018 | 0.027 | 0.031 | 0.03 | 0.034 | 0.033 | 0.033 | 0.019 | 0.029333 | 0.033333 | 0.026167 | 592.511 | 0.0031667 | 5.1549e-05
592.791 | 0.011 | 592.791 | 0.011 | 0.013 | 0.018 | 0.021 | 0.02 | 0.022 | 0.024 | 0.024 | 0.027 | 0.014 | 0.021 | 0.025 | 0.0195 | 592.791 | 0.0015 | 2.4418e-05
593.071 | 0.012 | 593.071 | 0.012 | 0.015 | 0.02 | 0.022 | 0.022 | 0.03 | 0.03 | 0.032 | 0.028 | 0.015667 | 0.024667 | 0.03 | 0.022833 | 593.071 | 0.0018333 | 2.9844e-05
593.352 | -0.003 | 593.352 | -0.003 | 0.001 | -0.004 | 0.009 | 0.011 | 0.012 | 0.012 | 0.018 | 0.011 | -0.002 | 0.010667 | 0.013667 | 0.0058333 | 593.352 | 0.0048333 | 7.868e-05
593.632 | 0.015 | 593.632 | 0.015 | 0.016 | 0.016 | 0.021 | 0.019 | 0.026 | 0.028 | 0.025 | 0.028 | 0.015667 | 0.022 | 0.027 | 0.021333 | 593.632 | 0.00066667 | 1.0852e-05
593.912 | 0.015 | 593.912 | 0.015 | 0.016 | 0.015 | 0.024 | 0.021 | 0.024 | 0.028 | 0.028 | 0.03 | 0.015333 | 0.023 | 0.028667 | 0.022 | 593.912 | 0.001 | 1.6279e-05
594.192 | 0.021 | 594.192 | 0.021 | 0.022 | 0.023 | 0.025 | 0.03 | 0.031 | 0.036 | 0.036 | 0.034 | 0.022 | 0.028667 | 0.035333 | 0.028667 | 594.192 | 0 | 0
594.472 | 0.008 | 594.472 | 0.008 | 0.012 | 0.015 | 0.019 | 0.023 | 0.026 | 0.027 | 0.028 | 0.027 | 0.011667 | 0.022667 | 0.027333 | 0.0195 | 594.472 | 0.0031667 | 5.1549e-05
594.752 | 0.017 | 594.752 | 0.017 | 0.018 | 0.016 | 0.022 | 0.026 | 0.025 | 0.024 | 0.028 | 0.03 | 0.017 | 0.024333 | 0.027333 | 0.022167 | 594.752 | 0.0021667 | 3.527e-05
595.032 | 0.01 | 595.032 | 0.01 | 0.013 | 0.013 | 0.021 | 0.022 | 0.023 | 0.023 | 0.022 | 0.027 | 0.012 | 0.022 | 0.024 | 0.018 | 595.032 | 0.004 | 6.5114e-05
595.312 | 0.001 | 595.312 | 0.001 | 0 | 0.003 | 0.011 | 0.003 | 0.014 | 0.013 | 0.017 | 0.015 | 0.0013333 | 0.0093333 | 0.015 | 0.0081667 | 595.312 | 0.0011667 | 1.8992e-05
595.592 | 0.011 | 595.592 | 0.011 | 0.014 | 0.013 | 0.019 | 0.019 | 0.025 | 0.027 | 0.025 | 0.026 | 0.012667 | 0.021 | 0.026 | 0.019333 | 595.592 | 0.0016667 | 2.7131e-05
595.871 | 0.01 | 595.871 | 0.01 | 0.012 | 0.014 | 0.018 | 0.02 | 0.023 | 0.024 | 0.022 | 0.02 | 0.012 | 0.020333 | 0.022 | 0.017 | 595.871 | 0.0033333 | 5.4262e-05
596.151 | 0.006 | 596.151 | 0.006 | 0.012 | 0.012 | 0.016 | 0.019 | 0.022 | 0.022 | 0.024 | 0.024 | 0.01 | 0.019 | 0.023333 | 0.016667 | 596.151 | 0.0023333 | 3.7983e-05
596.43 | -0.003 | 596.43 | -0.003 | -0.001 | 0.002 | 0.003 | 0.008 | 0.004 | 0.008 | 0.011 | 0.012 | -0.00066667 | 0.005 | 0.010333 | 0.0048333 | 596.43 | 0.00016667 | 2.7131e-06
596.71 | 0.005 | 596.71 | 0.005 | 0.007 | 0.01 | 0.015 | 0.017 | 0.014 | 0.016 | 0.023 | 0.023 | 0.0073333 | 0.015333 | 0.020667 | 0.014 | 596.71 | 0.0013333 | 2.1705e-05
596.989 | 0.024 | 596.989 | 0.024 | 0.023 | 0.025 | 0.027 | 0.026 | 0.034 | 0.03 | 0.037 | 0.038 | 0.024 | 0.029 | 0.035 | 0.0295 | 596.989 | -0.0005 | -8.1393e-06
597.269 | 0.003 | 597.269 | 0.003 | 0.009 | 0.015 | 0.017 | 0.02 | 0.021 | 0.026 | 0.021 | 0.023 | 0.009 | 0.019333 | 0.023333 | 0.016167 | 597.269 | 0.0031667 | 5.1549e-05
597.548 | 0.006 | 597.548 | 0.006 | 0.007 | 0.012 | 0.017 | 0.019 | 0.024 | 0.022 | 0.022 | 0.029 | 0.0083333 | 0.02 | 0.024333 | 0.016333 | 597.548 | 0.0036667 | 5.9688e-05
597.827 | 0.018 | 597.827 | 0.018 | 0.024 | 0.026 | 0.03 | 0.032 | 0.031 | 0.033 | 0.036 | 0.034 | 0.022667 | 0.031 | 0.034333 | 0.0285 | 597.827 | 0.0025 | 4.0696e-05
598.106 | 0.011 | 598.106 | 0.011 | 0.016 | 0.019 | 0.028 | 0.024 | 0.026 | 0.03 | 0.033 | 0.026 | 0.015333 | 0.026 | 0.029667 | 0.0225 | 598.106 | 0.0035 | 5.6975e-05
598.385 | 0.012 | 598.385 | 0.012 | 0.015 | 0.013 | 0.019 | 0.017 | 0.021 | 0.028 | 0.029 | 0.028 | 0.013333 | 0.019 | 0.028333 | 0.020833 | 598.385 | -0.0018333 | -2.9844e-05
598.665 | -0.007 | 598.665 | -0.007 | -0.006 | 0.001 | 0.002 | 0.002 | 0.007 | 0.002 | 0.009 | 0.008 | -0.004 | 0.0036667 | 0.0063333 | 0.0011667 | 598.665 | 0.0025 | 4.0696e-05
598.944 | 0 | 598.944 | 0 | 0 | 0.006 | 0.008 | 0.008 | 0.01 | 0.013 | 0.01 | 0.017 | 0.002 | 0.0086667 | 0.013333 | 0.0076667 | 598.944 | 0.001 | 1.6279e-05
599.222 | 0.01 | 599.222 | 0.01 | 0.01 | 0.012 | 0.011 | 0.016 | 0.017 | 0.017 | 0.02 | 0.022 | 0.010667 | 0.014667 | 0.019667 | 0.015167 | 599.222 | -0.0005 | -8.1393e-06
599.501 | 0.007 | 599.501 | 0.007 | 0.007 | 0.007 | 0.015 | 0.02 | 0.021 | 0.021 | 0.025 | 0.026 | 0.007 | 0.018667 | 0.024 | 0.0155 | 599.501 | 0.0031667 | 5.1549e-05
599.78 | 0.013 | 599.78 | 0.013 | 0.016 | 0.013 | 0.021 | 0.018 | 0.021 | 0.023 | 0.029 | 0.026 | 0.014 | 0.02 | 0.026 | 0.02 | 599.78 | 0 | 0
600.059 | 0.009 | 600.059 | 0.009 | 0.015 | 0.014 | 0.016 | 0.021 | 0.022 | 0.025 | 0.026 | 0.025 | 0.012667 | 0.019667 | 0.025333 | 0.019 | 600.059 | 0.00066667 | 1.0852e-05
600.337 | 0.006 | 600.337 | 0.006 | 0.008 | 0.007 | 0.014 | 0.01 | 0.02 | 0.022 | 0.022 | 0.025 | 0.007 | 0.014667 | 0.023 | 0.015 | 600.337 | -0.00033333 | -5.4262e-06
600.616 | 0.009 | 600.616 | 0.009 | 0.008 | 0.013 | 0.013 | 0.013 | 0.016 | 0.016 | 0.022 | 0.022 | 0.01 | 0.014 | 0.02 | 0.015 | 600.616 | -0.001 | -1.6279e-05
600.895 | 0.012 | 600.895 | 0.012 | 0.014 | 0.012 | 0.021 | 0.023 | 0.021 | 0.028 | 0.024 | 0.026 | 0.012667 | 0.021667 | 0.026 | 0.019333 | 600.895 | 0.0023333 | 3.7983e-05
601.173 | 0.022 | 601.173 | 0.022 | 0.027 | 0.03 | 0.042 | 0.044 | 0.045 | 0.051 | 0.052 | 0.05 | 0.026333 | 0.043667 | 0.051 | 0.038667 | 601.173 | 0.005 | 8.1393e-05
601.451 | 0.005 | 601.451 | 0.005 | 0.01 | 0.013 | 0.016 | 0.019 | 0.018 | 0.023 | 0.023 | 0.023 | 0.0093333 | 0.017667 | 0.023 | 0.016167 | 601.451 | 0.0015 | 2.4418e-05
601.73 | 0.008 | 601.73 | 0.008 | 0.011 | 0.007 | 0.014 | 0.013 | 0.019 | 0.015 | 0.024 | 0.02 | 0.0086667 | 0.015333 | 0.019667 | 0.014167 | 601.73 | 0.0011667 | 1.8992e-05
602.008 | 0.015 | 602.008 | 0.015 | 0.013 | 0.016 | 0.021 | 0.02 | 0.029 | 0.029 | 0.03 | 0.03 | 0.014667 | 0.023333 | 0.029667 | 0.022167 | 602.008 | 0.0011667 | 1.8992e-05
602.286 | 0 | 602.286 | 0 | 0.001 | 0.001 | 0.006 | 0.008 | 0.006 | 0.012 | 0.014 | 0.007 | 0.00066667 | 0.0066667 | 0.011 | 0.0058333 | 602.286 | 0.00083333 | 1.3565e-05
602.564 | 0.011 | 602.564 | 0.011 | 0.016 | 0.016 | 0.022 | 0.024 | 0.023 | 0.026 | 0.028 | 0.03 | 0.014333 | 0.023 | 0.028 | 0.021167 | 602.564 | 0.0018333 | 2.9844e-05
602.842 | 0.004 | 602.842 | 0.004 | 0.008 | 0.008 | 0.013 | 0.015 | 0.02 | 0.015 | 0.023 | 0.022 | 0.0066667 | 0.016 | 0.02 | 0.013333 | 602.842 | 0.0026667 | 4.3409e-05
603.12 | 0.003 | 603.12 | 0.003 | 0.006 | 0.005 | 0.007 | 0.011 | 0.015 | 0.014 | 0.016 | 0.013 | 0.0046667 | 0.011 | 0.014333 | 0.0095 | 603.12 | 0.0015 | 2.4418e-05
603.398 | 0.004 | 603.398 | 0.004 | 0.009 | 0.01 | 0.017 | 0.019 | 0.016 | 0.018 | 0.021 | 0.022 | 0.0076667 | 0.017333 | 0.020333 | 0.014 | 603.398 | 0.0033333 | 5.4262e-05
603.676 | 0.008 | 603.676 | 0.008 | 0.011 | 0.011 | 0.013 | 0.015 | 0.023 | 0.023 | 0.023 | 0.026 | 0.01 | 0.017 | 0.024 | 0.017 | 603.676 | 0 | 0
603.954 | 0.002 | 603.954 | 0.002 | 0.01 | 0.008 | 0.011 | 0.013 | 0.015 | 0.019 | 0.018 | 0.019 | 0.0066667 | 0.013 | 0.018667 | 0.012667 | 603.954 | 0.00033333 | 5.4262e-06
604.232 | 0.01 | 604.232 | 0.01 | 0.01 | 0.011 | 0.016 | 0.019 | 0.02 | 0.02 | 0.024 | 0.019 | 0.010333 | 0.018333 | 0.021 | 0.015667 | 604.232 | 0.0026667 | 4.3409e-05
604.51 | 0.003 | 604.51 | 0.003 | 0.005 | 0.004 | 0.011 | 0.013 | 0.014 | 0.013 | 0.02 | 0.018 | 0.004 | 0.012667 | 0.017 | 0.0105 | 604.51 | 0.0021667 | 3.527e-05
604.787 | -0.004 | 604.787 | -0.004 | 0.003 | 0.001 | 0.006 | 0.007 | 0.009 | 0.01 | 0.014 | 0.014 | 0 | 0.0073333 | 0.012667 | 0.0063333 | 604.787 | 0.001 | 1.6279e-05
605.065 | 0.01 | 605.065 | 0.01 | 0.014 | 0.018 | 0.013 | 0.017 | 0.022 | 0.022 | 0.024 | 0.023 | 0.014 | 0.017333 | 0.023 | 0.0185 | 605.065 | -0.0011667 | -1.8992e-05
605.342 | 0.002 | 605.342 | 0.002 | 0.005 | 0.008 | 0.011 | 0.014 | 0.013 | 0.015 | 0.019 | 0.02 | 0.005 | 0.012667 | 0.018 | 0.0115 | 605.342 | 0.0011667 | 1.8992e-05
605.62 | -0.005 | 605.62 | -0.005 | -0.005 | 0.003 | 0.013 | 0.01 | 0.011 | 0.011 | 0.012 | 0.012 | -0.0023333 | 0.011333 | 0.011667 | 0.0046667 | 605.62 | 0.0066667 | 0.00010852
605.897 | 0.007 | 605.897 | 0.007 | 0.014 | 0.009 | 0.014 | 0.019 | 0.019 | 0.019 | 0.019 | 0.021 | 0.01 | 0.017333 | 0.019667 | 0.014833 | 605.897 | 0.0025 | 4.0696e-05
606.174 | 0.017 | 606.174 | 0.017 | 0.015 | 0.018 | 0.02 | 0.027 | 0.024 | 0.031 | 0.033 | 0.029 | 0.016667 | 0.023667 | 0.031 | 0.023833 | 606.174 | -0.00016667 | -2.7131e-06
606.452 | 0.005 | 606.452 | 0.005 | 0.007 | 0.012 | 0.01 | 0.014 | 0.014 | 0.015 | 0.017 | 0.018 | 0.008 | 0.012667 | 0.016667 | 0.012333 | 606.452 | 0.00033333 | 5.4262e-06
606.729 | 0.015 | 606.729 | 0.015 | 0.019 | 0.017 | 0.024 | 0.027 | 0.026 | 0.03 | 0.032 | 0.033 | 0.017 | 0.025667 | 0.031667 | 0.024333 | 606.729 | 0.0013333 | 2.1705e-05
607.006 | 0.013 | 607.006 | 0.013 | 0.02 | 0.02 | 0.022 | 0.026 | 0.026 | 0.024 | 0.027 | 0.03 | 0.017667 | 0.024667 | 0.027 | 0.022333 | 607.006 | 0.0023333 | 3.7983e-05
607.283 | 0.013 | 607.283 | 0.013 | 0.014 | 0.014 | 0.016 | 0.019 | 0.022 | 0.019 | 0.023 | 0.025 | 0.013667 | 0.019 | 0.022333 | 0.018 | 607.283 | 0.001 | 1.6279e-05
607.56 | 0 | 607.56 | 0 | 0.005 | 0.009 | 0.012 | 0.013 | 0.012 | 0.015 | 0.018 | 0.019 | 0.0046667 | 0.012333 | 0.017333 | 0.011 | 607.56 | 0.0013333 | 2.1705e-05
607.837 | 0 | 607.837 | 0 | 0.005 | 0.003 | 0.005 | 0.006 | 0.008 | 0.011 | 0.011 | 0.014 | 0.0026667 | 0.0063333 | 0.012 | 0.0073333 | 607.837 | -0.001 | -1.6279e-05
608.114 | 0.009 | 608.114 | 0.009 | 0.009 | 0.013 | 0.016 | 0.019 | 0.021 | 0.019 | 0.023 | 0.022 | 0.010333 | 0.018667 | 0.021333 | 0.015833 | 608.114 | 0.0028333 | 4.6123e-05
608.391 | 0.03 | 608.391 | 0.03 | 0.032 | 0.036 | 0.041 | 0.042 | 0.046 | 0.045 | 0.055 | 0.048 | 0.032667 | 0.043 | 0.049333 | 0.041 | 608.391 | 0.002 | 3.2557e-05
608.667 | -0.01 | 608.667 | -0.01 | -0.009 | -0.002 | 0.004 | 0.002 | 0.005 | 0.006 | 0.016 | 0.01 | -0.007 | 0.0036667 | 0.010667 | 0.0018333 | 608.667 | 0.0018333 | 2.9844e-05
608.944 | 0.007 | 608.944 | 0.007 | 0.01 | 0.006 | 0.013 | 0.017 | 0.02 | 0.024 | 0.028 | 0.025 | 0.0076667 | 0.016667 | 0.025667 | 0.016667 | 608.944 | 0 | 0
609.221 | -0.002 | 609.221 | -0.002 | 0.003 | 0.007 | 0.009 | 0.008 | 0.012 | 0.012 | 0.015 | 0.014 | 0.0026667 | 0.0096667 | 0.013667 | 0.0081667 | 609.221 | 0.0015 | 2.4418e-05
609.497 | 0.005 | 609.497 | 0.005 | 0.008 | 0.007 | 0.011 | 0.011 | 0.012 | 0.019 | 0.022 | 0.019 | 0.0066667 | 0.011333 | 0.02 | 0.013333 | 609.497 | -0.002 | -3.2557e-05
609.774 | -0.005 | 609.774 | -0.005 | 0 | 0.003 | 0.006 | 0.007 | 0.006 | 0.012 | 0.011 | 0.013 | -0.00066667 | 0.0063333 | 0.012 | 0.0056667 | 609.774 | 0.00066667 | 1.0852e-05
610.05 | 0.019 | 610.05 | 0.019 | 0.019 | 0.024 | 0.022 | 0.021 | 0.024 | 0.03 | 0.031 | 0.028 | 0.020667 | 0.022333 | 0.029667 | 0.025167 | 610.05 | -0.0028333 | -4.6123e-05
610.326 | 0.01 | 610.326 | 0.01 | 0.012 | 0.016 | 0.019 | 0.02 | 0.02 | 0.023 | 0.027 | 0.022 | 0.012667 | 0.019667 | 0.024 | 0.018333 | 610.326 | 0.0013333 | 2.1705e-05
610.603 | 0.012 | 610.603 | 0.012 | 0.015 | 0.02 | 0.02 | 0.019 | 0.024 | 0.025 | 0.026 | 0.033 | 0.015667 | 0.021 | 0.028 | 0.021833 | 610.603 | -0.00083333 | -1.3565e-05
610.879 | 0.017 | 610.879 | 0.017 | 0.019 | 0.021 | 0.025 | 0.026 | 0.032 | 0.032 | 0.034 | 0.035 | 0.019 | 0.027667 | 0.033667 | 0.026333 | 610.879 | 0.0013333 | 2.1705e-05
611.155 | 0.012 | 611.155 | 0.012 | 0.015 | 0.017 | 0.021 | 0.024 | 0.023 | 0.026 | 0.025 | 0.025 | 0.014667 | 0.022667 | 0.025333 | 0.02 | 611.155 | 0.0026667 | 4.3409e-05
611.431 | 0.005 | 611.431 | 0.005 | 0.003 | 0.006 | 0.011 | 0.014 | 0.013 | 0.013 | 0.015 | 0.019 | 0.0046667 | 0.012667 | 0.015667 | 0.010167 | 611.431 | 0.0025 | 4.0696e-05
611.707 | 0.012 | 611.707 | 0.012 | 0.014 | 0.014 | 0.016 | 0.021 | 0.024 | 0.022 | 0.025 | 0.027 | 0.013333 | 0.020333 | 0.024667 | 0.019 | 611.707 | 0.0013333 | 2.1705e-05
611.983 | 0.002 | 611.983 | 0.002 | 0.006 | -0.001 | 0.007 | 0.011 | 0.013 | 0.014 | 0.019 | 0.016 | 0.0023333 | 0.010333 | 0.016333 | 0.0093333 | 611.983 | 0.001 | 1.6279e-05
612.259 | 0.016 | 612.259 | 0.016 | 0.017 | 0.021 | 0.024 | 0.029 | 0.026 | 0.025 | 0.033 | 0.027 | 0.018 | 0.026333 | 0.028333 | 0.023167 | 612.259 | 0.0031667 | 5.1549e-05
612.535 | 0.003 | 612.535 | 0.003 | 0.004 | 0.004 | 0.012 | 0.014 | 0.015 | 0.017 | 0.017 | 0.02 | 0.0036667 | 0.013667 | 0.018 | 0.010833 | 612.535 | 0.0028333 | 4.6123e-05
612.811 | 0.007 | 612.811 | 0.007 | 0.012 | 0.011 | 0.011 | 0.015 | 0.018 | 0.02 | 0.023 | 0.02 | 0.01 | 0.014667 | 0.021 | 0.0155 | 612.811 | -0.00083333 | -1.3565e-05
613.087 | 0.01 | 613.087 | 0.01 | 0.007 | 0.011 | 0.015 | 0.015 | 0.02 | 0.023 | 0.021 | 0.024 | 0.0093333 | 0.016667 | 0.022667 | 0.016 | 613.087 | 0.00066667 | 1.0852e-05
613.362 | 0.004 | 613.362 | 0.004 | 0.01 | 0.009 | 0.011 | 0.013 | 0.014 | 0.019 | 0.02 | 0.021 | 0.0076667 | 0.012667 | 0.02 | 0.013833 | 613.362 | -0.0011667 | -1.8992e-05
613.638 | 0.007 | 613.638 | 0.007 | 0.003 | 0.008 | 0.011 | 0.018 | 0.017 | 0.014 | 0.019 | 0.023 | 0.006 | 0.015333 | 0.018667 | 0.012333 | 613.638 | 0.003 | 4.8836e-05
613.913 | 0.009 | 613.913 | 0.009 | 0.007 | 0.013 | 0.016 | 0.017 | 0.015 | 0.02 | 0.021 | 0.023 | 0.0096667 | 0.016 | 0.021333 | 0.0155 | 613.913 | 0.0005 | 8.1393e-06
614.189 | 0.004 | 614.189 | 0.004 | 0.009 | 0.01 | 0.015 | 0.014 | 0.016 | 0.018 | 0.019 | 0.026 | 0.0076667 | 0.015 | 0.021 | 0.014333 | 614.189 | 0.00066667 | 1.0852e-05
614.464 | 0.014 | 614.464 | 0.014 | 0.01 | 0.015 | 0.017 | 0.016 | 0.025 | 0.022 | 0.023 | 0.026 | 0.013 | 0.019333 | 0.023667 | 0.018333 | 614.464 | 0.001 | 1.6279e-05
614.74 | 0.014 | 614.74 | 0.014 | 0.02 | 0.025 | 0.025 | 0.029 | 0.026 | 0.032 | 0.034 | 0.036 | 0.019667 | 0.026667 | 0.034 | 0.026833 | 614.74 | -0.00016667 | -2.7131e-06
615.015 | -0.004 | 615.015 | -0.004 | -0.002 | 0.002 | 0.002 | 0.003 | 0.005 | 0.005 | 0.011 | 0.01 | -0.0013333 | 0.0033333 | 0.0086667 | 0.0036667 | 615.015 | -0.00033333 | -5.4262e-06
615.29 | 0.01 | 615.29 | 0.01 | 0.013 | 0.017 | 0.017 | 0.019 | 0.021 | 0.024 | 0.024 | 0.024 | 0.013333 | 0.019 | 0.024 | 0.018667 | 615.29 | 0.00033333 | 5.4262e-06
615.565 | 0.012 | 615.565 | 0.012 | 0.017 | 0.012 | 0.018 | 0.024 | 0.023 | 0.022 | 0.028 | 0.023 | 0.013667 | 0.021667 | 0.024333 | 0.019 | 615.565 | 0.0026667 | 4.3409e-05
615.84 | -0.007 | 615.84 | -0.007 | -0.004 | -0.003 | 0.007 | 0.003 | 0.003 | 0.01 | 0.013 | 0.01 | -0.0046667 | 0.0043333 | 0.011 | 0.0031667 | 615.84 | 0.0011667 | 1.8992e-05
616.115 | 0.012 | 616.115 | 0.012 | 0.019 | 0.018 | 0.021 | 0.022 | 0.025 | 0.03 | 0.031 | 0.03 | 0.016333 | 0.022667 | 0.030333 | 0.023333 | 616.115 | -0.00066667 | -1.0852e-05
616.39 | 0.008 | 616.39 | 0.008 | 0.015 | 0.009 | 0.018 | 0.022 | 0.022 | 0.019 | 0.026 | 0.023 | 0.010667 | 0.020667 | 0.022667 | 0.016667 | 616.39 | 0.004 | 6.5114e-05
616.665 | 0.012 | 616.665 | 0.012 | 0.015 | 0.015 | 0.023 | 0.017 | 0.024 | 0.027 | 0.026 | 0.03 | 0.014 | 0.021333 | 0.027667 | 0.020833 | 616.665 | 0.0005 | 8.1393e-06
616.94 | 0.007 | 616.94 | 0.007 | 0.012 | 0.011 | 0.019 | 0.023 | 0.019 | 0.02 | 0.019 | 0.02 | 0.01 | 0.020333 | 0.019667 | 0.014833 | 616.94 | 0.0055 | 8.9532e-05
617.215 | 0.017 | 617.215 | 0.017 | 0.019 | 0.019 | 0.024 | 0.022 | 0.027 | 0.031 | 0.031 | 0.03 | 0.018333 | 0.024333 | 0.030667 | 0.0245 | 617.215 | -0.00016667 | -2.7131e-06
617.49 | 0.01 | 617.49 | 0.01 | 0.009 | 0.008 | 0.014 | 0.018 | 0.022 | 0.023 | 0.024 | 0.026 | 0.009 | 0.018 | 0.024333 | 0.016667 | 617.49 | 0.0013333 | 2.1705e-05
617.764 | 0.011 | 617.764 | 0.011 | 0.016 | 0.017 | 0.025 | 0.023 | 0.029 | 0.027 | 0.03 | 0.026 | 0.014667 | 0.025667 | 0.027667 | 0.021167 | 617.764 | 0.0045 | 7.3254e-05
618.039 | 0.017 | 618.039 | 0.017 | 0.016 | 0.018 | 0.021 | 0.027 | 0.023 | 0.026 | 0.031 | 0.028 | 0.017 | 0.023667 | 0.028333 | 0.022667 | 618.039 | 0.001 | 1.6279e-05
618.313 | 0.01 | 618.313 | 0.01 | 0.019 | 0.018 | 0.022 | 0.024 | 0.024 | 0.026 | 0.031 | 0.03 | 0.015667 | 0.023333 | 0.029 | 0.022333 | 618.313 | 0.001 | 1.6279e-05
618.588 | 0.021 | 618.588 | 0.021 | 0.021 | 0.023 | 0.03 | 0.03 | 0.03 | 0.031 | 0.036 | 0.036 | 0.021667 | 0.03 | 0.034333 | 0.028 | 618.588 | 0.002 | 3.2557e-05
618.862 | 0.01 | 618.862 | 0.01 | 0.011 | 0.012 | 0.017 | 0.017 | 0.02 | 0.024 | 0.023 | 0.024 | 0.011 | 0.018 | 0.023667 | 0.017333 | 618.862 | 0.00066667 | 1.0852e-05
619.137 | 0.019 | 619.137 | 0.019 | 0.019 | 0.024 | 0.028 | 0.03 | 0.029 | 0.031 | 0.036 | 0.035 | 0.020667 | 0.029 | 0.034 | 0.027333 | 619.137 | 0.0016667 | 2.7131e-05
619.411 | 0.009 | 619.411 | 0.009 | 0.011 | 0.011 | 0.012 | 0.019 | 0.019 | 0.016 | 0.021 | 0.023 | 0.010333 | 0.016667 | 0.02 | 0.015167 | 619.411 | 0.0015 | 2.4418e-05
619.685 | 0.011 | 619.685 | 0.011 | 0.015 | 0.013 | 0.016 | 0.02 | 0.023 | 0.024 | 0.024 | 0.027 | 0.013 | 0.019667 | 0.025 | 0.019 | 619.685 | 0.00066667 | 1.0852e-05
619.959 | 0.005 | 619.959 | 0.005 | 0.007 | 0.009 | 0.013 | 0.017 | 0.016 | 0.021 | 0.02 | 0.019 | 0.007 | 0.015333 | 0.02 | 0.0135 | 619.959 | 0.0018333 | 2.9844e-05
620.233 | 0.02 | 620.233 | 0.02 | 0.022 | 0.027 | 0.032 | 0.028 | 0.032 | 0.032 | 0.031 | 0.033 | 0.023 | 0.030667 | 0.032 | 0.0275 | 620.233 | 0.0031667 | 5.1549e-05
620.507 | 0.004 | 620.507 | 0.004 | 0.008 | 0.006 | 0.014 | 0.013 | 0.016 | 0.017 | 0.022 | 0.02 | 0.006 | 0.014333 | 0.019667 | 0.012833 | 620.507 | 0.0015 | 2.4418e-05
620.781 | 0.011 | 620.781 | 0.011 | 0.014 | 0.009 | 0.02 | 0.02 | 0.021 | 0.017 | 0.029 | 0.024 | 0.011333 | 0.020333 | 0.023333 | 0.017333 | 620.781 | 0.003 | 4.8836e-05
621.055 | 0.006 | 621.055 | 0.006 | 0.007 | 0.008 | 0.015 | 0.016 | 0.015 | 0.017 | 0.019 | 0.016 | 0.007 | 0.015333 | 0.017333 | 0.012167 | 621.055 | 0.0031667 | 5.1549e-05
621.329 | 0.01 | 621.329 | 0.01 | 0.015 | 0.013 | 0.016 | 0.019 | 0.021 | 0.024 | 0.023 | 0.025 | 0.012667 | 0.018667 | 0.024 | 0.018333 | 621.329 | 0.00033333 | 5.4262e-06
621.603 | 0.008 | 621.603 | 0.008 | 0.008 | 0.012 | 0.016 | 0.022 | 0.021 | 0.018 | 0.027 | 0.026 | 0.0093333 | 0.019667 | 0.023667 | 0.0165 | 621.603 | 0.0031667 | 5.1549e-05
621.876 | 0.015 | 621.876 | 0.015 | 0.019 | 0.02 | 0.027 | 0.025 | 0.026 | 0.029 | 0.032 | 0.033 | 0.018 | 0.026 | 0.031333 | 0.024667 | 621.876 | 0.0013333 | 2.1705e-05
622.15 | 0.014 | 622.15 | 0.014 | 0.016 | 0.014 | 0.021 | 0.021 | 0.02 | 0.022 | 0.026 | 0.028 | 0.014667 | 0.020667 | 0.025333 | 0.02 | 622.15 | 0.00066667 | 1.0852e-05
622.424 | 0.009 | 622.424 | 0.009 | 0.009 | 0.013 | 0.015 | 0.015 | 0.018 | 0.021 | 0.021 | 0.023 | 0.010333 | 0.016 | 0.021667 | 0.016 | 622.424 | 0 | 0
622.697 | 0.002 | 622.697 | 0.002 | 0.007 | 0.01 | 0.017 | 0.016 | 0.015 | 0.024 | 0.022 | 0.024 | 0.0063333 | 0.016 | 0.023333 | 0.014833 | 622.697 | 0.0011667 | 1.8992e-05
622.971 | 0.004 | 622.971 | 0.004 | 0.006 | 0.01 | 0.015 | 0.011 | 0.015 | 0.02 | 0.015 | 0.018 | 0.0066667 | 0.013667 | 0.017667 | 0.012167 | 622.971 | 0.0015 | 2.4418e-05
623.244 | 0 | 623.244 | 0 | 0.005 | 0.003 | 0.011 | 0.013 | 0.014 | 0.013 | 0.015 | 0.021 | 0.0026667 | 0.012667 | 0.016333 | 0.0095 | 623.244 | 0.0031667 | 5.1549e-05
623.517 | -0.001 | 623.517 | -0.001 | 0.003 | 0.007 | 0.01 | 0.01 | 0.012 | 0.015 | 0.017 | 0.017 | 0.003 | 0.010667 | 0.016333 | 0.0096667 | 623.517 | 0.001 | 1.6279e-05
623.791 | 0.008 | 623.791 | 0.008 | 0.013 | 0.012 | 0.015 | 0.018 | 0.023 | 0.021 | 0.027 | 0.025 | 0.011 | 0.018667 | 0.024333 | 0.017667 | 623.791 | 0.001 | 1.6279e-05
624.064 | 0.012 | 624.064 | 0.012 | 0.016 | 0.019 | 0.025 | 0.026 | 0.026 | 0.026 | 0.027 | 0.034 | 0.015667 | 0.025667 | 0.029 | 0.022333 | 624.064 | 0.0033333 | 5.4262e-05
624.337 | 0.019 | 624.337 | 0.019 | 0.018 | 0.023 | 0.026 | 0.026 | 0.029 | 0.029 | 0.029 | 0.032 | 0.02 | 0.027 | 0.03 | 0.025 | 624.337 | 0.002 | 3.2557e-05
624.61 | 0.002 | 624.61 | 0.002 | 0.005 | 0.01 | 0.009 | 0.014 | 0.017 | 0.017 | 0.018 | 0.025 | 0.0056667 | 0.013333 | 0.02 | 0.012833 | 624.61 | 0.0005 | 8.1393e-06
624.883 | 0.013 | 624.883 | 0.013 | 0.012 | 0.016 | 0.019 | 0.02 | 0.023 | 0.025 | 0.028 | 0.028 | 0.013667 | 0.020667 | 0.027 | 0.020333 | 624.883 | 0.00033333 | 5.4262e-06
625.156 | 0.002 | 625.156 | 0.002 | 0.006 | 0.006 | 0.012 | 0.009 | 0.011 | 0.014 | 0.017 | 0.019 | 0.0046667 | 0.010667 | 0.016667 | 0.010667 | 625.156 | 0 | 0
625.429 | 0.013 | 625.429 | 0.013 | 0.011 | 0.014 | 0.02 | 0.02 | 0.021 | 0.026 | 0.022 | 0.024 | 0.012667 | 0.020333 | 0.024 | 0.018333 | 625.429 | 0.002 | 3.2557e-05
625.702 | 0.005 | 625.702 | 0.005 | 0.003 | 0.01 | 0.009 | 0.018 | 0.012 | 0.014 | 0.019 | 0.018 | 0.006 | 0.013 | 0.017 | 0.0115 | 625.702 | 0.0015 | 2.4418e-05
625.974 | 0.02 | 625.974 | 0.02 | 0.022 | 0.023 | 0.025 | 0.027 | 0.03 | 0.031 | 0.033 | 0.034 | 0.021667 | 0.027333 | 0.032667 | 0.027167 | 625.974 | 0.00016667 | 2.7131e-06
626.247 | 0.011 | 626.247 | 0.011 | 0.012 | 0.014 | -0.001 | 0.002 | 0.006 | 0.012 | 0.024 | 0.026 | 0.012333 | 0.0023333 | 0.020667 | 0.0165 | 626.247 | -0.014167 | -0.00023061
626.52 | 0.013 | 626.52 | 0.013 | 0.018 | 0.021 | 0.02 | 0.02 | 0.023 | 0.026 | 0.024 | 0.027 | 0.017333 | 0.021 | 0.025667 | 0.0215 | 626.52 | -0.0005 | -8.1393e-06
626.792 | 0.011 | 626.792 | 0.011 | 0.013 | 0.011 | 0.017 | 0.021 | 0.02 | 0.026 | 0.024 | 0.026 | 0.011667 | 0.019333 | 0.025333 | 0.0185 | 626.792 | 0.00083333 | 1.3565e-05
627.065 | 0.01 | 627.065 | 0.01 | 0.014 | 0.013 | 0.021 | 0.017 | 0.019 | 0.026 | 0.026 | 0.029 | 0.012333 | 0.019 | 0.027 | 0.019667 | 627.065 | -0.00066667 | -1.0852e-05
627.337 | 0.01 | 627.337 | 0.01 | 0.012 | 0.014 | 0.02 | 0.024 | 0.023 | 0.022 | 0.026 | 0.029 | 0.012 | 0.022333 | 0.025667 | 0.018833 | 627.337 | 0.0035 | 5.6975e-05
627.61 | 0.006 | 627.61 | 0.006 | 0.01 | 0.015 | 0.017 | 0.02 | 0.022 | 0.022 | 0.024 | 0.029 | 0.010333 | 0.019667 | 0.025 | 0.017667 | 627.61 | 0.002 | 3.2557e-05
627.882 | 0.008 | 627.882 | 0.008 | 0.014 | 0.014 | 0.019 | 0.02 | 0.015 | 0.021 | 0.022 | 0.024 | 0.012 | 0.018 | 0.022333 | 0.017167 | 627.882 | 0.00083333 | 1.3565e-05
628.154 | 0.008 | 628.154 | 0.008 | 0.013 | 0.014 | 0.016 | 0.018 | 0.021 | 0.024 | 0.026 | 0.027 | 0.011667 | 0.018333 | 0.025667 | 0.018667 | 628.154 | -0.00033333 | -5.4262e-06
628.427 | 0.019 | 628.427 | 0.019 | 0.024 | 0.023 | 0.029 | 0.032 | 0.033 | 0.031 | 0.035 | 0.04 | 0.022 | 0.031333 | 0.035333 | 0.028667 | 628.427 | 0.0026667 | 4.3409e-05
628.699 | 0.014 | 628.699 | 0.014 | 0.015 | 0.019 | 0.018 | 0.021 | 0.027 | 0.021 | 0.028 | 0.027 | 0.016 | 0.022 | 0.025333 | 0.020667 | 628.699 | 0.0013333 | 2.1705e-05
628.971 | 0.008 | 628.971 | 0.008 | 0.01 | 0.012 | 0.019 | 0.021 | 0.021 | 0.021 | 0.023 | 0.026 | 0.01 | 0.020333 | 0.023333 | 0.016667 | 628.971 | 0.0036667 | 5.9688e-05
629.243 | 0 | 629.243 | 0 | 0 | 0.006 | 0.009 | 0.013 | 0.009 | 0.012 | 0.014 | 0.016 | 0.002 | 0.010333 | 0.014 | 0.008 | 629.243 | 0.0023333 | 3.7983e-05
629.515 | 0.018 | 629.515 | 0.018 | 0.016 | 0.02 | 0.026 | 0.029 | 0.029 | 0.035 | 0.034 | 0.033 | 0.018 | 0.028 | 0.034 | 0.026 | 629.515 | 0.002 | 3.2557e-05
629.787 | 0.007 | 629.787 | 0.007 | 0.005 | 0.007 | 0.018 | 0.018 | 0.021 | 0.022 | 0.022 | 0.025 | 0.0063333 | 0.019 | 0.023 | 0.014667 | 629.787 | 0.0043333 | 7.054e-05
630.058 | 0.021 | 630.058 | 0.021 | 0.022 | 0.022 | 0.027 | 0.033 | 0.029 | 0.034 | 0.03 | 0.038 | 0.021667 | 0.029667 | 0.034 | 0.027833 | 630.058 | 0.0018333 | 2.9844e-05
630.33 | 0.015 | 630.33 | 0.015 | 0.019 | 0.02 | 0.027 | 0.023 | 0.03 | 0.029 | 0.034 | 0.032 | 0.018 | 0.026667 | 0.031667 | 0.024833 | 630.33 | 0.0018333 | 2.9844e-05
630.602 | -0.004 | 630.602 | -0.004 | -0.001 | -0.001 | 0.005 | 0.009 | 0.006 | 0.007 | 0.01 | 0.012 | -0.002 | 0.0066667 | 0.0096667 | 0.0038333 | 630.602 | 0.0028333 | 4.6123e-05
630.874 | 0.014 | 630.874 | 0.014 | 0.02 | 0.021 | 0.025 | 0.023 | 0.027 | 0.027 | 0.032 | 0.032 | 0.018333 | 0.025 | 0.030333 | 0.024333 | 630.874 | 0.00066667 | 1.0852e-05
631.145 | 0.018 | 631.145 | 0.018 | 0.026 | 0.017 | 0.031 | 0.031 | 0.031 | 0.035 | 0.038 | 0.037 | 0.020333 | 0.031 | 0.036667 | 0.0285 | 631.145 | 0.0025 | 4.0696e-05
631.417 | 0.004 | 631.417 | 0.004 | 0 | 0.008 | 0.009 | 0.014 | 0.015 | 0.014 | 0.017 | 0.017 | 0.004 | 0.012667 | 0.016 | 0.01 | 631.417 | 0.0026667 | 4.3409e-05
631.688 | 0.015 | 631.688 | 0.015 | 0.02 | 0.023 | 0.021 | 0.029 | 0.025 | 0.026 | 0.033 | 0.035 | 0.019333 | 0.025 | 0.031333 | 0.025333 | 631.688 | -0.00033333 | -5.4262e-06
631.959 | 0.009 | 631.959 | 0.009 | 0.011 | 0.012 | 0.02 | 0.021 | 0.018 | 0.019 | 0.022 | 0.023 | 0.010667 | 0.019667 | 0.021333 | 0.016 | 631.959 | 0.0036667 | 5.9688e-05
632.231 | 0.007 | 632.231 | 0.007 | 0.007 | 0.013 | 0.016 | 0.022 | 0.018 | 0.022 | 0.022 | 0.025 | 0.009 | 0.018667 | 0.023 | 0.016 | 632.231 | 0.0026667 | 4.3409e-05
632.502 | 0.015 | 632.502 | 0.015 | 0.019 | 0.021 | 0.021 | 0.023 | 0.027 | 0.031 | 0.033 | 0.034 | 0.018333 | 0.023667 | 0.032667 | 0.0255 | 632.502 | -0.0018333 | -2.9844e-05
632.773 | 0.02 | 632.773 | 0.02 | 0.023 | 0.027 | 0.027 | 0.032 | 0.027 | 0.03 | 0.03 | 0.035 | 0.023333 | 0.028667 | 0.031667 | 0.0275 | 632.773 | 0.0011667 | 1.8992e-05
633.044 | 0.029 | 633.044 | 0.029 | 0.027 | 0.026 | 0.033 | 0.037 | 0.037 | 0.038 | 0.043 | 0.043 | 0.027333 | 0.035667 | 0.041333 | 0.034333 | 633.044 | 0.0013333 | 2.1705e-05
633.316 | 0.015 | 633.316 | 0.015 | 0.019 | 0.025 | 0.026 | 0.026 | 0.027 | 0.03 | 0.035 | 0.035 | 0.019667 | 0.026333 | 0.033333 | 0.0265 | 633.316 | -0.00016667 | -2.7131e-06
633.587 | 0 | 633.587 | 0 | -0.003 | -0.002 | 0.007 | 0.005 | 0.01 | 0.009 | 0.013 | 0.012 | -0.0016667 | 0.0073333 | 0.011333 | 0.0048333 | 633.587 | 0.0025 | 4.0696e-05
633.857 | 0.011 | 633.857 | 0.011 | 0.01 | 0.012 | 0.017 | 0.018 | 0.014 | 0.013 | 0.019 | 0.026 | 0.011 | 0.016333 | 0.019333 | 0.015167 | 633.857 | 0.0011667 | 1.8992e-05
634.128 | 0.018 | 634.128 | 0.018 | 0.019 | 0.021 | 0.032 | 0.029 | 0.033 | 0.034 | 0.041 | 0.033 | 0.019333 | 0.031333 | 0.036 | 0.027667 | 634.128 | 0.0036667 | 5.9688e-05
634.399 | 0.029 | 634.399 | 0.029 | 0.029 | 0.031 | 0.034 | 0.033 | 0.037 | 0.041 | 0.042 | 0.042 | 0.029667 | 0.034667 | 0.041667 | 0.035667 | 634.399 | -0.001 | -1.6279e-05
634.67 | 0.003 | 634.67 | 0.003 | 0.004 | 0.007 | 0.015 | 0.012 | 0.013 | 0.019 | 0.016 | 0.019 | 0.0046667 | 0.013333 | 0.018 | 0.011333 | 634.67 | 0.002 | 3.2557e-05
634.941 | 0.015 | 634.941 | 0.015 | 0.007 | 0.014 | 0.019 | 0.021 | 0.021 | 0.026 | 0.024 | 0.025 | 0.012 | 0.020333 | 0.025 | 0.0185 | 634.941 | 0.0018333 | 2.9844e-05
635.211 | 0.02 | 635.211 | 0.02 | 0.02 | 0.023 | 0.023 | 0.028 | 0.028 | 0.03 | 0.03 | 0.034 | 0.021 | 0.026333 | 0.031333 | 0.026167 | 635.211 | 0.00016667 | 2.7131e-06
635.482 | 0.008 | 635.482 | 0.008 | 0.01 | 0.011 | 0.016 | 0.017 | 0.014 | 0.02 | 0.02 | 0.018 | 0.0096667 | 0.015667 | 0.019333 | 0.0145 | 635.482 | 0.0011667 | 1.8992e-05
635.752 | 0.013 | 635.752 | 0.013 | 0.013 | 0.021 | 0.023 | 0.025 | 0.027 | 0.028 | 0.025 | 0.034 | 0.015667 | 0.025 | 0.029 | 0.022333 | 635.752 | 0.0026667 | 4.3409e-05
636.023 | -0.002 | 636.023 | -0.002 | -0.002 | 0.001 | 0.008 | 0.012 | 0.008 | 0.01 | 0.014 | 0.019 | -0.001 | 0.0093333 | 0.014333 | 0.0066667 | 636.023 | 0.0026667 | 4.3409e-05
636.293 | 0.006 | 636.293 | 0.006 | 0.008 | 0.008 | 0.014 | 0.012 | 0.017 | 0.022 | 0.027 | 0.028 | 0.0073333 | 0.014333 | 0.025667 | 0.0165 | 636.293 | -0.0021667 | -3.527e-05
636.564 | 0.026 | 636.564 | 0.026 | 0.027 | 0.027 | 0.03 | 0.037 | 0.035 | 0.038 | 0.037 | 0.041 | 0.026667 | 0.034 | 0.038667 | 0.032667 | 636.564 | 0.0013333 | 2.1705e-05
636.834 | 0.011 | 636.834 | 0.011 | 0.011 | 0.014 | 0.018 | 0.021 | 0.026 | 0.022 | 0.023 | 0.025 | 0.012 | 0.021667 | 0.023333 | 0.017667 | 636.834 | 0.004 | 6.5114e-05
637.104 | 0.02 | 637.104 | 0.02 | 0.019 | 0.016 | 0.027 | 0.029 | 0.031 | 0.033 | 0.033 | 0.039 | 0.018333 | 0.029 | 0.035 | 0.026667 | 637.104 | 0.0023333 | 3.7983e-05
637.374 | 0.012 | 637.374 | 0.012 | 0.01 | 0.016 | 0.018 | 0.014 | 0.021 | 0.021 | 0.021 | 0.026 | 0.012667 | 0.017667 | 0.022667 | 0.017667 | 637.374 | 0 | 0
637.644 | 0.016 | 637.644 | 0.016 | 0.019 | 0.02 | 0.02 | 0.024 | 0.024 | 0.031 | 0.032 | 0.031 | 0.018333 | 0.022667 | 0.031333 | 0.024833 | 637.644 | -0.0021667 | -3.527e-05
637.914 | 0.017 | 637.914 | 0.017 | 0.018 | 0.024 | 0.018 | 0.026 | 0.027 | 0.029 | 0.028 | 0.036 | 0.019667 | 0.023667 | 0.031 | 0.025333 | 637.914 | -0.0016667 | -2.7131e-05
638.184 | 0.014 | 638.184 | 0.014 | 0.015 | 0.016 | 0.024 | 0.027 | 0.027 | 0.028 | 0.026 | 0.03 | 0.015 | 0.026 | 0.028 | 0.0215 | 638.184 | 0.0045 | 7.3254e-05
638.454 | 0.01 | 638.454 | 0.01 | 0.009 | 0.013 | 0.016 | 0.021 | 0.019 | 0.021 | 0.024 | 0.019 | 0.010667 | 0.018667 | 0.021333 | 0.016 | 638.454 | 0.0026667 | 4.3409e-05
638.724 | 0.022 | 638.724 | 0.022 | 0.027 | 0.026 | 0.03 | 0.026 | 0.035 | 0.039 | 0.04 | 0.04 | 0.025 | 0.030333 | 0.039667 | 0.032333 | 638.724 | -0.002 | -3.2557e-05
638.994 | 0.008 | 638.994 | 0.008 | 0.012 | 0.012 | 0.02 | 0.02 | 0.025 | 0.026 | 0.027 | 0.027 | 0.010667 | 0.021667 | 0.026667 | 0.018667 | 638.994 | 0.003 | 4.8836e-05
639.264 | 0.017 | 639.264 | 0.017 | 0.021 | 0.021 | 0.03 | 0.025 | 0.033 | 0.033 | 0.034 | 0.033 | 0.019667 | 0.029333 | 0.033333 | 0.0265 | 639.264 | 0.0028333 | 4.6123e-05
639.533 | 0.01 | 639.533 | 0.01 | 0.013 | 0.014 | 0.015 | 0.018 | 0.02 | 0.024 | 0.023 | 0.026 | 0.012333 | 0.017667 | 0.024333 | 0.018333 | 639.533 | -0.00066667 | -1.0852e-05
639.803 | 0.013 | 639.803 | 0.013 | 0.012 | 0.008 | 0.013 | 0.017 | 0.018 | 0.024 | 0.023 | 0.022 | 0.011 | 0.016 | 0.023 | 0.017 | 639.803 | -0.001 | -1.6279e-05
640.072 | 0.015 | 640.072 | 0.015 | 0.016 | 0.014 | 0.02 | 0.022 | 0.023 | 0.024 | 0.026 | 0.028 | 0.015 | 0.021667 | 0.026 | 0.0205 | 640.072 | 0.0011667 | 1.8992e-05
640.342 | 0.005 | 640.342 | 0.005 | 0.013 | 0.015 | 0.016 | 0.02 | 0.021 | 0.023 | 0.023 | 0.022 | 0.011 | 0.019 | 0.022667 | 0.016833 | 640.342 | 0.0021667 | 3.527e-05
640.611 | 0.015 | 640.611 | 0.015 | 0.02 | 0.026 | 0.022 | 0.023 | 0.029 | 0.032 | 0.035 | 0.033 | 0.020333 | 0.024667 | 0.033333 | 0.026833 | 640.611 | -0.0021667 | -3.527e-05
640.881 | 0.009 | 640.881 | 0.009 | 0.011 | 0.016 | 0.02 | 0.022 | 0.02 | 0.025 | 0.026 | 0.029 | 0.012 | 0.020667 | 0.026667 | 0.019333 | 640.881 | 0.0013333 | 2.1705e-05
641.15 | 0.028 | 641.15 | 0.028 | 0.037 | 0.038 | 0.037 | 0.044 | 0.046 | 0.044 | 0.046 | 0.047 | 0.034333 | 0.042333 | 0.045667 | 0.04 | 641.15 | 0.0023333 | 3.7983e-05
641.419 | 0.001 | 641.419 | 0.001 | 0.001 | 0.009 | 0.004 | 0.013 | 0.013 | 0.015 | 0.017 | 0.015 | 0.0036667 | 0.01 | 0.015667 | 0.0096667 | 641.419 | 0.00033333 | 5.4262e-06
641.688 | 0.012 | 641.688 | 0.012 | 0.013 | 0.016 | 0.02 | 0.025 | 0.025 | 0.025 | 0.026 | 0.03 | 0.013667 | 0.023333 | 0.027 | 0.020333 | 641.688 | 0.003 | 4.8836e-05
641.957 | 0.009 | 641.957 | 0.009 | 0.011 | 0.013 | 0.017 | 0.021 | 0.016 | 0.022 | 0.024 | 0.027 | 0.011 | 0.018 | 0.024333 | 0.017667 | 641.957 | 0.00033333 | 5.4262e-06
642.227 | 0.014 | 642.227 | 0.014 | 0.016 | 0.02 | 0.023 | 0.024 | 0.027 | 0.027 | 0.032 | 0.034 | 0.016667 | 0.024667 | 0.031 | 0.023833 | 642.227 | 0.00083333 | 1.3565e-05
642.496 | 0.024 | 642.496 | 0.024 | 0.024 | 0.026 | 0.031 | 0.031 | 0.033 | 0.032 | 0.034 | 0.04 | 0.024667 | 0.031667 | 0.035333 | 0.03 | 642.496 | 0.0016667 | 2.7131e-05
642.764 | 0.018 | 642.764 | 0.018 | 0.017 | 0.017 | 0.029 | 0.028 | 0.025 | 0.03 | 0.033 | 0.033 | 0.017333 | 0.027333 | 0.032 | 0.024667 | 642.764 | 0.0026667 | 4.3409e-05
643.033 | 0.023 | 643.033 | 0.023 | 0.023 | 0.024 | 0.029 | 0.034 | 0.035 | 0.035 | 0.038 | 0.039 | 0.023333 | 0.032667 | 0.037333 | 0.030333 | 643.033 | 0.0023333 | 3.7983e-05
643.302 | 0.004 | 643.302 | 0.004 | 0.004 | 0.011 | 0.015 | 0.021 | 0.011 | 0.022 | 0.024 | 0.025 | 0.0063333 | 0.015667 | 0.023667 | 0.015 | 643.302 | 0.00066667 | 1.0852e-05
643.571 | 0.005 | 643.571 | 0.005 | 0.011 | 0.012 | 0.021 | 0.018 | 0.016 | 0.022 | 0.021 | 0.024 | 0.0093333 | 0.018333 | 0.022333 | 0.015833 | 643.571 | 0.0025 | 4.0696e-05
643.84 | 0.007 | 643.84 | 0.007 | 0.007 | 0.01 | 0.015 | 0.016 | 0.016 | 0.017 | 0.021 | 0.025 | 0.008 | 0.015667 | 0.021 | 0.0145 | 643.84 | 0.0011667 | 1.8992e-05
644.108 | 0.017 | 644.108 | 0.017 | 0.024 | 0.021 | 0.033 | 0.03 | 0.038 | 0.033 | 0.033 | 0.039 | 0.020667 | 0.033667 | 0.035 | 0.027833 | 644.108 | 0.0058333 | 9.4958e-05
644.377 | -0.007 | 644.377 | -0.007 | -0.006 | -0.004 | 0.002 | 0.003 | 0.004 | 0.006 | 0.012 | 0.01 | -0.0056667 | 0.003 | 0.0093333 | 0.0018333 | 644.377 | 0.0011667 | 1.8992e-05
644.645 | 0.007 | 644.645 | 0.007 | 0.006 | 0.011 | 0.011 | 0.014 | 0.013 | 0.017 | 0.021 | 0.017 | 0.008 | 0.012667 | 0.018333 | 0.013167 | 644.645 | -0.0005 | -8.1393e-06
644.914 | 0.009 | 644.914 | 0.009 | 0.012 | 0.011 | 0.012 | 0.018 | 0.021 | 0.021 | 0.022 | 0.026 | 0.010667 | 0.017 | 0.023 | 0.016833 | 644.914 | 0.00016667 | 2.7131e-06
645.182 | 0.019 | 645.182 | 0.019 | 0.022 | 0.024 | 0.027 | 0.03 | 0.031 | 0.034 | 0.036 | 0.034 | 0.021667 | 0.029333 | 0.034667 | 0.028167 | 645.182 | 0.0011667 | 1.8992e-05
645.45 | 0.018 | 645.45 | 0.018 | 0.022 | 0.024 | 0.024 | 0.029 | 0.027 | 0.032 | 0.032 | 0.033 | 0.021333 | 0.026667 | 0.032333 | 0.026833 | 645.45 | -0.00016667 | -2.7131e-06
645.719 | 0.026 | 645.719 | 0.026 | 0.026 | 0.029 | 0.032 | 0.031 | 0.036 | 0.041 | 0.041 | 0.04 | 0.027 | 0.033 | 0.040667 | 0.033833 | 645.719 | -0.00083333 | -1.3565e-05
645.987 | 0.015 | 645.987 | 0.015 | 0.013 | 0.018 | 0.031 | 0.031 | 0.027 | 0.036 | 0.034 | 0.035 | 0.015333 | 0.029667 | 0.035 | 0.025167 | 645.987 | 0.0045 | 7.3254e-05
646.255 | 0.006 | 646.255 | 0.006 | 0.012 | 0.015 | 0.013 | 0.015 | 0.019 | 0.024 | 0.025 | 0.03 | 0.011 | 0.015667 | 0.026333 | 0.018667 | 646.255 | -0.003 | -4.8836e-05
646.523 | 0.01 | 646.523 | 0.01 | 0.01 | 0.016 | 0.017 | 0.016 | 0.021 | 0.02 | 0.024 | 0.026 | 0.012 | 0.018 | 0.023333 | 0.017667 | 646.523 | 0.00033333 | 5.4262e-06
646.791 | 0.007 | 646.791 | 0.007 | 0.006 | 0.013 | 0.013 | 0.017 | 0.019 | 0.022 | 0.025 | 0.023 | 0.0086667 | 0.016333 | 0.023333 | 0.016 | 646.791 | 0.00033333 | 5.4262e-06
647.059 | 0.026 | 647.059 | 0.026 | 0.024 | 0.029 | 0.032 | 0.036 | 0.04 | 0.04 | 0.041 | 0.04 | 0.026333 | 0.036 | 0.040333 | 0.033333 | 647.059 | 0.0026667 | 4.3409e-05
647.327 | 0.011 | 647.327 | 0.011 | 0.015 | 0.009 | 0.018 | 0.019 | 0.021 | 0.021 | 0.03 | 0.028 | 0.011667 | 0.019333 | 0.026333 | 0.019 | 647.327 | 0.00033333 | 5.4262e-06
647.595 | 0.007 | 647.595 | 0.007 | 0.015 | 0.015 | 0.019 | 0.021 | 0.022 | 0.024 | 0.029 | 0.026 | 0.012333 | 0.020667 | 0.026333 | 0.019333 | 647.595 | 0.0013333 | 2.1705e-05
647.862 | -0.003 | 647.862 | -0.003 | -0.001 | 0.006 | 0.005 | 0.01 | 0.007 | 0.014 | 0.014 | 0.015 | 0.00066667 | 0.0073333 | 0.014333 | 0.0075 | 647.862 | -0.00016667 | -2.7131e-06
648.13 | 0.009 | 648.13 | 0.009 | 0.013 | 0.012 | 0.018 | 0.016 | 0.018 | 0.02 | 0.021 | 0.021 | 0.011333 | 0.017333 | 0.020667 | 0.016 | 648.13 | 0.0013333 | 2.1705e-05
648.398 | 0.017 | 648.398 | 0.017 | 0.018 | 0.022 | 0.032 | 0.03 | 0.029 | 0.029 | 0.037 | 0.039 | 0.019 | 0.030333 | 0.035 | 0.027 | 648.398 | 0.0033333 | 5.4262e-05
648.665 | 0.005 | 648.665 | 0.005 | 0.01 | 0.013 | 0.014 | 0.019 | 0.017 | 0.02 | 0.026 | 0.025 | 0.0093333 | 0.016667 | 0.023667 | 0.0165 | 648.665 | 0.00016667 | 2.7131e-06
648.933 | 0.023 | 648.933 | 0.023 | 0.029 | 0.031 | 0.031 | 0.041 | 0.039 | 0.04 | 0.043 | 0.046 | 0.027667 | 0.037 | 0.043 | 0.035333 | 648.933 | 0.0016667 | 2.7131e-05
649.2 | 0.011 | 649.2 | 0.011 | 0.015 | 0.018 | 0.021 | 0.022 | 0.027 | 0.03 | 0.031 | 0.032 | 0.014667 | 0.023333 | 0.031 | 0.022833 | 649.2 | 0.0005 | 8.1393e-06
649.468 | 0.01 | 649.468 | 0.01 | 0.015 | 0.02 | 0.017 | 0.025 | 0.023 | 0.024 | 0.026 | 0.031 | 0.015 | 0.021667 | 0.027 | 0.021 | 649.468 | 0.00066667 | 1.0852e-05
649.735 | 0.023 | 649.735 | 0.023 | 0.022 | 0.029 | 0.03 | 0.03 | 0.027 | 0.032 | 0.036 | 0.035 | 0.024667 | 0.029 | 0.034333 | 0.0295 | 649.735 | -0.0005 | -8.1393e-06
650.002 | 0.016 | 650.002 | 0.016 | 0.015 | 0.021 | 0.023 | 0.022 | 0.024 | 0.026 | 0.03 | 0.033 | 0.017333 | 0.023 | 0.029667 | 0.0235 | 650.002 | -0.0005 | -8.1393e-06
650.269 | 0.011 | 650.269 | 0.011 | 0.009 | 0.015 | 0.023 | 0.018 | 0.023 | 0.025 | 0.029 | 0.03 | 0.011667 | 0.021333 | 0.028 | 0.019833 | 650.269 | 0.0015 | 2.4418e-05
650.536 | 0.004 | 650.536 | 0.004 | 0.004 | 0.008 | 0.015 | 0.011 | 0.013 | 0.02 | 0.022 | 0.023 | 0.0053333 | 0.013 | 0.021667 | 0.0135 | 650.536 | -0.0005 | -8.1393e-06
650.804 | 0.007 | 650.804 | 0.007 | 0.007 | 0.008 | 0.015 | 0.015 | 0.017 | 0.02 | 0.022 | 0.019 | 0.0073333 | 0.015667 | 0.020333 | 0.013833 | 650.804 | 0.0018333 | 2.9844e-05
651.071 | 0.02 | 651.071 | 0.02 | 0.026 | 0.029 | 0.033 | 0.036 | 0.041 | 0.039 | 0.044 | 0.042 | 0.025 | 0.036667 | 0.041667 | 0.033333 | 651.071 | 0.0033333 | 5.4262e-05
651.338 | 0.014 | 651.338 | 0.014 | 0.015 | 0.014 | 0.021 | 0.018 | 0.019 | 0.021 | 0.024 | 0.024 | 0.014333 | 0.019333 | 0.023 | 0.018667 | 651.338 | 0.00066667 | 1.0852e-05
651.604 | 0.015 | 651.604 | 0.015 | 0.015 | 0.018 | 0.025 | 0.024 | 0.026 | 0.024 | 0.027 | 0.031 | 0.016 | 0.025 | 0.027333 | 0.021667 | 651.604 | 0.0033333 | 5.4262e-05
651.871 | 0.002 | 651.871 | 0.002 | 0.007 | 0.009 | 0.015 | 0.014 | 0.017 | 0.019 | 0.02 | 0.021 | 0.006 | 0.015333 | 0.02 | 0.013 | 651.871 | 0.0023333 | 3.7983e-05
652.138 | 0.008 | 652.138 | 0.008 | 0.014 | 0.013 | 0.018 | 0.021 | 0.024 | 0.021 | 0.028 | 0.026 | 0.011667 | 0.021 | 0.025 | 0.018333 | 652.138 | 0.0026667 | 4.3409e-05
652.405 | -0.002 | 652.405 | -0.002 | 0 | 0.002 | 0.009 | 0.007 | 0.007 | 0.011 | 0.013 | 0.014 | 0 | 0.0076667 | 0.012667 | 0.0063333 | 652.405 | 0.0013333 | 2.1705e-05
652.671 | 0.018 | 652.671 | 0.018 | 0.019 | 0.018 | 0.026 | 0.026 | 0.028 | 0.025 | 0.031 | 0.034 | 0.018333 | 0.026667 | 0.03 | 0.024167 | 652.671 | 0.0025 | 4.0696e-05
652.938 | 0.011 | 652.938 | 0.011 | 0.017 | 0.015 | 0.02 | 0.02 | 0.026 | 0.022 | 0.028 | 0.03 | 0.014333 | 0.022 | 0.026667 | 0.0205 | 652.938 | 0.0015 | 2.4418e-05
653.205 | 0.009 | 653.205 | 0.009 | 0.007 | 0.018 | 0.018 | 0.017 | 0.025 | 0.028 | 0.032 | 0.027 | 0.011333 | 0.02 | 0.029 | 0.020167 | 653.205 | -0.00016667 | -2.7131e-06
653.471 | 0.004 | 653.471 | 0.004 | 0.001 | 0.005 | 0.004 | 0.014 | 0.011 | 0.008 | 0.015 | 0.016 | 0.0033333 | 0.0096667 | 0.013 | 0.0081667 | 653.471 | 0.0015 | 2.4418e-05
653.737 | 0.011 | 653.737 | 0.011 | 0.011 | 0.019 | 0.02 | 0.017 | 0.027 | 0.023 | 0.03 | 0.028 | 0.013667 | 0.021333 | 0.027 | 0.020333 | 653.737 | 0.001 | 1.6279e-05
654.004 | 0.001 | 654.004 | 0.001 | 0.002 | 0.004 | 0.008 | 0.011 | 0.013 | 0.012 | 0.014 | 0.016 | 0.0023333 | 0.010667 | 0.014 | 0.0081667 | 654.004 | 0.0025 | 4.0696e-05
654.27 | 0.005 | 654.27 | 0.005 | 0.013 | 0.016 | 0.014 | 0.013 | 0.018 | 0.018 | 0.021 | 0.023 | 0.011333 | 0.015 | 0.020667 | 0.016 | 654.27 | -0.001 | -1.6279e-05
654.536 | 0.012 | 654.536 | 0.012 | 0.01 | 0.015 | 0.02 | 0.021 | 0.022 | 0.022 | 0.027 | 0.03 | 0.012333 | 0.021 | 0.026333 | 0.019333 | 654.536 | 0.0016667 | 2.7131e-05
654.803 | 0.005 | 654.803 | 0.005 | 0.006 | 0.009 | 0.013 | 0.014 | 0.017 | 0.02 | 0.023 | 0.024 | 0.0066667 | 0.014667 | 0.022333 | 0.0145 | 654.803 | 0.00016667 | 2.7131e-06
655.069 | 0.009 | 655.069 | 0.009 | 0.013 | 0.018 | 0.019 | 0.022 | 0.021 | 0.02 | 0.026 | 0.026 | 0.013333 | 0.020667 | 0.024 | 0.018667 | 655.069 | 0.002 | 3.2557e-05
655.335 | 0.011 | 655.335 | 0.011 | 0.012 | 0.013 | 0.02 | 0.016 | 0.018 | 0.023 | 0.023 | 0.023 | 0.012 | 0.018 | 0.023 | 0.0175 | 655.335 | 0.0005 | 8.1393e-06
655.601 | 0.017 | 655.601 | 0.017 | 0.019 | 0.021 | 0.024 | 0.026 | 0.027 | 0.025 | 0.034 | 0.03 | 0.019 | 0.025667 | 0.029667 | 0.024333 | 655.601 | 0.0013333 | 2.1705e-05
655.867 | 0.012 | 655.867 | 0.012 | 0.013 | 0.014 | 0.019 | 0.021 | 0.024 | 0.02 | 0.028 | 0.028 | 0.013 | 0.021333 | 0.025333 | 0.019167 | 655.867 | 0.0021667 | 3.527e-05
656.132 | 0.016 | 656.132 | 0.016 | 0.019 | 0.02 | 0.029 | 0.028 | 0.027 | 0.032 | 0.034 | 0.034 | 0.018333 | 0.028 | 0.033333 | 0.025833 | 656.132 | 0.0021667 | 3.527e-05
656.398 | 0.011 | 656.398 | 0.011 | 0.017 | 0.017 | 0.021 | 0.019 | 0.022 | 0.023 | 0.027 | 0.029 | 0.015 | 0.020667 | 0.026333 | 0.020667 | 656.398 | 0 | 0
656.664 | 0.014 | 656.664 | 0.014 | 0.012 | 0.018 | 0.023 | 0.021 | 0.023 | 0.024 | 0.03 | 0.024 | 0.014667 | 0.022333 | 0.026 | 0.020333 | 656.664 | 0.002 | 3.2557e-05
656.93 | 0.001 | 656.93 | 0.001 | 0.007 | 0.012 | 0.015 | 0.017 | 0.016 | 0.019 | 0.021 | 0.02 | 0.0066667 | 0.016 | 0.02 | 0.013333 | 656.93 | 0.0026667 | 4.3409e-05
657.195 | 0.004 | 657.195 | 0.004 | 0.01 | 0.011 | 0.018 | 0.011 | 0.014 | 0.018 | 0.02 | 0.021 | 0.0083333 | 0.014333 | 0.019667 | 0.014 | 657.195 | 0.00033333 | 5.4262e-06
657.461 | 0.014 | 657.461 | 0.014 | 0.014 | 0.017 | 0.019 | 0.017 | 0.021 | 0.023 | 0.025 | 0.027 | 0.015 | 0.019 | 0.025 | 0.02 | 657.461 | -0.001 | -1.6279e-05
657.726 | 0.012 | 657.726 | 0.012 | 0.014 | 0.014 | 0.017 | 0.023 | 0.021 | 0.024 | 0.025 | 0.032 | 0.013333 | 0.020333 | 0.027 | 0.020167 | 657.726 | 0.00016667 | 2.7131e-06
657.992 | 0.001 | 657.992 | 0.001 | 0 | 0.004 | 0.005 | 0.007 | 0.006 | 0.006 | 0.01 | 0.014 | 0.0016667 | 0.006 | 0.01 | 0.0058333 | 657.992 | 0.00016667 | 2.7131e-06
658.257 | 0.012 | 658.257 | 0.012 | 0.013 | 0.016 | 0.02 | 0.019 | 0.024 | 0.022 | 0.026 | 0.022 | 0.013667 | 0.021 | 0.023333 | 0.0185 | 658.257 | 0.0025 | 4.0696e-05
658.523 | 0.021 | 658.523 | 0.021 | 0.019 | 0.018 | 0.021 | 0.026 | 0.024 | 0.029 | 0.029 | 0.032 | 0.019333 | 0.023667 | 0.03 | 0.024667 | 658.523 | -0.001 | -1.6279e-05
658.788 | 0.001 | 658.788 | 0.001 | 0.004 | 0.004 | 0.009 | 0.01 | 0.014 | 0.015 | 0.017 | 0.017 | 0.003 | 0.011 | 0.016333 | 0.0096667 | 658.788 | 0.0013333 | 2.1705e-05
659.053 | 0.006 | 659.053 | 0.006 | 0.012 | 0.013 | 0.015 | 0.016 | 0.02 | 0.022 | 0.026 | 0.028 | 0.010333 | 0.017 | 0.025333 | 0.017833 | 659.053 | -0.00083333 | -1.3565e-05
659.318 | 0.004 | 659.318 | 0.004 | 0.007 | 0.012 | 0.015 | 0.017 | 0.023 | 0.021 | 0.023 | 0.025 | 0.0076667 | 0.018333 | 0.023 | 0.015333 | 659.318 | 0.003 | 4.8836e-05
659.583 | 0.009 | 659.583 | 0.009 | 0.013 | 0.016 | 0.019 | 0.017 | 0.019 | 0.018 | 0.02 | 0.023 | 0.012667 | 0.018333 | 0.020333 | 0.0165 | 659.583 | 0.0018333 | 2.9844e-05
659.848 | 0.002 | 659.848 | 0.002 | 0.006 | 0.008 | 0.014 | 0.012 | 0.012 | 0.015 | 0.02 | 0.021 | 0.0053333 | 0.012667 | 0.018667 | 0.012 | 659.848 | 0.00066667 | 1.0852e-05
660.113 | 0.012 | 660.113 | 0.012 | 0.014 | 0.015 | 0.02 | 0.022 | 0.027 | 0.024 | 0.026 | 0.027 | 0.013667 | 0.023 | 0.025667 | 0.019667 | 660.113 | 0.0033333 | 5.4262e-05
660.378 | 0.017 | 660.378 | 0.017 | 0.015 | 0.017 | 0.023 | 0.023 | 0.026 | 0.026 | 0.025 | 0.029 | 0.016333 | 0.024 | 0.026667 | 0.0215 | 660.378 | 0.0025 | 4.0696e-05
660.643 | 0.008 | 660.643 | 0.008 | 0.012 | 0.016 | 0.017 | 0.016 | 0.021 | 0.019 | 0.024 | 0.031 | 0.012 | 0.018 | 0.024667 | 0.018333 | 660.643 | -0.00033333 | -5.4262e-06
660.908 | -0.002 | 660.908 | -0.002 | 0 | 0.004 | 0.01 | 0.007 | 0.011 | 0.01 | 0.012 | 0.017 | 0.00066667 | 0.0093333 | 0.013 | 0.0068333 | 660.908 | 0.0025 | 4.0696e-05
661.172 | -0.008 | 661.172 | -0.008 | -0.001 | -0.001 | 0.002 | 0.004 | 0.002 | 0.012 | 0.011 | 0.008 | -0.0033333 | 0.0026667 | 0.010333 | 0.0035 | 661.172 | -0.00083333 | -1.3565e-05
661.437 | 0.003 | 661.437 | 0.003 | 0.008 | 0.007 | 0.014 | 0.013 | 0.016 | 0.017 | 0.021 | 0.024 | 0.006 | 0.014333 | 0.020667 | 0.013333 | 661.437 | 0.001 | 1.6279e-05
661.702 | 0.013 | 661.702 | 0.013 | 0.013 | 0.019 | 0.02 | 0.02 | 0.021 | 0.02 | 0.027 | 0.025 | 0.015 | 0.020333 | 0.024 | 0.0195 | 661.702 | 0.00083333 | 1.3565e-05
661.966 | 0.019 | 661.966 | 0.019 | 0.022 | 0.021 | 0.026 | 0.027 | 0.026 | 0.032 | 0.033 | 0.034 | 0.020667 | 0.026333 | 0.033 | 0.026833 | 661.966 | -0.0005 | -8.1393e-06
662.231 | 0.01 | 662.231 | 0.01 | 0.014 | 0.017 | 0.018 | 0.019 | 0.022 | 0.023 | 0.022 | 0.024 | 0.013667 | 0.019667 | 0.023 | 0.018333 | 662.231 | 0.0013333 | 2.1705e-05
662.495 | 0.015 | 662.495 | 0.015 | 0.013 | 0.016 | 0.022 | 0.023 | 0.023 | 0.025 | 0.033 | 0.027 | 0.014667 | 0.022667 | 0.028333 | 0.0215 | 662.495 | 0.0011667 | 1.8992e-05
662.76 | 0.029 | 662.76 | 0.029 | 0.033 | 0.033 | 0.035 | 0.035 | 0.04 | 0.043 | 0.044 | 0.052 | 0.031667 | 0.036667 | 0.046333 | 0.039 | 662.76 | -0.0023333 | -3.7983e-05
663.024 | 0.005 | 663.024 | 0.005 | 0.014 | 0.012 | 0.015 | 0.02 | 0.02 | 0.018 | 0.022 | 0.022 | 0.010333 | 0.018333 | 0.020667 | 0.0155 | 663.024 | 0.0028333 | 4.6123e-05
663.288 | 0.005 | 663.288 | 0.005 | 0.003 | 0.006 | 0.014 | 0.01 | 0.013 | 0.014 | 0.017 | 0.015 | 0.0046667 | 0.012333 | 0.015333 | 0.01 | 663.288 | 0.0023333 | 3.7983e-05
663.552 | 0 | 663.552 | 0 | 0.005 | 0.009 | 0.009 | 0.012 | 0.013 | 0.015 | 0.015 | 0.018 | 0.0046667 | 0.011333 | 0.016 | 0.010333 | 663.552 | 0.001 | 1.6279e-05
663.816 | 0.004 | 663.816 | 0.004 | 0.012 | 0.012 | 0.016 | 0.018 | 0.023 | 0.02 | 0.025 | 0.027 | 0.0093333 | 0.019 | 0.024 | 0.016667 | 663.816 | 0.0023333 | 3.7983e-05
664.08 | 0.022 | 664.08 | 0.022 | 0.019 | 0.024 | 0.03 | 0.027 | 0.033 | 0.039 | 0.034 | 0.03 | 0.021667 | 0.03 | 0.034333 | 0.028 | 664.08 | 0.002 | 3.2557e-05
664.344 | 0.013 | 664.344 | 0.013 | 0.014 | 0.015 | 0.022 | 0.017 | 0.018 | 0.023 | 0.025 | 0.028 | 0.014 | 0.019 | 0.025333 | 0.019667 | 664.344 | -0.00066667 | -1.0852e-05
664.608 | 0.01 | 664.608 | 0.01 | 0.015 | 0.017 | 0.018 | 0.025 | 0.026 | 0.025 | 0.03 | 0.027 | 0.014 | 0.023 | 0.027333 | 0.020667 | 664.608 | 0.0023333 | 3.7983e-05
664.872 | 0 | 664.872 | 0 | 0.003 | 0.007 | 0.013 | 0.015 | 0.015 | 0.016 | 0.021 | 0.022 | 0.0033333 | 0.014333 | 0.019667 | 0.0115 | 664.872 | 0.0028333 | 4.6123e-05
665.136 | 0.003 | 665.136 | 0.003 | 0.007 | 0.004 | 0.011 | 0.016 | 0.014 | 0.017 | 0.016 | 0.018 | 0.0046667 | 0.013667 | 0.017 | 0.010833 | 665.136 | 0.0028333 | 4.6123e-05
665.4 | 0.016 | 665.4 | 0.016 | 0.014 | 0.02 | 0.018 | 0.025 | 0.024 | 0.023 | 0.029 | 0.029 | 0.016667 | 0.022333 | 0.027 | 0.021833 | 665.4 | 0.0005 | 8.1393e-06
665.664 | 0.014 | 665.664 | 0.014 | 0.018 | 0.02 | 0.025 | 0.024 | 0.028 | 0.028 | 0.031 | 0.034 | 0.017333 | 0.025667 | 0.031 | 0.024167 | 665.664 | 0.0015 | 2.4418e-05
665.927 | 0.011 | 665.927 | 0.011 | 0.017 | 0.017 | 0.021 | 0.019 | 0.026 | 0.025 | 0.028 | 0.028 | 0.015 | 0.022 | 0.027 | 0.021 | 665.927 | 0.001 | 1.6279e-05
666.191 | 0.01 | 666.191 | 0.01 | 0.01 | 0.013 | 0.021 | 0.022 | 0.023 | 0.024 | 0.024 | 0.027 | 0.011 | 0.022 | 0.025 | 0.018 | 666.191 | 0.004 | 6.5114e-05
666.454 | 0.002 | 666.454 | 0.002 | 0.007 | 0.003 | 0.008 | 0.009 | 0.013 | 0.01 | 0.016 | 0.017 | 0.004 | 0.01 | 0.014333 | 0.0091667 | 666.454 | 0.00083333 | 1.3565e-05
666.718 | 0.008 | 666.718 | 0.008 | 0.008 | 0.011 | 0.016 | 0.016 | 0.021 | 0.023 | 0.022 | 0.027 | 0.009 | 0.017667 | 0.024 | 0.0165 | 666.718 | 0.0011667 | 1.8992e-05
666.981 | -0.004 | 666.981 | -0.004 | 0.003 | 0.001 | 0.005 | 0.006 | 0.009 | 0.01 | 0.013 | 0.011 | 0 | 0.0066667 | 0.011333 | 0.0056667 | 666.981 | 0.001 | 1.6279e-05
667.244 | 0.003 | 667.244 | 0.003 | 0.003 | 0.007 | 0.015 | 0.012 | 0.013 | 0.015 | 0.021 | 0.02 | 0.0043333 | 0.013333 | 0.018667 | 0.0115 | 667.244 | 0.0018333 | 2.9844e-05
667.508 | 0.021 | 667.508 | 0.021 | 0.024 | 0.023 | 0.03 | 0.033 | 0.031 | 0.032 | 0.039 | 0.038 | 0.022667 | 0.031333 | 0.036333 | 0.0295 | 667.508 | 0.0018333 | 2.9844e-05
667.771 | 0.005 | 667.771 | 0.005 | 0.011 | 0.014 | 0.017 | 0.017 | 0.019 | 0.017 | 0.024 | 0.023 | 0.01 | 0.017667 | 0.021333 | 0.015667 | 667.771 | 0.002 | 3.2557e-05
668.034 | 0.017 | 668.034 | 0.017 | 0.016 | 0.021 | 0.025 | 0.027 | 0.024 | 0.03 | 0.023 | 0.032 | 0.018 | 0.025333 | 0.028333 | 0.023167 | 668.034 | 0.0021667 | 3.527e-05
668.297 | 0.004 | 668.297 | 0.004 | 0.005 | 0.008 | 0.011 | 0.014 | 0.015 | 0.018 | 0.019 | 0.025 | 0.0056667 | 0.013333 | 0.020667 | 0.013167 | 668.297 | 0.00016667 | 2.7131e-06
668.56 | 0.012 | 668.56 | 0.012 | 0.016 | 0.015 | 0.018 | 0.022 | 0.022 | 0.027 | 0.028 | 0.029 | 0.014333 | 0.020667 | 0.028 | 0.021167 | 668.56 | -0.0005 | -8.1393e-06
668.823 | -0.009 | 668.823 | -0.009 | -0.005 | 0 | -0.001 | 0.007 | 0.007 | 0.008 | 0.013 | 0.01 | -0.0046667 | 0.0043333 | 0.010333 | 0.0028333 | 668.823 | 0.0015 | 2.4418e-05
669.086 | 0.018 | 669.086 | 0.018 | 0.023 | 0.021 | 0.031 | 0.029 | 0.033 | 0.033 | 0.036 | 0.031 | 0.020667 | 0.031 | 0.033333 | 0.027 | 669.086 | 0.004 | 6.5114e-05
669.349 | 0.004 | 669.349 | 0.004 | 0.009 | 0.005 | 0.016 | 0.014 | 0.015 | 0.017 | 0.019 | 0.024 | 0.006 | 0.015 | 0.02 | 0.013 | 669.349 | 0.002 | 3.2557e-05
669.612 | 0.01 | 669.612 | 0.01 | 0.01 | 0.016 | 0.011 | 0.017 | 0.02 | 0.018 | 0.022 | 0.022 | 0.012 | 0.016 | 0.020667 | 0.016333 | 669.612 | -0.00033333 | -5.4262e-06
669.875 | 0.003 | 669.875 | 0.003 | 0.003 | 0.008 | 0.013 | 0.012 | 0.015 | 0.011 | 0.016 | 0.017 | 0.0046667 | 0.013333 | 0.014667 | 0.0096667 | 669.875 | 0.0036667 | 5.9688e-05
670.137 | 0.014 | 670.137 | 0.014 | 0.017 | 0.017 | 0.022 | 0.026 | 0.026 | 0.027 | 0.031 | 0.031 | 0.016 | 0.024667 | 0.029667 | 0.022833 | 670.137 | 0.0018333 | 2.9844e-05
670.4 | 0.015 | 670.4 | 0.015 | 0.012 | 0.014 | 0.022 | 0.024 | 0.026 | 0.025 | 0.023 | 0.027 | 0.013667 | 0.024 | 0.025 | 0.019333 | 670.4 | 0.0046667 | 7.5967e-05
670.662 | 0.002 | 670.662 | 0.002 | 0.005 | 0.011 | 0.012 | 0.015 | 0.019 | 0.017 | 0.02 | 0.027 | 0.006 | 0.015333 | 0.021333 | 0.013667 | 670.662 | 0.0016667 | 2.7131e-05
670.925 | 0.014 | 670.925 | 0.014 | 0.021 | 0.024 | 0.025 | 0.02 | 0.024 | 0.025 | 0.026 | 0.029 | 0.019667 | 0.023 | 0.026667 | 0.023167 | 670.925 | -0.00016667 | -2.7131e-06
671.187 | 0.013 | 671.187 | 0.013 | 0.023 | 0.024 | 0.025 | 0.028 | 0.034 | 0.032 | 0.037 | 0.038 | 0.02 | 0.029 | 0.035667 | 0.027833 | 671.187 | 0.0011667 | 1.8992e-05
671.45 | -0.008 | 671.45 | -0.008 | -0.005 | -0.005 | 0.001 | 0.002 | 0.001 | 0.005 | 0.007 | 0.008 | -0.006 | 0.0013333 | 0.0066667 | 0.00033333 | 671.45 | 0.001 | 1.6279e-05
671.712 | 0.024 | 671.712 | 0.024 | 0.03 | 0.03 | 0.042 | 0.038 | 0.037 | 0.043 | 0.041 | 0.046 | 0.028 | 0.039 | 0.043333 | 0.035667 | 671.712 | 0.0033333 | 5.4262e-05
671.974 | 0.017 | 671.974 | 0.017 | 0.022 | 0.026 | 0.021 | 0.029 | 0.027 | 0.033 | 0.032 | 0.035 | 0.021667 | 0.025667 | 0.033333 | 0.0275 | 671.974 | -0.0018333 | -2.9844e-05
672.237 | 0.005 | 672.237 | 0.005 | 0.004 | 0.005 | 0.011 | 0.013 | 0.01 | 0.012 | 0.019 | 0.022 | 0.0046667 | 0.011333 | 0.017667 | 0.011167 | 672.237 | 0.00016667 | 2.7131e-06
672.499 | 0.009 | 672.499 | 0.009 | 0.005 | 0.009 | 0.017 | 0.017 | 0.023 | 0.02 | 0.028 | 0.024 | 0.0076667 | 0.019 | 0.024 | 0.015833 | 672.499 | 0.0031667 | 5.1549e-05
672.761 | 0.011 | 672.761 | 0.011 | 0.016 | 0.018 | 0.023 | 0.022 | 0.019 | 0.025 | 0.028 | 0.025 | 0.015 | 0.021333 | 0.026 | 0.0205 | 672.761 | 0.00083333 | 1.3565e-05
673.023 | 0.01 | 673.023 | 0.01 | 0.01 | 0.012 | 0.012 | 0.018 | 0.02 | 0.022 | 0.024 | 0.026 | 0.010667 | 0.016667 | 0.024 | 0.017333 | 673.023 | -0.00066667 | -1.0852e-05
673.285 | 0.014 | 673.285 | 0.014 | 0.013 | 0.015 | 0.018 | 0.023 | 0.025 | 0.025 | 0.027 | 0.031 | 0.014 | 0.022 | 0.027667 | 0.020833 | 673.285 | 0.0011667 | 1.8992e-05
673.547 | 0.014 | 673.547 | 0.014 | 0.018 | 0.023 | 0.022 | 0.026 | 0.026 | 0.028 | 0.03 | 0.033 | 0.018333 | 0.024667 | 0.030333 | 0.024333 | 673.547 | 0.00033333 | 5.4262e-06
673.808 | 0.005 | 673.808 | 0.005 | 0.009 | 0.009 | 0.019 | 0.022 | 0.025 | 0.024 | 0.024 | 0.025 | 0.0076667 | 0.022 | 0.024333 | 0.016 | 673.808 | 0.006 | 9.7671e-05
674.07 | 0.004 | 674.07 | 0.004 | 0.008 | 0.009 | 0.014 | 0.012 | 0.016 | 0.017 | 0.026 | 0.023 | 0.007 | 0.014 | 0.022 | 0.0145 | 674.07 | -0.0005 | -8.1393e-06
674.332 | 0.014 | 674.332 | 0.014 | 0.018 | 0.015 | 0.02 | 0.023 | 0.023 | 0.025 | 0.026 | 0.029 | 0.015667 | 0.022 | 0.026667 | 0.021167 | 674.332 | 0.00083333 | 1.3565e-05
674.594 | 0.007 | 674.594 | 0.007 | 0.011 | 0.008 | 0.013 | 0.018 | 0.019 | 0.018 | 0.018 | 0.024 | 0.0086667 | 0.016667 | 0.02 | 0.014333 | 674.594 | 0.0023333 | 3.7983e-05
674.855 | 0.009 | 674.855 | 0.009 | 0.014 | 0.013 | 0.018 | 0.019 | 0.023 | 0.024 | 0.027 | 0.022 | 0.012 | 0.02 | 0.024333 | 0.018167 | 674.855 | 0.0018333 | 2.9844e-05
675.117 | 0 | 675.117 | 0 | 0.006 | 0.009 | 0.008 | 0.015 | 0.015 | 0.014 | 0.018 | 0.018 | 0.005 | 0.012667 | 0.016667 | 0.010833 | 675.117 | 0.0018333 | 2.9844e-05
675.378 | 0.016 | 675.378 | 0.016 | 0.02 | 0.024 | 0.026 | 0.028 | 0.028 | 0.028 | 0.034 | 0.037 | 0.02 | 0.027333 | 0.033 | 0.0265 | 675.378 | 0.00083333 | 1.3565e-05
675.64 | 0.007 | 675.64 | 0.007 | 0.009 | 0.009 | 0.014 | 0.013 | 0.018 | 0.017 | 0.021 | 0.024 | 0.0083333 | 0.015 | 0.020667 | 0.0145 | 675.64 | 0.0005 | 8.1393e-06
675.901 | -0.006 | 675.901 | -0.006 | -0.005 | -0.003 | 0 | 0.008 | 0.003 | 0.008 | 0.004 | 0.007 | -0.0046667 | 0.0036667 | 0.0063333 | 0.00083333 | 675.901 | 0.0028333 | 4.6123e-05
676.162 | 0.013 | 676.162 | 0.013 | 0.017 | 0.017 | 0.021 | 0.024 | 0.026 | 0.027 | 0.028 | 0.03 | 0.015667 | 0.023667 | 0.028333 | 0.022 | 676.162 | 0.0016667 | 2.7131e-05
676.424 | 0.002 | 676.424 | 0.002 | 0.005 | 0.004 | 0.012 | 0.017 | 0.02 | 0.023 | 0.02 | 0.023 | 0.0036667 | 0.016333 | 0.022 | 0.012833 | 676.424 | 0.0035 | 5.6975e-05
676.685 | 0.016 | 676.685 | 0.016 | 0.017 | 0.024 | 0.026 | 0.028 | 0.027 | 0.029 | 0.031 | 0.035 | 0.019 | 0.027 | 0.031667 | 0.025333 | 676.685 | 0.0016667 | 2.7131e-05
676.946 | 0.016 | 676.946 | 0.016 | 0.02 | 0.022 | 0.022 | 0.025 | 0.027 | 0.026 | 0.029 | 0.031 | 0.019333 | 0.024667 | 0.028667 | 0.024 | 676.946 | 0.00066667 | 1.0852e-05
677.207 | 0.024 | 677.207 | 0.024 | 0.026 | 0.027 | 0.031 | 0.033 | 0.038 | 0.037 | 0.038 | 0.039 | 0.025667 | 0.034 | 0.038 | 0.031833 | 677.207 | 0.0021667 | 3.527e-05
677.468 | 0.017 | 677.468 | 0.017 | 0.026 | 0.029 | 0.037 | 0.031 | 0.032 | 0.038 | 0.039 | 0.042 | 0.024 | 0.033333 | 0.039667 | 0.031833 | 677.468 | 0.0015 | 2.4418e-05
677.729 | 0.001 | 677.729 | 0.001 | 0.006 | 0.006 | 0.008 | 0.012 | 0.014 | 0.017 | 0.015 | 0.016 | 0.0043333 | 0.011333 | 0.016 | 0.010167 | 677.729 | 0.0011667 | 1.8992e-05
677.99 | -0.001 | 677.99 | -0.001 | 0.003 | 0.007 | 0.01 | 0.01 | 0.01 | 0.012 | 0.018 | 0.016 | 0.003 | 0.01 | 0.015333 | 0.0091667 | 677.99 | 0.00083333 | 1.3565e-05
678.251 | 0.013 | 678.251 | 0.013 | 0.018 | 0.019 | 0.021 | 0.028 | 0.029 | 0.028 | 0.03 | 0.036 | 0.016667 | 0.026 | 0.031333 | 0.024 | 678.251 | 0.002 | 3.2557e-05
678.512 | 0.011 | 678.512 | 0.011 | 0.015 | 0.014 | 0.018 | 0.025 | 0.02 | 0.029 | 0.027 | 0.029 | 0.013333 | 0.021 | 0.028333 | 0.020833 | 678.512 | 0.00016667 | 2.7131e-06
678.772 | 0.008 | 678.772 | 0.008 | 0.014 | 0.019 | 0.021 | 0.019 | 0.03 | 0.027 | 0.029 | 0.035 | 0.013667 | 0.023333 | 0.030333 | 0.022 | 678.772 | 0.0013333 | 2.1705e-05
679.033 | 0.024 | 679.033 | 0.024 | 0.026 | 0.03 | 0.034 | 0.033 | 0.037 | 0.038 | 0.041 | 0.042 | 0.026667 | 0.034667 | 0.040333 | 0.0335 | 679.033 | 0.0011667 | 1.8992e-05
679.294 | 0.011 | 679.294 | 0.011 | 0.01 | 0.009 | 0.02 | 0.018 | 0.02 | 0.022 | 0.024 | 0.025 | 0.01 | 0.019333 | 0.023667 | 0.016833 | 679.294 | 0.0025 | 4.0696e-05
679.554 | 0.009 | 679.554 | 0.009 | 0.017 | 0.016 | 0.021 | 0.022 | 0.026 | 0.027 | 0.028 | 0.027 | 0.014 | 0.023 | 0.027333 | 0.020667 | 679.554 | 0.0023333 | 3.7983e-05
679.815 | 0.01 | 679.815 | 0.01 | 0.016 | 0.013 | 0.018 | 0.026 | 0.023 | 0.025 | 0.028 | 0.032 | 0.013 | 0.022333 | 0.028333 | 0.020667 | 679.815 | 0.0016667 | 2.7131e-05
680.075 | 0.007 | 680.075 | 0.007 | 0.01 | 0.012 | 0.012 | 0.014 | 0.017 | 0.018 | 0.02 | 0.021 | 0.0096667 | 0.014333 | 0.019667 | 0.014667 | 680.075 | -0.00033333 | -5.4262e-06
680.335 | 0.019 | 680.335 | 0.019 | 0.016 | 0.022 | 0.022 | 0.025 | 0.025 | 0.028 | 0.03 | 0.034 | 0.019 | 0.024 | 0.030667 | 0.024833 | 680.335 | -0.00083333 | -1.3565e-05
680.596 | 0 | 680.596 | 0 | 0.004 | 0.005 | 0.008 | 0.015 | 0.012 | 0.016 | 0.022 | 0.014 | 0.003 | 0.011667 | 0.017333 | 0.010167 | 680.596 | 0.0015 | 2.4418e-05
680.856 | 0.013 | 680.856 | 0.013 | 0.016 | 0.02 | 0.024 | 0.025 | 0.027 | 0.026 | 0.026 | 0.035 | 0.016333 | 0.025333 | 0.029 | 0.022667 | 680.856 | 0.0026667 | 4.3409e-05
681.116 | 0.032 | 681.116 | 0.032 | 0.032 | 0.035 | 0.036 | 0.042 | 0.042 | 0.039 | 0.046 | 0.045 | 0.033 | 0.04 | 0.043333 | 0.038167 | 681.116 | 0.0018333 | 2.9844e-05
681.376 | 0.015 | 681.376 | 0.015 | 0.019 | 0.018 | 0.022 | 0.028 | 0.03 | 0.029 | 0.032 | 0.034 | 0.017333 | 0.026667 | 0.031667 | 0.0245 | 681.376 | 0.0021667 | 3.527e-05
681.636 | 0.01 | 681.636 | 0.01 | 0.008 | 0.01 | 0.018 | 0.02 | 0.021 | 0.022 | 0.025 | 0.027 | 0.0093333 | 0.019667 | 0.024667 | 0.017 | 681.636 | 0.0026667 | 4.3409e-05
681.896 | 0.021 | 681.896 | 0.021 | 0.024 | 0.024 | 0.029 | 0.033 | 0.037 | 0.039 | 0.04 | 0.039 | 0.023 | 0.033 | 0.039333 | 0.031167 | 681.896 | 0.0018333 | 2.9844e-05
682.156 | 0.016 | 682.156 | 0.016 | 0.019 | 0.027 | 0.029 | 0.031 | 0.032 | 0.037 | 0.034 | 0.036 | 0.020667 | 0.030667 | 0.035667 | 0.028167 | 682.156 | 0.0025 | 4.0696e-05
682.416 | 0.008 | 682.416 | 0.008 | 0.013 | 0.015 | 0.017 | 0.019 | 0.018 | 0.024 | 0.025 | 0.03 | 0.012 | 0.018 | 0.026333 | 0.019167 | 682.416 | -0.0011667 | -1.8992e-05
682.676 | 0.005 | 682.676 | 0.005 | 0.009 | 0.012 | 0.017 | 0.019 | 0.02 | 0.023 | 0.023 | 0.026 | 0.0086667 | 0.018667 | 0.024 | 0.016333 | 682.676 | 0.0023333 | 3.7983e-05
682.936 | 0.022 | 682.936 | 0.022 | 0.024 | 0.023 | 0.029 | 0.031 | 0.031 | 0.035 | 0.033 | 0.034 | 0.023 | 0.030333 | 0.034 | 0.0285 | 682.936 | 0.0018333 | 2.9844e-05
683.196 | 0.012 | 683.196 | 0.012 | 0.012 | 0.017 | 0.017 | 0.023 | 0.022 | 0.023 | 0.027 | 0.027 | 0.013667 | 0.020667 | 0.025667 | 0.019667 | 683.196 | 0.001 | 1.6279e-05
683.455 | 0.008 | 683.455 | 0.008 | 0.013 | 0.014 | 0.017 | 0.021 | 0.023 | 0.023 | 0.021 | 0.029 | 0.011667 | 0.020333 | 0.024333 | 0.018 | 683.455 | 0.0023333 | 3.7983e-05
683.715 | 0.009 | 683.715 | 0.009 | 0.013 | 0.018 | 0.02 | 0.02 | 0.019 | 0.025 | 0.024 | 0.029 | 0.013333 | 0.019667 | 0.026 | 0.019667 | 683.715 | 0 | 0
683.974 | 0.014 | 683.974 | 0.014 | 0.017 | 0.017 | 0.022 | 0.027 | 0.024 | 0.028 | 0.027 | 0.028 | 0.016 | 0.024333 | 0.027667 | 0.021833 | 683.974 | 0.0025 | 4.0696e-05
684.234 | 0.009 | 684.234 | 0.009 | 0.009 | 0.012 | 0.017 | 0.017 | 0.02 | 0.023 | 0.021 | 0.025 | 0.01 | 0.018 | 0.023 | 0.0165 | 684.234 | 0.0015 | 2.4418e-05
684.493 | -0.001 | 684.493 | -0.001 | 0 | 0.008 | 0.012 | 0.008 | 0.013 | 0.011 | 0.014 | 0.017 | 0.0023333 | 0.011 | 0.014 | 0.0081667 | 684.493 | 0.0028333 | 4.6123e-05
684.753 | 0.006 | 684.753 | 0.006 | 0.014 | 0.021 | 0.025 | 0.026 | 0.026 | 0.024 | 0.031 | 0.031 | 0.013667 | 0.025667 | 0.028667 | 0.021167 | 684.753 | 0.0045 | 7.3254e-05
685.012 | 0.018 | 685.012 | 0.018 | 0.022 | 0.029 | 0.029 | 0.035 | 0.031 | 0.034 | 0.035 | 0.039 | 0.023 | 0.031667 | 0.036 | 0.0295 | 685.012 | 0.0021667 | 3.527e-05
685.271 | 0.007 | 685.271 | 0.007 | 0.007 | 0.011 | 0.017 | 0.012 | 0.015 | 0.017 | 0.025 | 0.022 | 0.0083333 | 0.014667 | 0.021333 | 0.014833 | 685.271 | -0.00016667 | -2.7131e-06
685.53 | 0.01 | 685.53 | 0.01 | 0.007 | 0.013 | 0.02 | 0.021 | 0.02 | 0.025 | 0.026 | 0.029 | 0.01 | 0.020333 | 0.026667 | 0.018333 | 685.53 | 0.002 | 3.2557e-05
685.789 | 0.025 | 685.789 | 0.025 | 0.025 | 0.024 | 0.028 | 0.034 | 0.037 | 0.035 | 0.035 | 0.036 | 0.024667 | 0.033 | 0.035333 | 0.03 | 685.789 | 0.003 | 4.8836e-05
686.049 | 0.015 | 686.049 | 0.015 | 0.015 | 0.02 | 0.021 | 0.026 | 0.027 | 0.029 | 0.031 | 0.039 | 0.016667 | 0.024667 | 0.033 | 0.024833 | 686.049 | -0.00016667 | -2.7131e-06
686.308 | 0.001 | 686.308 | 0.001 | 0.01 | 0.008 | 0.015 | 0.019 | 0.017 | 0.023 | 0.023 | 0.022 | 0.0063333 | 0.017 | 0.022667 | 0.0145 | 686.308 | 0.0025 | 4.0696e-05
686.567 | 0.013 | 686.567 | 0.013 | 0.016 | 0.019 | 0.02 | 0.018 | 0.023 | 0.024 | 0.027 | 0.026 | 0.016 | 0.020333 | 0.025667 | 0.020833 | 686.567 | -0.0005 | -8.1393e-06
686.825 | 0.021 | 686.825 | 0.021 | 0.022 | 0.023 | 0.027 | 0.033 | 0.034 | 0.038 | 0.041 | 0.04 | 0.022 | 0.031333 | 0.039667 | 0.030833 | 686.825 | 0.0005 | 8.1393e-06
687.084 | 0.013 | 687.084 | 0.013 | 0.017 | 0.016 | 0.026 | 0.025 | 0.026 | 0.028 | 0.035 | 0.031 | 0.015333 | 0.025667 | 0.031333 | 0.023333 | 687.084 | 0.0023333 | 3.7983e-05
687.343 | 0.006 | 687.343 | 0.006 | 0.016 | 0.018 | 0.021 | 0.023 | 0.026 | 0.032 | 0.036 | 0.037 | 0.013333 | 0.023333 | 0.035 | 0.024167 | 687.343 | -0.00083333 | -1.3565e-05
687.602 | 0.015 | 687.602 | 0.015 | 0.026 | 0.021 | 0.025 | 0.028 | 0.035 | 0.033 | 0.036 | 0.037 | 0.020667 | 0.029333 | 0.035333 | 0.028 | 687.602 | 0.0013333 | 2.1705e-05
687.86 | 0.015 | 687.86 | 0.015 | 0.015 | 0.026 | 0.026 | 0.031 | 0.03 | 0.033 | 0.029 | 0.029 | 0.018667 | 0.029 | 0.030333 | 0.0245 | 687.86 | 0.0045 | 7.3254e-05
688.119 | 0.008 | 688.119 | 0.008 | 0.016 | 0.021 | 0.022 | 0.027 | 0.024 | 0.026 | 0.026 | 0.035 | 0.015 | 0.024333 | 0.029 | 0.022 | 688.119 | 0.0023333 | 3.7983e-05
688.378 | 0.02 | 688.378 | 0.02 | 0.022 | 0.023 | 0.028 | 0.032 | 0.033 | 0.036 | 0.038 | 0.04 | 0.021667 | 0.031 | 0.038 | 0.029833 | 688.378 | 0.0011667 | 1.8992e-05
688.636 | 0.018 | 688.636 | 0.018 | 0.021 | 0.017 | 0.029 | 0.028 | 0.029 | 0.039 | 0.029 | 0.038 | 0.018667 | 0.028667 | 0.035333 | 0.027 | 688.636 | 0.0016667 | 2.7131e-05
688.894 | 0.013 | 688.894 | 0.013 | 0.009 | 0.013 | 0.017 | 0.022 | 0.025 | 0.028 | 0.028 | 0.031 | 0.011667 | 0.021333 | 0.029 | 0.020333 | 688.894 | 0.001 | 1.6279e-05
689.153 | 0.01 | 689.153 | 0.01 | 0.014 | 0.019 | 0.023 | 0.026 | 0.025 | 0.024 | 0.028 | 0.032 | 0.014333 | 0.024667 | 0.028 | 0.021167 | 689.153 | 0.0035 | 5.6975e-05
689.411 | 0.019 | 689.411 | 0.019 | 0.024 | 0.022 | 0.031 | 0.033 | 0.036 | 0.038 | 0.035 | 0.039 | 0.021667 | 0.033333 | 0.037333 | 0.0295 | 689.411 | 0.0038333 | 6.2401e-05
689.669 | 0.015 | 689.669 | 0.015 | 0.016 | 0.013 | 0.023 | 0.025 | 0.024 | 0.027 | 0.029 | 0.034 | 0.014667 | 0.024 | 0.03 | 0.022333 | 689.669 | 0.0016667 | 2.7131e-05
689.928 | 0.022 | 689.928 | 0.022 | 0.028 | 0.028 | 0.03 | 0.027 | 0.038 | 0.033 | 0.037 | 0.035 | 0.026 | 0.031667 | 0.035 | 0.0305 | 689.928 | 0.0011667 | 1.8992e-05
690.186 | -0.004 | 690.186 | -0.004 | -0.005 | 0.004 | 0.008 | 0.008 | 0.012 | 0.008 | 0.013 | 0.015 | -0.0016667 | 0.0093333 | 0.012 | 0.0051667 | 690.186 | 0.0041667 | 6.7827e-05
690.444 | 0.02 | 690.444 | 0.02 | 0.025 | 0.019 | 0.042 | 0.041 | 0.043 | 0.036 | 0.039 | 0.037 | 0.021333 | 0.042 | 0.037333 | 0.029333 | 690.444 | 0.012667 | 0.0002062
690.702 | 0.016 | 690.702 | 0.016 | 0.02 | 0.022 | 0.04 | 0.043 | 0.045 | 0.029 | 0.035 | 0.031 | 0.019333 | 0.042667 | 0.031667 | 0.0255 | 690.702 | 0.017167 | 0.00027945
690.96 | 0.022 | 690.96 | 0.022 | 0.031 | 0.027 | 0.034 | 0.041 | 0.041 | 0.04 | 0.04 | 0.04 | 0.026667 | 0.038667 | 0.04 | 0.033333 | 690.96 | 0.0053333 | 8.6819e-05
691.218 | 0.003 | 691.218 | 0.003 | 0.007 | 0.012 | 0.018 | 0.019 | 0.017 | 0.02 | 0.02 | 0.029 | 0.0073333 | 0.018 | 0.023 | 0.015167 | 691.218 | 0.0028333 | 4.6123e-05
691.476 | 0.016 | 691.476 | 0.016 | 0.019 | 0.027 | 0.025 | 0.028 | 0.032 | 0.031 | 0.037 | 0.035 | 0.020667 | 0.028333 | 0.034333 | 0.0275 | 691.476 | 0.00083333 | 1.3565e-05
691.733 | 0.016 | 691.733 | 0.016 | 0.019 | 0.021 | 0.025 | 0.03 | 0.031 | 0.031 | 0.03 | 0.034 | 0.018667 | 0.028667 | 0.031667 | 0.025167 | 691.733 | 0.0035 | 5.6975e-05
691.991 | 0.015 | 691.991 | 0.015 | 0.016 | 0.018 | 0.03 | 0.025 | 0.028 | 0.029 | 0.031 | 0.03 | 0.016333 | 0.027667 | 0.03 | 0.023167 | 691.991 | 0.0045 | 7.3254e-05
692.249 | 0.004 | 692.249 | 0.004 | 0.007 | 0.014 | 0.016 | 0.016 | 0.02 | 0.022 | 0.021 | 0.026 | 0.0083333 | 0.017333 | 0.023 | 0.015667 | 692.249 | 0.0016667 | 2.7131e-05
692.506 | 0.011 | 692.506 | 0.011 | 0.012 | 0.013 | 0.019 | 0.024 | 0.021 | 0.023 | 0.026 | 0.032 | 0.012 | 0.021333 | 0.027 | 0.0195 | 692.506 | 0.0018333 | 2.9844e-05
692.764 | 0.019 | 692.764 | 0.019 | 0.022 | 0.022 | 0.022 | 0.031 | 0.025 | 0.032 | 0.032 | 0.035 | 0.021 | 0.026 | 0.033 | 0.027 | 692.764 | -0.001 | -1.6279e-05
693.021 | 0.02 | 693.021 | 0.02 | 0.02 | 0.019 | 0.03 | 0.033 | 0.038 | 0.041 | 0.039 | 0.045 | 0.019667 | 0.033667 | 0.041667 | 0.030667 | 693.021 | 0.003 | 4.8836e-05
693.279 | 0.016 | 693.279 | 0.016 | 0.016 | 0.025 | 0.026 | 0.026 | 0.028 | 0.036 | 0.04 | 0.035 | 0.019 | 0.026667 | 0.037 | 0.028 | 693.279 | -0.0013333 | -2.1705e-05
693.536 | 0.018 | 693.536 | 0.018 | 0.02 | 0.022 | 0.029 | 0.032 | 0.027 | 0.034 | 0.034 | 0.035 | 0.02 | 0.029333 | 0.034333 | 0.027167 | 693.536 | 0.0021667 | 3.527e-05
693.794 | -0.007 | 693.794 | -0.007 | 0.001 | 0.001 | 0.002 | 0.013 | 0.008 | 0.009 | 0.014 | 0.014 | -0.0016667 | 0.0076667 | 0.012333 | 0.0053333 | 693.794 | 0.0023333 | 3.7983e-05
694.051 | 0.007 | 694.051 | 0.007 | 0.012 | 0.009 | 0.017 | 0.02 | 0.024 | 0.023 | 0.024 | 0.025 | 0.0093333 | 0.020333 | 0.024 | 0.016667 | 694.051 | 0.0036667 | 5.9688e-05
694.308 | 0.013 | 694.308 | 0.013 | 0.018 | 0.016 | 0.021 | 0.026 | 0.033 | 0.031 | 0.034 | 0.028 | 0.015667 | 0.026667 | 0.031 | 0.023333 | 694.308 | 0.0033333 | 5.4262e-05
694.565 | 0.018 | 694.565 | 0.018 | 0.026 | 0.03 | 0.031 | 0.03 | 0.035 | 0.038 | 0.042 | 0.044 | 0.024667 | 0.032 | 0.041333 | 0.033 | 694.565 | -0.001 | -1.6279e-05
694.822 | 0.025 | 694.822 | 0.025 | 0.031 | 0.031 | 0.036 | 0.039 | 0.044 | 0.039 | 0.045 | 0.046 | 0.029 | 0.039667 | 0.043333 | 0.036167 | 694.822 | 0.0035 | 5.6975e-05
695.079 | 0.007 | 695.079 | 0.007 | 0.003 | 0.009 | 0.009 | 0.013 | 0.017 | 0.015 | 0.018 | 0.017 | 0.0063333 | 0.013 | 0.016667 | 0.0115 | 695.079 | 0.0015 | 2.4418e-05
695.336 | 0.015 | 695.336 | 0.015 | 0.015 | 0.015 | 0.02 | 0.024 | 0.026 | 0.03 | 0.034 | 0.032 | 0.015 | 0.023333 | 0.032 | 0.0235 | 695.336 | -0.00016667 | -2.7131e-06
695.593 | 0 | 695.593 | 0 | 0.004 | 0.008 | 0.013 | 0.017 | 0.016 | 0.019 | 0.025 | 0.024 | 0.004 | 0.015333 | 0.022667 | 0.013333 | 695.593 | 0.002 | 3.2557e-05
695.85 | 0.005 | 695.85 | 0.005 | 0.011 | 0.014 | 0.018 | 0.015 | 0.021 | 0.029 | 0.029 | 0.03 | 0.01 | 0.018 | 0.029333 | 0.019667 | 695.85 | -0.0016667 | -2.7131e-05
696.107 | 0.008 | 696.107 | 0.008 | 0.012 | 0.01 | 0.014 | 0.017 | 0.022 | 0.022 | 0.022 | 0.026 | 0.01 | 0.017667 | 0.023333 | 0.016667 | 696.107 | 0.001 | 1.6279e-05
696.364 | 0.013 | 696.364 | 0.013 | 0.015 | 0.015 | 0.023 | 0.021 | 0.023 | 0.027 | 0.028 | 0.03 | 0.014333 | 0.022333 | 0.028333 | 0.021333 | 696.364 | 0.001 | 1.6279e-05
696.62 | 0.012 | 696.62 | 0.012 | 0.01 | 0.02 | 0.033 | 0.029 | 0.03 | 0.034 | 0.027 | 0.031 | 0.014 | 0.030667 | 0.030667 | 0.022333 | 696.62 | 0.0083333 | 0.00013565
696.877 | 0.013 | 696.877 | 0.013 | 0.017 | 0.021 | 0.026 | 0.029 | 0.031 | 0.031 | 0.035 | 0.04 | 0.017 | 0.028667 | 0.035333 | 0.026167 | 696.877 | 0.0025 | 4.0696e-05
697.133 | 0.005 | 697.133 | 0.005 | 0.006 | 0.013 | 0.011 | 0.014 | 0.011 | 0.023 | 0.023 | 0.025 | 0.008 | 0.012 | 0.023667 | 0.015833 | 697.133 | -0.0038333 | -6.2401e-05
697.39 | 0.015 | 697.39 | 0.015 | 0.015 | 0.021 | 0.024 | 0.028 | 0.035 | 0.037 | 0.042 | 0.039 | 0.017 | 0.029 | 0.039333 | 0.028167 | 697.39 | 0.00083333 | 1.3565e-05
697.646 | 0.017 | 697.646 | 0.017 | 0.022 | 0.019 | 0.023 | 0.029 | 0.031 | 0.035 | 0.036 | 0.031 | 0.019333 | 0.027667 | 0.034 | 0.026667 | 697.646 | 0.001 | 1.6279e-05
697.903 | 0.004 | 697.903 | 0.004 | 0.009 | 0.01 | 0.013 | 0.017 | 0.019 | 0.022 | 0.025 | 0.024 | 0.0076667 | 0.016333 | 0.023667 | 0.015667 | 697.903 | 0.00066667 | 1.0852e-05
698.159 | -0.002 | 698.159 | -0.002 | 0 | 0.004 | 0.008 | 0.006 | 0.011 | 0.015 | 0.012 | 0.014 | 0.00066667 | 0.0083333 | 0.013667 | 0.0071667 | 698.159 | 0.0011667 | 1.8992e-05
698.415 | 0.004 | 698.415 | 0.004 | 0.008 | 0.008 | 0.017 | 0.016 | 0.02 | 0.02 | 0.025 | 0.022 | 0.0066667 | 0.017667 | 0.022333 | 0.0145 | 698.415 | 0.0031667 | 5.1549e-05
698.672 | 0.018 | 698.672 | 0.018 | 0.025 | 0.023 | 0.027 | 0.031 | 0.03 | 0.035 | 0.036 | 0.043 | 0.022 | 0.029333 | 0.038 | 0.03 | 698.672 | -0.00066667 | -1.0852e-05
698.928 | 0.01 | 698.928 | 0.01 | 0.012 | 0.016 | 0.019 | 0.018 | 0.023 | 0.028 | 0.025 | 0.029 | 0.012667 | 0.02 | 0.027333 | 0.02 | 698.928 | 0 | 0
699.184 | 0.017 | 699.184 | 0.017 | 0.02 | 0.021 | 0.027 | 0.031 | 0.035 | 0.03 | 0.04 | 0.036 | 0.019333 | 0.031 | 0.035333 | 0.027333 | 699.184 | 0.0036667 | 5.9688e-05
699.44 | 0.008 | 699.44 | 0.008 | 0.014 | 0.016 | 0.02 | 0.02 | 0.025 | 0.028 | 0.024 | 0.026 | 0.012667 | 0.021667 | 0.026 | 0.019333 | 699.44 | 0.0023333 | 3.7983e-05
699.696 | 0.019 | 699.696 | 0.019 | 0.026 | 0.024 | 0.037 | 0.03 | 0.037 | 0.038 | 0.043 | 0.045 | 0.023 | 0.034667 | 0.042 | 0.0325 | 699.696 | 0.0021667 | 3.527e-05
699.952 | -0.004 | 699.952 | -0.004 | -0.002 | 0.003 | 0.001 | 0.008 | 0.006 | 0.011 | 0.015 | 0.012 | -0.001 | 0.005 | 0.012667 | 0.0058333 | 699.952 | -0.00083333 | -1.3565e-05
700.208 | 0.009 | 700.208 | 0.009 | 0.013 | 0.016 | 0.016 | 0.017 | 0.027 | 0.025 | 0.025 | 0.031 | 0.012667 | 0.02 | 0.027 | 0.019833 | 700.208 | 0.00016667 | 2.7131e-06
700.464 | 0.006 | 700.464 | 0.006 | 0.009 | 0.012 | 0.015 | 0.016 | 0.022 | 0.024 | 0.023 | 0.024 | 0.009 | 0.017667 | 0.023667 | 0.016333 | 700.464 | 0.0013333 | 2.1705e-05
700.719 | 0.026 | 700.719 | 0.026 | 0.024 | 0.032 | 0.04 | 0.039 | 0.043 | 0.044 | 0.042 | 0.048 | 0.027333 | 0.040667 | 0.044667 | 0.036 | 700.719 | 0.0046667 | 7.5967e-05
700.975 | 0.014 | 700.975 | 0.014 | 0.02 | 0.022 | 0.025 | 0.02 | 0.028 | 0.037 | 0.04 | 0.032 | 0.018667 | 0.024333 | 0.036333 | 0.0275 | 700.975 | -0.0031667 | -5.1549e-05
701.231 | 0.012 | 701.231 | 0.012 | 0.016 | 0.019 | 0.029 | 0.025 | 0.026 | 0.036 | 0.037 | 0.034 | 0.015667 | 0.026667 | 0.035667 | 0.025667 | 701.231 | 0.001 | 1.6279e-05
701.486 | 0.024 | 701.486 | 0.024 | 0.025 | 0.027 | 0.033 | 0.038 | 0.034 | 0.041 | 0.04 | 0.046 | 0.025333 | 0.035 | 0.042333 | 0.033833 | 701.486 | 0.0011667 | 1.8992e-05
701.742 | 0.011 | 701.742 | 0.011 | 0.012 | 0.014 | 0.021 | 0.027 | 0.028 | 0.027 | 0.034 | 0.035 | 0.012333 | 0.025333 | 0.032 | 0.022167 | 701.742 | 0.0031667 | 5.1549e-05
701.997 | 0.012 | 701.997 | 0.012 | 0.02 | 0.018 | 0.022 | 0.028 | 0.027 | 0.034 | 0.037 | 0.042 | 0.016667 | 0.025667 | 0.037667 | 0.027167 | 701.997 | -0.0015 | -2.4418e-05
702.253 | 0.013 | 702.253 | 0.013 | 0.02 | 0.022 | 0.025 | 0.025 | 0.034 | 0.03 | 0.039 | 0.039 | 0.018333 | 0.028 | 0.036 | 0.027167 | 702.253 | 0.00083333 | 1.3565e-05
702.508 | -0.002 | 702.508 | -0.002 | 0.001 | 0.005 | 0.005 | 0.003 | 0.012 | 0.01 | 0.013 | 0.016 | 0.0013333 | 0.0066667 | 0.013 | 0.0071667 | 702.508 | -0.0005 | -8.1393e-06
702.763 | 0.015 | 702.763 | 0.015 | 0.021 | 0.019 | 0.026 | 0.029 | 0.031 | 0.032 | 0.033 | 0.036 | 0.018333 | 0.028667 | 0.033667 | 0.026 | 702.763 | 0.0026667 | 4.3409e-05
703.019 | 0.019 | 703.019 | 0.019 | 0.019 | 0.026 | 0.03 | 0.025 | 0.03 | 0.036 | 0.034 | 0.04 | 0.021333 | 0.028333 | 0.036667 | 0.029 | 703.019 | -0.00066667 | -1.0852e-05
703.274 | 0.023 | 703.274 | 0.023 | 0.02 | 0.023 | 0.027 | 0.028 | 0.027 | 0.035 | 0.04 | 0.034 | 0.022 | 0.027333 | 0.036333 | 0.029167 | 703.274 | -0.0018333 | -2.9844e-05
703.529 | 0.02 | 703.529 | 0.02 | 0.018 | 0.021 | 0.033 | 0.027 | 0.028 | 0.036 | 0.032 | 0.035 | 0.019667 | 0.029333 | 0.034333 | 0.027 | 703.529 | 0.0023333 | 3.7983e-05
703.784 | 0.005 | 703.784 | 0.005 | 0.011 | 0.014 | 0.022 | 0.023 | 0.024 | 0.026 | 0.026 | 0.028 | 0.01 | 0.023 | 0.026667 | 0.018333 | 703.784 | 0.0046667 | 7.5967e-05
704.039 | 0.032 | 704.039 | 0.032 | 0.035 | 0.037 | 0.042 | 0.046 | 0.055 | 0.048 | 0.052 | 0.057 | 0.034667 | 0.047667 | 0.052333 | 0.0435 | 704.039 | 0.0041667 | 6.7827e-05
704.294 | 0.015 | 704.294 | 0.015 | 0.01 | 0.014 | 0.019 | 0.018 | 0.021 | 0.024 | 0.028 | 0.028 | 0.013 | 0.019333 | 0.026667 | 0.019833 | 704.294 | -0.0005 | -8.1393e-06
704.549 | 0.006 | 704.549 | 0.006 | 0.011 | 0.01 | 0.014 | 0.017 | 0.019 | 0.021 | 0.02 | 0.025 | 0.009 | 0.016667 | 0.022 | 0.0155 | 704.549 | 0.0011667 | 1.8992e-05
704.804 | 0.009 | 704.804 | 0.009 | 0.006 | 0.009 | 0.015 | 0.016 | 0.017 | 0.03 | 0.024 | 0.025 | 0.008 | 0.016 | 0.026333 | 0.017167 | 704.804 | -0.0011667 | -1.8992e-05
705.058 | 0.016 | 705.058 | 0.016 | 0.014 | 0.018 | 0.028 | 0.026 | 0.037 | 0.036 | 0.033 | 0.037 | 0.016 | 0.030333 | 0.035333 | 0.025667 | 705.058 | 0.0046667 | 7.5967e-05
705.313 | 0.001 | 705.313 | 0.001 | 0.006 | 0.005 | 0.009 | 0.009 | 0.02 | 0.02 | 0.025 | 0.025 | 0.004 | 0.012667 | 0.023333 | 0.013667 | 705.313 | -0.001 | -1.6279e-05
705.568 | 0.027 | 705.568 | 0.027 | 0.033 | 0.035 | 0.037 | 0.038 | 0.043 | 0.043 | 0.047 | 0.044 | 0.031667 | 0.039333 | 0.044667 | 0.038167 | 705.568 | 0.0011667 | 1.8992e-05
705.822 | 0.021 | 705.822 | 0.021 | 0.026 | 0.023 | 0.029 | 0.031 | 0.038 | 0.04 | 0.039 | 0.043 | 0.023333 | 0.032667 | 0.040667 | 0.032 | 705.822 | 0.00066667 | 1.0852e-05
706.077 | 0.011 | 706.077 | 0.011 | 0.014 | 0.014 | 0.016 | 0.018 | 0.028 | 0.021 | 0.027 | 0.027 | 0.013 | 0.020667 | 0.025 | 0.019 | 706.077 | 0.0016667 | 2.7131e-05
706.331 | 0.006 | 706.331 | 0.006 | 0.011 | 0.005 | 0.012 | 0.013 | 0.016 | 0.021 | 0.025 | 0.024 | 0.0073333 | 0.013667 | 0.023333 | 0.015333 | 706.331 | -0.0016667 | -2.7131e-05
706.586 | 0.006 | 706.586 | 0.006 | 0.008 | 0.012 | 0.022 | 0.022 | 0.023 | 0.024 | 0.029 | 0.025 | 0.0086667 | 0.022333 | 0.026 | 0.017333 | 706.586 | 0.005 | 8.1393e-05
706.84 | 0.019 | 706.84 | 0.019 | 0.021 | 0.027 | 0.035 | 0.038 | 0.039 | 0.042 | 0.041 | 0.039 | 0.022333 | 0.037333 | 0.040667 | 0.0315 | 706.84 | 0.0058333 | 9.4958e-05
707.095 | 0.011 | 707.095 | 0.011 | 0.017 | 0.019 | 0.023 | 0.026 | 0.031 | 0.027 | 0.028 | 0.028 | 0.015667 | 0.026667 | 0.027667 | 0.021667 | 707.095 | 0.005 | 8.1393e-05
707.349 | -0.016 | 707.349 | -0.016 | -0.013 | -0.01 | -0.007 | -0.004 | -0.003 | 0 | 0.004 | 0.01 | -0.013 | -0.0046667 | 0.0046667 | -0.0041667 | 707.349 | -0.0005 | -8.1393e-06
707.603 | 0.005 | 707.603 | 0.005 | 0.004 | 0.008 | 0.015 | 0.019 | 0.02 | 0.022 | 0.023 | 0.028 | 0.0056667 | 0.018 | 0.024333 | 0.015 | 707.603 | 0.003 | 4.8836e-05
707.857 | 0.014 | 707.857 | 0.014 | 0.016 | 0.012 | 0.021 | 0.025 | 0.027 | 0.026 | 0.037 | 0.028 | 0.014 | 0.024333 | 0.030333 | 0.022167 | 707.857 | 0.0021667 | 3.527e-05
708.111 | 0.003 | 708.111 | 0.003 | 0.009 | 0.004 | 0.016 | 0.015 | 0.017 | 0.019 | 0.023 | 0.022 | 0.0053333 | 0.016 | 0.021333 | 0.013333 | 708.111 | 0.0026667 | 4.3409e-05
708.365 | 0.01 | 708.365 | 0.01 | 0.009 | 0.014 | 0.014 | 0.023 | 0.024 | 0.023 | 0.024 | 0.026 | 0.011 | 0.020333 | 0.024333 | 0.017667 | 708.365 | 0.0026667 | 4.3409e-05
708.619 | 0.003 | 708.619 | 0.003 | 0.009 | 0.012 | 0.02 | 0.019 | 0.02 | 0.024 | 0.025 | 0.026 | 0.008 | 0.019667 | 0.025 | 0.0165 | 708.619 | 0.0031667 | 5.1549e-05
708.873 | 0.025 | 708.873 | 0.025 | 0.029 | 0.028 | 0.036 | 0.039 | 0.036 | 0.043 | 0.042 | 0.043 | 0.027333 | 0.037 | 0.042667 | 0.035 | 708.873 | 0.002 | 3.2557e-05
709.127 | 0.007 | 709.127 | 0.007 | 0.008 | 0.009 | 0.01 | 0.017 | 0.02 | 0.024 | 0.026 | 0.026 | 0.008 | 0.015667 | 0.025333 | 0.016667 | 709.127 | -0.001 | -1.6279e-05
709.381 | 0.004 | 709.381 | 0.004 | 0.005 | 0.007 | 0.009 | 0.011 | 0.019 | 0.02 | 0.018 | 0.02 | 0.0053333 | 0.013 | 0.019333 | 0.012333 | 709.381 | 0.00066667 | 1.0852e-05
709.635 | 0.007 | 709.635 | 0.007 | 0.006 | 0.021 | 0.018 | 0.016 | 0.023 | 0.026 | 0.031 | 0.025 | 0.011333 | 0.019 | 0.027333 | 0.019333 | 709.635 | -0.00033333 | -5.4262e-06
709.888 | 0.023 | 709.888 | 0.023 | 0.027 | 0.028 | 0.031 | 0.034 | 0.039 | 0.035 | 0.044 | 0.033 | 0.026 | 0.034667 | 0.037333 | 0.031667 | 709.888 | 0.003 | 4.8836e-05
710.142 | 0.007 | 710.142 | 0.007 | 0.011 | 0.016 | 0.02 | 0.021 | 0.017 | 0.025 | 0.026 | 0.027 | 0.011333 | 0.019333 | 0.026 | 0.018667 | 710.142 | 0.00066667 | 1.0852e-05
710.396 | 0.011 | 710.396 | 0.011 | 0.014 | 0.015 | 0.019 | 0.022 | 0.024 | 0.034 | 0.033 | 0.035 | 0.013333 | 0.021667 | 0.034 | 0.023667 | 710.396 | -0.002 | -3.2557e-05
710.649 | 0.008 | 710.649 | 0.008 | 0.01 | 0.02 | 0.022 | 0.022 | 0.029 | 0.024 | 0.032 | 0.03 | 0.012667 | 0.024333 | 0.028667 | 0.020667 | 710.649 | 0.0036667 | 5.9688e-05
710.903 | 0.007 | 710.903 | 0.007 | 0.011 | 0.008 | 0.016 | 0.012 | 0.016 | 0.019 | 0.022 | 0.024 | 0.0086667 | 0.014667 | 0.021667 | 0.015167 | 710.903 | -0.0005 | -8.1393e-06
711.156 | 0.001 | 711.156 | 0.001 | 0.01 | 0.017 | 0.018 | 0.022 | 0.025 | 0.027 | 0.034 | 0.024 | 0.0093333 | 0.021667 | 0.028333 | 0.018833 | 711.156 | 0.0028333 | 4.6123e-05
711.409 | 0.023 | 711.409 | 0.023 | 0.024 | 0.024 | 0.033 | 0.032 | 0.033 | 0.035 | 0.046 | 0.038 | 0.023667 | 0.032667 | 0.039667 | 0.031667 | 711.409 | 0.001 | 1.6279e-05
711.663 | 0.021 | 711.663 | 0.021 | 0.022 | 0.026 | 0.027 | 0.034 | 0.036 | 0.042 | 0.043 | 0.04 | 0.023 | 0.032333 | 0.041667 | 0.032333 | 711.663 | 0 | 0
711.916 | 0.013 | 711.916 | 0.013 | 0.008 | 0.015 | 0.015 | 0.021 | 0.02 | 0.023 | 0.025 | 0.026 | 0.012 | 0.018667 | 0.024667 | 0.018333 | 711.916 | 0.00033333 | 5.4262e-06
712.169 | 0.009 | 712.169 | 0.009 | 0.012 | 0.014 | 0.016 | 0.02 | 0.027 | 0.025 | 0.023 | 0.025 | 0.011667 | 0.021 | 0.024333 | 0.018 | 712.169 | 0.003 | 4.8836e-05
712.422 | 0.012 | 712.422 | 0.012 | 0.016 | 0.02 | 0.027 | 0.025 | 0.026 | 0.03 | 0.032 | 0.036 | 0.016 | 0.026 | 0.032667 | 0.024333 | 712.422 | 0.0016667 | 2.7131e-05
712.675 | 0.018 | 712.675 | 0.018 | 0.018 | 0.018 | 0.019 | 0.026 | 0.032 | 0.026 | 0.033 | 0.036 | 0.018 | 0.025667 | 0.031667 | 0.024833 | 712.675 | 0.00083333 | 1.3565e-05
712.928 | 0.026 | 712.928 | 0.026 | 0.028 | 0.036 | 0.042 | 0.042 | 0.045 | 0.047 | 0.052 | 0.047 | 0.03 | 0.043 | 0.048667 | 0.039333 | 712.928 | 0.0036667 | 5.9688e-05
713.181 | 0.021 | 713.181 | 0.021 | 0.021 | 0.022 | 0.03 | 0.037 | 0.035 | 0.034 | 0.039 | 0.04 | 0.021333 | 0.034 | 0.037667 | 0.0295 | 713.181 | 0.0045 | 7.3254e-05
713.434 | 0.012 | 713.434 | 0.012 | 0.018 | 0.018 | 0.027 | 0.02 | 0.031 | 0.034 | 0.034 | 0.027 | 0.016 | 0.026 | 0.031667 | 0.023833 | 713.434 | 0.0021667 | 3.527e-05
713.687 | 0.029 | 713.687 | 0.029 | 0.032 | 0.031 | 0.035 | 0.038 | 0.043 | 0.045 | 0.048 | 0.048 | 0.030667 | 0.038667 | 0.047 | 0.038833 | 713.687 | -0.00016667 | -2.7131e-06
713.94 | 0.006 | 713.94 | 0.006 | 0.008 | 0.013 | 0.02 | 0.022 | 0.025 | 0.028 | 0.029 | 0.027 | 0.009 | 0.022333 | 0.028 | 0.0185 | 713.94 | 0.0038333 | 6.2401e-05
714.192 | 0.017 | 714.192 | 0.017 | 0.019 | 0.017 | 0.024 | 0.027 | 0.031 | 0.035 | 0.034 | 0.039 | 0.017667 | 0.027333 | 0.036 | 0.026833 | 714.192 | 0.0005 | 8.1393e-06
714.445 | -0.005 | 714.445 | -0.005 | 0 | -0.007 | -0.001 | 0.005 | 0.004 | 0.013 | 0.002 | 0.007 | -0.004 | 0.0026667 | 0.0073333 | 0.0016667 | 714.445 | 0.001 | 1.6279e-05
714.698 | 0.023 | 714.698 | 0.023 | 0.021 | 0.024 | 0.034 | 0.032 | 0.034 | 0.03 | 0.041 | 0.043 | 0.022667 | 0.033333 | 0.038 | 0.030333 | 714.698 | 0.003 | 4.8836e-05
714.95 | 0.022 | 714.95 | 0.022 | 0.023 | 0.032 | 0.029 | 0.032 | 0.038 | 0.037 | 0.041 | 0.037 | 0.025667 | 0.033 | 0.038333 | 0.032 | 714.95 | 0.001 | 1.6279e-05
715.203 | 0.029 | 715.203 | 0.029 | 0.039 | 0.044 | 0.048 | 0.051 | 0.055 | 0.056 | 0.059 | 0.057 | 0.037333 | 0.051333 | 0.057333 | 0.047333 | 715.203 | 0.004 | 6.5114e-05
715.455 | 0.014 | 715.455 | 0.014 | 0.015 | 0.017 | 0.023 | 0.025 | 0.027 | 0.032 | 0.037 | 0.038 | 0.015333 | 0.025 | 0.035667 | 0.0255 | 715.455 | -0.0005 | -8.1393e-06
715.708 | 0.018 | 715.708 | 0.018 | 0.012 | 0.021 | 0.036 | 0.033 | 0.038 | 0.037 | 0.044 | 0.04 | 0.017 | 0.035667 | 0.040333 | 0.028667 | 715.708 | 0.007 | 0.00011395
715.96 | 0.026 | 715.96 | 0.026 | 0.035 | 0.036 | 0.041 | 0.046 | 0.047 | 0.05 | 0.051 | 0.056 | 0.032333 | 0.044667 | 0.052333 | 0.042333 | 715.96 | 0.0023333 | 3.7983e-05
716.212 | 0.012 | 716.212 | 0.012 | 0.019 | 0.016 | 0.02 | 0.026 | 0.025 | 0.025 | 0.03 | 0.034 | 0.015667 | 0.023667 | 0.029667 | 0.022667 | 716.212 | 0.001 | 1.6279e-05
716.464 | 0.028 | 716.464 | 0.028 | 0.028 | 0.039 | 0.038 | 0.038 | 0.045 | 0.044 | 0.047 | 0.043 | 0.031667 | 0.040333 | 0.044667 | 0.038167 | 716.464 | 0.0021667 | 3.527e-05
716.717 | 0.019 | 716.717 | 0.019 | 0.023 | 0.024 | 0.028 | 0.032 | 0.033 | 0.038 | 0.044 | 0.047 | 0.022 | 0.031 | 0.043 | 0.0325 | 716.717 | -0.0015 | -2.4418e-05
716.969 | 0.027 | 716.969 | 0.027 | 0.025 | 0.027 | 0.031 | 0.034 | 0.033 | 0.037 | 0.039 | 0.041 | 0.026333 | 0.032667 | 0.039 | 0.032667 | 716.969 | 0 | 0
717.221 | 0.016 | 717.221 | 0.016 | 0.026 | 0.028 | 0.03 | 0.029 | 0.042 | 0.04 | 0.045 | 0.044 | 0.023333 | 0.033667 | 0.043 | 0.033167 | 717.221 | 0.0005 | 8.1393e-06
717.473 | -0.01 | 717.473 | -0.01 | -0.002 | -0.001 | 0.014 | 0.006 | 0.005 | 0.015 | 0.018 | 0.017 | -0.0043333 | 0.0083333 | 0.016667 | 0.0061667 | 717.473 | 0.0021667 | 3.527e-05
717.725 | 0.016 | 717.725 | 0.016 | 0.021 | 0.017 | 0.024 | 0.028 | 0.032 | 0.033 | 0.036 | 0.036 | 0.018 | 0.028 | 0.035 | 0.0265 | 717.725 | 0.0015 | 2.4418e-05
717.977 | 0.017 | 717.977 | 0.017 | 0.017 | 0.024 | 0.031 | 0.027 | 0.028 | 0.027 | 0.039 | 0.036 | 0.019333 | 0.028667 | 0.034 | 0.026667 | 717.977 | 0.002 | 3.2557e-05
718.228 | 0.029 | 718.228 | 0.029 | 0.029 | 0.032 | 0.04 | 0.041 | 0.04 | 0.045 | 0.047 | 0.049 | 0.03 | 0.040333 | 0.047 | 0.0385 | 718.228 | 0.0018333 | 2.9844e-05
718.48 | 0.011 | 718.48 | 0.011 | 0.013 | 0.015 | 0.025 | 0.017 | 0.02 | 0.024 | 0.032 | 0.028 | 0.013 | 0.020667 | 0.028 | 0.0205 | 718.48 | 0.00016667 | 2.7131e-06
718.732 | 0.016 | 718.732 | 0.016 | 0.024 | 0.018 | 0.027 | 0.029 | 0.038 | 0.036 | 0.037 | 0.041 | 0.019333 | 0.031333 | 0.038 | 0.028667 | 718.732 | 0.0026667 | 4.3409e-05
718.984 | 0.027 | 718.984 | 0.027 | 0.032 | 0.031 | 0.038 | 0.039 | 0.039 | 0.041 | 0.05 | 0.04 | 0.03 | 0.038667 | 0.043667 | 0.036833 | 718.984 | 0.0018333 | 2.9844e-05
719.235 | 0.027 | 719.235 | 0.027 | 0.027 | 0.035 | 0.037 | 0.042 | 0.045 | 0.04 | 0.05 | 0.053 | 0.029667 | 0.041333 | 0.047667 | 0.038667 | 719.235 | 0.0026667 | 4.3409e-05
719.487 | 0.023 | 719.487 | 0.023 | 0.028 | 0.027 | 0.032 | 0.035 | 0.038 | 0.034 | 0.039 | 0.043 | 0.026 | 0.035 | 0.038667 | 0.032333 | 719.487 | 0.0026667 | 4.3409e-05
719.738 | 0.007 | 719.738 | 0.007 | 0.014 | 0.013 | 0.018 | 0.02 | 0.028 | 0.025 | 0.026 | 0.033 | 0.011333 | 0.022 | 0.028 | 0.019667 | 719.738 | 0.0023333 | 3.7983e-05
719.99 | 0.031 | 719.99 | 0.031 | 0.035 | 0.034 | 0.038 | 0.044 | 0.047 | 0.047 | 0.051 | 0.053 | 0.033333 | 0.043 | 0.050333 | 0.041833 | 719.99 | 0.0011667 | 1.8992e-05
720.241 | 0.013 | 720.241 | 0.013 | 0.012 | 0.024 | 0.026 | 0.019 | 0.023 | 0.03 | 0.033 | 0.032 | 0.016333 | 0.022667 | 0.031667 | 0.024 | 720.241 | -0.0013333 | -2.1705e-05
720.492 | 0.02 | 720.492 | 0.02 | 0.029 | 0.028 | 0.03 | 0.034 | 0.032 | 0.041 | 0.042 | 0.044 | 0.025667 | 0.032 | 0.042333 | 0.034 | 720.492 | -0.002 | -3.2557e-05
720.744 | 0.01 | 720.744 | 0.01 | 0.012 | 0.015 | 0.017 | 0.018 | 0.026 | 0.023 | 0.031 | 0.03 | 0.012333 | 0.020333 | 0.028 | 0.020167 | 720.744 | 0.00016667 | 2.7131e-06
720.995 | 0.008 | 720.995 | 0.008 | 0.007 | 0.008 | 0.015 | 0.02 | 0.02 | 0.019 | 0.029 | 0.022 | 0.0076667 | 0.018333 | 0.023333 | 0.0155 | 720.995 | 0.0028333 | 4.6123e-05
721.246 | 0.017 | 721.246 | 0.017 | 0.015 | 0.016 | 0.021 | 0.025 | 0.029 | 0.027 | 0.034 | 0.037 | 0.016 | 0.025 | 0.032667 | 0.024333 | 721.246 | 0.00066667 | 1.0852e-05
721.497 | 0.03 | 721.497 | 0.03 | 0.032 | 0.036 | 0.043 | 0.04 | 0.048 | 0.046 | 0.049 | 0.047 | 0.032667 | 0.043667 | 0.047333 | 0.04 | 721.497 | 0.0036667 | 5.9688e-05
721.748 | 0.02 | 721.748 | 0.02 | 0.02 | 0.022 | 0.025 | 0.029 | 0.028 | 0.03 | 0.033 | 0.039 | 0.020667 | 0.027333 | 0.034 | 0.027333 | 721.748 | 0 | 0
721.999 | 0.007 | 721.999 | 0.007 | 0.01 | 0.008 | 0.018 | 0.022 | 0.02 | 0.025 | 0.032 | 0.032 | 0.0083333 | 0.02 | 0.029667 | 0.019 | 721.999 | 0.001 | 1.6279e-05
722.25 | 0.038 | 722.25 | 0.038 | 0.046 | 0.041 | 0.057 | 0.05 | 0.055 | 0.052 | 0.06 | 0.06 | 0.041667 | 0.054 | 0.057333 | 0.0495 | 722.25 | 0.0045 | 7.3254e-05
722.501 | 0.026 | 722.501 | 0.026 | 0.036 | 0.03 | 0.05 | 0.046 | 0.047 | 0.046 | 0.053 | 0.053 | 0.030667 | 0.047667 | 0.050667 | 0.040667 | 722.501 | 0.007 | 0.00011395
722.752 | 0.02 | 722.752 | 0.02 | 0.022 | 0.027 | 0.026 | 0.028 | 0.03 | 0.039 | 0.045 | 0.041 | 0.023 | 0.028 | 0.041667 | 0.032333 | 722.752 | -0.0043333 | -7.054e-05
723.002 | 0.016 | 723.002 | 0.016 | 0.012 | 0.011 | 0.023 | 0.023 | 0.023 | 0.032 | 0.025 | 0.031 | 0.013 | 0.023 | 0.029333 | 0.021167 | 723.002 | 0.0018333 | 2.9844e-05
723.253 | 0.014 | 723.253 | 0.014 | 0.005 | 0.012 | 0.018 | 0.017 | 0.015 | 0.024 | 0.024 | 0.032 | 0.010333 | 0.016667 | 0.026667 | 0.0185 | 723.253 | -0.0018333 | -2.9844e-05
723.504 | 0.03 | 723.504 | 0.03 | 0.032 | 0.035 | 0.042 | 0.045 | 0.045 | 0.051 | 0.052 | 0.055 | 0.032333 | 0.044 | 0.052667 | 0.0425 | 723.504 | 0.0015 | 2.4418e-05
723.754 | 0.02 | 723.754 | 0.02 | 0.025 | 0.02 | 0.029 | 0.029 | 0.036 | 0.041 | 0.039 | 0.044 | 0.021667 | 0.031333 | 0.041333 | 0.0315 | 723.754 | -0.00016667 | -2.7131e-06
724.005 | 0.029 | 724.005 | 0.029 | 0.035 | 0.033 | 0.036 | 0.041 | 0.045 | 0.051 | 0.051 | 0.05 | 0.032333 | 0.040667 | 0.050667 | 0.0415 | 724.005 | -0.00083333 | -1.3565e-05
724.255 | 0.019 | 724.255 | 0.019 | 0.032 | 0.026 | 0.031 | 0.037 | 0.035 | 0.041 | 0.045 | 0.045 | 0.025667 | 0.034333 | 0.043667 | 0.034667 | 724.255 | -0.00033333 | -5.4262e-06
724.506 | 0.019 | 724.506 | 0.019 | 0.025 | 0.033 | 0.038 | 0.039 | 0.041 | 0.042 | 0.045 | 0.047 | 0.025667 | 0.039333 | 0.044667 | 0.035167 | 724.506 | 0.0041667 | 6.7827e-05
724.756 | 0.027 | 724.756 | 0.027 | 0.035 | 0.03 | 0.038 | 0.042 | 0.045 | 0.041 | 0.05 | 0.045 | 0.030667 | 0.041667 | 0.045333 | 0.038 | 724.756 | 0.0036667 | 5.9688e-05
725.006 | 0.005 | 725.006 | 0.005 | 0.009 | 0.015 | 0.021 | 0.018 | 0.024 | 0.029 | 0.032 | 0.033 | 0.0096667 | 0.021 | 0.031333 | 0.0205 | 725.006 | 0.0005 | 8.1393e-06
725.257 | 0.02 | 725.257 | 0.02 | 0.012 | 0.013 | 0.03 | 0.028 | 0.033 | 0.035 | 0.033 | 0.042 | 0.015 | 0.030333 | 0.036667 | 0.025833 | 725.257 | 0.0045 | 7.3254e-05
725.507 | 0.021 | 725.507 | 0.021 | 0.025 | 0.024 | 0.032 | 0.033 | 0.038 | 0.043 | 0.039 | 0.045 | 0.023333 | 0.034333 | 0.042333 | 0.032833 | 725.507 | 0.0015 | 2.4418e-05
725.757 | 0.012 | 725.757 | 0.012 | 0.012 | 0.018 | 0.023 | 0.028 | 0.028 | 0.033 | 0.029 | 0.037 | 0.014 | 0.026333 | 0.033 | 0.0235 | 725.757 | 0.0028333 | 4.6123e-05
726.007 | -0.002 | 726.007 | -0.002 | 0.011 | 0.007 | 0.014 | 0.009 | 0.016 | 0.022 | 0.031 | 0.028 | 0.0053333 | 0.013 | 0.027 | 0.016167 | 726.007 | -0.0031667 | -5.1549e-05
726.257 | 0.028 | 726.257 | 0.028 | 0.031 | 0.028 | 0.033 | 0.046 | 0.043 | 0.047 | 0.047 | 0.046 | 0.029 | 0.040667 | 0.046667 | 0.037833 | 726.257 | 0.0028333 | 4.6123e-05
726.507 | 0.009 | 726.507 | 0.009 | 0.006 | 0.009 | 0.015 | 0.011 | 0.021 | 0.025 | 0.027 | 0.024 | 0.008 | 0.015667 | 0.025333 | 0.016667 | 726.507 | -0.001 | -1.6279e-05
726.757 | 0.043 | 726.757 | 0.043 | 0.045 | 0.048 | 0.055 | 0.059 | 0.064 | 0.071 | 0.076 | 0.07 | 0.045333 | 0.059333 | 0.072333 | 0.058833 | 726.757 | 0.0005 | 8.1393e-06
727.007 | 0.028 | 727.007 | 0.028 | 0.029 | 0.035 | 0.042 | 0.045 | 0.042 | 0.045 | 0.051 | 0.052 | 0.030667 | 0.043 | 0.049333 | 0.04 | 727.007 | 0.003 | 4.8836e-05
727.257 | 0.005 | 727.257 | 0.005 | 0.003 | 0.009 | 0.013 | 0.019 | 0.02 | 0.024 | 0.03 | 0.025 | 0.0056667 | 0.017333 | 0.026333 | 0.016 | 727.257 | 0.0013333 | 2.1705e-05
727.506 | 0.027 | 727.506 | 0.027 | 0.027 | 0.033 | 0.036 | 0.038 | 0.043 | 0.046 | 0.045 | 0.045 | 0.029 | 0.039 | 0.045333 | 0.037167 | 727.506 | 0.0018333 | 2.9844e-05
727.756 | 0.016 | 727.756 | 0.016 | 0.018 | 0.016 | 0.023 | 0.026 | 0.032 | 0.022 | 0.035 | 0.035 | 0.016667 | 0.027 | 0.030667 | 0.023667 | 727.756 | 0.0033333 | 5.4262e-05
728.006 | 0.024 | 728.006 | 0.024 | 0.028 | 0.033 | 0.039 | 0.044 | 0.037 | 0.043 | 0.046 | 0.054 | 0.028333 | 0.04 | 0.047667 | 0.038 | 728.006 | 0.002 | 3.2557e-05
728.255 | 0.022 | 728.255 | 0.022 | 0.034 | 0.032 | 0.041 | 0.041 | 0.046 | 0.044 | 0.047 | 0.051 | 0.029333 | 0.042667 | 0.047333 | 0.038333 | 728.255 | 0.0043333 | 7.054e-05
728.505 | 0.001 | 728.505 | 0.001 | 0.002 | 0.005 | 0.012 | 0.017 | 0.017 | 0.021 | 0.023 | 0.018 | 0.0026667 | 0.015333 | 0.020667 | 0.011667 | 728.505 | 0.0036667 | 5.9688e-05
728.754 | 0.019 | 728.754 | 0.019 | 0.02 | 0.025 | 0.031 | 0.029 | 0.037 | 0.042 | 0.04 | 0.041 | 0.021333 | 0.032333 | 0.041 | 0.031167 | 728.754 | 0.0011667 | 1.8992e-05
729.004 | 0.03 | 729.004 | 0.03 | 0.037 | 0.035 | 0.051 | 0.042 | 0.053 | 0.051 | 0.053 | 0.061 | 0.034 | 0.048667 | 0.055 | 0.0445 | 729.004 | 0.0041667 | 6.7827e-05
729.253 | 0.035 | 729.253 | 0.035 | 0.044 | 0.049 | 0.054 | 0.056 | 0.06 | 0.064 | 0.077 | 0.073 | 0.042667 | 0.056667 | 0.071333 | 0.057 | 729.253 | -0.00033333 | -5.4262e-06
729.502 | 0.042 | 729.502 | 0.042 | 0.051 | 0.054 | 0.061 | 0.062 | 0.061 | 0.064 | 0.061 | 0.071 | 0.049 | 0.061333 | 0.065333 | 0.057167 | 729.502 | 0.0041667 | 6.7827e-05
729.751 | 0.021 | 729.751 | 0.021 | 0.019 | 0.029 | 0.03 | 0.034 | 0.036 | 0.042 | 0.044 | 0.039 | 0.023 | 0.033333 | 0.041667 | 0.032333 | 729.751 | 0.001 | 1.6279e-05
730.001 | -0.001 | 730.001 | -0.001 | 0.008 | 0.006 | 0.018 | 0.011 | 0.019 | 0.018 | 0.024 | 0.024 | 0.0043333 | 0.016 | 0.022 | 0.013167 | 730.001 | 0.0028333 | 4.6123e-05
730.25 | 0.01 | 730.25 | 0.01 | 0.014 | 0.014 | 0.024 | 0.025 | 0.029 | 0.028 | 0.036 | 0.037 | 0.012667 | 0.026 | 0.033667 | 0.023167 | 730.25 | 0.0028333 | 4.6123e-05
730.499 | 0.029 | 730.499 | 0.029 | 0.031 | 0.037 | 0.041 | 0.043 | 0.045 | 0.047 | 0.062 | 0.054 | 0.032333 | 0.043 | 0.054333 | 0.043333 | 730.499 | -0.00033333 | -5.4262e-06
730.748 | 0.016 | 730.748 | 0.016 | 0.024 | 0.028 | 0.034 | 0.032 | 0.041 | 0.036 | 0.039 | 0.047 | 0.022667 | 0.035667 | 0.040667 | 0.031667 | 730.748 | 0.004 | 6.5114e-05
730.997 | 0.018 | 730.997 | 0.018 | 0.014 | 0.022 | 0.03 | 0.026 | 0.02 | 0.034 | 0.034 | 0.037 | 0.018 | 0.025333 | 0.035 | 0.0265 | 730.997 | -0.0011667 | -1.8992e-05
731.246 | 0.029 | 731.246 | 0.029 | 0.026 | 0.037 | 0.034 | 0.044 | 0.048 | 0.05 | 0.05 | 0.05 | 0.030667 | 0.042 | 0.05 | 0.040333 | 731.246 | 0.0016667 | 2.7131e-05
731.495 | 0.022 | 731.495 | 0.022 | 0.018 | 0.017 | 0.029 | 0.034 | 0.037 | 0.036 | 0.038 | 0.041 | 0.019 | 0.033333 | 0.038333 | 0.028667 | 731.495 | 0.0046667 | 7.5967e-05
731.743 | 0.027 | 731.743 | 0.027 | 0.027 | 0.041 | 0.038 | 0.04 | 0.041 | 0.046 | 0.045 | 0.052 | 0.031667 | 0.039667 | 0.047667 | 0.039667 | 731.743 | 0 | 0
731.992 | 0.019 | 731.992 | 0.019 | 0.025 | 0.027 | 0.032 | 0.029 | 0.036 | 0.04 | 0.044 | 0.039 | 0.023667 | 0.032333 | 0.041 | 0.032333 | 731.992 | 0 | 0
732.241 | 0.005 | 732.241 | 0.005 | 0.009 | 0.007 | 0.016 | 0.018 | 0.024 | 0.025 | 0.025 | 0.027 | 0.007 | 0.019333 | 0.025667 | 0.016333 | 732.241 | 0.003 | 4.8836e-05
732.489 | 0.02 | 732.489 | 0.02 | 0.025 | 0.024 | 0.028 | 0.034 | 0.043 | 0.044 | 0.047 | 0.046 | 0.023 | 0.035 | 0.045667 | 0.034333 | 732.489 | 0.00066667 | 1.0852e-05
732.738 | 0.041 | 732.738 | 0.041 | 0.038 | 0.049 | 0.057 | 0.052 | 0.06 | 0.065 | 0.069 | 0.075 | 0.042667 | 0.056333 | 0.069667 | 0.056167 | 732.738 | 0.00016667 | 2.7131e-06
732.987 | 0.017 | 732.987 | 0.017 | 0.017 | 0.019 | 0.026 | 0.029 | 0.025 | 0.033 | 0.041 | 0.04 | 0.017667 | 0.026667 | 0.038 | 0.027833 | 732.987 | -0.0011667 | -1.8992e-05
733.235 | 0.016 | 733.235 | 0.016 | 0.024 | 0.022 | 0.033 | 0.031 | 0.037 | 0.04 | 0.038 | 0.041 | 0.020667 | 0.033667 | 0.039667 | 0.030167 | 733.235 | 0.0035 | 5.6975e-05
733.483 | 0.03 | 733.483 | 0.03 | 0.036 | 0.035 | 0.042 | 0.045 | 0.047 | 0.051 | 0.048 | 0.059 | 0.033667 | 0.044667 | 0.052667 | 0.043167 | 733.483 | 0.0015 | 2.4418e-05
733.732 | 0.026 | 733.732 | 0.026 | 0.034 | 0.033 | 0.042 | 0.044 | 0.043 | 0.044 | 0.056 | 0.048 | 0.031 | 0.043 | 0.049333 | 0.040167 | 733.732 | 0.0028333 | 4.6123e-05
733.98 | -0.003 | 733.98 | -0.003 | 0.009 | 0.007 | 0.019 | 0.016 | 0.018 | 0.023 | 0.024 | 0.024 | 0.0043333 | 0.017667 | 0.023667 | 0.014 | 733.98 | 0.0036667 | 5.9688e-05
734.228 | 0.012 | 734.228 | 0.012 | 0.008 | 0.01 | 0.024 | 0.019 | 0.02 | 0.028 | 0.032 | 0.029 | 0.01 | 0.021 | 0.029667 | 0.019833 | 734.228 | 0.0011667 | 1.8992e-05
734.477 | 0.012 | 734.477 | 0.012 | 0.007 | 0.019 | 0.022 | 0.018 | 0.025 | 0.03 | 0.033 | 0.037 | 0.012667 | 0.021667 | 0.033333 | 0.023 | 734.477 | -0.0013333 | -2.1705e-05
734.725 | 0.025 | 734.725 | 0.025 | 0.034 | 0.036 | 0.041 | 0.044 | 0.056 | 0.048 | 0.056 | 0.049 | 0.031667 | 0.047 | 0.051 | 0.041333 | 734.725 | 0.0056667 | 9.2245e-05
734.973 | 0.016 | 734.973 | 0.016 | 0.018 | 0.024 | 0.027 | 0.023 | 0.024 | 0.031 | 0.037 | 0.038 | 0.019333 | 0.024667 | 0.035333 | 0.027333 | 734.973 | -0.0026667 | -4.3409e-05
735.221 | 0.027 | 735.221 | 0.027 | 0.026 | 0.037 | 0.042 | 0.04 | 0.045 | 0.046 | 0.053 | 0.049 | 0.03 | 0.042333 | 0.049333 | 0.039667 | 735.221 | 0.0026667 | 4.3409e-05
735.469 | 0.031 | 735.469 | 0.031 | 0.035 | 0.028 | 0.044 | 0.043 | 0.047 | 0.042 | 0.053 | 0.054 | 0.031333 | 0.044667 | 0.049667 | 0.0405 | 735.469 | 0.0041667 | 6.7827e-05
735.717 | 0.028 | 735.717 | 0.028 | 0.032 | 0.035 | 0.039 | 0.048 | 0.045 | 0.05 | 0.049 | 0.055 | 0.031667 | 0.044 | 0.051333 | 0.0415 | 735.717 | 0.0025 | 4.0696e-05
735.965 | 0.003 | 735.965 | 0.003 | 0.013 | 0.016 | 0.027 | 0.026 | 0.025 | 0.03 | 0.04 | 0.035 | 0.010667 | 0.026 | 0.035 | 0.022833 | 735.965 | 0.0031667 | 5.1549e-05
736.212 | 0.007 | 736.212 | 0.007 | 0.012 | 0.02 | 0.021 | 0.029 | 0.026 | 0.027 | 0.039 | 0.033 | 0.013 | 0.025333 | 0.033 | 0.023 | 736.212 | 0.0023333 | 3.7983e-05
736.46 | 0.031 | 736.46 | 0.031 | 0.034 | 0.037 | 0.041 | 0.039 | 0.043 | 0.043 | 0.055 | 0.052 | 0.034 | 0.041 | 0.05 | 0.042 | 736.46 | -0.001 | -1.6279e-05
736.708 | 0.021 | 736.708 | 0.021 | 0.029 | 0.023 | 0.032 | 0.038 | 0.041 | 0.041 | 0.047 | 0.056 | 0.024333 | 0.037 | 0.048 | 0.036167 | 736.708 | 0.00083333 | 1.3565e-05
736.955 | 0.043 | 736.955 | 0.043 | 0.047 | 0.052 | 0.057 | 0.055 | 0.06 | 0.06 | 0.066 | 0.065 | 0.047333 | 0.057333 | 0.063667 | 0.0555 | 736.955 | 0.0018333 | 2.9844e-05
737.203 | 0.003 | 737.203 | 0.003 | 0.01 | 0.007 | 0.017 | 0.022 | 0.022 | 0.026 | 0.035 | 0.033 | 0.0066667 | 0.020333 | 0.031333 | 0.019 | 737.203 | 0.0013333 | 2.1705e-05
737.451 | 0.035 | 737.451 | 0.035 | 0.043 | 0.043 | 0.045 | 0.048 | 0.054 | 0.057 | 0.065 | 0.057 | 0.040333 | 0.049 | 0.059667 | 0.05 | 737.451 | -0.001 | -1.6279e-05
737.698 | 0.022 | 737.698 | 0.022 | 0.022 | 0.03 | 0.028 | 0.037 | 0.039 | 0.039 | 0.051 | 0.045 | 0.024667 | 0.034667 | 0.045 | 0.034833 | 737.698 | -0.00016667 | -2.7131e-06
737.945 | 0.017 | 737.945 | 0.017 | 0.019 | 0.024 | 0.031 | 0.03 | 0.034 | 0.036 | 0.042 | 0.035 | 0.02 | 0.031667 | 0.037667 | 0.028833 | 737.945 | 0.0028333 | 4.6123e-05
738.193 | 0.027 | 738.193 | 0.027 | 0.031 | 0.034 | 0.054 | 0.051 | 0.061 | 0.046 | 0.048 | 0.052 | 0.030667 | 0.055333 | 0.048667 | 0.039667 | 738.193 | 0.015667 | 0.00025503
738.44 | 0.014 | 738.44 | 0.014 | 0.015 | 0.024 | 0.039 | 0.044 | 0.039 | 0.032 | 0.04 | 0.036 | 0.017667 | 0.040667 | 0.036 | 0.026833 | 738.44 | 0.013833 | 0.00022519
738.687 | 0.024 | 738.687 | 0.024 | 0.023 | 0.032 | 0.036 | 0.041 | 0.036 | 0.038 | 0.043 | 0.043 | 0.026333 | 0.037667 | 0.041333 | 0.033833 | 738.687 | 0.0038333 | 6.2401e-05
738.935 | 0.015 | 738.935 | 0.015 | 0.023 | 0.031 | 0.031 | 0.033 | 0.039 | 0.034 | 0.044 | 0.043 | 0.023 | 0.034333 | 0.040333 | 0.031667 | 738.935 | 0.0026667 | 4.3409e-05
739.182 | 0.015 | 739.182 | 0.015 | 0.014 | 0.012 | 0.029 | 0.029 | 0.034 | 0.038 | 0.037 | 0.039 | 0.013667 | 0.030667 | 0.038 | 0.025833 | 739.182 | 0.0048333 | 7.868e-05
739.429 | 0.029 | 739.429 | 0.029 | 0.035 | 0.037 | 0.037 | 0.041 | 0.043 | 0.046 | 0.047 | 0.054 | 0.033667 | 0.040333 | 0.049 | 0.041333 | 739.429 | -0.001 | -1.6279e-05
739.676 | 0.015 | 739.676 | 0.015 | 0.008 | 0.016 | 0.025 | 0.022 | 0.026 | 0.031 | 0.032 | 0.035 | 0.013 | 0.024333 | 0.032667 | 0.022833 | 739.676 | 0.0015 | 2.4418e-05
739.923 | 0.021 | 739.923 | 0.021 | 0.029 | 0.028 | 0.036 | 0.037 | 0.039 | 0.043 | 0.051 | 0.048 | 0.026 | 0.037333 | 0.047333 | 0.036667 | 739.923 | 0.00066667 | 1.0852e-05
740.17 | 0.033 | 740.17 | 0.033 | 0.031 | 0.036 | 0.04 | 0.044 | 0.047 | 0.05 | 0.048 | 0.06 | 0.033333 | 0.043667 | 0.052667 | 0.043 | 740.17 | 0.00066667 | 1.0852e-05
740.417 | 0.009 | 740.417 | 0.009 | 0.012 | 0.022 | 0.023 | 0.032 | 0.03 | 0.034 | 0.039 | 0.043 | 0.014333 | 0.028333 | 0.038667 | 0.0265 | 740.417 | 0.0018333 | 2.9844e-05
740.663 | 0.008 | 740.663 | 0.008 | 0.013 | 0.01 | 0.024 | 0.026 | 0.029 | 0.027 | 0.029 | 0.032 | 0.010333 | 0.026333 | 0.029333 | 0.019833 | 740.663 | 0.0065 | 0.00010581
740.91 | 0.027 | 740.91 | 0.027 | 0.028 | 0.033 | 0.033 | 0.034 | 0.039 | 0.045 | 0.048 | 0.046 | 0.029333 | 0.035333 | 0.046333 | 0.037833 | 740.91 | -0.0025 | -4.0696e-05
741.157 | 0.03 | 741.157 | 0.03 | 0.035 | 0.041 | 0.043 | 0.045 | 0.049 | 0.051 | 0.053 | 0.055 | 0.035333 | 0.045667 | 0.053 | 0.044167 | 741.157 | 0.0015 | 2.4418e-05
741.403 | 0.015 | 741.403 | 0.015 | 0.014 | 0.023 | 0.028 | 0.03 | 0.032 | 0.036 | 0.042 | 0.036 | 0.017333 | 0.03 | 0.038 | 0.027667 | 741.403 | 0.0023333 | 3.7983e-05
741.65 | -0.006 | 741.65 | -0.006 | -0.007 | -0.001 | 0.006 | 0.008 | 0.009 | 0.022 | 0.014 | 0.024 | -0.0046667 | 0.0076667 | 0.02 | 0.0076667 | 741.65 | 0 | 0
741.897 | 0.021 | 741.897 | 0.021 | 0.022 | 0.021 | 0.03 | 0.04 | 0.039 | 0.043 | 0.047 | 0.044 | 0.021333 | 0.036333 | 0.044667 | 0.033 | 741.897 | 0.0033333 | 5.4262e-05
742.143 | 0.008 | 742.143 | 0.008 | 0.012 | 0.014 | 0.016 | 0.02 | 0.026 | 0.03 | 0.031 | 0.034 | 0.011333 | 0.020667 | 0.031667 | 0.0215 | 742.143 | -0.00083333 | -1.3565e-05
742.389 | 0.018 | 742.389 | 0.018 | 0.03 | 0.023 | 0.032 | 0.037 | 0.039 | 0.038 | 0.054 | 0.044 | 0.023667 | 0.036 | 0.045333 | 0.0345 | 742.389 | 0.0015 | 2.4418e-05
742.636 | 0 | 742.636 | 0 | 0.006 | 0.012 | 0.02 | 0.019 | 0.019 | 0.027 | 0.028 | 0.032 | 0.006 | 0.019333 | 0.029 | 0.0175 | 742.636 | 0.0018333 | 2.9844e-05
742.882 | 0.009 | 742.882 | 0.009 | 0.01 | 0.022 | 0.021 | 0.026 | 0.037 | 0.029 | 0.036 | 0.034 | 0.013667 | 0.028 | 0.033 | 0.023333 | 742.882 | 0.0046667 | 7.5967e-05
743.128 | 0.017 | 743.128 | 0.017 | 0.022 | 0.024 | 0.027 | 0.032 | 0.032 | 0.041 | 0.038 | 0.043 | 0.021 | 0.030333 | 0.040667 | 0.030833 | 743.128 | -0.0005 | -8.1393e-06
743.375 | 0.005 | 743.375 | 0.005 | 0.003 | 0.003 | 0.012 | 0.011 | 0.013 | 0.013 | 0.026 | 0.025 | 0.0036667 | 0.012 | 0.021333 | 0.0125 | 743.375 | -0.0005 | -8.1393e-06
743.621 | 0.04 | 743.621 | 0.04 | 0.045 | 0.045 | 0.054 | 0.055 | 0.054 | 0.063 | 0.052 | 0.066 | 0.043333 | 0.054333 | 0.060333 | 0.051833 | 743.621 | 0.0025 | 4.0696e-05
743.867 | 0.023 | 743.867 | 0.023 | 0.022 | 0.029 | 0.034 | 0.043 | 0.041 | 0.049 | 0.046 | 0.049 | 0.024667 | 0.039333 | 0.048 | 0.036333 | 743.867 | 0.003 | 4.8836e-05
744.113 | 0.025 | 744.113 | 0.025 | 0.022 | 0.024 | 0.026 | 0.032 | 0.03 | 0.037 | 0.047 | 0.048 | 0.023667 | 0.029333 | 0.044 | 0.033833 | 744.113 | -0.0045 | -7.3254e-05
744.359 | 0.007 | 744.359 | 0.007 | 0.016 | 0.017 | 0.022 | 0.028 | 0.029 | 0.027 | 0.032 | 0.038 | 0.013333 | 0.026333 | 0.032333 | 0.022833 | 744.359 | 0.0035 | 5.6975e-05
744.605 | 0.023 | 744.605 | 0.023 | 0.027 | 0.034 | 0.042 | 0.047 | 0.041 | 0.05 | 0.054 | 0.053 | 0.028 | 0.043333 | 0.052333 | 0.040167 | 744.605 | 0.0031667 | 5.1549e-05
744.851 | 0.046 | 744.851 | 0.046 | 0.049 | 0.052 | 0.06 | 0.067 | 0.062 | 0.072 | 0.072 | 0.074 | 0.049 | 0.063 | 0.072667 | 0.060833 | 744.851 | 0.0021667 | 3.527e-05
745.097 | 0.047 | 745.097 | 0.047 | 0.044 | 0.047 | 0.048 | 0.052 | 0.057 | 0.056 | 0.061 | 0.066 | 0.046 | 0.052333 | 0.061 | 0.0535 | 745.097 | -0.0011667 | -1.8992e-05
745.342 | 0.017 | 745.342 | 0.017 | 0.02 | 0.018 | 0.027 | 0.032 | 0.033 | 0.036 | 0.038 | 0.046 | 0.018333 | 0.030667 | 0.04 | 0.029167 | 745.342 | 0.0015 | 2.4418e-05
745.588 | 0.023 | 745.588 | 0.023 | 0.025 | 0.024 | 0.035 | 0.039 | 0.049 | 0.043 | 0.047 | 0.053 | 0.024 | 0.041 | 0.047667 | 0.035833 | 745.588 | 0.0051667 | 8.4106e-05
745.834 | 0.014 | 745.834 | 0.014 | 0.018 | 0.022 | 0.028 | 0.029 | 0.036 | 0.038 | 0.043 | 0.035 | 0.018 | 0.031 | 0.038667 | 0.028333 | 745.834 | 0.0026667 | 4.3409e-05
746.079 | 0.032 | 746.079 | 0.032 | 0.037 | 0.046 | 0.048 | 0.051 | 0.055 | 0.063 | 0.061 | 0.06 | 0.038333 | 0.051333 | 0.061333 | 0.049833 | 746.079 | 0.0015 | 2.4418e-05
746.325 | 0.025 | 746.325 | 0.025 | 0.028 | 0.023 | 0.04 | 0.041 | 0.048 | 0.045 | 0.05 | 0.059 | 0.025333 | 0.043 | 0.051333 | 0.038333 | 746.325 | 0.0046667 | 7.5967e-05
746.57 | 0.032 | 746.57 | 0.032 | 0.024 | 0.032 | 0.04 | 0.047 | 0.046 | 0.057 | 0.058 | 0.056 | 0.029333 | 0.044333 | 0.057 | 0.043167 | 746.57 | 0.0011667 | 1.8992e-05
746.816 | 0.003 | 746.816 | 0.003 | 0.009 | 0.009 | 0.009 | 0.008 | 0.018 | 0.021 | 0.023 | 0.018 | 0.007 | 0.011667 | 0.020667 | 0.013833 | 746.816 | -0.0021667 | -3.527e-05
747.061 | 0.02 | 747.061 | 0.02 | 0.021 | 0.032 | 0.039 | 0.04 | 0.045 | 0.053 | 0.044 | 0.044 | 0.024333 | 0.041333 | 0.047 | 0.035667 | 747.061 | 0.0056667 | 9.2245e-05
747.307 | 0.023 | 747.307 | 0.023 | 0.024 | 0.03 | 0.037 | 0.036 | 0.045 | 0.05 | 0.052 | 0.051 | 0.025667 | 0.039333 | 0.051 | 0.038333 | 747.307 | 0.001 | 1.6279e-05
747.552 | 0.013 | 747.552 | 0.013 | 0.012 | 0.024 | 0.026 | 0.032 | 0.037 | 0.034 | 0.043 | 0.037 | 0.016333 | 0.031667 | 0.038 | 0.027167 | 747.552 | 0.0045 | 7.3254e-05
747.797 | 0.031 | 747.797 | 0.031 | 0.037 | 0.042 | 0.042 | 0.052 | 0.053 | 0.054 | 0.058 | 0.053 | 0.036667 | 0.049 | 0.055 | 0.045833 | 747.797 | 0.0031667 | 5.1549e-05
748.042 | 0.018 | 748.042 | 0.018 | 0.023 | 0.024 | 0.038 | 0.027 | 0.037 | 0.043 | 0.046 | 0.051 | 0.021667 | 0.034 | 0.046667 | 0.034167 | 748.042 | -0.00016667 | -2.7131e-06
748.287 | 0.019 | 748.287 | 0.019 | 0.02 | 0.028 | 0.028 | 0.029 | 0.038 | 0.037 | 0.04 | 0.05 | 0.022333 | 0.031667 | 0.042333 | 0.032333 | 748.287 | -0.00066667 | -1.0852e-05
748.532 | 0.012 | 748.532 | 0.012 | 0.011 | 0.019 | 0.021 | 0.017 | 0.032 | 0.031 | 0.033 | 0.038 | 0.014 | 0.023333 | 0.034 | 0.024 | 748.532 | -0.00066667 | -1.0852e-05
748.777 | 0.027 | 748.777 | 0.027 | 0.034 | 0.026 | 0.038 | 0.039 | 0.04 | 0.046 | 0.045 | 0.051 | 0.029 | 0.039 | 0.047333 | 0.038167 | 748.777 | 0.00083333 | 1.3565e-05
749.022 | 0.033 | 749.022 | 0.033 | 0.031 | 0.034 | 0.042 | 0.038 | 0.041 | 0.049 | 0.056 | 0.053 | 0.032667 | 0.040333 | 0.052667 | 0.042667 | 749.022 | -0.0023333 | -3.7983e-05
749.267 | 0.021 | 749.267 | 0.021 | 0.028 | 0.039 | 0.036 | 0.046 | 0.049 | 0.05 | 0.055 | 0.06 | 0.029333 | 0.043667 | 0.055 | 0.042167 | 749.267 | 0.0015 | 2.4418e-05
749.512 | 0.011 | 749.512 | 0.011 | 0.019 | 0.025 | 0.03 | 0.028 | 0.028 | 0.028 | 0.036 | 0.04 | 0.018333 | 0.028667 | 0.034667 | 0.0265 | 749.512 | 0.0021667 | 3.527e-05
749.757 | 0.014 | 749.757 | 0.014 | 0.02 | 0.017 | 0.032 | 0.029 | 0.028 | 0.04 | 0.041 | 0.039 | 0.017 | 0.029667 | 0.04 | 0.0285 | 749.757 | 0.0011667 | 1.8992e-05
750.002 | 0.021 | 750.002 | 0.021 | 0.021 | 0.024 | 0.028 | 0.036 | 0.042 | 0.042 | 0.045 | 0.041 | 0.022 | 0.035333 | 0.042667 | 0.032333 | 750.002 | 0.003 | 4.8836e-05
750.246 | 0.03 | 750.246 | 0.03 | 0.02 | 0.03 | 0.038 | 0.044 | 0.054 | 0.037 | 0.045 | 0.047 | 0.026667 | 0.045333 | 0.043 | 0.034833 | 750.246 | 0.0105 | 0.00017092
750.491 | 0.025 | 750.491 | 0.025 | 0.03 | 0.035 | 0.056 | 0.048 | 0.05 | 0.044 | 0.052 | 0.053 | 0.03 | 0.051333 | 0.049667 | 0.039833 | 750.491 | 0.0115 | 0.0001872
750.736 | 0.01 | 750.736 | 0.01 | 0.015 | 0.02 | 0.029 | 0.03 | 0.033 | 0.034 | 0.035 | 0.039 | 0.015 | 0.030667 | 0.036 | 0.0255 | 750.736 | 0.0051667 | 8.4106e-05
750.98 | 0.024 | 750.98 | 0.024 | 0.024 | 0.034 | 0.042 | 0.052 | 0.043 | 0.042 | 0.058 | 0.057 | 0.027333 | 0.045667 | 0.052333 | 0.039833 | 750.98 | 0.0058333 | 9.4958e-05
751.225 | 0.018 | 751.225 | 0.018 | 0.022 | 0.023 | 0.04 | 0.038 | 0.05 | 0.042 | 0.044 | 0.055 | 0.021 | 0.042667 | 0.047 | 0.034 | 751.225 | 0.0086667 | 0.00014108
751.469 | 0.026 | 751.469 | 0.026 | 0.026 | 0.031 | 0.058 | 0.055 | 0.058 | 0.048 | 0.042 | 0.052 | 0.027667 | 0.057 | 0.047333 | 0.0375 | 751.469 | 0.0195 | 0.00031743
751.713 | 0.032 | 751.713 | 0.032 | 0.031 | 0.027 | 0.047 | 0.051 | 0.05 | 0.049 | 0.054 | 0.052 | 0.03 | 0.049333 | 0.051667 | 0.040833 | 751.713 | 0.0085 | 0.00013837
751.958 | 0.016 | 751.958 | 0.016 | 0.025 | 0.033 | 0.029 | 0.037 | 0.041 | 0.052 | 0.054 | 0.056 | 0.024667 | 0.035667 | 0.054 | 0.039333 | 751.958 | -0.0036667 | -5.9688e-05
752.202 | 0.008 | 752.202 | 0.008 | 0.019 | 0.013 | 0.027 | 0.03 | 0.029 | 0.036 | 0.034 | 0.036 | 0.013333 | 0.028667 | 0.035333 | 0.024333 | 752.202 | 0.0043333 | 7.054e-05
752.446 | 0.01 | 752.446 | 0.01 | 0.024 | 0.023 | 0.034 | 0.032 | 0.043 | 0.035 | 0.039 | 0.042 | 0.019 | 0.036333 | 0.038667 | 0.028833 | 752.446 | 0.0075 | 0.00012209
752.69 | 0.003 | 752.69 | 0.003 | 0.007 | 0.016 | 0.013 | 0.017 | 0.025 | 0.026 | 0.024 | 0.035 | 0.0086667 | 0.018333 | 0.028333 | 0.0185 | 752.69 | -0.00016667 | -2.7131e-06
752.934 | -0.004 | 752.934 | -0.004 | 0.01 | 0.01 | 0.015 | 0.022 | 0.016 | 0.024 | 0.032 | 0.028 | 0.0053333 | 0.017667 | 0.028 | 0.016667 | 752.934 | 0.001 | 1.6279e-05
753.178 | 0.028 | 753.178 | 0.028 | 0.026 | 0.036 | 0.04 | 0.033 | 0.047 | 0.046 | 0.052 | 0.048 | 0.03 | 0.04 | 0.048667 | 0.039333 | 753.178 | 0.00066667 | 1.0852e-05
753.422 | 0.004 | 753.422 | 0.004 | 0.013 | 0.021 | 0.021 | 0.03 | 0.033 | 0.028 | 0.042 | 0.05 | 0.012667 | 0.028 | 0.04 | 0.026333 | 753.422 | 0.0016667 | 2.7131e-05
753.666 | 0.035 | 753.666 | 0.035 | 0.033 | 0.045 | 0.048 | 0.053 | 0.057 | 0.068 | 0.059 | 0.065 | 0.037667 | 0.052667 | 0.064 | 0.050833 | 753.666 | 0.0018333 | 2.9844e-05
753.91 | 0.003 | 753.91 | 0.003 | 0.01 | 0.012 | 0.014 | 0.02 | 0.017 | 0.022 | 0.025 | 0.031 | 0.0083333 | 0.017 | 0.026 | 0.017167 | 753.91 | -0.00016667 | -2.7131e-06
754.154 | 0.029 | 754.154 | 0.029 | 0.034 | 0.034 | 0.047 | 0.04 | 0.053 | 0.051 | 0.061 | 0.057 | 0.032333 | 0.046667 | 0.056333 | 0.044333 | 754.154 | 0.0023333 | 3.7983e-05
754.398 | 0.056 | 754.398 | 0.056 | 0.058 | 0.061 | 0.074 | 0.067 | 0.072 | 0.08 | 0.078 | 0.082 | 0.058333 | 0.071 | 0.08 | 0.069167 | 754.398 | 0.0018333 | 2.9844e-05
754.641 | 0.001 | 754.641 | 0.001 | 0.005 | 0.011 | 0.012 | 0.017 | 0.025 | 0.027 | 0.03 | 0.027 | 0.0056667 | 0.018 | 0.028 | 0.016833 | 754.641 | 0.0011667 | 1.8992e-05
754.885 | 0.001 | 754.885 | 0.001 | 0.004 | 0.008 | 0.009 | 0.021 | 0.024 | 0.027 | 0.019 | 0.029 | 0.0043333 | 0.018 | 0.025 | 0.014667 | 754.885 | 0.0033333 | 5.4262e-05
755.129 | 0.021 | 755.129 | 0.021 | 0.021 | 0.032 | 0.035 | 0.034 | 0.048 | 0.048 | 0.048 | 0.047 | 0.024667 | 0.039 | 0.047667 | 0.036167 | 755.129 | 0.0028333 | 4.6123e-05
755.372 | -0.001 | 755.372 | -0.001 | -0.002 | 0.008 | 0.009 | 0.016 | 0.017 | 0.027 | 0.026 | 0.027 | 0.0016667 | 0.014 | 0.026667 | 0.014167 | 755.372 | -0.00016667 | -2.7131e-06
755.616 | 0.023 | 755.616 | 0.023 | 0.031 | 0.033 | 0.034 | 0.033 | 0.038 | 0.043 | 0.051 | 0.043 | 0.029 | 0.035 | 0.045667 | 0.037333 | 755.616 | -0.0023333 | -3.7983e-05
755.859 | 0.016 | 755.859 | 0.016 | 0.031 | 0.034 | 0.033 | 0.045 | 0.045 | 0.039 | 0.048 | 0.051 | 0.027 | 0.041 | 0.046 | 0.0365 | 755.859 | 0.0045 | 7.3254e-05
756.103 | 0.045 | 756.103 | 0.045 | 0.058 | 0.062 | 0.061 | 0.055 | 0.061 | 0.068 | 0.067 | 0.074 | 0.055 | 0.059 | 0.069667 | 0.062333 | 756.103 | -0.0033333 | -5.4262e-05
756.346 | -0.007 | 756.346 | -0.007 | -0.007 | 0.002 | 0.005 | 0.009 | 0.014 | 0.013 | 0.014 | 0.021 | -0.004 | 0.0093333 | 0.016 | 0.006 | 756.346 | 0.0033333 | 5.4262e-05
756.589 | 0.022 | 756.589 | 0.022 | 0.017 | 0.034 | 0.036 | 0.039 | 0.032 | 0.049 | 0.043 | 0.051 | 0.024333 | 0.035667 | 0.047667 | 0.036 | 756.589 | -0.00033333 | -5.4262e-06
756.832 | 0.014 | 756.832 | 0.014 | 0.019 | 0.026 | 0.027 | 0.033 | 0.037 | 0.036 | 0.043 | 0.054 | 0.019667 | 0.032333 | 0.044333 | 0.032 | 756.832 | 0.00033333 | 5.4262e-06
757.076 | 0.033 | 757.076 | 0.033 | 0.039 | 0.032 | 0.047 | 0.046 | 0.041 | 0.044 | 0.055 | 0.056 | 0.034667 | 0.044667 | 0.051667 | 0.043167 | 757.076 | 0.0015 | 2.4418e-05
757.319 | 0.029 | 757.319 | 0.029 | 0.036 | 0.037 | 0.04 | 0.04 | 0.038 | 0.049 | 0.054 | 0.051 | 0.034 | 0.039333 | 0.051333 | 0.042667 | 757.319 | -0.0033333 | -5.4262e-05
757.562 | 0.026 | 757.562 | 0.026 | 0.033 | 0.036 | 0.05 | 0.04 | 0.055 | 0.057 | 0.057 | 0.061 | 0.031667 | 0.048333 | 0.058333 | 0.045 | 757.562 | 0.0033333 | 5.4262e-05
757.805 | 0.029 | 757.805 | 0.029 | 0.032 | 0.04 | 0.044 | 0.056 | 0.045 | 0.056 | 0.055 | 0.06 | 0.033667 | 0.048333 | 0.057 | 0.045333 | 757.805 | 0.003 | 4.8836e-05
758.048 | 0.021 | 758.048 | 0.021 | 0.038 | 0.036 | 0.037 | 0.041 | 0.046 | 0.054 | 0.058 | 0.057 | 0.031667 | 0.041333 | 0.056333 | 0.044 | 758.048 | -0.0026667 | -4.3409e-05
758.291 | 0.022 | 758.291 | 0.022 | 0.02 | 0.026 | 0.026 | 0.034 | 0.033 | 0.043 | 0.045 | 0.049 | 0.022667 | 0.031 | 0.045667 | 0.034167 | 758.291 | -0.0031667 | -5.1549e-05
758.533 | 0.012 | 758.533 | 0.012 | 0.024 | 0.029 | 0.028 | 0.029 | 0.038 | 0.038 | 0.046 | 0.044 | 0.021667 | 0.031667 | 0.042667 | 0.032167 | 758.533 | -0.0005 | -8.1393e-06
758.776 | 0.009 | 758.776 | 0.009 | 0.016 | 0.014 | 0.033 | 0.037 | 0.034 | 0.037 | 0.044 | 0.042 | 0.013 | 0.034667 | 0.041 | 0.027 | 758.776 | 0.0076667 | 0.0001248
759.019 | 0.036 | 759.019 | 0.036 | 0.026 | 0.034 | 0.045 | 0.048 | 0.052 | 0.052 | 0.059 | 0.061 | 0.032 | 0.048333 | 0.057333 | 0.044667 | 759.019 | 0.0036667 | 5.9688e-05
759.262 | 0.029 | 759.262 | 0.029 | 0.04 | 0.036 | 0.041 | 0.039 | 0.044 | 0.049 | 0.058 | 0.055 | 0.035 | 0.041333 | 0.054 | 0.0445 | 759.262 | -0.0031667 | -5.1549e-05
759.504 | 0.021 | 759.504 | 0.021 | 0.032 | 0.034 | 0.038 | 0.038 | 0.041 | 0.05 | 0.048 | 0.053 | 0.029 | 0.039 | 0.050333 | 0.039667 | 759.504 | -0.00066667 | -1.0852e-05
759.747 | 0.021 | 759.747 | 0.021 | 0.027 | 0.029 | 0.039 | 0.033 | 0.042 | 0.037 | 0.042 | 0.052 | 0.025667 | 0.038 | 0.043667 | 0.034667 | 759.747 | 0.0033333 | 5.4262e-05
759.989 | 0.03 | 759.989 | 0.03 | 0.039 | 0.034 | 0.048 | 0.04 | 0.052 | 0.057 | 0.057 | 0.055 | 0.034333 | 0.046667 | 0.056333 | 0.045333 | 759.989 | 0.0013333 | 2.1705e-05
760.232 | 0.039 | 760.232 | 0.039 | 0.036 | 0.04 | 0.047 | 0.051 | 0.055 | 0.063 | 0.067 | 0.067 | 0.038333 | 0.051 | 0.065667 | 0.052 | 760.232 | -0.001 | -1.6279e-05
760.474 | 0.007 | 760.474 | 0.007 | 0.02 | 0.015 | 0.029 | 0.03 | 0.031 | 0.031 | 0.038 | 0.035 | 0.014 | 0.03 | 0.034667 | 0.024333 | 760.474 | 0.0056667 | 9.2245e-05
760.717 | 0.049 | 760.717 | 0.049 | 0.055 | 0.058 | 0.066 | 0.068 | 0.076 | 0.066 | 0.084 | 0.084 | 0.054 | 0.07 | 0.078 | 0.066 | 760.717 | 0.004 | 6.5114e-05
760.959 | 0.01 | 760.959 | 0.01 | 0.018 | 0.024 | 0.025 | 0.029 | 0.031 | 0.031 | 0.032 | 0.042 | 0.017333 | 0.028333 | 0.035 | 0.026167 | 760.959 | 0.0021667 | 3.527e-05
761.201 | 0.011 | 761.201 | 0.011 | 0.018 | 0.014 | 0.026 | 0.029 | 0.034 | 0.034 | 0.034 | 0.041 | 0.014333 | 0.029667 | 0.036333 | 0.025333 | 761.201 | 0.0043333 | 7.054e-05
761.444 | 0.02 | 761.444 | 0.02 | 0.027 | 0.027 | 0.033 | 0.037 | 0.037 | 0.049 | 0.042 | 0.05 | 0.024667 | 0.035667 | 0.047 | 0.035833 | 761.444 | -0.00016667 | -2.7131e-06
761.686 | 0.034 | 761.686 | 0.034 | 0.031 | 0.047 | 0.045 | 0.044 | 0.056 | 0.059 | 0.062 | 0.063 | 0.037333 | 0.048333 | 0.061333 | 0.049333 | 761.686 | -0.001 | -1.6279e-05
761.928 | 0.013 | 761.928 | 0.013 | 0.016 | 0.022 | 0.04 | 0.035 | 0.043 | 0.037 | 0.043 | 0.047 | 0.017 | 0.039333 | 0.042333 | 0.029667 | 761.928 | 0.0096667 | 0.00015736
762.17 | 0.014 | 762.17 | 0.014 | 0.02 | 0.033 | 0.029 | 0.036 | 0.034 | 0.046 | 0.051 | 0.051 | 0.022333 | 0.033 | 0.049333 | 0.035833 | 762.17 | -0.0028333 | -4.6123e-05
762.412 | 0.028 | 762.412 | 0.028 | 0.029 | 0.042 | 0.044 | 0.049 | 0.05 | 0.058 | 0.062 | 0.066 | 0.033 | 0.047667 | 0.062 | 0.0475 | 762.412 | 0.00016667 | 2.7131e-06
762.654 | 0.033 | 762.654 | 0.033 | 0.037 | 0.04 | 0.042 | 0.044 | 0.05 | 0.052 | 0.054 | 0.058 | 0.036667 | 0.045333 | 0.054667 | 0.045667 | 762.654 | -0.00033333 | -5.4262e-06
762.896 | 0.016 | 762.896 | 0.016 | 0.018 | 0.028 | 0.039 | 0.027 | 0.04 | 0.037 | 0.042 | 0.047 | 0.020667 | 0.035333 | 0.042 | 0.031333 | 762.896 | 0.004 | 6.5114e-05
763.137 | 0.027 | 763.137 | 0.027 | 0.024 | 0.029 | 0.046 | 0.044 | 0.049 | 0.056 | 0.049 | 0.065 | 0.026667 | 0.046333 | 0.056667 | 0.041667 | 763.137 | 0.0046667 | 7.5967e-05
763.379 | 0.011 | 763.379 | 0.011 | 0.013 | 0.021 | 0.058 | 0.061 | 0.064 | 0.029 | 0.035 | 0.035 | 0.015 | 0.061 | 0.033 | 0.024 | 763.379 | 0.037 | 0.00060231
763.621 | 0.031 | 763.621 | 0.031 | 0.035 | 0.041 | 0.096 | 0.088 | 0.099 | 0.052 | 0.063 | 0.072 | 0.035667 | 0.094333 | 0.062333 | 0.049 | 763.621 | 0.045333 | 0.00073796
763.863 | 0.02 | 763.863 | 0.02 | 0.029 | 0.031 | 0.047 | 0.054 | 0.059 | 0.041 | 0.044 | 0.049 | 0.026667 | 0.053333 | 0.044667 | 0.035667 | 763.863 | 0.017667 | 0.00028759
764.104 | 0.019 | 764.104 | 0.019 | 0.029 | 0.026 | 0.042 | 0.036 | 0.042 | 0.051 | 0.052 | 0.05 | 0.024667 | 0.04 | 0.051 | 0.037833 | 764.104 | 0.0021667 | 3.527e-05
764.346 | 0.033 | 764.346 | 0.033 | 0.05 | 0.043 | 0.066 | 0.061 | 0.068 | 0.066 | 0.069 | 0.083 | 0.042 | 0.065 | 0.072667 | 0.057333 | 764.346 | 0.0076667 | 0.0001248
764.587 | 0.024 | 764.587 | 0.024 | 0.021 | 0.024 | 0.029 | 0.033 | 0.041 | 0.037 | 0.037 | 0.044 | 0.023 | 0.034333 | 0.039333 | 0.031167 | 764.587 | 0.0031667 | 5.1549e-05
764.829 | 0.015 | 764.829 | 0.015 | 0.025 | 0.027 | 0.029 | 0.035 | 0.045 | 0.045 | 0.05 | 0.05 | 0.022333 | 0.036333 | 0.048333 | 0.035333 | 764.829 | 0.001 | 1.6279e-05
765.07 | 0.007 | 765.07 | 0.007 | 0.009 | 0.016 | 0.018 | 0.025 | 0.034 | 0.028 | 0.038 | 0.041 | 0.010667 | 0.025667 | 0.035667 | 0.023167 | 765.07 | 0.0025 | 4.0696e-05
765.312 | 0.026 | 765.312 | 0.026 | 0.051 | 0.045 | 0.051 | 0.057 | 0.057 | 0.067 | 0.066 | 0.074 | 0.040667 | 0.055 | 0.069 | 0.054833 | 765.312 | 0.00016667 | 2.7131e-06
765.553 | 0.033 | 765.553 | 0.033 | 0.05 | 0.051 | 0.055 | 0.054 | 0.063 | 0.064 | 0.072 | 0.088 | 0.044667 | 0.057333 | 0.074667 | 0.059667 | 765.553 | -0.0023333 | -3.7983e-05
765.794 | 0.039 | 765.794 | 0.039 | 0.033 | 0.041 | 0.044 | 0.049 | 0.058 | 0.066 | 0.057 | 0.056 | 0.037667 | 0.050333 | 0.059667 | 0.048667 | 765.794 | 0.0016667 | 2.7131e-05
766.035 | 0.035 | 766.035 | 0.035 | 0.04 | 0.036 | 0.042 | 0.053 | 0.044 | 0.059 | 0.057 | 0.057 | 0.037 | 0.046333 | 0.057667 | 0.047333 | 766.035 | -0.001 | -1.6279e-05
766.276 | 0.048 | 766.276 | 0.048 | 0.045 | 0.047 | 0.061 | 0.061 | 0.065 | 0.061 | 0.083 | 0.073 | 0.046667 | 0.062333 | 0.072333 | 0.0595 | 766.276 | 0.0028333 | 4.6123e-05
766.518 | 0.034 | 766.518 | 0.034 | 0.032 | 0.036 | 0.04 | 0.045 | 0.049 | 0.058 | 0.052 | 0.061 | 0.034 | 0.044667 | 0.057 | 0.0455 | 766.518 | -0.00083333 | -1.3565e-05
766.759 | 0.041 | 766.759 | 0.041 | 0.052 | 0.059 | 0.064 | 0.06 | 0.063 | 0.077 | 0.071 | 0.07 | 0.050667 | 0.062333 | 0.072667 | 0.061667 | 766.759 | 0.00066667 | 1.0852e-05
767 | 0.022 | 767 | 0.022 | 0.027 | 0.031 | 0.04 | 0.032 | 0.043 | 0.04 | 0.049 | 0.05 | 0.026667 | 0.038333 | 0.046333 | 0.0365 | 767 | 0.0018333 | 2.9844e-05
767.241 | 0.029 | 767.241 | 0.029 | 0.033 | 0.04 | 0.044 | 0.04 | 0.053 | 0.065 | 0.053 | 0.059 | 0.034 | 0.045667 | 0.059 | 0.0465 | 767.241 | -0.00083333 | -1.3565e-05
767.481 | 0.01 | 767.481 | 0.01 | 0.029 | 0.028 | 0.034 | 0.038 | 0.044 | 0.049 | 0.049 | 0.048 | 0.022333 | 0.038667 | 0.048667 | 0.0355 | 767.481 | 0.0031667 | 5.1549e-05
767.722 | 0.034 | 767.722 | 0.034 | 0.027 | 0.029 | 0.046 | 0.041 | 0.051 | 0.051 | 0.055 | 0.05 | 0.03 | 0.046 | 0.052 | 0.041 | 767.722 | 0.005 | 8.1393e-05
767.963 | 0.029 | 767.963 | 0.029 | 0.038 | 0.036 | 0.039 | 0.052 | 0.059 | 0.05 | 0.059 | 0.061 | 0.034333 | 0.05 | 0.056667 | 0.0455 | 767.963 | 0.0045 | 7.3254e-05
768.204 | 0.028 | 768.204 | 0.028 | 0.029 | 0.045 | 0.051 | 0.047 | 0.052 | 0.063 | 0.058 | 0.064 | 0.034 | 0.05 | 0.061667 | 0.047833 | 768.204 | 0.0021667 | 3.527e-05
768.444 | 0.053 | 768.444 | 0.053 | 0.048 | 0.055 | 0.064 | 0.059 | 0.064 | 0.079 | 0.072 | 0.084 | 0.052 | 0.062333 | 0.078333 | 0.065167 | 768.444 | -0.0028333 | -4.6123e-05
768.685 | 0.039 | 768.685 | 0.039 | 0.039 | 0.036 | 0.04 | 0.055 | 0.055 | 0.057 | 0.057 | 0.064 | 0.038 | 0.05 | 0.059333 | 0.048667 | 768.685 | 0.0013333 | 2.1705e-05
768.926 | 0.01 | 768.926 | 0.01 | 0.015 | 0.009 | 0.02 | 0.026 | 0.035 | 0.036 | 0.035 | 0.035 | 0.011333 | 0.027 | 0.035333 | 0.023333 | 768.926 | 0.0036667 | 5.9688e-05
769.166 | 0.023 | 769.166 | 0.023 | 0.026 | 0.026 | 0.035 | 0.028 | 0.04 | 0.044 | 0.042 | 0.053 | 0.025 | 0.034333 | 0.046333 | 0.035667 | 769.166 | -0.0013333 | -2.1705e-05
769.407 | 0.015 | 769.407 | 0.015 | 0.02 | 0.023 | 0.035 | 0.034 | 0.044 | 0.044 | 0.046 | 0.048 | 0.019333 | 0.037667 | 0.046 | 0.032667 | 769.407 | 0.005 | 8.1393e-05
769.647 | 0.048 | 769.647 | 0.048 | 0.054 | 0.056 | 0.072 | 0.068 | 0.074 | 0.077 | 0.078 | 0.087 | 0.052667 | 0.071333 | 0.080667 | 0.066667 | 769.647 | 0.0046667 | 7.5967e-05
769.887 | 0.01 | 769.887 | 0.01 | 0.005 | 0.015 | 0.014 | 0.023 | 0.029 | 0.029 | 0.037 | 0.03 | 0.01 | 0.022 | 0.032 | 0.021 | 769.887 | 0.001 | 1.6279e-05
770.128 | 0.033 | 770.128 | 0.033 | 0.035 | 0.041 | 0.045 | 0.049 | 0.064 | 0.057 | 0.057 | 0.064 | 0.036333 | 0.052667 | 0.059333 | 0.047833 | 770.128 | 0.0048333 | 7.868e-05
770.368 | 0.033 | 770.368 | 0.033 | 0.033 | 0.038 | 0.031 | 0.043 | 0.048 | 0.049 | 0.054 | 0.061 | 0.034667 | 0.040667 | 0.054667 | 0.044667 | 770.368 | -0.004 | -6.5114e-05
770.608 | 0.037 | 770.608 | 0.037 | 0.045 | 0.049 | 0.056 | 0.051 | 0.062 | 0.066 | 0.073 | 0.072 | 0.043667 | 0.056333 | 0.070333 | 0.057 | 770.608 | -0.00066667 | -1.0852e-05
770.848 | 0.032 | 770.848 | 0.032 | 0.04 | 0.041 | 0.042 | 0.052 | 0.052 | 0.057 | 0.057 | 0.069 | 0.037667 | 0.048667 | 0.061 | 0.049333 | 770.848 | -0.00066667 | -1.0852e-05
771.088 | 0.014 | 771.088 | 0.014 | 0.028 | 0.029 | 0.032 | 0.031 | 0.044 | 0.051 | 0.052 | 0.051 | 0.023667 | 0.035667 | 0.051333 | 0.0375 | 771.088 | -0.0018333 | -2.9844e-05
771.328 | 0.032 | 771.328 | 0.032 | 0.036 | 0.039 | 0.04 | 0.048 | 0.047 | 0.069 | 0.06 | 0.068 | 0.035667 | 0.045 | 0.065667 | 0.050667 | 771.328 | -0.0056667 | -9.2245e-05
771.568 | 0.018 | 771.568 | 0.018 | 0.021 | 0.028 | 0.032 | 0.025 | 0.04 | 0.041 | 0.039 | 0.045 | 0.022333 | 0.032333 | 0.041667 | 0.032 | 771.568 | 0.00033333 | 5.4262e-06
771.808 | 0.03 | 771.808 | 0.03 | 0.033 | 0.045 | 0.053 | 0.047 | 0.053 | 0.056 | 0.062 | 0.062 | 0.036 | 0.051 | 0.06 | 0.048 | 771.808 | 0.003 | 4.8836e-05
772.048 | 0.015 | 772.048 | 0.015 | 0.02 | 0.024 | 0.029 | 0.037 | 0.043 | 0.043 | 0.053 | 0.053 | 0.019667 | 0.036333 | 0.049667 | 0.034667 | 772.048 | 0.0016667 | 2.7131e-05
772.288 | 0.032 | 772.288 | 0.032 | 0.025 | 0.039 | 0.044 | 0.043 | 0.062 | 0.056 | 0.06 | 0.059 | 0.032 | 0.049667 | 0.058333 | 0.045167 | 772.288 | 0.0045 | 7.3254e-05
772.528 | 0.011 | 772.528 | 0.011 | 0.009 | 0.013 | 0.039 | 0.045 | 0.049 | 0.04 | 0.042 | 0.039 | 0.011 | 0.044333 | 0.040333 | 0.025667 | 772.528 | 0.018667 | 0.00030387
772.767 | 0.021 | 772.767 | 0.021 | 0.022 | 0.027 | 0.041 | 0.047 | 0.049 | 0.052 | 0.057 | 0.049 | 0.023333 | 0.045667 | 0.052667 | 0.038 | 772.767 | 0.0076667 | 0.0001248
773.007 | 0.03 | 773.007 | 0.03 | 0.033 | 0.032 | 0.051 | 0.048 | 0.052 | 0.056 | 0.061 | 0.063 | 0.031667 | 0.050333 | 0.06 | 0.045833 | 773.007 | 0.0045 | 7.3254e-05
773.247 | 0.037 | 773.247 | 0.037 | 0.04 | 0.035 | 0.045 | 0.048 | 0.057 | 0.064 | 0.066 | 0.064 | 0.037333 | 0.05 | 0.064667 | 0.051 | 773.247 | -0.001 | -1.6279e-05
773.486 | 0.023 | 773.486 | 0.023 | 0.03 | 0.027 | 0.032 | 0.042 | 0.042 | 0.047 | 0.054 | 0.05 | 0.026667 | 0.038667 | 0.050333 | 0.0385 | 773.486 | 0.00016667 | 2.7131e-06
773.726 | 0.01 | 773.726 | 0.01 | 0.003 | 0.016 | 0.022 | 0.027 | 0.031 | 0.039 | 0.039 | 0.044 | 0.0096667 | 0.026667 | 0.040667 | 0.025167 | 773.726 | 0.0015 | 2.4418e-05
773.965 | 0.036 | 773.965 | 0.036 | 0.043 | 0.043 | 0.056 | 0.054 | 0.066 | 0.069 | 0.072 | 0.076 | 0.040667 | 0.058667 | 0.072333 | 0.0565 | 773.965 | 0.0021667 | 3.527e-05
774.205 | 0.029 | 774.205 | 0.029 | 0.038 | 0.032 | 0.051 | 0.059 | 0.054 | 0.069 | 0.061 | 0.065 | 0.033 | 0.054667 | 0.065 | 0.049 | 774.205 | 0.0056667 | 9.2245e-05
774.444 | 0.029 | 774.444 | 0.029 | 0.036 | 0.044 | 0.045 | 0.057 | 0.061 | 0.057 | 0.063 | 0.07 | 0.036333 | 0.054333 | 0.063333 | 0.049833 | 774.444 | 0.0045 | 7.3254e-05
774.683 | 0.031 | 774.683 | 0.031 | 0.042 | 0.04 | 0.053 | 0.067 | 0.062 | 0.063 | 0.075 | 0.076 | 0.037667 | 0.060667 | 0.071333 | 0.0545 | 774.683 | 0.0061667 | 0.00010038
774.922 | 0.032 | 774.922 | 0.032 | 0.05 | 0.052 | 0.057 | 0.063 | 0.062 | 0.066 | 0.074 | 0.078 | 0.044667 | 0.060667 | 0.072667 | 0.058667 | 774.922 | 0.002 | 3.2557e-05
775.162 | 0.025 | 775.162 | 0.025 | 0.027 | 0.042 | 0.05 | 0.042 | 0.052 | 0.066 | 0.056 | 0.067 | 0.031333 | 0.048 | 0.063 | 0.047167 | 775.162 | 0.00083333 | 1.3565e-05
775.401 | 0.04 | 775.401 | 0.04 | 0.046 | 0.054 | 0.062 | 0.068 | 0.065 | 0.077 | 0.076 | 0.075 | 0.046667 | 0.065 | 0.076 | 0.061333 | 775.401 | 0.0036667 | 5.9688e-05
775.64 | 0.019 | 775.64 | 0.019 | 0.03 | 0.032 | 0.043 | 0.044 | 0.054 | 0.053 | 0.055 | 0.061 | 0.027 | 0.047 | 0.056333 | 0.041667 | 775.64 | 0.0053333 | 8.6819e-05
775.879 | 0.036 | 775.879 | 0.036 | 0.032 | 0.036 | 0.044 | 0.05 | 0.048 | 0.064 | 0.066 | 0.064 | 0.034667 | 0.047333 | 0.064667 | 0.049667 | 775.879 | -0.0023333 | -3.7983e-05
776.118 | 0.022 | 776.118 | 0.022 | 0.03 | 0.031 | 0.044 | 0.045 | 0.044 | 0.049 | 0.054 | 0.053 | 0.027667 | 0.044333 | 0.052 | 0.039833 | 776.118 | 0.0045 | 7.3254e-05
776.357 | 0.026 | 776.357 | 0.026 | 0.03 | 0.028 | 0.043 | 0.042 | 0.04 | 0.045 | 0.055 | 0.057 | 0.028 | 0.041667 | 0.052333 | 0.040167 | 776.357 | 0.0015 | 2.4418e-05
776.596 | 0.02 | 776.596 | 0.02 | 0.037 | 0.031 | 0.044 | 0.056 | 0.056 | 0.064 | 0.075 | 0.075 | 0.029333 | 0.052 | 0.071333 | 0.050333 | 776.596 | 0.0016667 | 2.7131e-05
776.834 | 0.048 | 776.834 | 0.048 | 0.05 | 0.057 | 0.051 | 0.061 | 0.059 | 0.07 | 0.069 | 0.079 | 0.051667 | 0.057 | 0.072667 | 0.062167 | 776.834 | -0.0051667 | -8.4106e-05
777.073 | 0.032 | 777.073 | 0.032 | 0.033 | 0.035 | 0.045 | 0.053 | 0.059 | 0.063 | 0.064 | 0.065 | 0.033333 | 0.052333 | 0.064 | 0.048667 | 777.073 | 0.0036667 | 5.9688e-05
777.312 | 0.041 | 777.312 | 0.041 | 0.034 | 0.04 | 0.062 | 0.055 | 0.064 | 0.061 | 0.068 | 0.069 | 0.038333 | 0.060333 | 0.066 | 0.052167 | 777.312 | 0.0081667 | 0.00013294
777.551 | 0.016 | 777.551 | 0.016 | 0.017 | 0.022 | 0.032 | 0.041 | 0.036 | 0.037 | 0.044 | 0.047 | 0.018333 | 0.036333 | 0.042667 | 0.0305 | 777.551 | 0.0058333 | 9.4958e-05
777.789 | 0.003 | 777.789 | 0.003 | 0.015 | 0.016 | 0.027 | 0.021 | 0.039 | 0.033 | 0.043 | 0.036 | 0.011333 | 0.029 | 0.037333 | 0.024333 | 777.789 | 0.0046667 | 7.5967e-05
778.028 | 0.032 | 778.028 | 0.032 | 0.025 | 0.041 | 0.041 | 0.044 | 0.052 | 0.049 | 0.056 | 0.055 | 0.032667 | 0.045667 | 0.053333 | 0.043 | 778.028 | 0.0026667 | 4.3409e-05
778.266 | 0.027 | 778.266 | 0.027 | 0.044 | 0.04 | 0.049 | 0.054 | 0.057 | 0.06 | 0.064 | 0.067 | 0.037 | 0.053333 | 0.063667 | 0.050333 | 778.266 | 0.003 | 4.8836e-05
778.505 | 0.012 | 778.505 | 0.012 | 0.023 | 0.025 | 0.025 | 0.042 | 0.032 | 0.043 | 0.04 | 0.04 | 0.02 | 0.033 | 0.041 | 0.0305 | 778.505 | 0.0025 | 4.0696e-05
778.743 | 0.016 | 778.743 | 0.016 | 0.018 | 0.026 | 0.03 | 0.032 | 0.032 | 0.035 | 0.045 | 0.036 | 0.02 | 0.031333 | 0.038667 | 0.029333 | 778.743 | 0.002 | 3.2557e-05
778.982 | -0.001 | 778.982 | -0.001 | 0.014 | 0.006 | 0.022 | 0.027 | 0.016 | 0.035 | 0.033 | 0.037 | 0.0063333 | 0.021667 | 0.035 | 0.020667 | 778.982 | 0.001 | 1.6279e-05
779.22 | 0.028 | 779.22 | 0.028 | 0.014 | 0.031 | 0.041 | 0.04 | 0.044 | 0.052 | 0.048 | 0.055 | 0.024333 | 0.041667 | 0.051667 | 0.038 | 779.22 | 0.0036667 | 5.9688e-05
779.458 | 0.031 | 779.458 | 0.031 | 0.037 | 0.039 | 0.051 | 0.052 | 0.051 | 0.055 | 0.068 | 0.064 | 0.035667 | 0.051333 | 0.062333 | 0.049 | 779.458 | 0.0023333 | 3.7983e-05
779.696 | 0.019 | 779.696 | 0.019 | 0.027 | 0.028 | 0.047 | 0.044 | 0.041 | 0.057 | 0.069 | 0.058 | 0.024667 | 0.044 | 0.061333 | 0.043 | 779.696 | 0.001 | 1.6279e-05
779.934 | 0.039 | 779.934 | 0.039 | 0.039 | 0.047 | 0.057 | 0.056 | 0.065 | 0.061 | 0.068 | 0.075 | 0.041667 | 0.059333 | 0.068 | 0.054833 | 779.934 | 0.0045 | 7.3254e-05
780.173 | 0.039 | 780.173 | 0.039 | 0.037 | 0.039 | 0.053 | 0.049 | 0.055 | 0.065 | 0.059 | 0.059 | 0.038333 | 0.052333 | 0.061 | 0.049667 | 780.173 | 0.0026667 | 4.3409e-05
780.411 | 0.016 | 780.411 | 0.016 | 0.037 | 0.025 | 0.049 | 0.048 | 0.043 | 0.061 | 0.056 | 0.053 | 0.026 | 0.046667 | 0.056667 | 0.041333 | 780.411 | 0.0053333 | 8.6819e-05
780.649 | 0.026 | 780.649 | 0.026 | 0.023 | 0.021 | 0.047 | 0.045 | 0.046 | 0.052 | 0.055 | 0.054 | 0.023333 | 0.046 | 0.053667 | 0.0385 | 780.649 | 0.0075 | 0.00012209
780.887 | 0.031 | 780.887 | 0.031 | 0.036 | 0.034 | 0.047 | 0.048 | 0.047 | 0.049 | 0.065 | 0.059 | 0.033667 | 0.047333 | 0.057667 | 0.045667 | 780.887 | 0.0016667 | 2.7131e-05
781.124 | 0.031 | 781.124 | 0.031 | 0.047 | 0.047 | 0.062 | 0.057 | 0.069 | 0.064 | 0.071 | 0.075 | 0.041667 | 0.062667 | 0.07 | 0.055833 | 781.124 | 0.0068333 | 0.00011124
781.362 | 0.015 | 781.362 | 0.015 | 0.014 | 0.017 | 0.038 | 0.028 | 0.035 | 0.049 | 0.05 | 0.053 | 0.015333 | 0.033667 | 0.050667 | 0.033 | 781.362 | 0.00066667 | 1.0852e-05
781.6 | 0.033 | 781.6 | 0.033 | 0.036 | 0.043 | 0.048 | 0.057 | 0.053 | 0.068 | 0.066 | 0.062 | 0.037333 | 0.052667 | 0.065333 | 0.051333 | 781.6 | 0.0013333 | 2.1705e-05
781.838 | 0.018 | 781.838 | 0.018 | 0.027 | 0.029 | 0.037 | 0.034 | 0.052 | 0.041 | 0.048 | 0.055 | 0.024667 | 0.041 | 0.048 | 0.036333 | 781.838 | 0.0046667 | 7.5967e-05
782.075 | 0.009 | 782.075 | 0.009 | 0.017 | 0.036 | 0.036 | 0.043 | 0.031 | 0.042 | 0.045 | 0.058 | 0.020667 | 0.036667 | 0.048333 | 0.0345 | 782.075 | 0.0021667 | 3.527e-05
782.313 | 0.015 | 782.313 | 0.015 | 0.027 | 0.024 | 0.037 | 0.039 | 0.043 | 0.044 | 0.049 | 0.06 | 0.022 | 0.039667 | 0.051 | 0.0365 | 782.313 | 0.0031667 | 5.1549e-05
782.551 | -0.004 | 782.551 | -0.004 | 0.007 | 0.005 | 0.024 | 0.016 | 0.025 | 0.031 | 0.03 | 0.034 | 0.0026667 | 0.021667 | 0.031667 | 0.017167 | 782.551 | 0.0045 | 7.3254e-05
782.788 | 0.015 | 782.788 | 0.015 | 0.019 | 0.022 | 0.022 | 0.021 | 0.023 | 0.041 | 0.048 | 0.052 | 0.018667 | 0.022 | 0.047 | 0.032833 | 782.788 | -0.010833 | -0.00017635
783.026 | 0.018 | 783.026 | 0.018 | 0.027 | 0.031 | 0.043 | 0.037 | 0.048 | 0.048 | 0.058 | 0.055 | 0.025333 | 0.042667 | 0.053667 | 0.0395 | 783.026 | 0.0031667 | 5.1549e-05
783.263 | 0.019 | 783.263 | 0.019 | 0.028 | 0.027 | 0.035 | 0.045 | 0.048 | 0.048 | 0.052 | 0.062 | 0.024667 | 0.042667 | 0.054 | 0.039333 | 783.263 | 0.0033333 | 5.4262e-05
783.5 | 0.015 | 783.5 | 0.015 | 0.026 | 0.031 | 0.045 | 0.035 | 0.048 | 0.04 | 0.048 | 0.051 | 0.024 | 0.042667 | 0.046333 | 0.035167 | 783.5 | 0.0075 | 0.00012209
783.738 | 0.028 | 783.738 | 0.028 | 0.025 | 0.04 | 0.044 | 0.049 | 0.051 | 0.055 | 0.062 | 0.067 | 0.031 | 0.048 | 0.061333 | 0.046167 | 783.738 | 0.0018333 | 2.9844e-05
783.975 | 0.011 | 783.975 | 0.011 | 0.013 | 0.006 | 0.026 | 0.027 | 0.029 | 0.032 | 0.043 | 0.048 | 0.01 | 0.027333 | 0.041 | 0.0255 | 783.975 | 0.0018333 | 2.9844e-05
784.212 | -0.017 | 784.212 | -0.017 | -0.009 | -0.015 | -0.001 | 0.006 | 0 | 0.018 | 0.009 | 0.02 | -0.013667 | 0.0016667 | 0.015667 | 0.001 | 784.212 | 0.00066667 | 1.0852e-05
784.449 | 0.017 | 784.449 | 0.017 | 0.018 | 0.032 | 0.044 | 0.042 | 0.044 | 0.045 | 0.052 | 0.052 | 0.022333 | 0.043333 | 0.049667 | 0.036 | 784.449 | 0.0073333 | 0.00011938
784.686 | 0.013 | 784.686 | 0.013 | 0.019 | 0.024 | 0.032 | 0.031 | 0.034 | 0.046 | 0.052 | 0.059 | 0.018667 | 0.032333 | 0.052333 | 0.0355 | 784.686 | -0.0031667 | -5.1549e-05
784.924 | 0.021 | 784.924 | 0.021 | 0.031 | 0.032 | 0.04 | 0.042 | 0.051 | 0.044 | 0.06 | 0.062 | 0.028 | 0.044333 | 0.055333 | 0.041667 | 784.924 | 0.0026667 | 4.3409e-05
785.161 | 0.005 | 785.161 | 0.005 | 0.017 | 0.014 | 0.024 | 0.032 | 0.03 | 0.027 | 0.042 | 0.043 | 0.012 | 0.028667 | 0.037333 | 0.024667 | 785.161 | 0.004 | 6.5114e-05
785.397 | 0.021 | 785.397 | 0.021 | 0.028 | 0.03 | 0.034 | 0.048 | 0.045 | 0.057 | 0.055 | 0.061 | 0.026333 | 0.042333 | 0.057667 | 0.042 | 785.397 | 0.00033333 | 5.4262e-06
785.634 | 0.037 | 785.634 | 0.037 | 0.043 | 0.039 | 0.048 | 0.053 | 0.065 | 0.057 | 0.063 | 0.07 | 0.039667 | 0.055333 | 0.063333 | 0.0515 | 785.634 | 0.0038333 | 6.2401e-05
785.871 | 0.045 | 785.871 | 0.045 | 0.053 | 0.053 | 0.06 | 0.073 | 0.059 | 0.071 | 0.092 | 0.083 | 0.050333 | 0.064 | 0.082 | 0.066167 | 785.871 | -0.0021667 | -3.527e-05
786.108 | 0.018 | 786.108 | 0.018 | 0.021 | 0.014 | 0.044 | 0.038 | 0.042 | 0.055 | 0.058 | 0.053 | 0.017667 | 0.041333 | 0.055333 | 0.0365 | 786.108 | 0.0048333 | 7.868e-05
786.345 | 0.045 | 786.345 | 0.045 | 0.044 | 0.063 | 0.072 | 0.078 | 0.073 | 0.087 | 0.088 | 0.086 | 0.050667 | 0.074333 | 0.087 | 0.068833 | 786.345 | 0.0055 | 8.9532e-05
786.581 | 0.011 | 786.581 | 0.011 | 0.024 | 0.022 | 0.029 | 0.032 | 0.042 | 0.049 | 0.046 | 0.062 | 0.019 | 0.034333 | 0.052333 | 0.035667 | 786.581 | -0.0013333 | -2.1705e-05
786.818 | 0.033 | 786.818 | 0.033 | 0.04 | 0.04 | 0.058 | 0.061 | 0.057 | 0.073 | 0.07 | 0.073 | 0.037667 | 0.058667 | 0.072 | 0.054833 | 786.818 | 0.0038333 | 6.2401e-05
787.055 | 0.031 | 787.055 | 0.031 | 0.03 | 0.039 | 0.056 | 0.039 | 0.051 | 0.059 | 0.066 | 0.061 | 0.033333 | 0.048667 | 0.062 | 0.047667 | 787.055 | 0.001 | 1.6279e-05
787.291 | 0.043 | 787.291 | 0.043 | 0.048 | 0.049 | 0.059 | 0.059 | 0.068 | 0.076 | 0.078 | 0.087 | 0.046667 | 0.062 | 0.080333 | 0.0635 | 787.291 | -0.0015 | -2.4418e-05
787.528 | 0.042 | 787.528 | 0.042 | 0.041 | 0.053 | 0.061 | 0.075 | 0.078 | 0.085 | 0.09 | 0.08 | 0.045333 | 0.071333 | 0.085 | 0.065167 | 787.528 | 0.0061667 | 0.00010038
787.764 | 0.018 | 787.764 | 0.018 | 0.026 | 0.034 | 0.048 | 0.045 | 0.047 | 0.06 | 0.059 | 0.066 | 0.026 | 0.046667 | 0.061667 | 0.043833 | 787.764 | 0.0028333 | 4.6123e-05
788.001 | 0.018 | 788.001 | 0.018 | 0.021 | 0.016 | 0.028 | 0.028 | 0.035 | 0.031 | 0.044 | 0.046 | 0.018333 | 0.030333 | 0.040333 | 0.029333 | 788.001 | 0.001 | 1.6279e-05
788.237 | 0.032 | 788.237 | 0.032 | 0.043 | 0.055 | 0.065 | 0.069 | 0.065 | 0.075 | 0.072 | 0.067 | 0.043333 | 0.066333 | 0.071333 | 0.057333 | 788.237 | 0.009 | 0.00014651
788.473 | -0.001 | 788.473 | -0.001 | 0.008 | 0.015 | 0.019 | 0.029 | 0.032 | 0.033 | 0.041 | 0.04 | 0.0073333 | 0.026667 | 0.038 | 0.022667 | 788.473 | 0.004 | 6.5114e-05
788.709 | 0.016 | 788.709 | 0.016 | 0.022 | 0.022 | 0.037 | 0.042 | 0.04 | 0.057 | 0.055 | 0.056 | 0.02 | 0.039667 | 0.056 | 0.038 | 788.709 | 0.0016667 | 2.7131e-05
788.946 | 0.024 | 788.946 | 0.024 | 0.025 | 0.032 | 0.05 | 0.041 | 0.058 | 0.041 | 0.065 | 0.071 | 0.027 | 0.049667 | 0.059 | 0.043 | 788.946 | 0.0066667 | 0.00010852
789.182 | 0.005 | 789.182 | 0.005 | 0.007 | 0.016 | 0.029 | 0.015 | 0.025 | 0.032 | 0.04 | 0.036 | 0.0093333 | 0.023 | 0.036 | 0.022667 | 789.182 | 0.00033333 | 5.4262e-06
789.418 | 0.005 | 789.418 | 0.005 | 0.008 | 0.006 | 0.023 | 0.027 | 0.019 | 0.026 | 0.032 | 0.042 | 0.0063333 | 0.023 | 0.033333 | 0.019833 | 789.418 | 0.0031667 | 5.1549e-05
789.654 | 0.035 | 789.654 | 0.035 | 0.043 | 0.046 | 0.058 | 0.055 | 0.073 | 0.075 | 0.075 | 0.073 | 0.041333 | 0.062 | 0.074333 | 0.057833 | 789.654 | 0.0041667 | 6.7827e-05
789.89 | 0.016 | 789.89 | 0.016 | 0.014 | 0.029 | 0.033 | 0.033 | 0.033 | 0.036 | 0.042 | 0.051 | 0.019667 | 0.033 | 0.043 | 0.031333 | 789.89 | 0.0016667 | 2.7131e-05
790.126 | 0.044 | 790.126 | 0.044 | 0.049 | 0.054 | 0.068 | 0.064 | 0.08 | 0.073 | 0.082 | 0.074 | 0.049 | 0.070667 | 0.076333 | 0.062667 | 790.126 | 0.008 | 0.00013023
790.362 | 0.045 | 790.362 | 0.045 | 0.059 | 0.052 | 0.076 | 0.067 | 0.07 | 0.067 | 0.095 | 0.088 | 0.052 | 0.071 | 0.083333 | 0.067667 | 790.362 | 0.0033333 | 5.4262e-05
790.597 | 0.001 | 790.597 | 0.001 | 0.009 | 0.025 | 0.027 | 0.031 | 0.032 | 0.034 | 0.041 | 0.051 | 0.011667 | 0.03 | 0.042 | 0.026833 | 790.597 | 0.0031667 | 5.1549e-05
790.833 | 0.023 | 790.833 | 0.023 | 0.028 | 0.03 | 0.043 | 0.051 | 0.053 | 0.064 | 0.062 | 0.067 | 0.027 | 0.049 | 0.064333 | 0.045667 | 790.833 | 0.0033333 | 5.4262e-05
791.069 | 0.011 | 791.069 | 0.011 | 0.016 | 0.017 | 0.025 | 0.026 | 0.04 | 0.041 | 0.043 | 0.039 | 0.014667 | 0.030333 | 0.041 | 0.027833 | 791.069 | 0.0025 | 4.0696e-05
791.305 | -0.007 | 791.305 | -0.007 | 0.015 | 0.021 | 0.029 | 0.025 | 0.028 | 0.034 | 0.038 | 0.05 | 0.0096667 | 0.027333 | 0.040667 | 0.025167 | 791.305 | 0.0021667 | 3.527e-05
791.54 | 0.031 | 791.54 | 0.031 | 0.036 | 0.035 | 0.049 | 0.038 | 0.056 | 0.052 | 0.064 | 0.063 | 0.034 | 0.047667 | 0.059667 | 0.046833 | 791.54 | 0.00083333 | 1.3565e-05
791.776 | 0.049 | 791.776 | 0.049 | 0.05 | 0.056 | 0.062 | 0.059 | 0.071 | 0.081 | 0.075 | 0.086 | 0.051667 | 0.064 | 0.080667 | 0.066167 | 791.776 | -0.0021667 | -3.527e-05
792.011 | 0.04 | 792.011 | 0.04 | 0.036 | 0.044 | 0.056 | 0.059 | 0.051 | 0.071 | 0.073 | 0.066 | 0.04 | 0.055333 | 0.07 | 0.055 | 792.011 | 0.00033333 | 5.4262e-06
792.247 | -0.001 | 792.247 | -0.001 | 0.003 | 0.015 | 0.016 | 0.019 | 0.021 | 0.023 | 0.026 | 0.03 | 0.0056667 | 0.018667 | 0.026333 | 0.016 | 792.247 | 0.0026667 | 4.3409e-05
792.482 | 0.016 | 792.482 | 0.016 | 0.022 | 0.032 | 0.035 | 0.04 | 0.041 | 0.054 | 0.061 | 0.062 | 0.023333 | 0.038667 | 0.059 | 0.041167 | 792.482 | -0.0025 | -4.0696e-05
792.717 | 0.032 | 792.717 | 0.032 | 0.042 | 0.046 | 0.045 | 0.063 | 0.052 | 0.069 | 0.067 | 0.068 | 0.04 | 0.053333 | 0.068 | 0.054 | 792.717 | -0.00066667 | -1.0852e-05
792.953 | 0.012 | 792.953 | 0.012 | 0.026 | 0.031 | 0.034 | 0.036 | 0.046 | 0.058 | 0.058 | 0.064 | 0.023 | 0.038667 | 0.06 | 0.0415 | 792.953 | -0.0028333 | -4.6123e-05
793.188 | 0.033 | 793.188 | 0.033 | 0.041 | 0.041 | 0.066 | 0.062 | 0.066 | 0.071 | 0.075 | 0.074 | 0.038333 | 0.064667 | 0.073333 | 0.055833 | 793.188 | 0.0088333 | 0.00014379
793.423 | 0.053 | 793.423 | 0.053 | 0.048 | 0.048 | 0.061 | 0.068 | 0.071 | 0.084 | 0.074 | 0.082 | 0.049667 | 0.066667 | 0.08 | 0.064833 | 793.423 | 0.0018333 | 2.9844e-05
793.658 | 0.04 | 793.658 | 0.04 | 0.051 | 0.049 | 0.058 | 0.07 | 0.071 | 0.08 | 0.077 | 0.079 | 0.046667 | 0.066333 | 0.078667 | 0.062667 | 793.658 | 0.0036667 | 5.9688e-05
793.894 | 0.035 | 793.894 | 0.035 | 0.057 | 0.065 | 0.064 | 0.074 | 0.075 | 0.078 | 0.085 | 0.087 | 0.052333 | 0.071 | 0.083333 | 0.067833 | 793.894 | 0.0031667 | 5.1549e-05
794.129 | 0.025 | 794.129 | 0.025 | 0.031 | 0.034 | 0.038 | 0.043 | 0.052 | 0.055 | 0.049 | 0.071 | 0.03 | 0.044333 | 0.058333 | 0.044167 | 794.129 | 0.00016667 | 2.7131e-06
794.364 | 0.033 | 794.364 | 0.033 | 0.038 | 0.04 | 0.043 | 0.053 | 0.063 | 0.048 | 0.06 | 0.066 | 0.037 | 0.053 | 0.058 | 0.0475 | 794.364 | 0.0055 | 8.9532e-05
794.599 | 0.03 | 794.599 | 0.03 | 0.041 | 0.041 | 0.037 | 0.058 | 0.052 | 0.066 | 0.07 | 0.074 | 0.037333 | 0.049 | 0.07 | 0.053667 | 794.599 | -0.0046667 | -7.5967e-05
794.833 | 0.02 | 794.833 | 0.02 | 0.034 | 0.022 | 0.048 | 0.055 | 0.052 | 0.053 | 0.045 | 0.051 | 0.025333 | 0.051667 | 0.049667 | 0.0375 | 794.833 | 0.014167 | 0.00023061
795.068 | 0.024 | 795.068 | 0.024 | 0.036 | 0.038 | 0.069 | 0.064 | 0.075 | 0.065 | 0.063 | 0.066 | 0.032667 | 0.069333 | 0.064667 | 0.048667 | 795.068 | 0.020667 | 0.00033642
795.303 | 0.036 | 795.303 | 0.036 | 0.041 | 0.034 | 0.053 | 0.058 | 0.06 | 0.054 | 0.067 | 0.067 | 0.037 | 0.057 | 0.062667 | 0.049833 | 795.303 | 0.0071667 | 0.00011666
795.538 | 0.006 | 795.538 | 0.006 | 0.01 | 0.005 | 0.027 | 0.023 | 0.027 | 0.029 | 0.037 | 0.038 | 0.007 | 0.025667 | 0.034667 | 0.020833 | 795.538 | 0.0048333 | 7.868e-05
795.772 | 0.026 | 795.772 | 0.026 | 0.037 | 0.035 | 0.054 | 0.05 | 0.056 | 0.047 | 0.068 | 0.056 | 0.032667 | 0.053333 | 0.057 | 0.044833 | 795.772 | 0.0085 | 0.00013837
796.007 | 0.034 | 796.007 | 0.034 | 0.035 | 0.04 | 0.061 | 0.05 | 0.056 | 0.056 | 0.066 | 0.06 | 0.036333 | 0.055667 | 0.060667 | 0.0485 | 796.007 | 0.0071667 | 0.00011666
796.242 | 0.071 | 796.242 | 0.071 | 0.051 | 0.065 | 0.069 | 0.069 | 0.071 | 0.086 | 0.085 | 0.085 | 0.062333 | 0.069667 | 0.085333 | 0.073833 | 796.242 | -0.0041667 | -6.7827e-05
796.476 | 0.031 | 796.476 | 0.031 | 0.04 | 0.047 | 0.05 | 0.063 | 0.056 | 0.07 | 0.068 | 0.066 | 0.039333 | 0.056333 | 0.068 | 0.053667 | 796.476 | 0.0026667 | 4.3409e-05
796.711 | 0.014 | 796.711 | 0.014 | 0.014 | 0.024 | 0.038 | 0.03 | 0.044 | 0.043 | 0.054 | 0.044 | 0.017333 | 0.037333 | 0.047 | 0.032167 | 796.711 | 0.0051667 | 8.4106e-05
796.945 | 0.011 | 796.945 | 0.011 | 0.02 | 0.034 | 0.048 | 0.04 | 0.05 | 0.046 | 0.053 | 0.054 | 0.021667 | 0.046 | 0.051 | 0.036333 | 796.945 | 0.0096667 | 0.00015736
797.18 | 0.012 | 797.18 | 0.012 | 0.015 | 0.016 | 0.024 | 0.034 | 0.037 | 0.033 | 0.042 | 0.053 | 0.014333 | 0.031667 | 0.042667 | 0.0285 | 797.18 | 0.0031667 | 5.1549e-05
797.414 | -0.004 | 797.414 | -0.004 | 0.002 | 0.009 | 0.017 | 0.017 | 0.025 | 0.028 | 0.035 | 0.03 | 0.0023333 | 0.019667 | 0.031 | 0.016667 | 797.414 | 0.003 | 4.8836e-05
797.648 | 0.046 | 797.648 | 0.046 | 0.057 | 0.052 | 0.076 | 0.068 | 0.07 | 0.077 | 0.09 | 0.083 | 0.051667 | 0.071333 | 0.083333 | 0.0675 | 797.648 | 0.0038333 | 6.2401e-05
797.882 | 0.019 | 797.882 | 0.019 | 0.036 | 0.025 | 0.034 | 0.035 | 0.034 | 0.042 | 0.053 | 0.062 | 0.026667 | 0.034333 | 0.052333 | 0.0395 | 797.882 | -0.0051667 | -8.4106e-05
798.117 | 0.014 | 798.117 | 0.014 | 0.024 | 0.02 | 0.038 | 0.041 | 0.053 | 0.046 | 0.048 | 0.047 | 0.019333 | 0.044 | 0.047 | 0.033167 | 798.117 | 0.010833 | 0.00017635
798.351 | 0.001 | 798.351 | 0.001 | 0.018 | 0.018 | 0.033 | 0.03 | 0.025 | 0.035 | 0.045 | 0.039 | 0.012333 | 0.029333 | 0.039667 | 0.026 | 798.351 | 0.0033333 | 5.4262e-05
798.585 | 0.072 | 798.585 | 0.072 | 0.071 | 0.068 | 0.072 | 0.102 | 0.093 | 0.087 | 0.108 | 0.104 | 0.070333 | 0.089 | 0.099667 | 0.085 | 798.585 | 0.004 | 6.5114e-05
798.819 | -0.003 | 798.819 | -0.003 | 0.014 | 0.012 | 0.017 | 0.041 | 0.035 | 0.034 | 0.041 | 0.045 | 0.0076667 | 0.031 | 0.04 | 0.023833 | 798.819 | 0.0071667 | 0.00011666
799.053 | 0.037 | 799.053 | 0.037 | 0.036 | 0.047 | 0.059 | 0.065 | 0.063 | 0.066 | 0.073 | 0.074 | 0.04 | 0.062333 | 0.071 | 0.0555 | 799.053 | 0.0068333 | 0.00011124
799.287 | 0.027 | 799.287 | 0.027 | 0.031 | 0.035 | 0.046 | 0.053 | 0.046 | 0.053 | 0.053 | 0.075 | 0.031 | 0.048333 | 0.060333 | 0.045667 | 799.287 | 0.0026667 | 4.3409e-05
799.521 | 0.054 | 799.521 | 0.054 | 0.06 | 0.052 | 0.066 | 0.072 | 0.083 | 0.084 | 0.08 | 0.083 | 0.055333 | 0.073667 | 0.082333 | 0.068833 | 799.521 | 0.0048333 | 7.868e-05
799.755 | 0.038 | 799.755 | 0.038 | 0.044 | 0.042 | 0.053 | 0.06 | 0.061 | 0.073 | 0.081 | 0.077 | 0.041333 | 0.058 | 0.077 | 0.059167 | 799.755 | -0.0011667 | -1.8992e-05
799.988 | 0.034 | 799.988 | 0.034 | 0.042 | 0.044 | 0.057 | 0.067 | 0.067 | 0.059 | 0.064 | 0.073 | 0.04 | 0.063667 | 0.065333 | 0.052667 | 799.988 | 0.011 | 0.00017906
800.222 | 0.024 | 800.222 | 0.024 | 0.036 | 0.036 | 0.053 | 0.04 | 0.04 | 0.056 | 0.068 | 0.057 | 0.032 | 0.044333 | 0.060333 | 0.046167 | 800.222 | -0.0018333 | -2.9844e-05
800.456 | 0.016 | 800.456 | 0.016 | 0.03 | 0.03 | 0.05 | 0.036 | 0.044 | 0.046 | 0.056 | 0.059 | 0.025333 | 0.043333 | 0.053667 | 0.0395 | 800.456 | 0.0038333 | 6.2401e-05
800.689 | -0.008 | 800.689 | -0.008 | -0.012 | 0.002 | 0.021 | 0.023 | 0.017 | 0.018 | 0.029 | 0.03 | -0.006 | 0.020333 | 0.025667 | 0.0098333 | 800.689 | 0.0105 | 0.00017092
800.923 | 0.009 | 800.923 | 0.009 | 0.028 | 0.028 | 0.034 | 0.036 | 0.047 | 0.049 | 0.04 | 0.043 | 0.021667 | 0.039 | 0.044 | 0.032833 | 800.923 | 0.0061667 | 0.00010038
801.156 | 0.035 | 801.156 | 0.035 | 0.039 | 0.049 | 0.072 | 0.061 | 0.078 | 0.077 | 0.078 | 0.075 | 0.041 | 0.070333 | 0.076667 | 0.058833 | 801.156 | 0.0115 | 0.0001872
801.39 | 0.021 | 801.39 | 0.021 | 0.02 | 0.02 | 0.035 | 0.045 | 0.042 | 0.044 | 0.054 | 0.048 | 0.020333 | 0.040667 | 0.048667 | 0.0345 | 801.39 | 0.0061667 | 0.00010038
801.623 | 0.038 | 801.623 | 0.038 | 0.044 | 0.045 | 0.067 | 0.065 | 0.065 | 0.056 | 0.058 | 0.071 | 0.042333 | 0.065667 | 0.061667 | 0.052 | 801.623 | 0.013667 | 0.00022247
801.857 | 0.026 | 801.857 | 0.026 | 0.03 | 0.027 | 0.068 | 0.071 | 0.07 | 0.05 | 0.055 | 0.076 | 0.027667 | 0.069667 | 0.060333 | 0.044 | 801.857 | 0.025667 | 0.00041782
802.09 | 0.027 | 802.09 | 0.027 | 0.035 | 0.049 | 0.056 | 0.071 | 0.074 | 0.061 | 0.068 | 0.061 | 0.037 | 0.067 | 0.063333 | 0.050167 | 802.09 | 0.016833 | 0.00027402
802.323 | 0.03 | 802.323 | 0.03 | 0.051 | 0.041 | 0.048 | 0.062 | 0.06 | 0.06 | 0.071 | 0.066 | 0.040667 | 0.056667 | 0.065667 | 0.053167 | 802.323 | 0.0035 | 5.6975e-05
802.557 | 0.025 | 802.557 | 0.025 | 0.029 | 0.037 | 0.05 | 0.061 | 0.054 | 0.061 | 0.055 | 0.066 | 0.030333 | 0.055 | 0.060667 | 0.0455 | 802.557 | 0.0095 | 0.00015465
802.79 | 0.029 | 802.79 | 0.029 | 0.035 | 0.03 | 0.044 | 0.046 | 0.049 | 0.043 | 0.049 | 0.052 | 0.031333 | 0.046333 | 0.048 | 0.039667 | 802.79 | 0.0066667 | 0.00010852
803.023 | 0.017 | 803.023 | 0.017 | 0.008 | 0.023 | 0.031 | 0.027 | 0.034 | 0.042 | 0.053 | 0.056 | 0.016 | 0.030667 | 0.050333 | 0.033167 | 803.023 | -0.0025 | -4.0696e-05
803.256 | 0.018 | 803.256 | 0.018 | 0.018 | 0.027 | 0.035 | 0.036 | 0.043 | 0.048 | 0.042 | 0.052 | 0.021 | 0.038 | 0.047333 | 0.034167 | 803.256 | 0.0038333 | 6.2401e-05
803.489 | 0.025 | 803.489 | 0.025 | 0.026 | 0.041 | 0.047 | 0.061 | 0.05 | 0.054 | 0.055 | 0.065 | 0.030667 | 0.052667 | 0.058 | 0.044333 | 803.489 | 0.0083333 | 0.00013565
803.722 | 0.01 | 803.722 | 0.01 | 0.017 | 0.034 | 0.038 | 0.023 | 0.042 | 0.029 | 0.042 | 0.04 | 0.020333 | 0.034333 | 0.037 | 0.028667 | 803.722 | 0.0056667 | 9.2245e-05
803.955 | 0.011 | 803.955 | 0.011 | 0.016 | 0.023 | 0.035 | 0.027 | 0.033 | 0.04 | 0.047 | 0.036 | 0.016667 | 0.031667 | 0.041 | 0.028833 | 803.955 | 0.0028333 | 4.6123e-05
804.188 | 0.059 | 804.188 | 0.059 | 0.057 | 0.062 | 0.08 | 0.076 | 0.072 | 0.082 | 0.085 | 0.084 | 0.059333 | 0.076 | 0.083667 | 0.0715 | 804.188 | 0.0045 | 7.3254e-05
804.421 | 0.017 | 804.421 | 0.017 | 0.025 | 0.021 | 0.047 | 0.043 | 0.046 | 0.045 | 0.054 | 0.054 | 0.021 | 0.045333 | 0.051 | 0.036 | 804.421 | 0.0093333 | 0.00015193
804.653 | 0.031 | 804.653 | 0.031 | 0.045 | 0.037 | 0.051 | 0.056 | 0.058 | 0.06 | 0.064 | 0.066 | 0.037667 | 0.055 | 0.063333 | 0.0505 | 804.653 | 0.0045 | 7.3254e-05
804.886 | 0.026 | 804.886 | 0.026 | 0.034 | 0.035 | 0.036 | 0.042 | 0.042 | 0.043 | 0.06 | 0.045 | 0.031667 | 0.04 | 0.049333 | 0.0405 | 804.886 | -0.0005 | -8.1393e-06
805.119 | 0 | 805.119 | 0 | 0.009 | 0.014 | 0.017 | 0.02 | 0.021 | 0.026 | 0.03 | 0.033 | 0.0076667 | 0.019333 | 0.029667 | 0.018667 | 805.119 | 0.00066667 | 1.0852e-05
805.351 | 0.019 | 805.351 | 0.019 | 0.027 | 0.022 | 0.043 | 0.044 | 0.047 | 0.039 | 0.049 | 0.054 | 0.022667 | 0.044667 | 0.047333 | 0.035 | 805.351 | 0.0096667 | 0.00015736
805.584 | 0.035 | 805.584 | 0.035 | 0.047 | 0.053 | 0.062 | 0.054 | 0.052 | 0.072 | 0.07 | 0.07 | 0.045 | 0.056 | 0.070667 | 0.057833 | 805.584 | -0.0018333 | -2.9844e-05
805.816 | 0.005 | 805.816 | 0.005 | -0.003 | 0.014 | 0.02 | 0.029 | 0.023 | 0.028 | 0.032 | 0.035 | 0.0053333 | 0.024 | 0.031667 | 0.0185 | 805.816 | 0.0055 | 8.9532e-05
806.049 | 0.027 | 806.049 | 0.027 | 0.025 | 0.036 | 0.03 | 0.034 | 0.048 | 0.051 | 0.057 | 0.05 | 0.029333 | 0.037333 | 0.052667 | 0.041 | 806.049 | -0.0036667 | -5.9688e-05
806.281 | 0.013 | 806.281 | 0.013 | 0.018 | 0.047 | 0.046 | 0.045 | 0.052 | 0.051 | 0.068 | 0.069 | 0.026 | 0.047667 | 0.062667 | 0.044333 | 806.281 | 0.0033333 | 5.4262e-05
806.514 | 0.039 | 806.514 | 0.039 | 0.038 | 0.046 | 0.086 | 0.075 | 0.077 | 0.085 | 0.102 | 0.09 | 0.041 | 0.079333 | 0.092333 | 0.066667 | 806.514 | 0.012667 | 0.0002062
806.746 | 0.045 | 806.746 | 0.045 | 0.045 | 0.055 | 0.079 | 0.083 | 0.089 | 0.099 | 0.086 | 0.097 | 0.048333 | 0.083667 | 0.094 | 0.071167 | 806.746 | 0.0125 | 0.00020348
806.978 | 0.024 | 806.978 | 0.024 | 0.02 | 0.03 | 0.029 | 0.023 | 0.021 | 0.027 | 0.033 | 0.031 | 0.024667 | 0.024333 | 0.030333 | 0.0275 | 806.978 | -0.0031667 | -5.1549e-05
807.211 | 0.007 | 807.211 | 0.007 | 0.003 | 0.01 | 0.019 | 0.017 | 0.017 | 0.013 | 0.015 | 0.019 | 0.0066667 | 0.017667 | 0.015667 | 0.011167 | 807.211 | 0.0065 | 0.00010581
807.443 | 0.026 | 807.443 | 0.026 | 0.028 | 0.027 | 0.039 | 0.021 | 0.033 | 0.028 | 0.039 | 0.037 | 0.027 | 0.031 | 0.034667 | 0.030833 | 807.443 | 0.00016667 | 2.7131e-06
807.675 | 0.042 | 807.675 | 0.042 | 0.051 | 0.031 | 0.053 | 0.038 | 0.054 | 0.044 | 0.061 | 0.039 | 0.041333 | 0.048333 | 0.048 | 0.044667 | 807.675 | 0.0036667 | 5.9688e-05
background 2
AbsoluteIrradiance:211
195.562 | 0
195.934 | 0
196.307 | 0
196.679 | 0
197.051 | 0
197.423 | 0
197.795 | 0
198.167 | 0
198.539 | 0
198.911 | 0
199.283 | 0
199.654 | 0
200.026 | 0
200.397 | 0
200.769 | 0
201.14 | 0
201.511 | 0
201.882 | 0
202.253 | 0
202.624 | 0
202.995 | 0
203.366 | 0
203.736 | 0
204.107 | 0
204.477 | 0
204.848 | 2.213
205.218 | 4.959
205.588 | 0.504
205.959 | 1.38
206.329 | 1.817
206.699 | 0.943
207.068 | 1.535
207.438 | 1.887
207.808 | 0.52
208.178 | 1.501
208.547 | 1.087
208.917 | 0.776
209.286 | 2.026
209.655 | 1.489
210.025 | 0.637
210.394 | 0.906
210.763 | 1.424
211.132 | 0.398
211.501 | 0.374
211.869 | 0.12
212.238 | 0.737
212.607 | 1.193
212.975 | 0.322
213.344 | 0.231
213.712 | 0.245
214.08 | 0.584
214.448 | 0.527
214.817 | 0.454
215.185 | 1.028
215.553 | 0.689
215.92 | 0.729
216.288 | 0.585
216.656 | 0.7
217.023 | 0.35
217.391 | 0.335
217.758 | 0.631
218.126 | 0.556
218.493 | 0.392
218.86 | 0.394
219.227 | 0.22
219.594 | 0.24
219.961 | 0.647
220.328 | 0.198
220.695 | 0.395
221.061 | 0.544
221.428 | 0.413
221.795 | 0.532
222.161 | 0.58
222.527 | 0.271
222.893 | 0.589
223.26 | 0.161
223.626 | 0.646
223.992 | 0.377
224.358 | 0.275
224.723 | 0.481
225.089 | 0.014
225.455 | 0.329
225.82 | 0.207
226.186 | 0.179
226.551 | 0.014
226.917 | 0.134
227.282 | 0.145
227.647 | 0.326
228.012 | 0.229
228.377 | 0.175
228.742 | 0.205
229.107 | 0.145
229.471 | 0.243
229.836 | 0.11
230.201 | 0.193
230.565 | 0.2
230.929 | 0.113
231.294 | 0.044
231.658 | 0.094
232.022 | 0.099
232.386 | 0.052
232.75 | -0.099
233.114 | 0.055
233.478 | 0.084
233.842 | 0.101
234.205 | 0.116
234.569 | 0.056
234.932 | 0.048
235.296 | 0.032
235.659 | 0.06
236.022 | 0.045
236.385 | 0.039
236.748 | 0.088
237.111 | 0.016
237.474 | 0.096
237.837 | 0.083
238.2 | 0.056
238.562 | 0.048
238.925 | 0.109
239.287 | 0.077
239.65 | 0.045
240.012 | 0.088
240.374 | 0.052
240.736 | 0.097
241.098 | 0.04
241.46 | 0.06
241.822 | 0.095
242.184 | 0.013
242.546 | 0.091
242.907 | 0.038
243.269 | 0.033
243.63 | 0.007
243.992 | 0.082
244.353 | 0.046
244.714 | 0.042
245.075 | 0.042
245.436 | 0.053
245.797 | 0.028
246.158 | -0.009
246.519 | 0.066
246.879 | 0.069
247.24 | 0.041
247.601 | 0.055
247.961 | 0.063
248.321 | 0.001
248.682 | 0.029
249.042 | 0.018
249.402 | 0.035
249.762 | 0.074
250.122 | 0.03
250.482 | 0.072
250.842 | 0.034
251.201 | 0.043
251.561 | 0.043
251.921 | 0.042
252.28 | 0.046
252.639 | 0.045
252.999 | 0.051
253.358 | 0.035
253.717 | 0.037
254.076 | 0.044
254.435 | 0.027
254.794 | 0.005
255.153 | 0.072
255.511 | 0.039
255.87 | 0
256.228 | 0.037
256.587 | 0.03
256.945 | 0.033
257.304 | 0
257.662 | 0.014
258.02 | 0.033
258.378 | 0.049
258.736 | 0.002
259.094 | 0.072
259.452 | 0.06
259.809 | 0.041
260.167 | 0.044
260.524 | 0.026
260.882 | 0.001
261.239 | 0.023
261.597 | 0.029
261.954 | 0.017
262.311 | 0.03
262.668 | 0.035
263.025 | 0.03
263.382 | 0.037
263.739 | 0.022
264.095 | 0.037
264.452 | 0.014
264.809 | 0.028
265.165 | 0.024
265.521 | 0.047
265.878 | 0.021
266.234 | 0.021
266.59 | 0.02
266.946 | 0.064
267.302 | 0.025
267.658 | 0.045
268.014 | 0.027
268.37 | 0.026
268.725 | 0.03
269.081 | 0.031
269.436 | 0.035
269.792 | 0.001
270.147 | 0.022
270.502 | 0.048
270.857 | 0.032
271.212 | 0.022
271.567 | -0.007
271.922 | 0.061
272.277 | 0.021
272.632 | 0.043
272.987 | 0.044
273.341 | 0.05
273.696 | 0.033
274.05 | 0.003
274.404 | 0.022
274.759 | 0.045
275.113 | 0.057
275.467 | 0.005
275.821 | 0.014
276.175 | 0.017
276.529 | 0.046
276.882 | 0.078
277.236 | 0.03
277.59 | 0.046
277.943 | 0.001
278.297 | 0.082
278.65 | 0.004
279.003 | 0.018
279.356 | 0.03
279.709 | 0.037
280.062 | 0.021
280.415 | 0.036
280.768 | 0.023
281.121 | 0.057
281.474 | 0.035
281.826 | 0.018
282.179 | 0.036
282.531 | 0.011
282.884 | 0.038
283.236 | 0.026
283.588 | 0.064
283.94 | 0.013
284.292 | 0.03
284.644 | 0.016
284.996 | 0.033
285.348 | 0.029
285.7 | 0.043
286.051 | 0.032
286.403 | 0.029
286.754 | 0.033
287.106 | 0.053
287.457 | 0.05
287.808 | 0.041
288.159 | 0.032
288.51 | 0.04
288.861 | 0.027
289.212 | -0.007
289.563 | 0.015
289.914 | 0.042
290.264 | 0.011
290.615 | 0.03
290.966 | 0.045
291.316 | 0.027
291.666 | 0.041
292.017 | 0.032
292.367 | 0.015
292.717 | 0.015
293.067 | 0.027
293.417 | 0.026
293.767 | 0.021
294.116 | 0.046
294.466 | 0.025
294.816 | 0.04
295.165 | 0.018
295.515 | 0.016
295.864 | 0.034
296.213 | 0.015
296.562 | 0.01
296.911 | 0.03
297.261 | 0.022
297.609 | 0.02
297.958 | -0.001
298.307 | 0.021
298.656 | 0.021
299.004 | 0.016
299.353 | -0.013
299.701 | 0.021
300.05 | 0.01
300.398 | 0.026
300.746 | 0.023
301.095 | 0.018
301.443 | 0.015
301.791 | 0.012
302.138 | 0.018
302.486 | 0.01
302.834 | 0.005
303.182 | 0.006
303.529 | 0.02
303.877 | 0.022
304.224 | 0.021
304.572 | 0.022
304.919 | 0.004
305.266 | 0.021
305.613 | 0.007
305.96 | 0.03
306.307 | 0.005
306.654 | 0.031
307.001 | 0.025
307.347 | 0.02
307.694 | 0.003
308.04 | 0.012
308.387 | 0.024
308.733 | 0.016
309.08 | 0.016
309.426 | 0.026
309.772 | 0.019
310.118 | 0.027
310.464 | 0.008
310.81 | 0.017
311.156 | 0.016
311.501 | 0.021
311.847 | 0.017
312.193 | 0.011
312.538 | 0.021
312.883 | 0.015
313.229 | 0.015
313.574 | 0.012
313.919 | 0.016
314.264 | 0.024
314.609 | 0.024
314.954 | 0.012
315.299 | 0.014
315.644 | 0.007
315.988 | 0.023
316.333 | 0.02
316.678 | 0.016
317.022 | 0.016
317.366 | 0.013
317.711 | 0.015
318.055 | 0.006
318.399 | 0.013
318.743 | 0.016
319.087 | 0.016
319.431 | 0.023
319.775 | 0.017
320.118 | 0.023
320.462 | 0.015
320.806 | 0.014
321.149 | 0.022
321.493 | 0.013
321.836 | 0.007
322.179 | 0
322.522 | 0.014
322.865 | 0.014
323.208 | 0.015
323.551 | 0.025
323.894 | 0.011
324.237 | 0.006
324.58 | 0.017
324.922 | 0.014
325.265 | 0.02
325.607 | 0.021
325.95 | 0.023
326.292 | 0.003
326.634 | 0.017
326.976 | 0.006
327.318 | 0.011
327.66 | 0.016
328.002 | 0.002
328.344 | 0.017
328.686 | 0.015
329.027 | 0.007
329.369 | 0.006
329.711 | 0.017
330.052 | 0.012
330.393 | 0.005
330.735 | 0.005
331.076 | 0.015
331.417 | 0.007
331.758 | 0.016
332.099 | 0.002
332.44 | 0.017
332.78 | 0.006
333.121 | 0.017
333.462 | 0.007
333.802 | 0.015
334.143 | 0.009
334.483 | 0.011
334.823 | 0.003
335.164 | 0.014
335.504 | 0.004
335.844 | 0.017
336.184 | 0.005
336.524 | 0.006
336.864 | 0.009
337.203 | 0.011
337.543 | 0.022
337.883 | 0.016
338.222 | 0.022
338.562 | 0.009
338.901 | 0.016
339.24 | -0.002
339.58 | 0.009
339.919 | 0.007
340.258 | 0.016
340.597 | 0.01
340.936 | 0.003
341.274 | 0.012
341.613 | 0.016
341.952 | 0.012
342.29 | 0.011
342.629 | 0.021
342.967 | 0.011
343.306 | 0.007
343.644 | 0.011
343.982 | 0.005
344.32 | 0.014
344.658 | 0.006
344.996 | 0.012
345.334 | 0.007
345.672 | 0.013
346.01 | 0.007
346.347 | 0.024
346.685 | 0.01
347.022 | 0.023
347.36 | 0.015
347.697 | 0.015
348.034 | 0.012
348.371 | 0.01
348.708 | 0.011
349.045 | 0.016
349.382 | 0.017
349.719 | 0.004
350.056 | 0.009
350.393 | 0.011
350.729 | 0.019
351.066 | 0.007
351.402 | 0.011
351.739 | 0.021
352.075 | 0.016
352.411 | 0.01
352.747 | 0.009
353.083 | 0.013
353.419 | 0.011
353.755 | 0.008
354.091 | 0.017
354.427 | 0.011
354.763 | 0.015
355.098 | 0.018
355.434 | 0.004
355.769 | 0.025
356.104 | 0
356.44 | 0.001
356.775 | 0.015
357.11 | 0.013
357.445 | 0.009
357.78 | 0.006
358.115 | 0.016
358.45 | 0.015
358.785 | 0.014
359.119 | 0.005
359.454 | 0.009
359.788 | 0.013
360.123 | 0.015
360.457 | 0.003
360.791 | 0.021
361.126 | 0.005
361.46 | 0.005
361.794 | 0.012
362.128 | 0.017
362.462 | 0.025
362.795 | 0.023
363.129 | 0.014
363.463 | 0.001
363.796 | 0.016
364.13 | 0.022
364.463 | 0.016
364.797 | 0.017
365.13 | 0.02
365.463 | 0.005
365.796 | 0.009
366.129 | 0.015
366.462 | 0.008
366.795 | 0.008
367.128 | 0.021
367.461 | 0.008
367.793 | 0.018
368.126 | 0.011
368.458 | 0.016
368.791 | 0.014
369.123 | 0.013
369.455 | 0.007
369.787 | 0.007
370.12 | 0.005
370.452 | 0.003
370.784 | 0.008
371.115 | 0.007
371.447 | 0
371.779 | 0.011
372.111 | 0.011
372.442 | 0.007
372.774 | 0.018
373.105 | 0.012
373.436 | 0.009
373.768 | 0.005
374.099 | 0.001
374.43 | 0.017
374.761 | 0.006
375.092 | 0.01
375.423 | 0.005
375.754 | 0.006
376.084 | 0.01
376.415 | 0.007
376.745 | 0.008
377.076 | 0.015
377.406 | 0.013
377.737 | 0.025
378.067 | 0.013
378.397 | 0.008
378.727 | 0.007
379.057 | 0.004
379.387 | 0.006
379.717 | 0.011
380.047 | 0.013
380.377 | 0.009
380.706 | 0.01
381.036 | 0.011
381.365 | 0.023
381.695 | 0.006
382.024 | 0.006
382.353 | -0.002
382.682 | 0.008
383.012 | 0.002
383.341 | 0.014
383.67 | 0.013
383.998 | 0.004
384.327 | 0.017
384.656 | 0.005
384.985 | 0
385.313 | 0
385.642 | 0.004
385.97 | 0.008
386.298 | 0.005
386.627 | 0.004
386.955 | 0.014
387.283 | 0.003
387.611 | 0.004
387.939 | 0.008
388.267 | 0.006
388.595 | 0.018
388.922 | -0.003
389.25 | 0.019
389.578 | 0.007
389.905 | 0.011
390.233 | 0.005
390.56 | 0.009
390.887 | 0.015
391.214 | 0.014
391.542 | 0.002
391.869 | 0.007
392.196 | 0.002
392.522 | 0.001
392.849 | 0.006
393.176 | 0.001
393.503 | 0.008
393.829 | 0.007
394.156 | 0.005
394.482 | 0.009
394.809 | 0.01
395.135 | 0.011
395.461 | 0.003
395.787 | 0.014
396.113 | 0
396.439 | 0.009
396.765 | 0.005
397.091 | 0.008
397.417 | 0.002
397.742 | 0.007
398.068 | 0.014
398.394 | 0.012
398.719 | 0.007
399.044 | 0.005
399.37 | 0.013
399.695 | 0.007
400.02 | 0.013
400.345 | 0.016
400.67 | 0.012
400.995 | 0.015
401.32 | 0.01
401.645 | 0.013
401.969 | 0.006
402.294 | 0.006
402.618 | 0.02
402.943 | 0.002
403.267 | 0.002
403.592 | 0.005
403.916 | 0.02
404.24 | 0.009
404.564 | 0.005
404.888 | 0.02
405.212 | 0.014
405.536 | 0.002
405.86 | 0.013
406.183 | 0.007
406.507 | 0.01
406.831 | 0.013
407.154 | 0.006
407.478 | 0.009
407.801 | 0.007
408.124 | 0.001
408.447 | 0.007
408.77 | 0.014
409.093 | 0.011
409.416 | 0.005
409.739 | 0.011
410.062 | 0.001
410.385 | 0.013
410.708 | 0.014
411.03 | 0.012
411.353 | 0.004
411.675 | 0.01
411.997 | 0.008
412.32 | 0.004
412.642 | 0.01
412.964 | 0.007
413.286 | 0.006
413.608 | 0.011
413.93 | 0.012
414.252 | 0.001
414.574 | 0.002
414.895 | 0.004
415.217 | 0.014
415.538 | 0.005
415.86 | 0.006
416.181 | 0.011
416.503 | 0.011
416.824 | 0.006
417.145 | 0.004
417.466 | 0.011
417.787 | 0.017
418.108 | 0.004
418.429 | 0.006
418.75 | 0.007
419.07 | 0.006
419.391 | 0.013
419.712 | 0.011
420.032 | 0.007
420.353 | 0.005
420.673 | 0.005
420.993 | 0.01
421.313 | 0.015
421.633 | 0.012
421.953 | 0.002
422.273 | 0.007
422.593 | 0.01
422.913 | 0.006
423.233 | 0.01
423.553 | 0.005
423.872 | 0.004
424.192 | 0.007
424.511 | 0.005
424.83 | 0.012
425.15 | 0.014
425.469 | 0.01
425.788 | 0.015
426.107 | 0.009
426.426 | 0.018
426.745 | 0.019
427.064 | 0.012
427.383 | 0.012
427.702 | 0.001
428.02 | 0.013
428.339 | 0.009
428.657 | 0.01
428.976 | 0.012
429.294 | 0.007
429.612 | 0.004
429.93 | 0.012
430.248 | 0.017
430.567 | 0.007
430.884 | 0.009
431.202 | 0.006
431.52 | 0.007
431.838 | 0.005
432.156 | 0.009
432.473 | 0.009
432.791 | 0.011
433.108 | 0.005
433.426 | 0.018
433.743 | 0.014
434.06 | 0.009
434.377 | 0.013
434.694 | 0.015
435.011 | 0.006
435.328 | 0.008
435.645 | 0.008
435.962 | 0.008
436.279 | 0.007
436.595 | 0.004
436.912 | 0.007
437.228 | 0.007
437.545 | 0.009
437.861 | 0.005
438.177 | 0.002
438.493 | 0.003
438.81 | 0.011
439.126 | 0.005
439.442 | 0.013
439.757 | -0.001
440.073 | 0.009
440.389 | 0.005
440.705 | 0.015
441.02 | 0.008
441.336 | 0.004
441.651 | 0.013
441.967 | 0.003
442.282 | -0.001
442.597 | 0.01
442.912 | 0.012
443.227 | 0.01
443.543 | 0.012
443.857 | 0.012
444.172 | 0.005
444.487 | 0.012
444.802 | 0.015
445.116 | 0.015
445.431 | 0
445.746 | 0.002
446.06 | 0.008
446.374 | 0.007
446.689 | 0.014
447.003 | 0.007
447.317 | 0.011
447.631 | 0.001
447.945 | 0.01
448.259 | 0.013
448.573 | 0.012
448.887 | 0.004
449.2 | 0.018
449.514 | 0.01
449.827 | 0.01
450.141 | 0.007
450.454 | 0.006
450.768 | 0.007
451.081 | 0.012
451.394 | 0.007
451.707 | 0.008
452.02 | 0.015
452.333 | 0.01
452.646 | 0.007
452.959 | 0.005
453.272 | 0.007
453.584 | 0.007
453.897 | 0.007
454.21 | 0.011
454.522 | 0.006
454.834 | 0.005
455.147 | 0.007
455.459 | 0.001
455.771 | -0.004
456.083 | 0.002
456.395 | 0.006
456.707 | 0.007
457.019 | 0.011
457.331 | 0.01
457.643 | 0.007
457.954 | 0.009
458.266 | 0.004
458.577 | 0.006
458.889 | 0.003
459.2 | 0.009
459.512 | 0.003
459.823 | 0.006
460.134 | 0.009
460.445 | 0.013
460.756 | 0.008
461.067 | 0.009
461.378 | 0.005
461.689 | 0.005
461.999 | 0.016
462.31 | 0.016
462.621 | 0.003
462.931 | 0.008
463.242 | 0.012
463.552 | 0.013
463.862 | 0.002
464.172 | 0.008
464.483 | 0.018
464.793 | 0.007
465.103 | 0.008
465.413 | 0.011
465.722 | 0.009
466.032 | 0.009
466.342 | 0.007
466.652 | 0.008
466.961 | 0.012
467.271 | 0.012
467.58 | 0.009
467.889 | 0.01
468.199 | 0.011
468.508 | 0.007
468.817 | 0.012
469.126 | 0.01
469.435 | 0.017
469.744 | 0.009
470.053 | 0.009
470.362 | 0.006
470.67 | 0.006
470.979 | 0.01
471.287 | 0.003
471.596 | 0.007
471.904 | 0.015
472.213 | 0.002
472.521 | 0.015
472.829 | 0.01
473.137 | 0.004
473.445 | 0.003
473.753 | 0.021
474.061 | 0.005
474.369 | 0.007
474.677 | 0.015
474.985 | 0.003
475.292 | 0.005
475.6 | 0.012
475.907 | 0.007
476.215 | 0.009
476.522 | 0.018
476.829 | 0.008
477.137 | 0.016
477.444 | 0.004
477.751 | 0.016
478.058 | 0.003
478.365 | 0.01
478.671 | 0.014
478.978 | 0.014
479.285 | 0.005
479.592 | 0.006
479.898 | 0.005
480.205 | 0.01
480.511 | 0.001
480.817 | 0.012
481.124 | 0.009
481.43 | 0.012
481.736 | 0.02
482.042 | 0.008
482.348 | 0.006
482.654 | 0.017
482.96 | 0.015
483.266 | 0.009
483.571 | 0.012
483.877 | 0.001
484.182 | 0.005
484.488 | 0.01
484.793 | 0.003
485.099 | 0.008
485.404 | 0.011
485.709 | 0.015
486.014 | 0.008
486.319 | 0.017
486.624 | 0.006
486.929 | 0.004
487.234 | -0.001
487.539 | 0.01
487.844 | 0.012
488.148 | 0.007
488.453 | 0.017
488.757 | 0.01
489.062 | 0.005
489.366 | 0.011
489.67 | 0.004
489.975 | 0.002
490.279 | 0.008
490.583 | 0.011
490.887 | 0.005
491.191 | 0.011
491.495 | 0.012
491.798 | 0.016
492.102 | 0.013
492.406 | 0.007
492.709 | 0.012
493.013 | 0.017
493.316 | 0.01
493.62 | 0.016
493.923 | 0.011
494.226 | 0.018
494.529 | 0.01
494.832 | 0.007
495.135 | 0.007
495.438 | 0.014
495.741 | 0.022
496.044 | 0.007
496.347 | 0.004
496.649 | 0.014
496.952 | 0.009
497.255 | 0.008
497.557 | 0.007
497.859 | 0.009
498.162 | 0
498.464 | 0.016
498.766 | 0.013
499.068 | 0.006
499.37 | 0.009
499.672 | 0.017
499.974 | 0.009
500.276 | 0.01
500.578 | 0.012
500.879 | 0.003
501.181 | 0.014
501.483 | 0.01
501.784 | 0.013
502.085 | 0.007
502.387 | 0.003
502.688 | 0.003
502.989 | 0.01
503.29 | 0.001
503.591 | 0.004
503.892 | 0
504.193 | 0.017
504.494 | 0.007
504.795 | 0.001
505.096 | 0.016
505.396 | 0.001
505.697 | -0.003
505.997 | 0.006
506.298 | 0.01
506.598 | 0
506.898 | 0.009
507.199 | 0.008
507.499 | 0.009
507.799 | 0.004
508.099 | 0.005
508.399 | 0.011
508.699 | 0.011
508.998 | 0.008
509.298 | 0.014
509.598 | 0.008
509.897 | 0.011
510.197 | 0.004
510.496 | 0.007
510.796 | 0
511.095 | 0.009
511.394 | 0.008
511.693 | 0.015
511.992 | 0.012
512.291 | 0.011
512.59 | 0.014
512.889 | 0.014
513.188 | 0.011
513.487 | 0.008
513.786 | 0.011
514.084 | 0.012
514.383 | 0.014
514.681 | 0.018
514.98 | 0.009
515.278 | 0.013
515.576 | 0.016
515.874 | 0.011
516.173 | 0.004
516.471 | 0.017
516.769 | 0.011
517.067 | 0.009
517.364 | 0.003
517.662 | 0.003
517.96 | 0.007
518.257 | 0.006
518.555 | 0.002
518.853 | 0.005
519.15 | 0.005
519.447 | 0.008
519.745 | 0.009
520.042 | 0.011
520.339 | 0.003
520.636 | 0.013
520.933 | 0.02
521.23 | 0.005
521.527 | 0.003
521.824 | 0.014
522.121 | 0.005
522.417 | 0.004
522.714 | 0.001
523.01 | 0.009
523.307 | 0.008
523.603 | 0.011
523.9 | 0.011
524.196 | 0.012
524.492 | 0.002
524.788 | 0.018
525.084 | 0.012
525.38 | 0.016
525.676 | 0.002
525.972 | 0.007
526.268 | 0.011
526.564 | 0.009
526.859 | 0.009
527.155 | 0.005
527.45 | 0.012
527.746 | 0.013
528.041 | 0.015
528.336 | 0.013
528.632 | 0.005
528.927 | 0.004
529.222 | 0.008
529.517 | 0.015
529.812 | 0.007
530.107 | 0.005
530.402 | 0.01
530.696 | 0.014
530.991 | 0.016
531.286 | 0.012
531.58 | 0.013
531.875 | 0.009
532.169 | 0.011
532.463 | 0.007
532.758 | 0.001
533.052 | 0.01
533.346 | 0.004
533.64 | 0.013
533.934 | 0.009
534.228 | 0.016
534.522 | 0.014
534.816 | 0.013
535.11 | 0.007
535.403 | 0.021
535.697 | 0.014
535.99 | 0.014
536.284 | 0.005
536.577 | 0.021
536.871 | 0.006
537.164 | 0.009
537.457 | 0.016
537.75 | 0.007
538.043 | 0.008
538.336 | 0.013
538.629 | 0.008
538.922 | 0.011
539.215 | 0.001
539.507 | 0.007
539.8 | 0.005
540.093 | 0.017
540.385 | 0.006
540.678 | 0.004
540.97 | -0.001
541.262 | 0.009
541.554 | 0.008
541.847 | 0.011
542.139 | 0.01
542.431 | 0.014
542.723 | 0.009
543.015 | 0.014
543.307 | 0.016
543.598 | 0.014
543.89 | 0.017
544.182 | 0.01
544.473 | 0.013
544.765 | 0.015
545.056 | -0.001
545.347 | 0.009
545.639 | 0.003
545.93 | 0.003
546.221 | 0.01
546.512 | 0.011
546.803 | 0.003
547.094 | 0.003
547.385 | 0.007
547.676 | 0.014
547.967 | 0.021
548.257 | 0.006
548.548 | 0.013
548.839 | 0.003
549.129 | 0.005
549.419 | 0.015
549.71 | 0.009
550 | 0.013
550.29 | 0.011
550.58 | 0.001
550.871 | 0.003
551.161 | 0.003
551.45 | 0
551.74 | 0.007
552.03 | 0.005
552.32 | 0.004
552.61 | 0.007
552.899 | 0.018
553.189 | 0.008
553.478 | 0.01
553.768 | 0.004
554.057 | 0.018
554.346 | 0.016
554.635 | 0.006
554.925 | 0.009
555.214 | 0.015
555.503 | 0.003
555.792 | 0.011
556.081 | 0
556.369 | 0.005
556.658 | 0.005
556.947 | 0.015
557.235 | 0.003
557.524 | 0.021
557.812 | 0.007
558.101 | 0.009
558.389 | 0.015
558.677 | 0.011
558.966 | 0.014
559.254 | 0.019
559.542 | 0.01
559.83 | 0.014
560.118 | 0.011
560.406 | 0.005
560.693 | 0.025
560.981 | 0.018
561.269 | 0.006
561.556 | 0.01
561.844 | 0.014
562.131 | 0.013
562.419 | 0.015
562.706 | 0.012
562.994 | 0.006
563.281 | 0.017
563.568 | 0.012
563.855 | 0.021
564.142 | 0.011
564.429 | 0.016
564.716 | 0.016
565.003 | 0.021
565.289 | 0.013
565.576 | 0.019
565.863 | 0.014
566.149 | 0.016
566.436 | 0.018
566.722 | 0.01
567.008 | 0.009
567.295 | 0.018
567.581 | 0.011
567.867 | 0.009
568.153 | 0.007
568.439 | 0.016
568.725 | 0.009
569.011 | 0.016
569.297 | 0.013
569.583 | 0.013
569.868 | 0.01
570.154 | 0.011
570.439 | 0.007
570.725 | 0.014
571.01 | 0.013
571.296 | 0.02
571.581 | 0.012
571.866 | 0.019
572.151 | 0.017
572.436 | 0.008
572.721 | 0.017
573.006 | 0.013
573.291 | 0.02
573.576 | 0.014
573.861 | 0.018
574.146 | 0.007
574.43 | 0.015
574.715 | 0.018
574.999 | 0.019
575.284 | 0.025
575.568 | 0.009
575.852 | 0.013
576.137 | 0.021
576.421 | 0.026
576.705 | 0.01
576.989 | 0.016
577.273 | 0.003
577.557 | 0.011
577.841 | 0.011
578.124 | 0.02
578.408 | 0.02
578.692 | 0.002
578.975 | 0.015
579.259 | 0.01
579.542 | 0.007
579.826 | 0.007
580.109 | 0.015
580.392 | 0.019
580.676 | 0.007
580.959 | 0.014
581.242 | 0.013
581.525 | 0.012
581.808 | 0.014
582.091 | 0.01
582.373 | 0.007
582.656 | 0.017
582.939 | 0.014
583.221 | 0.004
583.504 | 0.015
583.786 | 0.012
584.069 | 0.005
584.351 | 0.01
584.633 | 0.004
584.916 | 0.001
585.198 | 0.014
585.48 | 0.013
585.762 | -0.003
586.044 | 0.004
586.326 | 0.015
586.607 | 0.008
586.889 | 0.023
587.171 | 0.016
587.453 | 0.007
587.734 | 0.014
588.016 | 0.016
588.297 | 0.003
588.578 | 0.012
588.86 | 0.014
589.141 | 0.01
589.422 | 0.013
589.703 | 0.02
589.984 | 0.005
590.265 | 0.009
590.546 | 0.012
590.827 | 0.004
591.108 | 0.007
591.389 | 0.018
591.669 | 0.005
591.95 | 0.014
592.23 | 0.01
592.511 | 0.02
592.791 | 0.013
593.071 | 0.015
593.352 | 0.001
593.632 | 0.016
593.912 | 0.016
594.192 | 0.022
594.472 | 0.012
594.752 | 0.018
595.032 | 0.013
595.312 | 0
595.592 | 0.014
595.871 | 0.012
596.151 | 0.012
596.43 | -0.001
596.71 | 0.007
596.989 | 0.023
597.269 | 0.009
597.548 | 0.007
597.827 | 0.024
598.106 | 0.016
598.385 | 0.015
598.665 | -0.006
598.944 | 0
599.222 | 0.01
599.501 | 0.007
599.78 | 0.016
600.059 | 0.015
600.337 | 0.008
600.616 | 0.008
600.895 | 0.014
601.173 | 0.027
601.451 | 0.01
601.73 | 0.011
602.008 | 0.013
602.286 | 0.001
602.564 | 0.016
602.842 | 0.008
603.12 | 0.006
603.398 | 0.009
603.676 | 0.011
603.954 | 0.01
604.232 | 0.01
604.51 | 0.005
604.787 | 0.003
605.065 | 0.014
605.342 | 0.005
605.62 | -0.005
605.897 | 0.014
606.174 | 0.015
606.452 | 0.007
606.729 | 0.019
607.006 | 0.02
607.283 | 0.014
607.56 | 0.005
607.837 | 0.005
608.114 | 0.009
608.391 | 0.032
608.667 | -0.009
608.944 | 0.01
609.221 | 0.003
609.497 | 0.008
609.774 | 0
610.05 | 0.019
610.326 | 0.012
610.603 | 0.015
610.879 | 0.019
611.155 | 0.015
611.431 | 0.003
611.707 | 0.014
611.983 | 0.006
612.259 | 0.017
612.535 | 0.004
612.811 | 0.012
613.087 | 0.007
613.362 | 0.01
613.638 | 0.003
613.913 | 0.007
614.189 | 0.009
614.464 | 0.01
614.74 | 0.02
615.015 | -0.002
615.29 | 0.013
615.565 | 0.017
615.84 | -0.004
616.115 | 0.019
616.39 | 0.015
616.665 | 0.015
616.94 | 0.012
617.215 | 0.019
617.49 | 0.009
617.764 | 0.016
618.039 | 0.016
618.313 | 0.019
618.588 | 0.021
618.862 | 0.011
619.137 | 0.019
619.411 | 0.011
619.685 | 0.015
619.959 | 0.007
620.233 | 0.022
620.507 | 0.008
620.781 | 0.014
621.055 | 0.007
621.329 | 0.015
621.603 | 0.008
621.876 | 0.019
622.15 | 0.016
622.424 | 0.009
622.697 | 0.007
622.971 | 0.006
623.244 | 0.005
623.517 | 0.003
623.791 | 0.013
624.064 | 0.016
624.337 | 0.018
624.61 | 0.005
624.883 | 0.012
625.156 | 0.006
625.429 | 0.011
625.702 | 0.003
625.974 | 0.022
626.247 | 0.012
626.52 | 0.018
626.792 | 0.013
627.065 | 0.014
627.337 | 0.012
627.61 | 0.01
627.882 | 0.014
628.154 | 0.013
628.427 | 0.024
628.699 | 0.015
628.971 | 0.01
629.243 | 0
629.515 | 0.016
629.787 | 0.005
630.058 | 0.022
630.33 | 0.019
630.602 | -0.001
630.874 | 0.02
631.145 | 0.026
631.417 | 0
631.688 | 0.02
631.959 | 0.011
632.231 | 0.007
632.502 | 0.019
632.773 | 0.023
633.044 | 0.027
633.316 | 0.019
633.587 | -0.003
633.857 | 0.01
634.128 | 0.019
634.399 | 0.029
634.67 | 0.004
634.941 | 0.007
635.211 | 0.02
635.482 | 0.01
635.752 | 0.013
636.023 | -0.002
636.293 | 0.008
636.564 | 0.027
636.834 | 0.011
637.104 | 0.019
637.374 | 0.01
637.644 | 0.019
637.914 | 0.018
638.184 | 0.015
638.454 | 0.009
638.724 | 0.027
638.994 | 0.012
639.264 | 0.021
639.533 | 0.013
639.803 | 0.012
640.072 | 0.016
640.342 | 0.013
640.611 | 0.02
640.881 | 0.011
641.15 | 0.037
641.419 | 0.001
641.688 | 0.013
641.957 | 0.011
642.227 | 0.016
642.496 | 0.024
642.764 | 0.017
643.033 | 0.023
643.302 | 0.004
643.571 | 0.011
643.84 | 0.007
644.108 | 0.024
644.377 | -0.006
644.645 | 0.006
644.914 | 0.012
645.182 | 0.022
645.45 | 0.022
645.719 | 0.026
645.987 | 0.013
646.255 | 0.012
646.523 | 0.01
646.791 | 0.006
647.059 | 0.024
647.327 | 0.015
647.595 | 0.015
647.862 | -0.001
648.13 | 0.013
648.398 | 0.018
648.665 | 0.01
648.933 | 0.029
649.2 | 0.015
649.468 | 0.015
649.735 | 0.022
650.002 | 0.015
650.269 | 0.009
650.536 | 0.004
650.804 | 0.007
651.071 | 0.026
651.338 | 0.015
651.604 | 0.015
651.871 | 0.007
652.138 | 0.014
652.405 | 0
652.671 | 0.019
652.938 | 0.017
653.205 | 0.007
653.471 | 0.001
653.737 | 0.011
654.004 | 0.002
654.27 | 0.013
654.536 | 0.01
654.803 | 0.006
655.069 | 0.013
655.335 | 0.012
655.601 | 0.019
655.867 | 0.013
656.132 | 0.019
656.398 | 0.017
656.664 | 0.012
656.93 | 0.007
657.195 | 0.01
657.461 | 0.014
657.726 | 0.014
657.992 | 0
658.257 | 0.013
658.523 | 0.019
658.788 | 0.004
659.053 | 0.012
659.318 | 0.007
659.583 | 0.013
659.848 | 0.006
660.113 | 0.014
660.378 | 0.015
660.643 | 0.012
660.908 | 0
661.172 | -0.001
661.437 | 0.008
661.702 | 0.013
661.966 | 0.022
662.231 | 0.014
662.495 | 0.013
662.76 | 0.033
663.024 | 0.014
663.288 | 0.003
663.552 | 0.005
663.816 | 0.012
664.08 | 0.019
664.344 | 0.014
664.608 | 0.015
664.872 | 0.003
665.136 | 0.007
665.4 | 0.014
665.664 | 0.018
665.927 | 0.017
666.191 | 0.01
666.454 | 0.007
666.718 | 0.008
666.981 | 0.003
667.244 | 0.003
667.508 | 0.024
667.771 | 0.011
668.034 | 0.016
668.297 | 0.005
668.56 | 0.016
668.823 | -0.005
669.086 | 0.023
669.349 | 0.009
669.612 | 0.01
669.875 | 0.003
670.137 | 0.017
670.4 | 0.012
670.662 | 0.005
670.925 | 0.021
671.187 | 0.023
671.45 | -0.005
671.712 | 0.03
671.974 | 0.022
672.237 | 0.004
672.499 | 0.005
672.761 | 0.016
673.023 | 0.01
673.285 | 0.013
673.547 | 0.018
673.808 | 0.009
674.07 | 0.008
674.332 | 0.018
674.594 | 0.011
674.855 | 0.014
675.117 | 0.006
675.378 | 0.02
675.64 | 0.009
675.901 | -0.005
676.162 | 0.017
676.424 | 0.005
676.685 | 0.017
676.946 | 0.02
677.207 | 0.026
677.468 | 0.026
677.729 | 0.006
677.99 | 0.003
678.251 | 0.018
678.512 | 0.015
678.772 | 0.014
679.033 | 0.026
679.294 | 0.01
679.554 | 0.017
679.815 | 0.016
680.075 | 0.01
680.335 | 0.016
680.596 | 0.004
680.856 | 0.016
681.116 | 0.032
681.376 | 0.019
681.636 | 0.008
681.896 | 0.024
682.156 | 0.019
682.416 | 0.013
682.676 | 0.009
682.936 | 0.024
683.196 | 0.012
683.455 | 0.013
683.715 | 0.013
683.974 | 0.017
684.234 | 0.009
684.493 | 0
684.753 | 0.014
685.012 | 0.022
685.271 | 0.007
685.53 | 0.007
685.789 | 0.025
686.049 | 0.015
686.308 | 0.01
686.567 | 0.016
686.825 | 0.022
687.084 | 0.017
687.343 | 0.016
687.602 | 0.026
687.86 | 0.015
688.119 | 0.016
688.378 | 0.022
688.636 | 0.021
688.894 | 0.009
689.153 | 0.014
689.411 | 0.024
689.669 | 0.016
689.928 | 0.028
690.186 | -0.005
690.444 | 0.025
690.702 | 0.02
690.96 | 0.031
691.218 | 0.007
691.476 | 0.019
691.733 | 0.019
691.991 | 0.016
692.249 | 0.007
692.506 | 0.012
692.764 | 0.022
693.021 | 0.02
693.279 | 0.016
693.536 | 0.02
693.794 | 0.001
694.051 | 0.012
694.308 | 0.018
694.565 | 0.026
694.822 | 0.031
695.079 | 0.003
695.336 | 0.015
695.593 | 0.004
695.85 | 0.011
696.107 | 0.012
696.364 | 0.015
696.62 | 0.01
696.877 | 0.017
697.133 | 0.006
697.39 | 0.015
697.646 | 0.022
697.903 | 0.009
698.159 | 0
698.415 | 0.008
698.672 | 0.025
698.928 | 0.012
699.184 | 0.02
699.44 | 0.014
699.696 | 0.026
699.952 | -0.002
700.208 | 0.013
700.464 | 0.009
700.719 | 0.024
700.975 | 0.02
701.231 | 0.016
701.486 | 0.025
701.742 | 0.012
701.997 | 0.02
702.253 | 0.02
702.508 | 0.001
702.763 | 0.021
703.019 | 0.019
703.274 | 0.02
703.529 | 0.018
703.784 | 0.011
704.039 | 0.035
704.294 | 0.01
704.549 | 0.011
704.804 | 0.006
705.058 | 0.014
705.313 | 0.006
705.568 | 0.033
705.822 | 0.026
706.077 | 0.014
706.331 | 0.011
706.586 | 0.008
706.84 | 0.021
707.095 | 0.017
707.349 | -0.013
707.603 | 0.004
707.857 | 0.016
708.111 | 0.009
708.365 | 0.009
708.619 | 0.009
708.873 | 0.029
709.127 | 0.008
709.381 | 0.005
709.635 | 0.006
709.888 | 0.027
710.142 | 0.011
710.396 | 0.014
710.649 | 0.01
710.903 | 0.011
711.156 | 0.01
711.409 | 0.024
711.663 | 0.022
711.916 | 0.008
712.169 | 0.012
712.422 | 0.016
712.675 | 0.018
712.928 | 0.028
713.181 | 0.021
713.434 | 0.018
713.687 | 0.032
713.94 | 0.008
714.192 | 0.019
714.445 | 0
714.698 | 0.021
714.95 | 0.023
715.203 | 0.039
715.455 | 0.015
715.708 | 0.012
715.96 | 0.035
716.212 | 0.019
716.464 | 0.028
716.717 | 0.023
716.969 | 0.025
717.221 | 0.026
717.473 | -0.002
717.725 | 0.021
717.977 | 0.017
718.228 | 0.029
718.48 | 0.013
718.732 | 0.024
718.984 | 0.032
719.235 | 0.027
719.487 | 0.028
719.738 | 0.014
719.99 | 0.035
720.241 | 0.012
720.492 | 0.029
720.744 | 0.012
720.995 | 0.007
721.246 | 0.015
721.497 | 0.032
721.748 | 0.02
721.999 | 0.01
722.25 | 0.046
722.501 | 0.036
722.752 | 0.022
723.002 | 0.012
723.253 | 0.005
723.504 | 0.032
723.754 | 0.025
724.005 | 0.035
724.255 | 0.032
724.506 | 0.025
724.756 | 0.035
725.006 | 0.009
725.257 | 0.012
725.507 | 0.025
725.757 | 0.012
726.007 | 0.011
726.257 | 0.031
726.507 | 0.006
726.757 | 0.045
727.007 | 0.029
727.257 | 0.003
727.506 | 0.027
727.756 | 0.018
728.006 | 0.028
728.255 | 0.034
728.505 | 0.002
728.754 | 0.02
729.004 | 0.037
729.253 | 0.044
729.502 | 0.051
729.751 | 0.019
730.001 | 0.008
730.25 | 0.014
730.499 | 0.031
730.748 | 0.024
730.997 | 0.014
731.246 | 0.026
731.495 | 0.018
731.743 | 0.027
731.992 | 0.025
732.241 | 0.009
732.489 | 0.025
732.738 | 0.038
732.987 | 0.017
733.235 | 0.024
733.483 | 0.036
733.732 | 0.034
733.98 | 0.009
734.228 | 0.008
734.477 | 0.007
734.725 | 0.034
734.973 | 0.018
735.221 | 0.026
735.469 | 0.035
735.717 | 0.032
735.965 | 0.013
736.212 | 0.012
736.46 | 0.034
736.708 | 0.029
736.955 | 0.047
737.203 | 0.01
737.451 | 0.043
737.698 | 0.022
737.945 | 0.019
738.193 | 0.031
738.44 | 0.015
738.687 | 0.023
738.935 | 0.023
739.182 | 0.014
739.429 | 0.035
739.676 | 0.008
739.923 | 0.029
740.17 | 0.031
740.417 | 0.012
740.663 | 0.013
740.91 | 0.028
741.157 | 0.035
741.403 | 0.014
741.65 | -0.007
741.897 | 0.022
742.143 | 0.012
742.389 | 0.03
742.636 | 0.006
742.882 | 0.01
743.128 | 0.022
743.375 | 0.003
743.621 | 0.045
743.867 | 0.022
744.113 | 0.022
744.359 | 0.016
744.605 | 0.027
744.851 | 0.049
745.097 | 0.044
745.342 | 0.02
745.588 | 0.025
745.834 | 0.018
746.079 | 0.037
746.325 | 0.028
746.57 | 0.024
746.816 | 0.009
747.061 | 0.021
747.307 | 0.024
747.552 | 0.012
747.797 | 0.037
748.042 | 0.023
748.287 | 0.02
748.532 | 0.011
748.777 | 0.034
749.022 | 0.031
749.267 | 0.028
749.512 | 0.019
749.757 | 0.02
750.002 | 0.021
750.246 | 0.02
750.491 | 0.03
750.736 | 0.015
750.98 | 0.024
751.225 | 0.022
751.469 | 0.026
751.713 | 0.031
751.958 | 0.025
752.202 | 0.019
752.446 | 0.024
752.69 | 0.007
752.934 | 0.01
753.178 | 0.026
753.422 | 0.013
753.666 | 0.033
753.91 | 0.01
754.154 | 0.034
754.398 | 0.058
754.641 | 0.005
754.885 | 0.004
755.129 | 0.021
755.372 | -0.002
755.616 | 0.031
755.859 | 0.031
756.103 | 0.058
756.346 | -0.007
756.589 | 0.017
756.832 | 0.019
757.076 | 0.039
757.319 | 0.036
757.562 | 0.033
757.805 | 0.032
758.048 | 0.038
758.291 | 0.02
758.533 | 0.024
758.776 | 0.016
759.019 | 0.026
759.262 | 0.04
759.504 | 0.032
759.747 | 0.027
759.989 | 0.039
760.232 | 0.036
760.474 | 0.02
760.717 | 0.055
760.959 | 0.018
761.201 | 0.018
761.444 | 0.027
761.686 | 0.031
761.928 | 0.016
762.17 | 0.02
762.412 | 0.029
762.654 | 0.037
762.896 | 0.018
763.137 | 0.024
763.379 | 0.013
763.621 | 0.035
763.863 | 0.029
764.104 | 0.029
764.346 | 0.05
764.587 | 0.021
764.829 | 0.025
765.07 | 0.009
765.312 | 0.051
765.553 | 0.05
765.794 | 0.033
766.035 | 0.04
766.276 | 0.045
766.518 | 0.032
766.759 | 0.052
767 | 0.027
767.241 | 0.033
767.481 | 0.029
767.722 | 0.027
767.963 | 0.038
768.204 | 0.029
768.444 | 0.048
768.685 | 0.039
768.926 | 0.015
769.166 | 0.026
769.407 | 0.02
769.647 | 0.054
769.887 | 0.005
770.128 | 0.035
770.368 | 0.033
770.608 | 0.045
770.848 | 0.04
771.088 | 0.028
771.328 | 0.036
771.568 | 0.021
771.808 | 0.033
772.048 | 0.02
772.288 | 0.025
772.528 | 0.009
772.767 | 0.022
773.007 | 0.033
773.247 | 0.04
773.486 | 0.03
773.726 | 0.003
773.965 | 0.043
774.205 | 0.038
774.444 | 0.036
774.683 | 0.042
774.922 | 0.05
775.162 | 0.027
775.401 | 0.046
775.64 | 0.03
775.879 | 0.032
776.118 | 0.03
776.357 | 0.03
776.596 | 0.037
776.834 | 0.05
777.073 | 0.033
777.312 | 0.034
777.551 | 0.017
777.789 | 0.015
778.028 | 0.025
778.266 | 0.044
778.505 | 0.023
778.743 | 0.018
778.982 | 0.014
779.22 | 0.014
779.458 | 0.037
779.696 | 0.027
779.934 | 0.039
780.173 | 0.037
780.411 | 0.037
780.649 | 0.023
780.887 | 0.036
781.124 | 0.047
781.362 | 0.014
781.6 | 0.036
781.838 | 0.027
782.075 | 0.017
782.313 | 0.027
782.551 | 0.007
782.788 | 0.019
783.026 | 0.027
783.263 | 0.028
783.5 | 0.026
783.738 | 0.025
783.975 | 0.013
784.212 | -0.009
784.449 | 0.018
784.686 | 0.019
784.924 | 0.031
785.161 | 0.017
785.397 | 0.028
785.634 | 0.043
785.871 | 0.053
786.108 | 0.021
786.345 | 0.044
786.581 | 0.024
786.818 | 0.04
787.055 | 0.03
787.291 | 0.048
787.528 | 0.041
787.764 | 0.026
788.001 | 0.021
788.237 | 0.043
788.473 | 0.008
788.709 | 0.022
788.946 | 0.025
789.182 | 0.007
789.418 | 0.008
789.654 | 0.043
789.89 | 0.014
790.126 | 0.049
790.362 | 0.059
790.597 | 0.009
790.833 | 0.028
791.069 | 0.016
791.305 | 0.015
791.54 | 0.036
791.776 | 0.05
792.011 | 0.036
792.247 | 0.003
792.482 | 0.022
792.717 | 0.042
792.953 | 0.026
793.188 | 0.041
793.423 | 0.048
793.658 | 0.051
793.894 | 0.057
794.129 | 0.031
794.364 | 0.038
794.599 | 0.041
794.833 | 0.034
795.068 | 0.036
795.303 | 0.041
795.538 | 0.01
795.772 | 0.037
796.007 | 0.035
796.242 | 0.051
796.476 | 0.04
796.711 | 0.014
796.945 | 0.02
797.18 | 0.015
797.414 | 0.002
797.648 | 0.057
797.882 | 0.036
798.117 | 0.024
798.351 | 0.018
798.585 | 0.071
798.819 | 0.014
799.053 | 0.036
799.287 | 0.031
799.521 | 0.06
799.755 | 0.044
799.988 | 0.042
800.222 | 0.036
800.456 | 0.03
800.689 | -0.012
800.923 | 0.028
801.156 | 0.039
801.39 | 0.02
801.623 | 0.044
801.857 | 0.03
802.09 | 0.035
802.323 | 0.051
802.557 | 0.029
802.79 | 0.035
803.023 | 0.008
803.256 | 0.018
803.489 | 0.026
803.722 | 0.017
803.955 | 0.016
804.188 | 0.057
804.421 | 0.025
804.653 | 0.045
804.886 | 0.034
805.119 | 0.009
805.351 | 0.027
805.584 | 0.047
805.816 | -0.003
806.049 | 0.025
806.281 | 0.018
806.514 | 0.038
806.746 | 0.045
806.978 | 0.02
807.211 | 0.003
807.443 | 0.028
807.675 | 0.051
background 3
AbsoluteIrradiance:211
195.562 | 0
195.934 | 0
196.307 | 0
196.679 | 0
197.051 | 0
197.423 | 0
197.795 | 0
198.167 | 0
198.539 | 0
198.911 | 0
199.283 | 0
199.654 | 0
200.026 | 0
200.397 | 0
200.769 | 0
201.14 | 0
201.511 | 0
201.882 | 0
202.253 | 0
202.624 | 0
202.995 | 0
203.366 | 0
203.736 | 0
204.107 | 0
204.477 | 0
204.848 | 2.213
205.218 | 4.72
205.588 | 0.704
205.959 | 0.918
206.329 | 1.445
206.699 | 0.904
207.068 | 1.499
207.438 | 1.847
207.808 | 0.776
208.178 | 1.533
208.547 | 1.099
208.917 | 1.082
209.286 | 1.799
209.655 | 1.597
210.025 | 0.73
210.394 | 0.966
210.763 | 1.176
211.132 | 0.58
211.501 | 0.168
211.869 | 0.463
212.238 | 1.037
212.607 | 1.272
212.975 | 0.416
213.344 | 0.404
213.712 | 0.438
214.08 | 0.599
214.448 | 0.627
214.817 | 0.531
215.185 | 1.22
215.553 | 0.816
215.92 | 0.782
216.288 | 0.687
216.656 | 0.747
217.023 | 0.412
217.391 | 0.35
217.758 | 0.897
218.126 | 0.55
218.493 | 0.475
218.86 | 0.529
219.227 | 0.228
219.594 | 0.236
219.961 | 0.57
220.328 | 0.334
220.695 | 0.511
221.061 | 0.545
221.428 | 0.47
221.795 | 0.615
222.161 | 0.553
222.527 | 0.295
222.893 | 0.552
223.26 | 0.123
223.626 | 0.707
223.992 | 0.404
224.358 | 0.219
224.723 | 0.462
225.089 | 0.025
225.455 | 0.353
225.82 | 0.213
226.186 | 0.106
226.551 | 0.04
226.917 | 0.179
227.282 | 0.246
227.647 | 0.332
228.012 | 0.249
228.377 | 0.185
228.742 | 0.234
229.107 | 0.094
229.471 | 0.228
229.836 | 0.107
230.201 | 0.222
230.565 | 0.199
230.929 | 0.165
231.294 | 0.066
231.658 | 0.103
232.022 | 0.079
232.386 | 0.07
232.75 | -0.069
233.114 | 0.08
233.478 | 0.092
233.842 | 0.118
234.205 | 0.119
234.569 | 0.056
234.932 | 0.049
235.296 | 0.036
235.659 | 0.075
236.022 | 0.064
236.385 | 0.033
236.748 | 0.111
237.111 | 0.019
237.474 | 0.116
237.837 | 0.092
238.2 | 0.073
238.562 | 0.05
238.925 | 0.085
239.287 | 0.098
239.65 | 0.038
240.012 | 0.096
240.374 | 0.07
240.736 | 0.083
241.098 | 0.06
241.46 | 0.103
241.822 | 0.095
242.184 | 0.036
242.546 | 0.071
242.907 | 0.053
243.269 | 0.028
243.63 | 0.014
243.992 | 0.084
244.353 | 0.05
244.714 | 0.047
245.075 | 0.067
245.436 | 0.06
245.797 | 0.032
246.158 | -0.007
246.519 | 0.058
246.879 | 0.074
247.24 | 0.045
247.601 | 0.064
247.961 | 0.062
248.321 | 0.011
248.682 | 0.034
249.042 | 0.028
249.402 | 0.02
249.762 | 0.077
250.122 | 0.034
250.482 | 0.075
250.842 | 0.028
251.201 | 0.033
251.561 | 0.051
251.921 | 0.05
252.28 | 0.057
252.639 | 0.039
252.999 | 0.054
253.358 | 0.038
253.717 | 0.033
254.076 | 0.038
254.435 | 0.026
254.794 | 0.003
255.153 | 0.066
255.511 | 0.047
255.87 | 0.017
256.228 | 0.044
256.587 | 0.034
256.945 | 0.025
257.304 | 0.008
257.662 | 0.032
258.02 | 0.032
258.378 | 0.048
258.736 | 0.015
259.094 | 0.068
259.452 | 0.047
259.809 | 0.04
260.167 | 0.046
260.524 | 0.018
260.882 | 0.016
261.239 | 0.029
261.597 | 0.034
261.954 | 0.026
262.311 | 0.041
262.668 | 0.038
263.025 | 0.032
263.382 | 0.033
263.739 | 0.014
264.095 | 0.051
264.452 | 0.002
264.809 | 0.032
265.165 | 0.026
265.521 | 0.058
265.878 | 0.033
266.234 | 0.032
266.59 | 0.029
266.946 | 0.068
267.302 | 0.033
267.658 | 0.044
268.014 | 0.034
268.37 | 0.032
268.725 | 0.029
269.081 | 0.03
269.436 | 0.029
269.792 | 0.006
270.147 | 0.025
270.502 | 0.053
270.857 | 0.046
271.212 | 0.016
271.567 | 0.002
271.922 | 0.065
272.277 | 0.018
272.632 | 0.039
272.987 | 0.059
273.341 | 0.035
273.696 | 0.043
274.05 | 0.004
274.404 | 0.021
274.759 | 0.054
275.113 | 0.047
275.467 | 0.014
275.821 | 0.013
276.175 | 0.043
276.529 | 0.053
276.882 | 0.075
277.236 | 0.044
277.59 | 0.05
277.943 | 0.004
278.297 | 0.081
278.65 | 0.02
279.003 | 0.008
279.356 | 0.027
279.709 | 0.034
280.062 | 0.019
280.415 | 0.04
280.768 | 0.044
281.121 | 0.067
281.474 | 0.031
281.826 | 0.038
282.179 | 0.032
282.531 | 0.017
282.884 | 0.04
283.236 | 0.032
283.588 | 0.066
283.94 | 0.024
284.292 | 0.039
284.644 | 0.016
284.996 | 0.035
285.348 | 0.034
285.7 | 0.056
286.051 | 0.028
286.403 | 0.036
286.754 | 0.033
287.106 | 0.057
287.457 | 0.051
287.808 | 0.036
288.159 | 0.048
288.51 | 0.046
288.861 | 0.027
289.212 | 0.009
289.563 | 0.014
289.914 | 0.046
290.264 | 0.019
290.615 | 0.039
290.966 | 0.047
291.316 | 0.019
291.666 | 0.048
292.017 | 0.035
292.367 | 0.017
292.717 | 0.013
293.067 | 0.025
293.417 | 0.027
293.767 | 0.024
294.116 | 0.048
294.466 | 0.026
294.816 | 0.04
295.165 | 0.021
295.515 | 0.014
295.864 | 0.032
296.213 | 0.018
296.562 | 0.011
296.911 | 0.039
297.261 | 0.024
297.609 | 0.026
297.958 | -0.005
298.307 | 0.023
298.656 | 0.03
299.004 | 0.023
299.353 | -0.005
299.701 | 0.023
300.05 | 0.015
300.398 | 0.027
300.746 | 0.022
301.095 | 0.019
301.443 | 0.02
301.791 | 0.012
302.138 | 0.02
302.486 | 0.006
302.834 | 0.009
303.182 | 0.005
303.529 | 0.021
303.877 | 0.02
304.224 | 0.025
304.572 | 0.022
304.919 | 0.005
305.266 | 0.019
305.613 | 0.012
305.96 | 0.029
306.307 | 0.015
306.654 | 0.03
307.001 | 0.028
307.347 | 0.016
307.694 | 0.011
308.04 | 0.014
308.387 | 0.027
308.733 | 0.018
309.08 | 0.014
309.426 | 0.031
309.772 | 0.022
310.118 | 0.028
310.464 | 0.013
310.81 | 0.022
311.156 | 0.014
311.501 | 0.023
311.847 | 0.018
312.193 | 0.014
312.538 | 0.016
312.883 | 0.019
313.229 | 0.018
313.574 | 0.013
313.919 | 0.018
314.264 | 0.027
314.609 | 0.023
314.954 | 0.012
315.299 | 0.015
315.644 | 0.009
315.988 | 0.026
316.333 | 0.028
316.678 | 0.016
317.022 | 0.017
317.366 | 0.014
317.711 | 0.021
318.055 | 0.009
318.399 | 0.014
318.743 | 0.019
319.087 | 0.019
319.431 | 0.021
319.775 | 0.017
320.118 | 0.024
320.462 | 0.016
320.806 | 0.013
321.149 | 0.02
321.493 | 0.021
321.836 | 0.013
322.179 | 0.004
322.522 | 0.016
322.865 | 0.019
323.208 | 0.018
323.551 | 0.023
323.894 | 0.013
324.237 | 0.01
324.58 | 0.015
324.922 | 0.017
325.265 | 0.021
325.607 | 0.024
325.95 | 0.024
326.292 | 0.001
326.634 | 0.014
326.976 | 0.01
327.318 | 0.012
327.66 | 0.014
328.002 | -0.002
328.344 | 0.016
328.686 | 0.019
329.027 | 0.006
329.369 | 0.008
329.711 | 0.018
330.052 | 0.016
330.393 | 0.008
330.735 | 0.009
331.076 | 0.016
331.417 | 0.009
331.758 | 0.015
332.099 | 0.001
332.44 | 0.018
332.78 | 0.006
333.121 | 0.014
333.462 | 0.013
333.802 | 0.013
334.143 | 0.011
334.483 | 0.012
334.823 | 0.007
335.164 | 0.015
335.504 | 0.008
335.844 | 0.017
336.184 | 0.006
336.524 | 0.011
336.864 | 0.014
337.203 | 0.011
337.543 | 0.022
337.883 | 0.014
338.222 | 0.024
338.562 | 0.005
338.901 | 0.018
339.24 | 0.001
339.58 | 0.013
339.919 | 0.005
340.258 | 0.017
340.597 | 0.013
340.936 | 0.005
341.274 | 0.011
341.613 | 0.021
341.952 | 0.017
342.29 | 0.012
342.629 | 0.024
342.967 | 0.009
343.306 | 0.014
343.644 | 0.015
343.982 | 0.011
344.32 | 0.015
344.658 | 0.007
344.996 | 0.013
345.334 | 0.008
345.672 | 0.02
346.01 | 0.006
346.347 | 0.03
346.685 | 0.014
347.022 | 0.022
347.36 | 0.023
347.697 | 0.015
348.034 | 0.012
348.371 | 0.016
348.708 | 0.011
349.045 | 0.02
349.382 | 0.018
349.719 | 0.008
350.056 | 0.011
350.393 | 0.009
350.729 | 0.021
351.066 | 0.013
351.402 | 0.015
351.739 | 0.023
352.075 | 0.019
352.411 | 0.012
352.747 | 0.011
353.083 | 0.014
353.419 | 0.01
353.755 | 0.011
354.091 | 0.013
354.427 | 0.012
354.763 | 0.015
355.098 | 0.019
355.434 | 0.004
355.769 | 0.025
356.104 | 0.004
356.44 | 0.007
356.775 | 0.015
357.11 | 0.014
357.445 | 0.011
357.78 | 0.007
358.115 | 0.018
358.45 | 0.016
358.785 | 0.015
359.119 | 0.007
359.454 | 0.014
359.788 | 0.013
360.123 | 0.011
360.457 | 0.004
360.791 | 0.019
361.126 | 0.007
361.46 | 0.008
361.794 | 0.014
362.128 | 0.018
362.462 | 0.024
362.795 | 0.021
363.129 | 0.015
363.463 | 0.004
363.796 | 0.019
364.13 | 0.024
364.463 | 0.018
364.797 | 0.016
365.13 | 0.02
365.463 | 0.006
365.796 | 0.007
366.129 | 0.015
366.462 | 0.006
366.795 | 0.011
367.128 | 0.021
367.461 | 0.004
367.793 | 0.025
368.126 | 0.01
368.458 | 0.021
368.791 | 0.014
369.123 | 0.015
369.455 | 0.015
369.787 | 0.009
370.12 | 0.006
370.452 | 0.003
370.784 | 0.011
371.115 | 0.007
371.447 | 0.003
371.779 | 0.009
372.111 | 0.017
372.442 | 0.011
372.774 | 0.021
373.105 | 0.016
373.436 | 0.011
373.768 | 0.009
374.099 | 0.003
374.43 | 0.019
374.761 | 0.01
375.092 | 0.014
375.423 | 0.012
375.754 | 0.01
376.084 | 0.009
376.415 | 0.01
376.745 | 0.013
377.076 | 0.014
377.406 | 0.012
377.737 | 0.024
378.067 | 0.018
378.397 | 0.008
378.727 | 0.011
379.057 | 0.009
379.387 | 0.012
379.717 | 0.011
380.047 | 0.012
380.377 | 0.012
380.706 | 0.019
381.036 | 0.013
381.365 | 0.023
381.695 | 0.011
382.024 | 0.008
382.353 | -0.001
382.682 | 0.01
383.012 | 0.005
383.341 | 0.015
383.67 | 0.015
383.998 | 0.005
384.327 | 0.017
384.656 | 0.006
384.985 | 0.002
385.313 | 0.002
385.642 | 0.006
385.97 | 0.009
386.298 | 0.005
386.627 | 0.005
386.955 | 0.02
387.283 | 0.006
387.611 | 0.006
387.939 | 0.009
388.267 | 0.008
388.595 | 0.02
388.922 | -0.002
389.25 | 0.021
389.578 | 0.011
389.905 | 0.011
390.233 | 0.01
390.56 | 0.01
390.887 | 0.022
391.214 | 0.014
391.542 | 0.001
391.869 | 0.009
392.196 | 0.004
392.522 | 0.004
392.849 | 0.006
393.176 | 0.003
393.503 | 0.011
393.829 | 0.009
394.156 | 0.007
394.482 | 0.014
394.809 | 0.009
395.135 | 0.014
395.461 | 0.005
395.787 | 0.016
396.113 | 0.002
396.439 | 0.009
396.765 | 0.004
397.091 | 0.007
397.417 | 0.005
397.742 | 0.01
398.068 | 0.013
398.394 | 0.009
398.719 | 0.005
399.044 | 0.007
399.37 | 0.015
399.695 | 0.01
400.02 | 0.015
400.345 | 0.016
400.67 | 0.014
400.995 | 0.021
401.32 | 0.013
401.645 | 0.008
401.969 | 0.006
402.294 | 0.007
402.618 | 0.02
402.943 | 0.002
403.267 | 0.004
403.592 | 0.008
403.916 | 0.018
404.24 | 0.009
404.564 | 0.006
404.888 | 0.023
405.212 | 0.013
405.536 | 0.002
405.86 | 0.014
406.183 | 0.01
406.507 | 0.011
406.831 | 0.015
407.154 | 0.01
407.478 | 0.009
407.801 | 0.009
408.124 | 0.001
408.447 | 0.009
408.77 | 0.013
409.093 | 0.014
409.416 | 0.004
409.739 | 0.015
410.062 | 0.001
410.385 | 0.013
410.708 | 0.014
411.03 | 0.016
411.353 | 0.002
411.675 | 0.012
411.997 | 0.008
412.32 | 0.006
412.642 | 0.01
412.964 | 0.009
413.286 | 0.009
413.608 | 0.011
413.93 | 0.01
414.252 | 0.001
414.574 | 0.004
414.895 | 0.008
415.217 | 0.014
415.538 | 0.008
415.86 | 0.008
416.181 | 0.014
416.503 | 0.009
416.824 | 0.01
417.145 | 0.007
417.466 | 0.008
417.787 | 0.016
418.108 | 0.004
418.429 | 0.006
418.75 | 0.01
419.07 | 0.006
419.391 | 0.013
419.712 | 0.01
420.032 | 0.007
420.353 | 0.005
420.673 | 0.009
420.993 | 0.012
421.313 | 0.012
421.633 | 0.01
421.953 | 0.007
422.273 | 0.009
422.593 | 0.011
422.913 | 0.006
423.233 | 0.009
423.553 | 0.007
423.872 | 0.006
424.192 | 0.008
424.511 | 0.006
424.83 | 0.012
425.15 | 0.014
425.469 | 0.014
425.788 | 0.016
426.107 | 0.011
426.426 | 0.015
426.745 | 0.018
427.064 | 0.012
427.383 | 0.013
427.702 | 0.001
428.02 | 0.012
428.339 | 0.012
428.657 | 0.01
428.976 | 0.013
429.294 | 0.01
429.612 | 0.005
429.93 | 0.012
430.248 | 0.018
430.567 | 0.011
430.884 | 0.009
431.202 | 0.006
431.52 | 0.005
431.838 | 0.006
432.156 | 0.007
432.473 | 0.01
432.791 | 0.013
433.108 | 0.007
433.426 | 0.017
433.743 | 0.013
434.06 | 0.009
434.377 | 0.013
434.694 | 0.015
435.011 | 0.005
435.328 | 0.01
435.645 | 0.011
435.962 | 0.007
436.279 | 0.01
436.595 | 0.003
436.912 | 0.005
437.228 | 0.009
437.545 | 0.003
437.861 | 0.007
438.177 | 0.004
438.493 | 0.003
438.81 | 0.013
439.126 | 0.007
439.442 | 0.012
439.757 | -0.002
440.073 | 0.005
440.389 | 0.006
440.705 | 0.018
441.02 | 0.008
441.336 | 0.005
441.651 | 0.012
441.967 | 0.002
442.282 | 0.001
442.597 | 0.009
442.912 | 0.015
443.227 | 0.011
443.543 | 0.014
443.857 | 0.013
444.172 | 0.006
444.487 | 0.01
444.802 | 0.014
445.116 | 0.015
445.431 | 0.003
445.746 | 0.002
446.06 | 0.008
446.374 | 0.007
446.689 | 0.015
447.003 | 0.009
447.317 | 0.011
447.631 | -0.002
447.945 | 0.011
448.259 | 0.014
448.573 | 0.016
448.887 | 0.003
449.2 | 0.017
449.514 | 0.014
449.827 | 0.012
450.141 | 0.01
450.454 | 0.009
450.768 | 0.008
451.081 | 0.013
451.394 | 0.006
451.707 | 0.006
452.02 | 0.015
452.333 | 0.009
452.646 | 0.006
452.959 | 0.005
453.272 | 0.008
453.584 | 0.007
453.897 | 0.007
454.21 | 0.013
454.522 | 0.007
454.834 | 0.007
455.147 | 0.009
455.459 | -0.001
455.771 | -0.002
456.083 | 0.004
456.395 | 0.008
456.707 | 0.01
457.019 | 0.013
457.331 | 0.009
457.643 | 0.004
457.954 | 0.011
458.266 | 0.008
458.577 | 0.005
458.889 | 0.005
459.2 | 0.009
459.512 | 0.005
459.823 | 0.007
460.134 | 0.01
460.445 | 0.012
460.756 | 0.007
461.067 | 0.011
461.378 | 0.007
461.689 | 0.005
461.999 | 0.017
462.31 | 0.011
462.621 | 0.003
462.931 | 0.009
463.242 | 0.011
463.552 | 0.015
463.862 | 0.005
464.172 | 0.01
464.483 | 0.017
464.793 | 0.007
465.103 | 0.012
465.413 | 0.016
465.722 | 0.008
466.032 | 0.011
466.342 | 0.008
466.652 | 0.012
466.961 | 0.012
467.271 | 0.013
467.58 | 0.012
467.889 | 0.013
468.199 | 0.012
468.508 | 0.008
468.817 | 0.011
469.126 | 0.01
469.435 | 0.016
469.744 | 0.011
470.053 | 0.01
470.362 | 0.008
470.67 | 0.002
470.979 | 0.011
471.287 | 0.006
471.596 | 0.007
471.904 | 0.017
472.213 | 0.004
472.521 | 0.012
472.829 | 0.01
473.137 | 0.008
473.445 | 0.003
473.753 | 0.023
474.061 | 0.006
474.369 | 0.011
474.677 | 0.014
474.985 | 0.005
475.292 | 0.003
475.6 | 0.015
475.907 | 0.011
476.215 | 0.01
476.522 | 0.016
476.829 | 0.011
477.137 | 0.014
477.444 | 0.005
477.751 | 0.014
478.058 | 0.005
478.365 | 0.009
478.671 | 0.016
478.978 | 0.013
479.285 | 0.004
479.592 | 0.007
479.898 | 0.004
480.205 | 0.012
480.511 | 0
480.817 | 0.016
481.124 | 0.008
481.43 | 0.01
481.736 | 0.021
482.042 | 0.01
482.348 | 0.008
482.654 | 0.016
482.96 | 0.013
483.266 | 0.012
483.571 | 0.015
483.877 | 0
484.182 | 0.006
484.488 | 0.009
484.793 | 0.004
485.099 | 0.005
485.404 | 0.012
485.709 | 0.014
486.014 | 0.011
486.319 | 0.018
486.624 | 0.008
486.929 | 0.006
487.234 | -0.003
487.539 | 0.01
487.844 | 0.016
488.148 | 0.009
488.453 | 0.018
488.757 | 0.012
489.062 | 0.006
489.366 | 0.014
489.67 | 0.009
489.975 | 0.003
490.279 | 0.008
490.583 | 0.011
490.887 | 0.008
491.191 | 0.008
491.495 | 0.014
491.798 | 0.014
492.102 | 0.013
492.406 | 0.007
492.709 | 0.013
493.013 | 0.018
493.316 | 0.012
493.62 | 0.019
493.923 | 0.011
494.226 | 0.015
494.529 | 0.012
494.832 | 0.01
495.135 | 0.01
495.438 | 0.015
495.741 | 0.023
496.044 | 0.008
496.347 | 0.004
496.649 | 0.018
496.952 | 0.009
497.255 | 0.008
497.557 | 0.009
497.859 | 0.011
498.162 | 0.002
498.464 | 0.017
498.766 | 0.012
499.068 | 0.008
499.37 | 0.011
499.672 | 0.016
499.974 | 0.009
500.276 | 0.012
500.578 | 0.016
500.879 | 0.002
501.181 | 0.019
501.483 | 0.008
501.784 | 0.015
502.085 | 0.006
502.387 | 0.005
502.688 | 0.004
502.989 | 0.014
503.29 | 0.002
503.591 | 0.003
503.892 | 0.006
504.193 | 0.017
504.494 | 0.008
504.795 | 0.003
505.096 | 0.022
505.396 | 0.004
505.697 | -0.001
505.997 | 0.009
506.298 | 0.012
506.598 | 0.008
506.898 | 0.011
507.199 | 0.006
507.499 | 0.01
507.799 | 0.009
508.099 | 0.008
508.399 | 0.012
508.699 | 0.012
508.998 | 0.011
509.298 | 0.017
509.598 | 0.004
509.897 | 0.013
510.197 | 0.005
510.496 | 0.009
510.796 | 0.005
511.095 | 0.008
511.394 | 0.012
511.693 | 0.016
511.992 | 0.011
512.291 | 0.011
512.59 | 0.014
512.889 | 0.015
513.188 | 0.014
513.487 | 0.011
513.786 | 0.014
514.084 | 0.011
514.383 | 0.015
514.681 | 0.02
514.98 | 0.011
515.278 | 0.017
515.576 | 0.019
515.874 | 0.009
516.173 | 0.005
516.471 | 0.018
516.769 | 0.012
517.067 | 0.011
517.364 | 0.006
517.662 | 0.001
517.96 | 0.005
518.257 | 0.007
518.555 | 0.007
518.853 | 0.007
519.15 | 0.005
519.447 | 0.008
519.745 | 0.01
520.042 | 0.017
520.339 | 0.004
520.636 | 0.014
520.933 | 0.02
521.23 | 0.006
521.527 | 0.005
521.824 | 0.012
522.121 | 0.007
522.417 | 0.008
522.714 | 0.003
523.01 | 0.01
523.307 | 0.011
523.603 | 0.012
523.9 | 0.012
524.196 | 0.013
524.492 | 0
524.788 | 0.02
525.084 | 0.01
525.38 | 0.016
525.676 | 0.003
525.972 | 0.01
526.268 | 0.01
526.564 | 0.01
526.859 | 0.015
527.155 | 0.005
527.45 | 0.015
527.746 | 0.01
528.041 | 0.015
528.336 | 0.015
528.632 | 0.007
528.927 | 0.007
529.222 | 0.011
529.517 | 0.015
529.812 | 0.009
530.107 | 0.007
530.402 | 0.009
530.696 | 0.014
530.991 | 0.018
531.286 | 0.013
531.58 | 0.015
531.875 | 0.007
532.169 | 0.012
532.463 | 0.006
532.758 | 0.002
533.052 | 0.009
533.346 | 0.005
533.64 | 0.011
533.934 | 0.011
534.228 | 0.017
534.522 | 0.016
534.816 | 0.01
535.11 | 0.004
535.403 | 0.019
535.697 | 0.015
535.99 | 0.014
536.284 | 0.008
536.577 | 0.025
536.871 | 0.009
537.164 | 0.009
537.457 | 0.015
537.75 | 0.011
538.043 | 0.011
538.336 | 0.013
538.629 | 0.01
538.922 | 0.013
539.215 | 0.004
539.507 | 0.009
539.8 | 0.001
540.093 | 0.016
540.385 | 0.005
540.678 | 0.006
540.97 | -0.001
541.262 | 0.011
541.554 | 0.01
541.847 | 0.011
542.139 | 0.012
542.431 | 0.013
542.723 | 0.012
543.015 | 0.013
543.307 | 0.016
543.598 | 0.015
543.89 | 0.016
544.182 | 0.011
544.473 | 0.017
544.765 | 0.021
545.056 | 0.003
545.347 | 0.007
545.639 | 0.007
545.93 | 0.005
546.221 | 0.007
546.512 | 0.011
546.803 | 0.003
547.094 | 0.005
547.385 | 0.007
547.676 | 0.013
547.967 | 0.026
548.257 | 0.004
548.548 | 0.017
548.839 | 0.004
549.129 | 0.009
549.419 | 0.016
549.71 | 0.012
550 | 0.017
550.29 | 0.011
550.58 | 0.002
550.871 | 0.004
551.161 | 0.004
551.45 | 0
551.74 | 0.009
552.03 | 0.004
552.32 | 0.01
552.61 | 0.01
552.899 | 0.021
553.189 | 0.006
553.478 | 0.007
553.768 | 0.006
554.057 | 0.02
554.346 | 0.016
554.635 | 0.011
554.925 | 0.01
555.214 | 0.016
555.503 | 0.004
555.792 | 0.017
556.081 | 0.001
556.369 | 0.01
556.658 | 0.006
556.947 | 0.019
557.235 | 0.004
557.524 | 0.026
557.812 | 0.011
558.101 | 0.009
558.389 | 0.019
558.677 | 0.015
558.966 | 0.014
559.254 | 0.02
559.542 | 0.014
559.83 | 0.013
560.118 | 0.017
560.406 | 0.005
560.693 | 0.027
560.981 | 0.017
561.269 | 0.011
561.556 | 0.016
561.844 | 0.015
562.131 | 0.014
562.419 | 0.021
562.706 | 0.011
562.994 | 0.01
563.281 | 0.021
563.568 | 0.016
563.855 | 0.024
564.142 | 0.014
564.429 | 0.019
564.716 | 0.016
565.003 | 0.023
565.289 | 0.016
565.576 | 0.021
565.863 | 0.019
566.149 | 0.017
566.436 | 0.021
566.722 | 0.009
567.008 | 0.012
567.295 | 0.021
567.581 | 0.012
567.867 | 0.012
568.153 | 0.006
568.439 | 0.016
568.725 | 0.011
569.011 | 0.015
569.297 | 0.012
569.583 | 0.016
569.868 | 0.015
570.154 | 0.006
570.439 | 0.008
570.725 | 0.013
571.01 | 0.015
571.296 | 0.018
571.581 | 0.012
571.866 | 0.023
572.151 | 0.022
572.436 | 0.007
572.721 | 0.012
573.006 | 0.013
573.291 | 0.02
573.576 | 0.015
573.861 | 0.018
574.146 | 0.011
574.43 | 0.017
574.715 | 0.024
574.999 | 0.022
575.284 | 0.021
575.568 | 0.008
575.852 | 0.012
576.137 | 0.024
576.421 | 0.032
576.705 | 0.01
576.989 | 0.016
577.273 | 0.003
577.557 | 0.015
577.841 | 0.013
578.124 | 0.023
578.408 | 0.02
578.692 | 0.002
578.975 | 0.018
579.259 | 0.011
579.542 | 0.012
579.826 | 0.008
580.109 | 0.014
580.392 | 0.02
580.676 | 0.007
580.959 | 0.015
581.242 | 0.015
581.525 | 0.01
581.808 | 0.013
582.091 | 0.011
582.373 | 0.008
582.656 | 0.015
582.939 | 0.011
583.221 | 0.005
583.504 | 0.012
583.786 | 0.016
584.069 | 0.008
584.351 | 0.013
584.633 | 0.007
584.916 | 0
585.198 | 0.013
585.48 | 0.009
585.762 | 0
586.044 | 0.004
586.326 | 0.017
586.607 | 0.01
586.889 | 0.026
587.171 | 0.019
587.453 | 0.005
587.734 | 0.013
588.016 | 0.014
588.297 | 0.005
588.578 | 0.02
588.86 | 0.017
589.141 | 0.009
589.422 | 0.017
589.703 | 0.019
589.984 | 0.004
590.265 | 0.01
590.546 | 0.017
590.827 | 0.006
591.108 | 0.008
591.389 | 0.021
591.669 | 0.005
591.95 | 0.009
592.23 | 0.016
592.511 | 0.018
592.791 | 0.018
593.071 | 0.02
593.352 | -0.004
593.632 | 0.016
593.912 | 0.015
594.192 | 0.023
594.472 | 0.015
594.752 | 0.016
595.032 | 0.013
595.312 | 0.003
595.592 | 0.013
595.871 | 0.014
596.151 | 0.012
596.43 | 0.002
596.71 | 0.01
596.989 | 0.025
597.269 | 0.015
597.548 | 0.012
597.827 | 0.026
598.106 | 0.019
598.385 | 0.013
598.665 | 0.001
598.944 | 0.006
599.222 | 0.012
599.501 | 0.007
599.78 | 0.013
600.059 | 0.014
600.337 | 0.007
600.616 | 0.013
600.895 | 0.012
601.173 | 0.03
601.451 | 0.013
601.73 | 0.007
602.008 | 0.016
602.286 | 0.001
602.564 | 0.016
602.842 | 0.008
603.12 | 0.005
603.398 | 0.01
603.676 | 0.011
603.954 | 0.008
604.232 | 0.011
604.51 | 0.004
604.787 | 0.001
605.065 | 0.018
605.342 | 0.008
605.62 | 0.003
605.897 | 0.009
606.174 | 0.018
606.452 | 0.012
606.729 | 0.017
607.006 | 0.02
607.283 | 0.014
607.56 | 0.009
607.837 | 0.003
608.114 | 0.013
608.391 | 0.036
608.667 | -0.002
608.944 | 0.006
609.221 | 0.007
609.497 | 0.007
609.774 | 0.003
610.05 | 0.024
610.326 | 0.016
610.603 | 0.02
610.879 | 0.021
611.155 | 0.017
611.431 | 0.006
611.707 | 0.014
611.983 | -0.001
612.259 | 0.021
612.535 | 0.004
612.811 | 0.011
613.087 | 0.011
613.362 | 0.009
613.638 | 0.008
613.913 | 0.013
614.189 | 0.01
614.464 | 0.015
614.74 | 0.025
615.015 | 0.002
615.29 | 0.017
615.565 | 0.012
615.84 | -0.003
616.115 | 0.018
616.39 | 0.009
616.665 | 0.015
616.94 | 0.011
617.215 | 0.019
617.49 | 0.008
617.764 | 0.017
618.039 | 0.018
618.313 | 0.018
618.588 | 0.023
618.862 | 0.012
619.137 | 0.024
619.411 | 0.011
619.685 | 0.013
619.959 | 0.009
620.233 | 0.027
620.507 | 0.006
620.781 | 0.009
621.055 | 0.008
621.329 | 0.013
621.603 | 0.012
621.876 | 0.02
622.15 | 0.014
622.424 | 0.013
622.697 | 0.01
622.971 | 0.01
623.244 | 0.003
623.517 | 0.007
623.791 | 0.012
624.064 | 0.019
624.337 | 0.023
624.61 | 0.01
624.883 | 0.016
625.156 | 0.006
625.429 | 0.014
625.702 | 0.01
625.974 | 0.023
626.247 | 0.014
626.52 | 0.021
626.792 | 0.011
627.065 | 0.013
627.337 | 0.014
627.61 | 0.015
627.882 | 0.014
628.154 | 0.014
628.427 | 0.023
628.699 | 0.019
628.971 | 0.012
629.243 | 0.006
629.515 | 0.02
629.787 | 0.007
630.058 | 0.022
630.33 | 0.02
630.602 | -0.001
630.874 | 0.021
631.145 | 0.017
631.417 | 0.008
631.688 | 0.023
631.959 | 0.012
632.231 | 0.013
632.502 | 0.021
632.773 | 0.027
633.044 | 0.026
633.316 | 0.025
633.587 | -0.002
633.857 | 0.012
634.128 | 0.021
634.399 | 0.031
634.67 | 0.007
634.941 | 0.014
635.211 | 0.023
635.482 | 0.011
635.752 | 0.021
636.023 | 0.001
636.293 | 0.008
636.564 | 0.027
636.834 | 0.014
637.104 | 0.016
637.374 | 0.016
637.644 | 0.02
637.914 | 0.024
638.184 | 0.016
638.454 | 0.013
638.724 | 0.026
638.994 | 0.012
639.264 | 0.021
639.533 | 0.014
639.803 | 0.008
640.072 | 0.014
640.342 | 0.015
640.611 | 0.026
640.881 | 0.016
641.15 | 0.038
641.419 | 0.009
641.688 | 0.016
641.957 | 0.013
642.227 | 0.02
642.496 | 0.026
642.764 | 0.017
643.033 | 0.024
643.302 | 0.011
643.571 | 0.012
643.84 | 0.01
644.108 | 0.021
644.377 | -0.004
644.645 | 0.011
644.914 | 0.011
645.182 | 0.024
645.45 | 0.024
645.719 | 0.029
645.987 | 0.018
646.255 | 0.015
646.523 | 0.016
646.791 | 0.013
647.059 | 0.029
647.327 | 0.009
647.595 | 0.015
647.862 | 0.006
648.13 | 0.012
648.398 | 0.022
648.665 | 0.013
648.933 | 0.031
649.2 | 0.018
649.468 | 0.02
649.735 | 0.029
650.002 | 0.021
650.269 | 0.015
650.536 | 0.008
650.804 | 0.008
651.071 | 0.029
651.338 | 0.014
651.604 | 0.018
651.871 | 0.009
652.138 | 0.013
652.405 | 0.002
652.671 | 0.018
652.938 | 0.015
653.205 | 0.018
653.471 | 0.005
653.737 | 0.019
654.004 | 0.004
654.27 | 0.016
654.536 | 0.015
654.803 | 0.009
655.069 | 0.018
655.335 | 0.013
655.601 | 0.021
655.867 | 0.014
656.132 | 0.02
656.398 | 0.017
656.664 | 0.018
656.93 | 0.012
657.195 | 0.011
657.461 | 0.017
657.726 | 0.014
657.992 | 0.004
658.257 | 0.016
658.523 | 0.018
658.788 | 0.004
659.053 | 0.013
659.318 | 0.012
659.583 | 0.016
659.848 | 0.008
660.113 | 0.015
660.378 | 0.017
660.643 | 0.016
660.908 | 0.004
661.172 | -0.001
661.437 | 0.007
661.702 | 0.019
661.966 | 0.021
662.231 | 0.017
662.495 | 0.016
662.76 | 0.033
663.024 | 0.012
663.288 | 0.006
663.552 | 0.009
663.816 | 0.012
664.08 | 0.024
664.344 | 0.015
664.608 | 0.017
664.872 | 0.007
665.136 | 0.004
665.4 | 0.02
665.664 | 0.02
665.927 | 0.017
666.191 | 0.013
666.454 | 0.003
666.718 | 0.011
666.981 | 0.001
667.244 | 0.007
667.508 | 0.023
667.771 | 0.014
668.034 | 0.021
668.297 | 0.008
668.56 | 0.015
668.823 | 0
669.086 | 0.021
669.349 | 0.005
669.612 | 0.016
669.875 | 0.008
670.137 | 0.017
670.4 | 0.014
670.662 | 0.011
670.925 | 0.024
671.187 | 0.024
671.45 | -0.005
671.712 | 0.03
671.974 | 0.026
672.237 | 0.005
672.499 | 0.009
672.761 | 0.018
673.023 | 0.012
673.285 | 0.015
673.547 | 0.023
673.808 | 0.009
674.07 | 0.009
674.332 | 0.015
674.594 | 0.008
674.855 | 0.013
675.117 | 0.009
675.378 | 0.024
675.64 | 0.009
675.901 | -0.003
676.162 | 0.017
676.424 | 0.004
676.685 | 0.024
676.946 | 0.022
677.207 | 0.027
677.468 | 0.029
677.729 | 0.006
677.99 | 0.007
678.251 | 0.019
678.512 | 0.014
678.772 | 0.019
679.033 | 0.03
679.294 | 0.009
679.554 | 0.016
679.815 | 0.013
680.075 | 0.012
680.335 | 0.022
680.596 | 0.005
680.856 | 0.02
681.116 | 0.035
681.376 | 0.018
681.636 | 0.01
681.896 | 0.024
682.156 | 0.027
682.416 | 0.015
682.676 | 0.012
682.936 | 0.023
683.196 | 0.017
683.455 | 0.014
683.715 | 0.018
683.974 | 0.017
684.234 | 0.012
684.493 | 0.008
684.753 | 0.021
685.012 | 0.029
685.271 | 0.011
685.53 | 0.013
685.789 | 0.024
686.049 | 0.02
686.308 | 0.008
686.567 | 0.019
686.825 | 0.023
687.084 | 0.016
687.343 | 0.018
687.602 | 0.021
687.86 | 0.026
688.119 | 0.021
688.378 | 0.023
688.636 | 0.017
688.894 | 0.013
689.153 | 0.019
689.411 | 0.022
689.669 | 0.013
689.928 | 0.028
690.186 | 0.004
690.444 | 0.019
690.702 | 0.022
690.96 | 0.027
691.218 | 0.012
691.476 | 0.027
691.733 | 0.021
691.991 | 0.018
692.249 | 0.014
692.506 | 0.013
692.764 | 0.022
693.021 | 0.019
693.279 | 0.025
693.536 | 0.022
693.794 | 0.001
694.051 | 0.009
694.308 | 0.016
694.565 | 0.03
694.822 | 0.031
695.079 | 0.009
695.336 | 0.015
695.593 | 0.008
695.85 | 0.014
696.107 | 0.01
696.364 | 0.015
696.62 | 0.02
696.877 | 0.021
697.133 | 0.013
697.39 | 0.021
697.646 | 0.019
697.903 | 0.01
698.159 | 0.004
698.415 | 0.008
698.672 | 0.023
698.928 | 0.016
699.184 | 0.021
699.44 | 0.016
699.696 | 0.024
699.952 | 0.003
700.208 | 0.016
700.464 | 0.012
700.719 | 0.032
700.975 | 0.022
701.231 | 0.019
701.486 | 0.027
701.742 | 0.014
701.997 | 0.018
702.253 | 0.022
702.508 | 0.005
702.763 | 0.019
703.019 | 0.026
703.274 | 0.023
703.529 | 0.021
703.784 | 0.014
704.039 | 0.037
704.294 | 0.014
704.549 | 0.01
704.804 | 0.009
705.058 | 0.018
705.313 | 0.005
705.568 | 0.035
705.822 | 0.023
706.077 | 0.014
706.331 | 0.005
706.586 | 0.012
706.84 | 0.027
707.095 | 0.019
707.349 | -0.01
707.603 | 0.008
707.857 | 0.012
708.111 | 0.004
708.365 | 0.014
708.619 | 0.012
708.873 | 0.028
709.127 | 0.009
709.381 | 0.007
709.635 | 0.021
709.888 | 0.028
710.142 | 0.016
710.396 | 0.015
710.649 | 0.02
710.903 | 0.008
711.156 | 0.017
711.409 | 0.024
711.663 | 0.026
711.916 | 0.015
712.169 | 0.014
712.422 | 0.02
712.675 | 0.018
712.928 | 0.036
713.181 | 0.022
713.434 | 0.018
713.687 | 0.031
713.94 | 0.013
714.192 | 0.017
714.445 | -0.007
714.698 | 0.024
714.95 | 0.032
715.203 | 0.044
715.455 | 0.017
715.708 | 0.021
715.96 | 0.036
716.212 | 0.016
716.464 | 0.039
716.717 | 0.024
716.969 | 0.027
717.221 | 0.028
717.473 | -0.001
717.725 | 0.017
717.977 | 0.024
718.228 | 0.032
718.48 | 0.015
718.732 | 0.018
718.984 | 0.031
719.235 | 0.035
719.487 | 0.027
719.738 | 0.013
719.99 | 0.034
720.241 | 0.024
720.492 | 0.028
720.744 | 0.015
720.995 | 0.008
721.246 | 0.016
721.497 | 0.036
721.748 | 0.022
721.999 | 0.008
722.25 | 0.041
722.501 | 0.03
722.752 | 0.027
723.002 | 0.011
723.253 | 0.012
723.504 | 0.035
723.754 | 0.02
724.005 | 0.033
724.255 | 0.026
724.506 | 0.033
724.756 | 0.03
725.006 | 0.015
725.257 | 0.013
725.507 | 0.024
725.757 | 0.018
726.007 | 0.007
726.257 | 0.028
726.507 | 0.009
726.757 | 0.048
727.007 | 0.035
727.257 | 0.009
727.506 | 0.033
727.756 | 0.016
728.006 | 0.033
728.255 | 0.032
728.505 | 0.005
728.754 | 0.025
729.004 | 0.035
729.253 | 0.049
729.502 | 0.054
729.751 | 0.029
730.001 | 0.006
730.25 | 0.014
730.499 | 0.037
730.748 | 0.028
730.997 | 0.022
731.246 | 0.037
731.495 | 0.017
731.743 | 0.041
731.992 | 0.027
732.241 | 0.007
732.489 | 0.024
732.738 | 0.049
732.987 | 0.019
733.235 | 0.022
733.483 | 0.035
733.732 | 0.033
733.98 | 0.007
734.228 | 0.01
734.477 | 0.019
734.725 | 0.036
734.973 | 0.024
735.221 | 0.037
735.469 | 0.028
735.717 | 0.035
735.965 | 0.016
736.212 | 0.02
736.46 | 0.037
736.708 | 0.023
736.955 | 0.052
737.203 | 0.007
737.451 | 0.043
737.698 | 0.03
737.945 | 0.024
738.193 | 0.034
738.44 | 0.024
738.687 | 0.032
738.935 | 0.031
739.182 | 0.012
739.429 | 0.037
739.676 | 0.016
739.923 | 0.028
740.17 | 0.036
740.417 | 0.022
740.663 | 0.01
740.91 | 0.033
741.157 | 0.041
741.403 | 0.023
741.65 | -0.001
741.897 | 0.021
742.143 | 0.014
742.389 | 0.023
742.636 | 0.012
742.882 | 0.022
743.128 | 0.024
743.375 | 0.003
743.621 | 0.045
743.867 | 0.029
744.113 | 0.024
744.359 | 0.017
744.605 | 0.034
744.851 | 0.052
745.097 | 0.047
745.342 | 0.018
745.588 | 0.024
745.834 | 0.022
746.079 | 0.046
746.325 | 0.023
746.57 | 0.032
746.816 | 0.009
747.061 | 0.032
747.307 | 0.03
747.552 | 0.024
747.797 | 0.042
748.042 | 0.024
748.287 | 0.028
748.532 | 0.019
748.777 | 0.026
749.022 | 0.034
749.267 | 0.039
749.512 | 0.025
749.757 | 0.017
750.002 | 0.024
750.246 | 0.03
750.491 | 0.035
750.736 | 0.02
750.98 | 0.034
751.225 | 0.023
751.469 | 0.031
751.713 | 0.027
751.958 | 0.033
752.202 | 0.013
752.446 | 0.023
752.69 | 0.016
752.934 | 0.01
753.178 | 0.036
753.422 | 0.021
753.666 | 0.045
753.91 | 0.012
754.154 | 0.034
754.398 | 0.061
754.641 | 0.011
754.885 | 0.008
755.129 | 0.032
755.372 | 0.008
755.616 | 0.033
755.859 | 0.034
756.103 | 0.062
756.346 | 0.002
756.589 | 0.034
756.832 | 0.026
757.076 | 0.032
757.319 | 0.037
757.562 | 0.036
757.805 | 0.04
758.048 | 0.036
758.291 | 0.026
758.533 | 0.029
758.776 | 0.014
759.019 | 0.034
759.262 | 0.036
759.504 | 0.034
759.747 | 0.029
759.989 | 0.034
760.232 | 0.04
760.474 | 0.015
760.717 | 0.058
760.959 | 0.024
761.201 | 0.014
761.444 | 0.027
761.686 | 0.047
761.928 | 0.022
762.17 | 0.033
762.412 | 0.042
762.654 | 0.04
762.896 | 0.028
763.137 | 0.029
763.379 | 0.021
763.621 | 0.041
763.863 | 0.031
764.104 | 0.026
764.346 | 0.043
764.587 | 0.024
764.829 | 0.027
765.07 | 0.016
765.312 | 0.045
765.553 | 0.051
765.794 | 0.041
766.035 | 0.036
766.276 | 0.047
766.518 | 0.036
766.759 | 0.059
767 | 0.031
767.241 | 0.04
767.481 | 0.028
767.722 | 0.029
767.963 | 0.036
768.204 | 0.045
768.444 | 0.055
768.685 | 0.036
768.926 | 0.009
769.166 | 0.026
769.407 | 0.023
769.647 | 0.056
769.887 | 0.015
770.128 | 0.041
770.368 | 0.038
770.608 | 0.049
770.848 | 0.041
771.088 | 0.029
771.328 | 0.039
771.568 | 0.028
771.808 | 0.045
772.048 | 0.024
772.288 | 0.039
772.528 | 0.013
772.767 | 0.027
773.007 | 0.032
773.247 | 0.035
773.486 | 0.027
773.726 | 0.016
773.965 | 0.043
774.205 | 0.032
774.444 | 0.044
774.683 | 0.04
774.922 | 0.052
775.162 | 0.042
775.401 | 0.054
775.64 | 0.032
775.879 | 0.036
776.118 | 0.031
776.357 | 0.028
776.596 | 0.031
776.834 | 0.057
777.073 | 0.035
777.312 | 0.04
777.551 | 0.022
777.789 | 0.016
778.028 | 0.041
778.266 | 0.04
778.505 | 0.025
778.743 | 0.026
778.982 | 0.006
779.22 | 0.031
779.458 | 0.039
779.696 | 0.028
779.934 | 0.047
780.173 | 0.039
780.411 | 0.025
780.649 | 0.021
780.887 | 0.034
781.124 | 0.047
781.362 | 0.017
781.6 | 0.043
781.838 | 0.029
782.075 | 0.036
782.313 | 0.024
782.551 | 0.005
782.788 | 0.022
783.026 | 0.031
783.263 | 0.027
783.5 | 0.031
783.738 | 0.04
783.975 | 0.006
784.212 | -0.015
784.449 | 0.032
784.686 | 0.024
784.924 | 0.032
785.161 | 0.014
785.397 | 0.03
785.634 | 0.039
785.871 | 0.053
786.108 | 0.014
786.345 | 0.063
786.581 | 0.022
786.818 | 0.04
787.055 | 0.039
787.291 | 0.049
787.528 | 0.053
787.764 | 0.034
788.001 | 0.016
788.237 | 0.055
788.473 | 0.015
788.709 | 0.022
788.946 | 0.032
789.182 | 0.016
789.418 | 0.006
789.654 | 0.046
789.89 | 0.029
790.126 | 0.054
790.362 | 0.052
790.597 | 0.025
790.833 | 0.03
791.069 | 0.017
791.305 | 0.021
791.54 | 0.035
791.776 | 0.056
792.011 | 0.044
792.247 | 0.015
792.482 | 0.032
792.717 | 0.046
792.953 | 0.031
793.188 | 0.041
793.423 | 0.048
793.658 | 0.049
793.894 | 0.065
794.129 | 0.034
794.364 | 0.04
794.599 | 0.041
794.833 | 0.022
795.068 | 0.038
795.303 | 0.034
795.538 | 0.005
795.772 | 0.035
796.007 | 0.04
796.242 | 0.065
796.476 | 0.047
796.711 | 0.024
796.945 | 0.034
797.18 | 0.016
797.414 | 0.009
797.648 | 0.052
797.882 | 0.025
798.117 | 0.02
798.351 | 0.018
798.585 | 0.068
798.819 | 0.012
799.053 | 0.047
799.287 | 0.035
799.521 | 0.052
799.755 | 0.042
799.988 | 0.044
800.222 | 0.036
800.456 | 0.03
800.689 | 0.002
800.923 | 0.028
801.156 | 0.049
801.39 | 0.02
801.623 | 0.045
801.857 | 0.027
802.09 | 0.049
802.323 | 0.041
802.557 | 0.037
802.79 | 0.03
803.023 | 0.023
803.256 | 0.027
803.489 | 0.041
803.722 | 0.034
803.955 | 0.023
804.188 | 0.062
804.421 | 0.021
804.653 | 0.037
804.886 | 0.035
805.119 | 0.014
805.351 | 0.022
805.584 | 0.053
805.816 | 0.014
806.049 | 0.036
806.281 | 0.047
806.514 | 0.046
806.746 | 0.055
806.978 | 0.03
807.211 | 0.01
807.443 | 0.027
807.675 | 0.031
254
AbsoluteIrradiance:211
195.562 | 0
195.934 | 0
196.307 | 0
196.679 | 0
197.051 | 0
197.423 | 0
197.795 | 0
198.167 | 0
198.539 | 0
198.911 | 0
199.283 | 0
199.654 | 0
200.026 | 0
200.397 | 0
200.769 | 0
201.14 | 0
201.511 | 0
201.882 | 0
202.253 | 0
202.624 | 0
202.995 | 0
203.366 | 0
203.736 | 0
204.107 | 0
204.477 | 0
204.848 | 3.052
205.218 | 6.453
205.588 | 2.319
205.959 | 1.759
206.329 | 2.398
206.699 | 1.499
207.068 | 2.342
207.438 | 2.513
207.808 | 1.534
208.178 | 2.366
208.547 | 1.16
208.917 | 1.364
209.286 | 2.651
209.655 | 2.153
210.025 | 1.28
210.394 | 1.625
210.763 | 1.935
211.132 | 0.939
211.501 | 0.903
211.869 | 0.862
212.238 | 1.381
212.607 | 1.726
212.975 | 0.712
213.344 | 0.605
213.712 | 0.61
214.08 | 0.813
214.448 | 0.967
214.817 | 0.614
215.185 | 1.235
215.553 | 0.899
215.92 | 0.94
216.288 | 0.83
216.656 | 0.97
217.023 | 0.677
217.391 | 0.714
217.758 | 0.983
218.126 | 0.77
218.493 | 0.615
218.86 | 0.666
219.227 | 0.406
219.594 | 0.463
219.961 | 0.833
220.328 | 0.429
220.695 | 0.595
221.061 | 0.754
221.428 | 0.647
221.795 | 0.73
222.161 | 0.884
222.527 | 0.324
222.893 | 0.808
223.26 | 0.329
223.626 | 0.755
223.992 | 0.522
224.358 | 0.396
224.723 | 0.616
225.089 | 0.197
225.455 | 0.469
225.82 | 0.354
226.186 | 0.268
226.551 | 0.131
226.917 | 0.248
227.282 | 0.352
227.647 | 0.469
228.012 | 0.356
228.377 | 0.293
228.742 | 0.318
229.107 | 0.211
229.471 | 0.287
229.836 | 0.198
230.201 | 0.266
230.565 | 0.305
230.929 | 0.225
231.294 | 0.159
231.658 | 0.128
232.022 | 0.168
232.386 | 0.117
232.75 | -0.053
233.114 | 0.122
233.478 | 0.137
233.842 | 0.179
234.205 | 0.18
234.569 | 0.127
234.932 | 0.096
235.296 | 0.075
235.659 | 0.135
236.022 | 0.103
236.385 | 0.079
236.748 | 0.134
237.111 | 0.067
237.474 | 0.172
237.837 | 0.165
238.2 | 0.124
238.562 | 0.105
238.925 | 0.163
239.287 | 0.133
239.65 | 0.082
240.012 | 0.151
240.374 | 0.124
240.736 | 0.15
241.098 | 0.122
241.46 | 0.125
241.822 | 0.156
242.184 | 0.071
242.546 | 0.138
242.907 | 0.101
243.269 | 0.087
243.63 | 0.088
243.992 | 0.152
244.353 | 0.142
244.714 | 0.132
245.075 | 0.156
245.436 | 0.155
245.797 | 0.138
246.158 | 0.113
246.519 | 0.178
246.879 | 0.183
247.24 | 0.17
247.601 | 0.218
247.961 | 0.246
248.321 | 0.21
248.682 | 0.248
249.042 | 0.224
249.402 | 0.266
249.762 | 0.343
250.122 | 0.352
250.482 | 0.459
250.842 | 0.502
251.201 | 0.824
251.561 | 2.84
251.921 | 8.632
252.28 | 11.274
252.639 | 13.195
252.999 | 18.092
253.358 | 33.694
253.717 | 61.873
254.076 | 28.301
254.435 | 3.729
254.794 | 1.215
255.153 | 0.867
255.511 | 0.723
255.87 | 0.623
256.228 | 0.538
256.587 | 0.474
256.945 | 0.415
257.304 | 0.366
257.662 | 0.374
258.02 | 0.35
258.378 | 0.35
258.736 | 0.288
259.094 | 0.36
259.452 | 0.333
259.809 | 0.283
260.167 | 0.282
260.524 | 0.255
260.882 | 0.235
261.239 | 0.279
261.597 | 0.248
261.954 | 0.239
262.311 | 0.258
262.668 | 0.27
263.025 | 0.248
263.382 | 0.257
263.739 | 0.261
264.095 | 0.304
264.452 | 0.241
264.809 | 0.263
265.165 | 0.269
265.521 | 0.273
265.878 | 0.17
266.234 | 0.154
266.59 | 0.117
266.946 | 0.195
267.302 | 0.138
267.658 | 0.153
268.014 | 0.122
268.37 | 0.14
268.725 | 0.12
269.081 | 0.125
269.436 | 0.131
269.792 | 0.104
270.147 | 0.122
270.502 | 0.151
270.857 | 0.143
271.212 | 0.08
271.567 | 0.04
271.922 | 0.101
272.277 | 0.071
272.632 | 0.077
272.987 | 0.087
273.341 | 0.077
273.696 | 0.08
274.05 | 0.045
274.404 | 0.071
274.759 | 0.092
275.113 | 0.128
275.467 | 0.081
275.821 | 0.046
276.175 | 0.065
276.529 | 0.09
276.882 | 0.113
277.236 | 0.076
277.59 | 0.082
277.943 | 0.032
278.297 | 0.107
278.65 | 0.045
279.003 | 0.045
279.356 | 0.058
279.709 | 0.077
280.062 | 0.073
280.415 | 0.086
280.768 | 0.068
281.121 | 0.088
281.474 | 0.075
281.826 | 0.062
282.179 | 0.051
282.531 | 0.043
282.884 | 0.063
283.236 | 0.057
283.588 | 0.083
283.94 | 0.039
284.292 | 0.059
284.644 | 0.042
284.996 | 0.051
285.348 | 0.05
285.7 | 0.078
286.051 | 0.056
286.403 | 0.055
286.754 | 0.05
287.106 | 0.078
287.457 | 0.069
287.808 | 0.056
288.159 | 0.069
288.51 | 0.082
288.861 | 0.085
289.212 | 0.083
289.563 | 0.07
289.914 | 0.067
290.264 | 0.037
290.615 | 0.053
290.966 | 0.07
291.316 | 0.033
291.666 | 0.064
292.017 | 0.052
292.367 | 0.037
292.717 | 0.035
293.067 | 0.041
293.417 | 0.045
293.767 | 0.038
294.116 | 0.065
294.466 | 0.037
294.816 | 0.062
295.165 | 0.049
295.515 | 0.065
295.864 | 0.106
296.213 | 0.154
296.562 | 0.236
296.911 | 0.165
297.261 | 0.051
297.609 | 0.047
297.958 | 0.014
298.307 | 0.039
298.656 | 0.033
299.004 | 0.028
299.353 | 0.004
299.701 | 0.031
300.05 | 0.021
300.398 | 0.039
300.746 | 0.047
301.095 | 0.063
301.443 | 0.081
301.791 | 0.133
302.138 | 0.161
302.486 | 0.062
302.834 | 0.028
303.182 | 0.016
303.529 | 0.029
303.877 | 0.031
304.224 | 0.034
304.572 | 0.036
304.919 | 0.011
305.266 | 0.039
305.613 | 0.023
305.96 | 0.04
306.307 | 0.02
306.654 | 0.042
307.001 | 0.039
307.347 | 0.027
307.694 | 0.018
308.04 | 0.025
308.387 | 0.034
308.733 | 0.027
309.08 | 0.025
309.426 | 0.045
309.772 | 0.032
310.118 | 0.036
310.464 | 0.024
310.81 | 0.038
311.156 | 0.052
311.501 | 0.137
311.847 | 0.281
312.193 | 0.598
312.538 | 0.799
312.883 | 0.796
313.229 | 0.521
313.574 | 0.107
313.919 | 0.04
314.264 | 0.034
314.609 | 0.038
314.954 | 0.026
315.299 | 0.024
315.644 | 0.017
315.988 | 0.03
316.333 | 0.034
316.678 | 0.028
317.022 | 0.023
317.366 | 0.023
317.711 | 0.029
318.055 | 0.02
318.399 | 0.018
318.743 | 0.03
319.087 | 0.024
319.431 | 0.033
319.775 | 0.02
320.118 | 0.038
320.462 | 0.024
320.806 | 0.021
321.149 | 0.032
321.493 | 0.027
321.836 | 0.018
322.179 | 0.009
322.522 | 0.025
322.865 | 0.025
323.208 | 0.024
323.551 | 0.033
323.894 | 0.018
324.237 | 0.013
324.58 | 0.017
324.922 | 0.025
325.265 | 0.029
325.607 | 0.029
325.95 | 0.03
326.292 | 0.009
326.634 | 0.022
326.976 | 0.016
327.318 | 0.019
327.66 | 0.028
328.002 | 0.008
328.344 | 0.024
328.686 | 0.024
329.027 | 0.007
329.369 | 0.012
329.711 | 0.022
330.052 | 0.021
330.393 | 0.012
330.735 | 0.011
331.076 | 0.023
331.417 | 0.014
331.758 | 0.023
332.099 | 0.007
332.44 | 0.025
332.78 | 0.013
333.121 | 0.032
333.462 | 0.048
333.802 | 0.086
334.143 | 0.07
334.483 | 0.028
334.823 | 0.01
335.164 | 0.022
335.504 | 0.011
335.844 | 0.024
336.184 | 0.01
336.524 | 0.014
336.864 | 0.013
337.203 | 0.015
337.543 | 0.025
337.883 | 0.018
338.222 | 0.029
338.562 | 0.012
338.901 | 0.02
339.24 | 0.009
339.58 | 0.018
339.919 | 0.014
340.258 | 0.023
340.597 | 0.019
340.936 | 0.012
341.274 | 0.017
341.613 | 0.025
341.952 | 0.018
342.29 | 0.018
342.629 | 0.026
342.967 | 0.01
343.306 | 0.018
343.644 | 0.019
343.982 | 0.014
344.32 | 0.023
344.658 | 0.012
344.996 | 0.016
345.334 | 0.015
345.672 | 0.027
346.01 | 0.007
346.347 | 0.032
346.685 | 0.019
347.022 | 0.025
347.36 | 0.025
347.697 | 0.017
348.034 | 0.019
348.371 | 0.017
348.708 | 0.015
349.045 | 0.022
349.382 | 0.023
349.719 | 0.011
350.056 | 0.018
350.393 | 0.013
350.729 | 0.025
351.066 | 0.015
351.402 | 0.018
351.739 | 0.028
352.075 | 0.024
352.411 | 0.019
352.747 | 0.013
353.083 | 0.018
353.419 | 0.017
353.755 | 0.019
354.091 | 0.02
354.427 | 0.016
354.763 | 0.02
355.098 | 0.021
355.434 | 0.009
355.769 | 0.032
356.104 | 0.008
356.44 | 0.012
356.775 | 0.025
357.11 | 0.021
357.445 | 0.017
357.78 | 0.014
358.115 | 0.019
358.45 | 0.022
358.785 | 0.026
359.119 | 0.013
359.454 | 0.016
359.788 | 0.021
360.123 | 0.021
360.457 | 0.013
360.791 | 0.028
361.126 | 0.011
361.46 | 0.016
361.794 | 0.025
362.128 | 0.024
362.462 | 0.037
362.795 | 0.03
363.129 | 0.025
363.463 | 0.016
363.796 | 0.053
364.13 | 0.187
364.463 | 0.524
364.797 | 0.746
365.13 | 0.597
365.463 | 0.256
365.796 | 0.165
366.129 | 0.147
366.462 | 0.087
366.795 | 0.03
367.128 | 0.031
367.461 | 0.018
367.793 | 0.029
368.126 | 0.023
368.458 | 0.027
368.791 | 0.025
369.123 | 0.024
369.455 | 0.023
369.787 | 0.019
370.12 | 0.019
370.452 | 0.013
370.784 | 0.015
371.115 | 0.021
371.447 | 0.011
371.779 | 0.021
372.111 | 0.02
372.442 | 0.019
372.774 | 0.034
373.105 | 0.023
373.436 | 0.02
373.768 | 0.019
374.099 | 0.007
374.43 | 0.028
374.761 | 0.016
375.092 | 0.023
375.423 | 0.019
375.754 | 0.019
376.084 | 0.015
376.415 | 0.016
376.745 | 0.015
377.076 | 0.022
377.406 | 0.021
377.737 | 0.037
378.067 | 0.024
378.397 | 0.015
378.727 | 0.017
379.057 | 0.015
379.387 | 0.014
379.717 | 0.018
380.047 | 0.02
380.377 | 0.021
380.706 | 0.023
381.036 | 0.017
381.365 | 0.032
381.695 | 0.016
382.024 | 0.012
382.353 | 0.007
382.682 | 0.019
383.012 | 0.016
383.341 | 0.021
383.67 | 0.021
383.998 | 0.012
384.327 | 0.025
384.656 | 0.012
384.985 | 0.01
385.313 | 0.01
385.642 | 0.014
385.97 | 0.014
386.298 | 0.014
386.627 | 0.016
386.955 | 0.024
387.283 | 0.014
387.611 | 0.012
387.939 | 0.014
388.267 | 0.017
388.595 | 0.026
388.922 | 0.004
389.25 | 0.029
389.578 | 0.017
389.905 | 0.021
390.233 | 0.017
390.56 | 0.025
390.887 | 0.029
391.214 | 0.021
391.542 | 0.013
391.869 | 0.017
392.196 | 0.009
392.522 | 0.013
392.849 | 0.014
393.176 | 0.011
393.503 | 0.016
393.829 | 0.014
394.156 | 0.014
394.482 | 0.018
394.809 | 0.018
395.135 | 0.021
395.461 | 0.013
395.787 | 0.02
396.113 | 0.008
396.439 | 0.017
396.765 | 0.014
397.091 | 0.016
397.417 | 0.013
397.742 | 0.016
398.068 | 0.021
398.394 | 0.018
398.719 | 0.017
399.044 | 0.013
399.37 | 0.023
399.695 | 0.016
400.02 | 0.021
400.345 | 0.021
400.67 | 0.023
400.995 | 0.03
401.32 | 0.019
401.645 | 0.018
401.969 | 0.016
402.294 | 0.015
402.618 | 0.031
402.943 | 0.016
403.267 | 0.029
403.592 | 0.111
403.916 | 0.38
404.24 | 0.791
404.564 | 1.038
404.888 | 0.615
405.212 | 0.116
405.536 | 0.019
405.86 | 0.028
406.183 | 0.018
406.507 | 0.025
406.831 | 0.035
407.154 | 0.056
407.478 | 0.091
407.801 | 0.094
408.124 | 0.042
408.447 | 0.018
408.77 | 0.022
409.093 | 0.021
409.416 | 0.015
409.739 | 0.022
410.062 | 0.012
410.385 | 0.023
410.708 | 0.021
411.03 | 0.023
411.353 | 0.013
411.675 | 0.021
411.997 | 0.018
412.32 | 0.014
412.642 | 0.022
412.964 | 0.018
413.286 | 0.017
413.608 | 0.02
413.93 | 0.019
414.252 | 0.012
414.574 | 0.011
414.895 | 0.013
415.217 | 0.023
415.538 | 0.013
415.86 | 0.017
416.181 | 0.024
416.503 | 0.015
416.824 | 0.018
417.145 | 0.011
417.466 | 0.02
417.787 | 0.024
418.108 | 0.008
418.429 | 0.016
418.75 | 0.014
419.07 | 0.014
419.391 | 0.018
419.712 | 0.018
420.032 | 0.015
420.353 | 0.016
420.673 | 0.014
420.993 | 0.016
421.313 | 0.022
421.633 | 0.017
421.953 | 0.009
422.273 | 0.016
422.593 | 0.018
422.913 | 0.015
423.233 | 0.012
423.553 | 0.012
423.872 | 0.01
424.192 | 0.013
424.511 | 0.007
424.83 | 0.015
425.15 | 0.02
425.469 | 0.016
425.788 | 0.02
426.107 | 0.017
426.426 | 0.021
426.745 | 0.024
427.064 | 0.02
427.383 | 0.022
427.702 | 0.007
428.02 | 0.021
428.339 | 0.015
428.657 | 0.017
428.976 | 0.019
429.294 | 0.011
429.612 | 0.011
429.93 | 0.016
430.248 | 0.022
430.567 | 0.016
430.884 | 0.016
431.202 | 0.011
431.52 | 0.012
431.838 | 0.013
432.156 | 0.018
432.473 | 0.018
432.791 | 0.024
433.108 | 0.016
433.426 | 0.033
433.743 | 0.035
434.06 | 0.045
434.377 | 0.073
434.694 | 0.199
435.011 | 0.646
435.328 | 1.259
435.645 | 2.167
435.962 | 2.374
436.279 | 0.76
436.595 | 0.085
436.912 | 0.029
437.228 | 0.024
437.545 | 0.017
437.861 | 0.018
438.177 | 0.019
438.493 | 0.017
438.81 | 0.025
439.126 | 0.02
439.442 | 0.027
439.757 | 0.013
440.073 | 0.02
440.389 | 0.02
440.705 | 0.033
441.02 | 0.02
441.336 | 0.017
441.651 | 0.024
441.967 | 0.017
442.282 | 0.011
442.597 | 0.02
442.912 | 0.028
443.227 | 0.022
443.543 | 0.024
443.857 | 0.028
444.172 | 0.021
444.487 | 0.024
444.802 | 0.032
445.116 | 0.028
445.431 | 0.013
445.746 | 0.014
446.06 | 0.017
446.374 | 0.017
446.689 | 0.021
447.003 | 0.016
447.317 | 0.015
447.631 | 0.008
447.945 | 0.016
448.259 | 0.021
448.573 | 0.018
448.887 | 0.011
449.2 | 0.022
449.514 | 0.017
449.827 | 0.015
450.141 | 0.015
450.454 | 0.016
450.768 | 0.014
451.081 | 0.023
451.394 | 0.012
451.707 | 0.011
452.02 | 0.02
452.333 | 0.016
452.646 | 0.012
452.959 | 0.011
453.272 | 0.011
453.584 | 0.013
453.897 | 0.014
454.21 | 0.016
454.522 | 0.011
454.834 | 0.008
455.147 | 0.012
455.459 | 0.003
455.771 | 0.001
456.083 | 0.008
456.395 | 0.013
456.707 | 0.013
457.019 | 0.016
457.331 | 0.014
457.643 | 0.01
457.954 | 0.014
458.266 | 0.01
458.577 | 0.013
458.889 | 0.01
459.2 | 0.013
459.512 | 0.01
459.823 | 0.016
460.134 | 0.016
460.445 | 0.018
460.756 | 0.014
461.067 | 0.014
461.378 | 0.009
461.689 | 0.008
461.999 | 0.022
462.31 | 0.018
462.621 | 0.011
462.931 | 0.013
463.242 | 0.015
463.552 | 0.02
463.862 | 0.008
464.172 | 0.013
464.483 | 0.021
464.793 | 0.011
465.103 | 0.014
465.413 | 0.017
465.722 | 0.012
466.032 | 0.019
466.342 | 0.012
466.652 | 0.014
466.961 | 0.015
467.271 | 0.015
467.58 | 0.015
467.889 | 0.015
468.199 | 0.016
468.508 | 0.012
468.817 | 0.018
469.126 | 0.016
469.435 | 0.021
469.744 | 0.014
470.053 | 0.012
470.362 | 0.009
470.67 | 0.01
470.979 | 0.015
471.287 | 0.008
471.596 | 0.012
471.904 | 0.02
472.213 | 0.01
472.521 | 0.018
472.829 | 0.015
473.137 | 0.014
473.445 | 0.01
473.753 | 0.023
474.061 | 0.007
474.369 | 0.016
474.677 | 0.021
474.985 | 0.007
475.292 | 0.007
475.6 | 0.019
475.907 | 0.015
476.215 | 0.01
476.522 | 0.022
476.829 | 0.014
477.137 | 0.017
477.444 | 0.008
477.751 | 0.019
478.058 | 0.006
478.365 | 0.013
478.671 | 0.019
478.978 | 0.016
479.285 | 0.01
479.592 | 0.011
479.898 | 0.008
480.205 | 0.014
480.511 | 0.002
480.817 | 0.019
481.124 | 0.011
481.43 | 0.017
481.736 | 0.023
482.042 | 0.014
482.348 | 0.011
482.654 | 0.021
482.96 | 0.019
483.266 | 0.015
483.571 | 0.019
483.877 | 0.007
484.182 | 0.006
484.488 | 0.016
484.793 | 0.005
485.099 | 0.013
485.404 | 0.014
485.709 | 0.017
486.014 | 0.012
486.319 | 0.022
486.624 | 0.01
486.929 | 0.008
487.234 | 0.002
487.539 | 0.016
487.844 | 0.018
488.148 | 0.011
488.453 | 0.021
488.757 | 0.013
489.062 | 0.011
489.366 | 0.015
489.67 | 0.014
489.975 | 0.011
490.279 | 0.012
490.583 | 0.015
490.887 | 0.014
491.191 | 0.015
491.495 | 0.023
491.798 | 0.027
492.102 | 0.023
492.406 | 0.011
492.709 | 0.015
493.013 | 0.02
493.316 | 0.014
493.62 | 0.02
493.923 | 0.013
494.226 | 0.022
494.529 | 0.012
494.832 | 0.01
495.135 | 0.012
495.438 | 0.021
495.741 | 0.027
496.044 | 0.012
496.347 | 0.01
496.649 | 0.019
496.952 | 0.012
497.255 | 0.012
497.557 | 0.012
497.859 | 0.015
498.162 | 0.003
498.464 | 0.024
498.766 | 0.014
499.068 | 0.011
499.37 | 0.013
499.672 | 0.019
499.974 | 0.012
500.276 | 0.015
500.578 | 0.018
500.879 | 0.006
501.181 | 0.022
501.483 | 0.01
501.784 | 0.018
502.085 | 0.006
502.387 | 0.007
502.688 | 0.01
502.989 | 0.02
503.29 | 0.007
503.591 | 0.009
503.892 | 0.008
504.193 | 0.02
504.494 | 0.013
504.795 | 0.009
505.096 | 0.02
505.396 | 0.01
505.697 | 0.002
505.997 | 0.011
506.298 | 0.017
506.598 | 0.012
506.898 | 0.02
507.199 | 0.012
507.499 | 0.013
507.799 | 0.01
508.099 | 0.011
508.399 | 0.017
508.699 | 0.014
508.998 | 0.01
509.298 | 0.022
509.598 | 0.011
509.897 | 0.015
510.197 | 0.008
510.496 | 0.012
510.796 | 0.007
511.095 | 0.013
511.394 | 0.011
511.693 | 0.019
511.992 | 0.016
512.291 | 0.016
512.59 | 0.017
512.889 | 0.019
513.188 | 0.018
513.487 | 0.014
513.786 | 0.016
514.084 | 0.015
514.383 | 0.018
514.681 | 0.021
514.98 | 0.016
515.278 | 0.021
515.576 | 0.02
515.874 | 0.014
516.173 | 0.012
516.471 | 0.016
516.769 | 0.013
517.067 | 0.012
517.364 | 0.008
517.662 | 0.007
517.96 | 0.007
518.257 | 0.011
518.555 | 0.014
518.853 | 0.007
519.15 | 0.009
519.447 | 0.01
519.745 | 0.014
520.042 | 0.018
520.339 | 0.009
520.636 | 0.015
520.933 | 0.022
521.23 | 0.011
521.527 | 0.014
521.824 | 0.017
522.121 | 0.008
522.417 | 0.011
522.714 | 0.005
523.01 | 0.013
523.307 | 0.009
523.603 | 0.018
523.9 | 0.017
524.196 | 0.019
524.492 | 0.008
524.788 | 0.022
525.084 | 0.015
525.38 | 0.02
525.676 | 0.009
525.972 | 0.012
526.268 | 0.014
526.564 | 0.018
526.859 | 0.014
527.155 | 0.013
527.45 | 0.021
527.746 | 0.015
528.041 | 0.023
528.336 | 0.018
528.632 | 0.009
528.927 | 0.008
529.222 | 0.014
529.517 | 0.017
529.812 | 0.013
530.107 | 0.01
530.402 | 0.015
530.696 | 0.018
530.991 | 0.022
531.286 | 0.014
531.58 | 0.015
531.875 | 0.01
532.169 | 0.014
532.463 | 0.011
532.758 | 0.01
533.052 | 0.015
533.346 | 0.008
533.64 | 0.014
533.934 | 0.018
534.228 | 0.018
534.522 | 0.019
534.816 | 0.019
535.11 | 0.008
535.403 | 0.022
535.697 | 0.023
535.99 | 0.017
536.284 | 0.014
536.577 | 0.027
536.871 | 0.014
537.164 | 0.014
537.457 | 0.021
537.75 | 0.014
538.043 | 0.017
538.336 | 0.016
538.629 | 0.015
538.922 | 0.017
539.215 | 0.003
539.507 | 0.015
539.8 | 0.006
540.093 | 0.025
540.385 | 0.01
540.678 | 0.015
540.97 | 0.004
541.262 | 0.016
541.554 | 0.013
541.847 | 0.018
542.139 | 0.02
542.431 | 0.02
542.723 | 0.014
543.015 | 0.019
543.307 | 0.023
543.598 | 0.023
543.89 | 0.022
544.182 | 0.018
544.473 | 0.026
544.765 | 0.042
545.056 | 0.082
545.347 | 0.336
545.639 | 0.708
545.93 | 0.959
546.221 | 1.487
546.512 | 0.773
546.803 | 0.122
547.094 | 0.034
547.385 | 0.027
547.676 | 0.026
547.967 | 0.034
548.257 | 0.018
548.548 | 0.024
548.839 | 0.007
549.129 | 0.014
549.419 | 0.023
549.71 | 0.019
550 | 0.026
550.29 | 0.018
550.58 | 0.014
550.871 | 0.016
551.161 | 0.014
551.45 | 0.008
551.74 | 0.014
552.03 | 0.013
552.32 | 0.019
552.61 | 0.016
552.899 | 0.027
553.189 | 0.015
553.478 | 0.017
553.768 | 0.012
554.057 | 0.027
554.346 | 0.025
554.635 | 0.019
554.925 | 0.019
555.214 | 0.023
555.503 | 0.014
555.792 | 0.021
556.081 | 0.011
556.369 | 0.01
556.658 | 0.012
556.947 | 0.019
557.235 | 0.01
557.524 | 0.03
557.812 | 0.016
558.101 | 0.015
558.389 | 0.021
558.677 | 0.013
558.966 | 0.022
559.254 | 0.026
559.542 | 0.019
559.83 | 0.016
560.118 | 0.017
560.406 | 0.01
560.693 | 0.031
560.981 | 0.024
561.269 | 0.015
561.556 | 0.016
561.844 | 0.021
562.131 | 0.022
562.419 | 0.023
562.706 | 0.017
562.994 | 0.012
563.281 | 0.025
563.568 | 0.021
563.855 | 0.026
564.142 | 0.011
564.429 | 0.023
564.716 | 0.019
565.003 | 0.025
565.289 | 0.023
565.576 | 0.019
565.863 | 0.019
566.149 | 0.024
566.436 | 0.022
566.722 | 0.011
567.008 | 0.017
567.295 | 0.024
567.581 | 0.015
567.867 | 0.015
568.153 | 0.01
568.439 | 0.017
568.725 | 0.017
569.011 | 0.021
569.297 | 0.02
569.583 | 0.019
569.868 | 0.014
570.154 | 0.02
570.439 | 0.011
570.725 | 0.021
571.01 | 0.018
571.296 | 0.027
571.581 | 0.017
571.866 | 0.021
572.151 | 0.023
572.436 | 0.015
572.721 | 0.018
573.006 | 0.015
573.291 | 0.025
573.576 | 0.018
573.861 | 0.024
574.146 | 0.012
574.43 | 0.027
574.715 | 0.029
574.999 | 0.022
575.284 | 0.029
575.568 | 0.015
575.852 | 0.025
576.137 | 0.052
576.421 | 0.13
576.705 | 0.159
576.989 | 0.249
577.273 | 0.208
577.557 | 0.06
577.841 | 0.027
578.124 | 0.041
578.408 | 0.089
578.692 | 0.136
578.975 | 0.205
579.259 | 0.246
579.542 | 0.098
579.826 | 0.026
580.109 | 0.027
580.392 | 0.024
580.676 | 0.018
580.959 | 0.02
581.242 | 0.02
581.525 | 0.017
581.808 | 0.02
582.091 | 0.017
582.373 | 0.012
582.656 | 0.021
582.939 | 0.019
583.221 | 0.01
583.504 | 0.016
583.786 | 0.019
584.069 | 0.013
584.351 | 0.015
584.633 | 0.015
584.916 | 0.008
585.198 | 0.018
585.48 | 0.019
585.762 | 0.006
586.044 | 0.014
586.326 | 0.018
586.607 | 0.016
586.889 | 0.035
587.171 | 0.021
587.453 | 0.013
587.734 | 0.021
588.016 | 0.023
588.297 | 0.011
588.578 | 0.024
588.86 | 0.024
589.141 | 0.014
589.422 | 0.017
589.703 | 0.029
589.984 | 0.01
590.265 | 0.02
590.546 | 0.02
590.827 | 0.012
591.108 | 0.012
591.389 | 0.026
591.669 | 0.012
591.95 | 0.018
592.23 | 0.018
592.511 | 0.027
592.791 | 0.021
593.071 | 0.022
593.352 | 0.009
593.632 | 0.021
593.912 | 0.024
594.192 | 0.025
594.472 | 0.019
594.752 | 0.022
595.032 | 0.021
595.312 | 0.011
595.592 | 0.019
595.871 | 0.018
596.151 | 0.016
596.43 | 0.003
596.71 | 0.015
596.989 | 0.027
597.269 | 0.017
597.548 | 0.017
597.827 | 0.03
598.106 | 0.028
598.385 | 0.019
598.665 | 0.002
598.944 | 0.008
599.222 | 0.011
599.501 | 0.015
599.78 | 0.021
600.059 | 0.016
600.337 | 0.014
600.616 | 0.013
600.895 | 0.021
601.173 | 0.042
601.451 | 0.016
601.73 | 0.014
602.008 | 0.021
602.286 | 0.006
602.564 | 0.022
602.842 | 0.013
603.12 | 0.007
603.398 | 0.017
603.676 | 0.013
603.954 | 0.011
604.232 | 0.016
604.51 | 0.011
604.787 | 0.006
605.065 | 0.013
605.342 | 0.011
605.62 | 0.013
605.897 | 0.014
606.174 | 0.02
606.452 | 0.01
606.729 | 0.024
607.006 | 0.022
607.283 | 0.016
607.56 | 0.012
607.837 | 0.005
608.114 | 0.016
608.391 | 0.041
608.667 | 0.004
608.944 | 0.013
609.221 | 0.009
609.497 | 0.011
609.774 | 0.006
610.05 | 0.022
610.326 | 0.019
610.603 | 0.02
610.879 | 0.025
611.155 | 0.021
611.431 | 0.011
611.707 | 0.016
611.983 | 0.007
612.259 | 0.024
612.535 | 0.012
612.811 | 0.011
613.087 | 0.015
613.362 | 0.011
613.638 | 0.011
613.913 | 0.016
614.189 | 0.015
614.464 | 0.017
614.74 | 0.025
615.015 | 0.002
615.29 | 0.017
615.565 | 0.018
615.84 | 0.007
616.115 | 0.021
616.39 | 0.018
616.665 | 0.023
616.94 | 0.019
617.215 | 0.024
617.49 | 0.014
617.764 | 0.025
618.039 | 0.021
618.313 | 0.022
618.588 | 0.03
618.862 | 0.017
619.137 | 0.028
619.411 | 0.012
619.685 | 0.016
619.959 | 0.013
620.233 | 0.032
620.507 | 0.014
620.781 | 0.02
621.055 | 0.015
621.329 | 0.016
621.603 | 0.016
621.876 | 0.027
622.15 | 0.021
622.424 | 0.015
622.697 | 0.017
622.971 | 0.015
623.244 | 0.011
623.517 | 0.01
623.791 | 0.015
624.064 | 0.025
624.337 | 0.026
624.61 | 0.009
624.883 | 0.019
625.156 | 0.012
625.429 | 0.02
625.702 | 0.009
625.974 | 0.025
626.247 | -0.001
626.52 | 0.02
626.792 | 0.017
627.065 | 0.021
627.337 | 0.02
627.61 | 0.017
627.882 | 0.019
628.154 | 0.016
628.427 | 0.029
628.699 | 0.018
628.971 | 0.019
629.243 | 0.009
629.515 | 0.026
629.787 | 0.018
630.058 | 0.027
630.33 | 0.027
630.602 | 0.005
630.874 | 0.025
631.145 | 0.031
631.417 | 0.009
631.688 | 0.021
631.959 | 0.02
632.231 | 0.016
632.502 | 0.021
632.773 | 0.027
633.044 | 0.033
633.316 | 0.026
633.587 | 0.007
633.857 | 0.017
634.128 | 0.032
634.399 | 0.034
634.67 | 0.015
634.941 | 0.019
635.211 | 0.023
635.482 | 0.016
635.752 | 0.023
636.023 | 0.008
636.293 | 0.014
636.564 | 0.03
636.834 | 0.018
637.104 | 0.027
637.374 | 0.018
637.644 | 0.02
637.914 | 0.018
638.184 | 0.024
638.454 | 0.016
638.724 | 0.03
638.994 | 0.02
639.264 | 0.03
639.533 | 0.015
639.803 | 0.013
640.072 | 0.02
640.342 | 0.016
640.611 | 0.022
640.881 | 0.02
641.15 | 0.037
641.419 | 0.004
641.688 | 0.02
641.957 | 0.017
642.227 | 0.023
642.496 | 0.031
642.764 | 0.029
643.033 | 0.029
643.302 | 0.015
643.571 | 0.021
643.84 | 0.015
644.108 | 0.033
644.377 | 0.002
644.645 | 0.011
644.914 | 0.012
645.182 | 0.027
645.45 | 0.024
645.719 | 0.032
645.987 | 0.031
646.255 | 0.013
646.523 | 0.017
646.791 | 0.013
647.059 | 0.032
647.327 | 0.018
647.595 | 0.019
647.862 | 0.005
648.13 | 0.018
648.398 | 0.032
648.665 | 0.014
648.933 | 0.031
649.2 | 0.021
649.468 | 0.017
649.735 | 0.03
650.002 | 0.023
650.269 | 0.023
650.536 | 0.015
650.804 | 0.015
651.071 | 0.033
651.338 | 0.021
651.604 | 0.025
651.871 | 0.015
652.138 | 0.018
652.405 | 0.009
652.671 | 0.026
652.938 | 0.02
653.205 | 0.018
653.471 | 0.004
653.737 | 0.02
654.004 | 0.008
654.27 | 0.014
654.536 | 0.02
654.803 | 0.013
655.069 | 0.019
655.335 | 0.02
655.601 | 0.024
655.867 | 0.019
656.132 | 0.029
656.398 | 0.021
656.664 | 0.023
656.93 | 0.015
657.195 | 0.018
657.461 | 0.019
657.726 | 0.017
657.992 | 0.005
658.257 | 0.02
658.523 | 0.021
658.788 | 0.009
659.053 | 0.015
659.318 | 0.015
659.583 | 0.019
659.848 | 0.014
660.113 | 0.02
660.378 | 0.023
660.643 | 0.017
660.908 | 0.01
661.172 | 0.002
661.437 | 0.014
661.702 | 0.02
661.966 | 0.026
662.231 | 0.018
662.495 | 0.022
662.76 | 0.035
663.024 | 0.015
663.288 | 0.014
663.552 | 0.009
663.816 | 0.016
664.08 | 0.03
664.344 | 0.022
664.608 | 0.018
664.872 | 0.013
665.136 | 0.011
665.4 | 0.018
665.664 | 0.025
665.927 | 0.021
666.191 | 0.021
666.454 | 0.008
666.718 | 0.016
666.981 | 0.005
667.244 | 0.015
667.508 | 0.03
667.771 | 0.017
668.034 | 0.025
668.297 | 0.011
668.56 | 0.018
668.823 | -0.001
669.086 | 0.031
669.349 | 0.016
669.612 | 0.011
669.875 | 0.013
670.137 | 0.022
670.4 | 0.022
670.662 | 0.012
670.925 | 0.025
671.187 | 0.025
671.45 | 0.001
671.712 | 0.042
671.974 | 0.021
672.237 | 0.011
672.499 | 0.017
672.761 | 0.023
673.023 | 0.012
673.285 | 0.018
673.547 | 0.022
673.808 | 0.019
674.07 | 0.014
674.332 | 0.02
674.594 | 0.013
674.855 | 0.018
675.117 | 0.008
675.378 | 0.026
675.64 | 0.014
675.901 | 0
676.162 | 0.021
676.424 | 0.012
676.685 | 0.026
676.946 | 0.022
677.207 | 0.031
677.468 | 0.037
677.729 | 0.008
677.99 | 0.01
678.251 | 0.021
678.512 | 0.018
678.772 | 0.021
679.033 | 0.034
679.294 | 0.02
679.554 | 0.021
679.815 | 0.018
680.075 | 0.012
680.335 | 0.022
680.596 | 0.008
680.856 | 0.024
681.116 | 0.036
681.376 | 0.022
681.636 | 0.018
681.896 | 0.029
682.156 | 0.029
682.416 | 0.017
682.676 | 0.017
682.936 | 0.029
683.196 | 0.017
683.455 | 0.017
683.715 | 0.02
683.974 | 0.022
684.234 | 0.017
684.493 | 0.012
684.753 | 0.025
685.012 | 0.029
685.271 | 0.017
685.53 | 0.02
685.789 | 0.028
686.049 | 0.021
686.308 | 0.015
686.567 | 0.02
686.825 | 0.027
687.084 | 0.026
687.343 | 0.021
687.602 | 0.025
687.86 | 0.026
688.119 | 0.022
688.378 | 0.028
688.636 | 0.029
688.894 | 0.017
689.153 | 0.023
689.411 | 0.031
689.669 | 0.023
689.928 | 0.03
690.186 | 0.008
690.444 | 0.042
690.702 | 0.04
690.96 | 0.034
691.218 | 0.018
691.476 | 0.025
691.733 | 0.025
691.991 | 0.03
692.249 | 0.016
692.506 | 0.019
692.764 | 0.022
693.021 | 0.03
693.279 | 0.026
693.536 | 0.029
693.794 | 0.002
694.051 | 0.017
694.308 | 0.021
694.565 | 0.031
694.822 | 0.036
695.079 | 0.009
695.336 | 0.02
695.593 | 0.013
695.85 | 0.018
696.107 | 0.014
696.364 | 0.023
696.62 | 0.033
696.877 | 0.026
697.133 | 0.011
697.39 | 0.024
697.646 | 0.023
697.903 | 0.013
698.159 | 0.008
698.415 | 0.017
698.672 | 0.027
698.928 | 0.019
699.184 | 0.027
699.44 | 0.02
699.696 | 0.037
699.952 | 0.001
700.208 | 0.016
700.464 | 0.015
700.719 | 0.04
700.975 | 0.025
701.231 | 0.029
701.486 | 0.033
701.742 | 0.021
701.997 | 0.022
702.253 | 0.025
702.508 | 0.005
702.763 | 0.026
703.019 | 0.03
703.274 | 0.027
703.529 | 0.033
703.784 | 0.022
704.039 | 0.042
704.294 | 0.019
704.549 | 0.014
704.804 | 0.015
705.058 | 0.028
705.313 | 0.009
705.568 | 0.037
705.822 | 0.029
706.077 | 0.016
706.331 | 0.012
706.586 | 0.022
706.84 | 0.035
707.095 | 0.023
707.349 | -0.007
707.603 | 0.015
707.857 | 0.021
708.111 | 0.016
708.365 | 0.014
708.619 | 0.02
708.873 | 0.036
709.127 | 0.01
709.381 | 0.009
709.635 | 0.018
709.888 | 0.031
710.142 | 0.02
710.396 | 0.019
710.649 | 0.022
710.903 | 0.016
711.156 | 0.018
711.409 | 0.033
711.663 | 0.027
711.916 | 0.015
712.169 | 0.016
712.422 | 0.027
712.675 | 0.019
712.928 | 0.042
713.181 | 0.03
713.434 | 0.027
713.687 | 0.035
713.94 | 0.02
714.192 | 0.024
714.445 | -0.001
714.698 | 0.034
714.95 | 0.029
715.203 | 0.048
715.455 | 0.023
715.708 | 0.036
715.96 | 0.041
716.212 | 0.02
716.464 | 0.038
716.717 | 0.028
716.969 | 0.031
717.221 | 0.03
717.473 | 0.014
717.725 | 0.024
717.977 | 0.031
718.228 | 0.04
718.48 | 0.025
718.732 | 0.027
718.984 | 0.038
719.235 | 0.037
719.487 | 0.032
719.738 | 0.018
719.99 | 0.038
720.241 | 0.026
720.492 | 0.03
720.744 | 0.017
720.995 | 0.015
721.246 | 0.021
721.497 | 0.043
721.748 | 0.025
721.999 | 0.018
722.25 | 0.057
722.501 | 0.05
722.752 | 0.026
723.002 | 0.023
723.253 | 0.018
723.504 | 0.042
723.754 | 0.029
724.005 | 0.036
724.255 | 0.031
724.506 | 0.038
724.756 | 0.038
725.006 | 0.021
725.257 | 0.03
725.507 | 0.032
725.757 | 0.023
726.007 | 0.014
726.257 | 0.033
726.507 | 0.015
726.757 | 0.055
727.007 | 0.042
727.257 | 0.013
727.506 | 0.036
727.756 | 0.023
728.006 | 0.039
728.255 | 0.041
728.505 | 0.012
728.754 | 0.031
729.004 | 0.051
729.253 | 0.054
729.502 | 0.061
729.751 | 0.03
730.001 | 0.018
730.25 | 0.024
730.499 | 0.041
730.748 | 0.034
730.997 | 0.03
731.246 | 0.034
731.495 | 0.029
731.743 | 0.038
731.992 | 0.032
732.241 | 0.016
732.489 | 0.028
732.738 | 0.057
732.987 | 0.026
733.235 | 0.033
733.483 | 0.042
733.732 | 0.042
733.98 | 0.019
734.228 | 0.024
734.477 | 0.022
734.725 | 0.041
734.973 | 0.027
735.221 | 0.042
735.469 | 0.044
735.717 | 0.039
735.965 | 0.027
736.212 | 0.021
736.46 | 0.041
736.708 | 0.032
736.955 | 0.057
737.203 | 0.017
737.451 | 0.045
737.698 | 0.028
737.945 | 0.031
738.193 | 0.054
738.44 | 0.039
738.687 | 0.036
738.935 | 0.031
739.182 | 0.029
739.429 | 0.037
739.676 | 0.025
739.923 | 0.036
740.17 | 0.04
740.417 | 0.023
740.663 | 0.024
740.91 | 0.033
741.157 | 0.043
741.403 | 0.028
741.65 | 0.006
741.897 | 0.03
742.143 | 0.016
742.389 | 0.032
742.636 | 0.02
742.882 | 0.021
743.128 | 0.027
743.375 | 0.012
743.621 | 0.054
743.867 | 0.034
744.113 | 0.026
744.359 | 0.022
744.605 | 0.042
744.851 | 0.06
745.097 | 0.048
745.342 | 0.027
745.588 | 0.035
745.834 | 0.028
746.079 | 0.048
746.325 | 0.04
746.57 | 0.04
746.816 | 0.009
747.061 | 0.039
747.307 | 0.037
747.552 | 0.026
747.797 | 0.042
748.042 | 0.038
748.287 | 0.028
748.532 | 0.021
748.777 | 0.038
749.022 | 0.042
749.267 | 0.036
749.512 | 0.03
749.757 | 0.032
750.002 | 0.028
750.246 | 0.038
750.491 | 0.056
750.736 | 0.029
750.98 | 0.042
751.225 | 0.04
751.469 | 0.058
751.713 | 0.047
751.958 | 0.029
752.202 | 0.027
752.446 | 0.034
752.69 | 0.013
752.934 | 0.015
753.178 | 0.04
753.422 | 0.021
753.666 | 0.048
753.91 | 0.014
754.154 | 0.047
754.398 | 0.074
754.641 | 0.012
754.885 | 0.009
755.129 | 0.035
755.372 | 0.009
755.616 | 0.034
755.859 | 0.033
756.103 | 0.061
756.346 | 0.005
756.589 | 0.036
756.832 | 0.027
757.076 | 0.047
757.319 | 0.04
757.562 | 0.05
757.805 | 0.044
758.048 | 0.037
758.291 | 0.026
758.533 | 0.028
758.776 | 0.033
759.019 | 0.045
759.262 | 0.041
759.504 | 0.038
759.747 | 0.039
759.989 | 0.048
760.232 | 0.047
760.474 | 0.029
760.717 | 0.066
760.959 | 0.025
761.201 | 0.026
761.444 | 0.033
761.686 | 0.045
761.928 | 0.04
762.17 | 0.029
762.412 | 0.044
762.654 | 0.042
762.896 | 0.039
763.137 | 0.046
763.379 | 0.058
763.621 | 0.096
763.863 | 0.047
764.104 | 0.042
764.346 | 0.066
764.587 | 0.029
764.829 | 0.029
765.07 | 0.018
765.312 | 0.051
765.553 | 0.055
765.794 | 0.044
766.035 | 0.042
766.276 | 0.061
766.518 | 0.04
766.759 | 0.064
767 | 0.04
767.241 | 0.044
767.481 | 0.034
767.722 | 0.046
767.963 | 0.039
768.204 | 0.051
768.444 | 0.064
768.685 | 0.04
768.926 | 0.02
769.166 | 0.035
769.407 | 0.035
769.647 | 0.072
769.887 | 0.014
770.128 | 0.045
770.368 | 0.031
770.608 | 0.056
770.848 | 0.042
771.088 | 0.032
771.328 | 0.04
771.568 | 0.032
771.808 | 0.053
772.048 | 0.029
772.288 | 0.044
772.528 | 0.039
772.767 | 0.041
773.007 | 0.051
773.247 | 0.045
773.486 | 0.032
773.726 | 0.022
773.965 | 0.056
774.205 | 0.051
774.444 | 0.045
774.683 | 0.053
774.922 | 0.057
775.162 | 0.05
775.401 | 0.062
775.64 | 0.043
775.879 | 0.044
776.118 | 0.044
776.357 | 0.043
776.596 | 0.044
776.834 | 0.051
777.073 | 0.045
777.312 | 0.062
777.551 | 0.032
777.789 | 0.027
778.028 | 0.041
778.266 | 0.049
778.505 | 0.025
778.743 | 0.03
778.982 | 0.022
779.22 | 0.041
779.458 | 0.051
779.696 | 0.047
779.934 | 0.057
780.173 | 0.053
780.411 | 0.049
780.649 | 0.047
780.887 | 0.047
781.124 | 0.062
781.362 | 0.038
781.6 | 0.048
781.838 | 0.037
782.075 | 0.036
782.313 | 0.037
782.551 | 0.024
782.788 | 0.022
783.026 | 0.043
783.263 | 0.035
783.5 | 0.045
783.738 | 0.044
783.975 | 0.026
784.212 | -0.001
784.449 | 0.044
784.686 | 0.032
784.924 | 0.04
785.161 | 0.024
785.397 | 0.034
785.634 | 0.048
785.871 | 0.06
786.108 | 0.044
786.345 | 0.072
786.581 | 0.029
786.818 | 0.058
787.055 | 0.056
787.291 | 0.059
787.528 | 0.061
787.764 | 0.048
788.001 | 0.028
788.237 | 0.065
788.473 | 0.019
788.709 | 0.037
788.946 | 0.05
789.182 | 0.029
789.418 | 0.023
789.654 | 0.058
789.89 | 0.033
790.126 | 0.068
790.362 | 0.076
790.597 | 0.027
790.833 | 0.043
791.069 | 0.025
791.305 | 0.029
791.54 | 0.049
791.776 | 0.062
792.011 | 0.056
792.247 | 0.016
792.482 | 0.035
792.717 | 0.045
792.953 | 0.034
793.188 | 0.066
793.423 | 0.061
793.658 | 0.058
793.894 | 0.064
794.129 | 0.038
794.364 | 0.043
794.599 | 0.037
794.833 | 0.048
795.068 | 0.069
795.303 | 0.053
795.538 | 0.027
795.772 | 0.054
796.007 | 0.061
796.242 | 0.069
796.476 | 0.05
796.711 | 0.038
796.945 | 0.048
797.18 | 0.024
797.414 | 0.017
797.648 | 0.076
797.882 | 0.034
798.117 | 0.038
798.351 | 0.033
798.585 | 0.072
798.819 | 0.017
799.053 | 0.059
799.287 | 0.046
799.521 | 0.066
799.755 | 0.053
799.988 | 0.057
800.222 | 0.053
800.456 | 0.05
800.689 | 0.021
800.923 | 0.034
801.156 | 0.072
801.39 | 0.035
801.623 | 0.067
801.857 | 0.068
802.09 | 0.056
802.323 | 0.048
802.557 | 0.05
802.79 | 0.044
803.023 | 0.031
803.256 | 0.035
803.489 | 0.047
803.722 | 0.038
803.955 | 0.035
804.188 | 0.08
804.421 | 0.047
804.653 | 0.051
804.886 | 0.036
805.119 | 0.017
805.351 | 0.043
805.584 | 0.062
805.816 | 0.02
806.049 | 0.03
806.281 | 0.046
806.514 | 0.086
806.746 | 0.079
806.978 | 0.029
807.211 | 0.019
807.443 | 0.039
807.675 | 0.053
254 2
AbsoluteIrradiance:211
195.562 | 0
195.934 | 0
196.307 | 0
196.679 | 0
197.051 | 0
197.423 | 0
197.795 | 0
198.167 | 0
198.539 | 0
198.911 | 0
199.283 | 0
199.654 | 0
200.026 | 0
200.397 | 0
200.769 | 0
201.14 | 0
201.511 | 0
201.882 | 0
202.253 | 0
202.624 | 0
202.995 | 0
203.366 | 0
203.736 | 0
204.107 | 0
204.477 | 0
204.848 | 3.052
205.218 | 6.453
205.588 | 2.216
205.959 | 1.98
206.329 | 2.638
206.699 | 2.105
207.068 | 2.154
207.438 | 2.701
207.808 | 1.378
208.178 | 2.253
208.547 | 1.49
208.917 | 1.575
209.286 | 2.826
209.655 | 2.283
210.025 | 1.323
210.394 | 1.483
210.763 | 1.645
211.132 | 1.102
211.501 | 0.652
211.869 | 0.954
212.238 | 1.628
212.607 | 1.748
212.975 | 0.705
213.344 | 0.807
213.712 | 0.673
214.08 | 0.763
214.448 | 0.805
214.817 | 0.663
215.185 | 1.452
215.553 | 0.893
215.92 | 1.028
216.288 | 0.864
216.656 | 0.995
217.023 | 0.733
217.391 | 0.457
217.758 | 0.948
218.126 | 0.846
218.493 | 0.83
218.86 | 0.738
219.227 | 0.437
219.594 | 0.59
219.961 | 0.849
220.328 | 0.47
220.695 | 0.533
221.061 | 0.701
221.428 | 0.682
221.795 | 0.72
222.161 | 0.775
222.527 | 0.4
222.893 | 0.869
223.26 | 0.247
223.626 | 0.789
223.992 | 0.579
224.358 | 0.4
224.723 | 0.633
225.089 | 0.216
225.455 | 0.522
225.82 | 0.428
226.186 | 0.309
226.551 | 0.136
226.917 | 0.289
227.282 | 0.334
227.647 | 0.47
228.012 | 0.339
228.377 | 0.273
228.742 | 0.306
229.107 | 0.24
229.471 | 0.304
229.836 | 0.217
230.201 | 0.291
230.565 | 0.325
230.929 | 0.231
231.294 | 0.173
231.658 | 0.163
232.022 | 0.169
232.386 | 0.106
232.75 | -0.04
233.114 | 0.1
233.478 | 0.153
233.842 | 0.191
234.205 | 0.174
234.569 | 0.149
234.932 | 0.119
235.296 | 0.104
235.659 | 0.122
236.022 | 0.104
236.385 | 0.091
236.748 | 0.168
237.111 | 0.072
237.474 | 0.158
237.837 | 0.17
238.2 | 0.128
238.562 | 0.105
238.925 | 0.173
239.287 | 0.155
239.65 | 0.088
240.012 | 0.15
240.374 | 0.149
240.736 | 0.153
241.098 | 0.136
241.46 | 0.139
241.822 | 0.164
242.184 | 0.082
242.546 | 0.133
242.907 | 0.128
243.269 | 0.088
243.63 | 0.088
243.992 | 0.154
244.353 | 0.144
244.714 | 0.136
245.075 | 0.141
245.436 | 0.156
245.797 | 0.138
246.158 | 0.105
246.519 | 0.182
246.879 | 0.205
247.24 | 0.183
247.601 | 0.217
247.961 | 0.259
248.321 | 0.212
248.682 | 0.242
249.042 | 0.242
249.402 | 0.266
249.762 | 0.355
250.122 | 0.348
250.482 | 0.468
250.842 | 0.507
251.201 | 0.847
251.561 | 2.818
251.921 | 8.547
252.28 | 11.174
252.639 | 13.064
252.999 | 17.959
253.358 | 33.301
253.717 | 61.263
254.076 | 28.02
254.435 | 3.696
254.794 | 1.235
255.153 | 0.861
255.511 | 0.715
255.87 | 0.606
256.228 | 0.543
256.587 | 0.465
256.945 | 0.416
257.304 | 0.354
257.662 | 0.387
258.02 | 0.344
258.378 | 0.342
258.736 | 0.304
259.094 | 0.364
259.452 | 0.329
259.809 | 0.295
260.167 | 0.29
260.524 | 0.278
260.882 | 0.244
261.239 | 0.265
261.597 | 0.25
261.954 | 0.243
262.311 | 0.272
262.668 | 0.271
263.025 | 0.25
263.382 | 0.259
263.739 | 0.263
264.095 | 0.292
264.452 | 0.26
264.809 | 0.267
265.165 | 0.275
265.521 | 0.274
265.878 | 0.164
266.234 | 0.149
266.59 | 0.126
266.946 | 0.182
267.302 | 0.149
267.658 | 0.145
268.014 | 0.127
268.37 | 0.149
268.725 | 0.139
269.081 | 0.132
269.436 | 0.138
269.792 | 0.107
270.147 | 0.133
270.502 | 0.154
270.857 | 0.144
271.212 | 0.089
271.567 | 0.05
271.922 | 0.113
272.277 | 0.07
272.632 | 0.089
272.987 | 0.101
273.341 | 0.093
273.696 | 0.073
274.05 | 0.047
274.404 | 0.079
274.759 | 0.108
275.113 | 0.117
275.467 | 0.063
275.821 | 0.057
276.175 | 0.064
276.529 | 0.079
276.882 | 0.109
277.236 | 0.08
277.59 | 0.085
277.943 | 0.041
278.297 | 0.116
278.65 | 0.046
279.003 | 0.046
279.356 | 0.058
279.709 | 0.067
280.062 | 0.067
280.415 | 0.1
280.768 | 0.069
281.121 | 0.089
281.474 | 0.076
281.826 | 0.058
282.179 | 0.06
282.531 | 0.043
282.884 | 0.071
283.236 | 0.061
283.588 | 0.097
283.94 | 0.045
284.292 | 0.056
284.644 | 0.038
284.996 | 0.06
285.348 | 0.053
285.7 | 0.069
286.051 | 0.059
286.403 | 0.052
286.754 | 0.056
287.106 | 0.083
287.457 | 0.069
287.808 | 0.065
288.159 | 0.083
288.51 | 0.083
288.861 | 0.087
289.212 | 0.087
289.563 | 0.064
289.914 | 0.067
290.264 | 0.03
290.615 | 0.054
290.966 | 0.07
291.316 | 0.034
291.666 | 0.068
292.017 | 0.055
292.367 | 0.046
292.717 | 0.046
293.067 | 0.042
293.417 | 0.052
293.767 | 0.04
294.116 | 0.064
294.466 | 0.045
294.816 | 0.057
295.165 | 0.044
295.515 | 0.074
295.864 | 0.109
296.213 | 0.152
296.562 | 0.239
296.911 | 0.155
297.261 | 0.047
297.609 | 0.04
297.958 | 0.015
298.307 | 0.033
298.656 | 0.03
299.004 | 0.029
299.353 | 0.003
299.701 | 0.036
300.05 | 0.024
300.398 | 0.037
300.746 | 0.045
301.095 | 0.063
301.443 | 0.082
301.791 | 0.13
302.138 | 0.16
302.486 | 0.068
302.834 | 0.032
303.182 | 0.016
303.529 | 0.032
303.877 | 0.027
304.224 | 0.04
304.572 | 0.035
304.919 | 0.023
305.266 | 0.038
305.613 | 0.02
305.96 | 0.043
306.307 | 0.028
306.654 | 0.041
307.001 | 0.036
307.347 | 0.03
307.694 | 0.02
308.04 | 0.023
308.387 | 0.034
308.733 | 0.026
309.08 | 0.028
309.426 | 0.041
309.772 | 0.033
310.118 | 0.041
310.464 | 0.021
310.81 | 0.039
311.156 | 0.055
311.501 | 0.142
311.847 | 0.281
312.193 | 0.594
312.538 | 0.8
312.883 | 0.786
313.229 | 0.524
313.574 | 0.108
313.919 | 0.036
314.264 | 0.036
314.609 | 0.035
314.954 | 0.025
315.299 | 0.02
315.644 | 0.023
315.988 | 0.029
316.333 | 0.035
316.678 | 0.023
317.022 | 0.023
317.366 | 0.024
317.711 | 0.026
318.055 | 0.024
318.399 | 0.021
318.743 | 0.027
319.087 | 0.023
319.431 | 0.032
319.775 | 0.025
320.118 | 0.036
320.462 | 0.022
320.806 | 0.023
321.149 | 0.034
321.493 | 0.025
321.836 | 0.025
322.179 | 0.012
322.522 | 0.027
322.865 | 0.027
323.208 | 0.024
323.551 | 0.032
323.894 | 0.02
324.237 | 0.014
324.58 | 0.023
324.922 | 0.025
325.265 | 0.033
325.607 | 0.03
325.95 | 0.033
326.292 | 0.016
326.634 | 0.02
326.976 | 0.019
327.318 | 0.015
327.66 | 0.028
328.002 | 0.004
328.344 | 0.022
328.686 | 0.023
329.027 | 0.01
329.369 | 0.015
329.711 | 0.022
330.052 | 0.022
330.393 | 0.015
330.735 | 0.014
331.076 | 0.023
331.417 | 0.013
331.758 | 0.022
332.099 | 0.008
332.44 | 0.025
332.78 | 0.012
333.121 | 0.031
333.462 | 0.051
333.802 | 0.086
334.143 | 0.068
334.483 | 0.028
334.823 | 0.011
335.164 | 0.021
335.504 | 0.012
335.844 | 0.026
336.184 | 0.013
336.524 | 0.012
336.864 | 0.017
337.203 | 0.018
337.543 | 0.028
337.883 | 0.022
338.222 | 0.035
338.562 | 0.011
338.901 | 0.027
339.24 | 0.009
339.58 | 0.014
339.919 | 0.015
340.258 | 0.022
340.597 | 0.016
340.936 | 0.012
341.274 | 0.021
341.613 | 0.026
341.952 | 0.022
342.29 | 0.02
342.629 | 0.031
342.967 | 0.01
343.306 | 0.017
343.644 | 0.026
343.982 | 0.01
344.32 | 0.023
344.658 | 0.013
344.996 | 0.018
345.334 | 0.014
345.672 | 0.023
346.01 | 0.012
346.347 | 0.032
346.685 | 0.023
347.022 | 0.027
347.36 | 0.027
347.697 | 0.017
348.034 | 0.021
348.371 | 0.021
348.708 | 0.019
349.045 | 0.026
349.382 | 0.021
349.719 | 0.014
350.056 | 0.014
350.393 | 0.017
350.729 | 0.025
351.066 | 0.015
351.402 | 0.021
351.739 | 0.032
352.075 | 0.024
352.411 | 0.021
352.747 | 0.016
353.083 | 0.016
353.419 | 0.015
353.755 | 0.022
354.091 | 0.024
354.427 | 0.021
354.763 | 0.022
355.098 | 0.026
355.434 | 0.012
355.769 | 0.034
356.104 | 0.008
356.44 | 0.014
356.775 | 0.024
357.11 | 0.021
357.445 | 0.02
357.78 | 0.01
358.115 | 0.025
358.45 | 0.024
358.785 | 0.025
359.119 | 0.015
359.454 | 0.018
359.788 | 0.017
360.123 | 0.021
360.457 | 0.011
360.791 | 0.03
361.126 | 0.016
361.46 | 0.017
361.794 | 0.026
362.128 | 0.031
362.462 | 0.032
362.795 | 0.032
363.129 | 0.03
363.463 | 0.019
363.796 | 0.051
364.13 | 0.191
364.463 | 0.526
364.797 | 0.737
365.13 | 0.59
365.463 | 0.254
365.796 | 0.166
366.129 | 0.148
366.462 | 0.083
366.795 | 0.037
367.128 | 0.031
367.461 | 0.017
367.793 | 0.033
368.126 | 0.02
368.458 | 0.033
368.791 | 0.025
369.123 | 0.025
369.455 | 0.023
369.787 | 0.02
370.12 | 0.021
370.452 | 0.014
370.784 | 0.021
371.115 | 0.022
371.447 | 0.011
371.779 | 0.021
372.111 | 0.026
372.442 | 0.021
372.774 | 0.032
373.105 | 0.026
373.436 | 0.02
373.768 | 0.019
374.099 | 0.014
374.43 | 0.027
374.761 | 0.016
375.092 | 0.023
375.423 | 0.015
375.754 | 0.018
376.084 | 0.019
376.415 | 0.02
376.745 | 0.02
377.076 | 0.022
377.406 | 0.021
377.737 | 0.034
378.067 | 0.026
378.397 | 0.017
378.727 | 0.017
379.057 | 0.013
379.387 | 0.016
379.717 | 0.02
380.047 | 0.024
380.377 | 0.022
380.706 | 0.027
381.036 | 0.018
381.365 | 0.036
381.695 | 0.02
382.024 | 0.016
382.353 | 0.007
382.682 | 0.02
383.012 | 0.016
383.341 | 0.026
383.67 | 0.023
383.998 | 0.011
384.327 | 0.025
384.656 | 0.017
384.985 | 0.009
385.313 | 0.012
385.642 | 0.015
385.97 | 0.016
386.298 | 0.016
386.627 | 0.015
386.955 | 0.027
387.283 | 0.014
387.611 | 0.014
387.939 | 0.018
388.267 | 0.018
388.595 | 0.028
388.922 | 0.006
389.25 | 0.031
389.578 | 0.018
389.905 | 0.024
390.233 | 0.019
390.56 | 0.024
390.887 | 0.029
391.214 | 0.022
391.542 | 0.011
391.869 | 0.017
392.196 | 0.011
392.522 | 0.013
392.849 | 0.016
393.176 | 0.015
393.503 | 0.016
393.829 | 0.017
394.156 | 0.017
394.482 | 0.016
394.809 | 0.022
395.135 | 0.022
395.461 | 0.011
395.787 | 0.025
396.113 | 0.01
396.439 | 0.015
396.765 | 0.016
397.091 | 0.017
397.417 | 0.014
397.742 | 0.016
398.068 | 0.027
398.394 | 0.017
398.719 | 0.015
399.044 | 0.017
399.37 | 0.024
399.695 | 0.017
400.02 | 0.023
400.345 | 0.026
400.67 | 0.021
400.995 | 0.029
401.32 | 0.019
401.645 | 0.021
401.969 | 0.015
402.294 | 0.018
402.618 | 0.033
402.943 | 0.019
403.267 | 0.032
403.592 | 0.112
403.916 | 0.377
404.24 | 0.783
404.564 | 1.031
404.888 | 0.606
405.212 | 0.114
405.536 | 0.024
405.86 | 0.028
406.183 | 0.021
406.507 | 0.027
406.831 | 0.038
407.154 | 0.056
407.478 | 0.094
407.801 | 0.098
408.124 | 0.045
408.447 | 0.022
408.77 | 0.023
409.093 | 0.024
409.416 | 0.011
409.739 | 0.025
410.062 | 0.013
410.385 | 0.023
410.708 | 0.023
411.03 | 0.025
411.353 | 0.011
411.675 | 0.025
411.997 | 0.016
412.32 | 0.015
412.642 | 0.024
412.964 | 0.017
413.286 | 0.018
413.608 | 0.02
413.93 | 0.023
414.252 | 0.01
414.574 | 0.013
414.895 | 0.016
415.217 | 0.024
415.538 | 0.012
415.86 | 0.018
416.181 | 0.025
416.503 | 0.016
416.824 | 0.014
417.145 | 0.017
417.466 | 0.017
417.787 | 0.027
418.108 | 0.01
418.429 | 0.015
418.75 | 0.017
419.07 | 0.016
419.391 | 0.02
419.712 | 0.02
420.032 | 0.016
420.353 | 0.016
420.673 | 0.013
420.993 | 0.019
421.313 | 0.02
421.633 | 0.021
421.953 | 0.012
422.273 | 0.014
422.593 | 0.017
422.913 | 0.016
423.233 | 0.014
423.553 | 0.015
423.872 | 0.013
424.192 | 0.014
424.511 | 0.012
424.83 | 0.022
425.15 | 0.02
425.469 | 0.021
425.788 | 0.022
426.107 | 0.021
426.426 | 0.024
426.745 | 0.027
427.064 | 0.02
427.383 | 0.023
427.702 | 0.009
428.02 | 0.022
428.339 | 0.018
428.657 | 0.021
428.976 | 0.021
429.294 | 0.016
429.612 | 0.012
429.93 | 0.02
430.248 | 0.025
430.567 | 0.018
430.884 | 0.017
431.202 | 0.013
431.52 | 0.014
431.838 | 0.015
432.156 | 0.019
432.473 | 0.019
432.791 | 0.025
433.108 | 0.019
433.426 | 0.035
433.743 | 0.036
434.06 | 0.049
434.377 | 0.074
434.694 | 0.197
435.011 | 0.637
435.328 | 1.25
435.645 | 2.162
435.962 | 2.342
436.279 | 0.747
436.595 | 0.09
436.912 | 0.03
437.228 | 0.023
437.545 | 0.02
437.861 | 0.02
438.177 | 0.018
438.493 | 0.015
438.81 | 0.024
439.126 | 0.025
439.442 | 0.03
439.757 | 0.014
440.073 | 0.024
440.389 | 0.023
440.705 | 0.034
441.02 | 0.022
441.336 | 0.02
441.651 | 0.024
441.967 | 0.017
442.282 | 0.015
442.597 | 0.019
442.912 | 0.03
443.227 | 0.026
443.543 | 0.026
443.857 | 0.025
444.172 | 0.021
444.487 | 0.024
444.802 | 0.027
445.116 | 0.03
445.431 | 0.015
445.746 | 0.013
446.06 | 0.018
446.374 | 0.017
446.689 | 0.021
447.003 | 0.017
447.317 | 0.017
447.631 | 0.008
447.945 | 0.018
448.259 | 0.019
448.573 | 0.022
448.887 | 0.012
449.2 | 0.024
449.514 | 0.021
449.827 | 0.018
450.141 | 0.019
450.454 | 0.012
450.768 | 0.015
451.081 | 0.021
451.394 | 0.013
451.707 | 0.012
452.02 | 0.024
452.333 | 0.015
452.646 | 0.012
452.959 | 0.013
453.272 | 0.015
453.584 | 0.013
453.897 | 0.016
454.21 | 0.017
454.522 | 0.013
454.834 | 0.01
455.147 | 0.011
455.459 | 0.006
455.771 | 0.004
456.083 | 0.008
456.395 | 0.013
456.707 | 0.013
457.019 | 0.017
457.331 | 0.013
457.643 | 0.011
457.954 | 0.018
458.266 | 0.012
458.577 | 0.012
458.889 | 0.009
459.2 | 0.015
459.512 | 0.01
459.823 | 0.015
460.134 | 0.013
460.445 | 0.018
460.756 | 0.016
461.067 | 0.014
461.378 | 0.01
461.689 | 0.01
461.999 | 0.025
462.31 | 0.019
462.621 | 0.011
462.931 | 0.013
463.242 | 0.018
463.552 | 0.021
463.862 | 0.01
464.172 | 0.014
464.483 | 0.023
464.793 | 0.009
465.103 | 0.016
465.413 | 0.016
465.722 | 0.014
466.032 | 0.018
466.342 | 0.014
466.652 | 0.015
466.961 | 0.015
467.271 | 0.019
467.58 | 0.016
467.889 | 0.017
468.199 | 0.017
468.508 | 0.012
468.817 | 0.016
469.126 | 0.015
469.435 | 0.021
469.744 | 0.013
470.053 | 0.012
470.362 | 0.016
470.67 | 0.011
470.979 | 0.015
471.287 | 0.009
471.596 | 0.011
471.904 | 0.021
472.213 | 0.01
472.521 | 0.019
472.829 | 0.018
473.137 | 0.014
473.445 | 0.01
473.753 | 0.025
474.061 | 0.011
474.369 | 0.014
474.677 | 0.014
474.985 | 0.009
475.292 | 0.007
475.6 | 0.019
475.907 | 0.016
476.215 | 0.011
476.522 | 0.022
476.829 | 0.016
477.137 | 0.02
477.444 | 0.009
477.751 | 0.02
478.058 | 0.012
478.365 | 0.015
478.671 | 0.021
478.978 | 0.019
479.285 | 0.011
479.592 | 0.012
479.898 | 0.007
480.205 | 0.017
480.511 | 0.004
480.817 | 0.019
481.124 | 0.015
481.43 | 0.018
481.736 | 0.027
482.042 | 0.012
482.348 | 0.013
482.654 | 0.02
482.96 | 0.019
483.266 | 0.019
483.571 | 0.022
483.877 | 0.007
484.182 | 0.008
484.488 | 0.015
484.793 | 0.01
485.099 | 0.012
485.404 | 0.014
485.709 | 0.019
486.014 | 0.015
486.319 | 0.024
486.624 | 0.012
486.929 | 0.007
487.234 | 0.004
487.539 | 0.018
487.844 | 0.018
488.148 | 0.014
488.453 | 0.024
488.757 | 0.018
489.062 | 0.013
489.366 | 0.017
489.67 | 0.012
489.975 | 0.009
490.279 | 0.014
490.583 | 0.02
490.887 | 0.015
491.191 | 0.016
491.495 | 0.028
491.798 | 0.029
492.102 | 0.021
492.406 | 0.01
492.709 | 0.016
493.013 | 0.02
493.316 | 0.016
493.62 | 0.022
493.923 | 0.013
494.226 | 0.021
494.529 | 0.015
494.832 | 0.016
495.135 | 0.013
495.438 | 0.02
495.741 | 0.028
496.044 | 0.011
496.347 | 0.008
496.649 | 0.025
496.952 | 0.014
497.255 | 0.013
497.557 | 0.014
497.859 | 0.012
498.162 | 0.006
498.464 | 0.022
498.766 | 0.019
499.068 | 0.013
499.37 | 0.012
499.672 | 0.02
499.974 | 0.013
500.276 | 0.017
500.578 | 0.018
500.879 | 0.006
501.181 | 0.024
501.483 | 0.013
501.784 | 0.021
502.085 | 0.009
502.387 | 0.009
502.688 | 0.009
502.989 | 0.018
503.29 | 0.009
503.591 | 0.011
503.892 | 0.008
504.193 | 0.019
504.494 | 0.014
504.795 | 0.006
505.096 | 0.023
505.396 | 0.008
505.697 | 0.005
505.997 | 0.015
506.298 | 0.016
506.598 | 0.009
506.898 | 0.02
507.199 | 0.012
507.499 | 0.017
507.799 | 0.01
508.099 | 0.014
508.399 | 0.016
508.699 | 0.018
508.998 | 0.014
509.298 | 0.021
509.598 | 0.009
509.897 | 0.017
510.197 | 0.01
510.496 | 0.013
510.796 | 0.01
511.095 | 0.017
511.394 | 0.015
511.693 | 0.024
511.992 | 0.015
512.291 | 0.019
512.59 | 0.02
512.889 | 0.018
513.188 | 0.02
513.487 | 0.019
513.786 | 0.016
514.084 | 0.017
514.383 | 0.021
514.681 | 0.025
514.98 | 0.014
515.278 | 0.02
515.576 | 0.023
515.874 | 0.018
516.173 | 0.013
516.471 | 0.023
516.769 | 0.017
517.067 | 0.016
517.364 | 0.011
517.662 | 0.011
517.96 | 0.013
518.257 | 0.013
518.555 | 0.015
518.853 | 0.011
519.15 | 0.013
519.447 | 0.017
519.745 | 0.013
520.042 | 0.02
520.339 | 0.009
520.636 | 0.018
520.933 | 0.023
521.23 | 0.012
521.527 | 0.013
521.824 | 0.02
522.121 | 0.011
522.417 | 0.008
522.714 | 0.006
523.01 | 0.015
523.307 | 0.012
523.603 | 0.015
523.9 | 0.013
524.196 | 0.019
524.492 | 0.008
524.788 | 0.024
525.084 | 0.018
525.38 | 0.023
525.676 | 0.011
525.972 | 0.017
526.268 | 0.016
526.564 | 0.017
526.859 | 0.018
527.155 | 0.01
527.45 | 0.017
527.746 | 0.018
528.041 | 0.02
528.336 | 0.016
528.632 | 0.01
528.927 | 0.012
529.222 | 0.017
529.517 | 0.019
529.812 | 0.014
530.107 | 0.014
530.402 | 0.016
530.696 | 0.021
530.991 | 0.025
531.286 | 0.018
531.58 | 0.017
531.875 | 0.014
532.169 | 0.017
532.463 | 0.012
532.758 | 0.009
533.052 | 0.018
533.346 | 0.01
533.64 | 0.017
533.934 | 0.019
534.228 | 0.021
534.522 | 0.018
534.816 | 0.017
535.11 | 0.008
535.403 | 0.019
535.697 | 0.023
535.99 | 0.018
536.284 | 0.015
536.577 | 0.027
536.871 | 0.015
537.164 | 0.015
537.457 | 0.018
537.75 | 0.015
538.043 | 0.017
538.336 | 0.018
538.629 | 0.016
538.922 | 0.016
539.215 | 0.003
539.507 | 0.011
539.8 | 0.007
540.093 | 0.025
540.385 | 0.011
540.678 | 0.015
540.97 | 0.006
541.262 | 0.017
541.554 | 0.014
541.847 | 0.018
542.139 | 0.021
542.431 | 0.021
542.723 | 0.015
543.015 | 0.02
543.307 | 0.023
543.598 | 0.026
543.89 | 0.024
544.182 | 0.018
544.473 | 0.028
544.765 | 0.044
545.056 | 0.079
545.347 | 0.34
545.639 | 0.696
545.93 | 0.952
546.221 | 1.464
546.512 | 0.748
546.803 | 0.121
547.094 | 0.039
547.385 | 0.019
547.676 | 0.026
547.967 | 0.035
548.257 | 0.015
548.548 | 0.023
548.839 | 0.008
549.129 | 0.015
549.419 | 0.028
549.71 | 0.022
550 | 0.031
550.29 | 0.022
550.58 | 0.009
550.871 | 0.02
551.161 | 0.016
551.45 | 0.012
551.74 | 0.018
552.03 | 0.011
552.32 | 0.019
552.61 | 0.022
552.899 | 0.029
553.189 | 0.018
553.478 | 0.019
553.768 | 0.02
554.057 | 0.031
554.346 | 0.026
554.635 | 0.024
554.925 | 0.024
555.214 | 0.024
555.503 | 0.013
555.792 | 0.024
556.081 | 0.009
556.369 | 0.017
556.658 | 0.013
556.947 | 0.023
557.235 | 0.017
557.524 | 0.027
557.812 | 0.019
558.101 | 0.016
558.389 | 0.022
558.677 | 0.018
558.966 | 0.023
559.254 | 0.028
559.542 | 0.021
559.83 | 0.024
560.118 | 0.021
560.406 | 0.013
560.693 | 0.032
560.981 | 0.032
561.269 | 0.015
561.556 | 0.017
561.844 | 0.022
562.131 | 0.022
562.419 | 0.027
562.706 | 0.017
562.994 | 0.015
563.281 | 0.027
563.568 | 0.025
563.855 | 0.029
564.142 | 0.015
564.429 | 0.026
564.716 | 0.025
565.003 | 0.03
565.289 | 0.022
565.576 | 0.022
565.863 | 0.021
566.149 | 0.023
566.436 | 0.025
566.722 | 0.017
567.008 | 0.017
567.295 | 0.026
567.581 | 0.016
567.867 | 0.019
568.153 | 0.014
568.439 | 0.021
568.725 | 0.017
569.011 | 0.024
569.297 | 0.019
569.583 | 0.021
569.868 | 0.017
570.154 | 0.02
570.439 | 0.012
570.725 | 0.02
571.01 | 0.019
571.296 | 0.024
571.581 | 0.015
571.866 | 0.027
572.151 | 0.026
572.436 | 0.016
572.721 | 0.017
573.006 | 0.017
573.291 | 0.026
573.576 | 0.023
573.861 | 0.025
574.146 | 0.012
574.43 | 0.023
574.715 | 0.028
574.999 | 0.029
575.284 | 0.03
575.568 | 0.017
575.852 | 0.023
576.137 | 0.054
576.421 | 0.127
576.705 | 0.162
576.989 | 0.247
577.273 | 0.205
577.557 | 0.063
577.841 | 0.03
578.124 | 0.039
578.408 | 0.09
578.692 | 0.138
578.975 | 0.197
579.259 | 0.244
579.542 | 0.093
579.826 | 0.03
580.109 | 0.025
580.392 | 0.029
580.676 | 0.019
580.959 | 0.025
581.242 | 0.028
581.525 | 0.021
581.808 | 0.021
582.091 | 0.018
582.373 | 0.016
582.656 | 0.026
582.939 | 0.021
583.221 | 0.01
583.504 | 0.021
583.786 | 0.02
584.069 | 0.014
584.351 | 0.02
584.633 | 0.013
584.916 | 0.012
585.198 | 0.021
585.48 | 0.021
585.762 | 0.005
586.044 | 0.014
586.326 | 0.022
586.607 | 0.021
586.889 | 0.036
587.171 | 0.025
587.453 | 0.015
587.734 | 0.022
588.016 | 0.028
588.297 | 0.011
588.578 | 0.023
588.86 | 0.024
589.141 | 0.018
589.422 | 0.022
589.703 | 0.026
589.984 | 0.011
590.265 | 0.02
590.546 | 0.021
590.827 | 0.012
591.108 | 0.011
591.389 | 0.024
591.669 | 0.013
591.95 | 0.025
592.23 | 0.023
592.511 | 0.031
592.791 | 0.02
593.071 | 0.022
593.352 | 0.011
593.632 | 0.019
593.912 | 0.021
594.192 | 0.03
594.472 | 0.023
594.752 | 0.026
595.032 | 0.022
595.312 | 0.003
595.592 | 0.019
595.871 | 0.02
596.151 | 0.019
596.43 | 0.008
596.71 | 0.017
596.989 | 0.026
597.269 | 0.02
597.548 | 0.019
597.827 | 0.032
598.106 | 0.024
598.385 | 0.017
598.665 | 0.002
598.944 | 0.008
599.222 | 0.016
599.501 | 0.02
599.78 | 0.018
600.059 | 0.021
600.337 | 0.01
600.616 | 0.013
600.895 | 0.023
601.173 | 0.044
601.451 | 0.019
601.73 | 0.013
602.008 | 0.02
602.286 | 0.008
602.564 | 0.024
602.842 | 0.015
603.12 | 0.011
603.398 | 0.019
603.676 | 0.015
603.954 | 0.013
604.232 | 0.019
604.51 | 0.013
604.787 | 0.007
605.065 | 0.017
605.342 | 0.014
605.62 | 0.01
605.897 | 0.019
606.174 | 0.027
606.452 | 0.014
606.729 | 0.027
607.006 | 0.026
607.283 | 0.019
607.56 | 0.013
607.837 | 0.006
608.114 | 0.019
608.391 | 0.042
608.667 | 0.002
608.944 | 0.017
609.221 | 0.008
609.497 | 0.011
609.774 | 0.007
610.05 | 0.021
610.326 | 0.02
610.603 | 0.019
610.879 | 0.026
611.155 | 0.024
611.431 | 0.014
611.707 | 0.021
611.983 | 0.011
612.259 | 0.029
612.535 | 0.014
612.811 | 0.015
613.087 | 0.015
613.362 | 0.013
613.638 | 0.018
613.913 | 0.017
614.189 | 0.014
614.464 | 0.016
614.74 | 0.029
615.015 | 0.003
615.29 | 0.019
615.565 | 0.024
615.84 | 0.003
616.115 | 0.022
616.39 | 0.022
616.665 | 0.017
616.94 | 0.023
617.215 | 0.022
617.49 | 0.018
617.764 | 0.023
618.039 | 0.027
618.313 | 0.024
618.588 | 0.03
618.862 | 0.017
619.137 | 0.03
619.411 | 0.019
619.685 | 0.02
619.959 | 0.017
620.233 | 0.028
620.507 | 0.013
620.781 | 0.02
621.055 | 0.016
621.329 | 0.019
621.603 | 0.022
621.876 | 0.025
622.15 | 0.021
622.424 | 0.015
622.697 | 0.016
622.971 | 0.011
623.244 | 0.013
623.517 | 0.01
623.791 | 0.018
624.064 | 0.026
624.337 | 0.026
624.61 | 0.014
624.883 | 0.02
625.156 | 0.009
625.429 | 0.02
625.702 | 0.018
625.974 | 0.027
626.247 | 0.002
626.52 | 0.02
626.792 | 0.021
627.065 | 0.017
627.337 | 0.024
627.61 | 0.02
627.882 | 0.02
628.154 | 0.018
628.427 | 0.032
628.699 | 0.021
628.971 | 0.021
629.243 | 0.013
629.515 | 0.029
629.787 | 0.018
630.058 | 0.033
630.33 | 0.023
630.602 | 0.009
630.874 | 0.023
631.145 | 0.031
631.417 | 0.014
631.688 | 0.029
631.959 | 0.021
632.231 | 0.022
632.502 | 0.023
632.773 | 0.032
633.044 | 0.037
633.316 | 0.026
633.587 | 0.005
633.857 | 0.018
634.128 | 0.029
634.399 | 0.033
634.67 | 0.012
634.941 | 0.021
635.211 | 0.028
635.482 | 0.017
635.752 | 0.025
636.023 | 0.012
636.293 | 0.012
636.564 | 0.037
636.834 | 0.021
637.104 | 0.029
637.374 | 0.014
637.644 | 0.024
637.914 | 0.026
638.184 | 0.027
638.454 | 0.021
638.724 | 0.026
638.994 | 0.02
639.264 | 0.025
639.533 | 0.018
639.803 | 0.017
640.072 | 0.022
640.342 | 0.02
640.611 | 0.023
640.881 | 0.022
641.15 | 0.044
641.419 | 0.013
641.688 | 0.025
641.957 | 0.021
642.227 | 0.024
642.496 | 0.031
642.764 | 0.028
643.033 | 0.034
643.302 | 0.021
643.571 | 0.018
643.84 | 0.016
644.108 | 0.03
644.377 | 0.003
644.645 | 0.014
644.914 | 0.018
645.182 | 0.03
645.45 | 0.029
645.719 | 0.031
645.987 | 0.031
646.255 | 0.015
646.523 | 0.016
646.791 | 0.017
647.059 | 0.036
647.327 | 0.019
647.595 | 0.021
647.862 | 0.01
648.13 | 0.016
648.398 | 0.03
648.665 | 0.019
648.933 | 0.041
649.2 | 0.022
649.468 | 0.025
649.735 | 0.03
650.002 | 0.022
650.269 | 0.018
650.536 | 0.011
650.804 | 0.015
651.071 | 0.036
651.338 | 0.018
651.604 | 0.024
651.871 | 0.014
652.138 | 0.021
652.405 | 0.007
652.671 | 0.026
652.938 | 0.02
653.205 | 0.017
653.471 | 0.014
653.737 | 0.017
654.004 | 0.011
654.27 | 0.013
654.536 | 0.021
654.803 | 0.014
655.069 | 0.022
655.335 | 0.016
655.601 | 0.026
655.867 | 0.021
656.132 | 0.028
656.398 | 0.019
656.664 | 0.021
656.93 | 0.017
657.195 | 0.011
657.461 | 0.017
657.726 | 0.023
657.992 | 0.007
658.257 | 0.019
658.523 | 0.026
658.788 | 0.01
659.053 | 0.016
659.318 | 0.017
659.583 | 0.017
659.848 | 0.012
660.113 | 0.022
660.378 | 0.023
660.643 | 0.016
660.908 | 0.007
661.172 | 0.004
661.437 | 0.013
661.702 | 0.02
661.966 | 0.027
662.231 | 0.019
662.495 | 0.023
662.76 | 0.035
663.024 | 0.02
663.288 | 0.01
663.552 | 0.012
663.816 | 0.018
664.08 | 0.027
664.344 | 0.017
664.608 | 0.025
664.872 | 0.015
665.136 | 0.016
665.4 | 0.025
665.664 | 0.024
665.927 | 0.019
666.191 | 0.022
666.454 | 0.009
666.718 | 0.016
666.981 | 0.006
667.244 | 0.012
667.508 | 0.033
667.771 | 0.017
668.034 | 0.027
668.297 | 0.014
668.56 | 0.022
668.823 | 0.007
669.086 | 0.029
669.349 | 0.014
669.612 | 0.017
669.875 | 0.012
670.137 | 0.026
670.4 | 0.024
670.662 | 0.015
670.925 | 0.02
671.187 | 0.028
671.45 | 0.002
671.712 | 0.038
671.974 | 0.029
672.237 | 0.013
672.499 | 0.017
672.761 | 0.022
673.023 | 0.018
673.285 | 0.023
673.547 | 0.026
673.808 | 0.022
674.07 | 0.012
674.332 | 0.023
674.594 | 0.018
674.855 | 0.019
675.117 | 0.015
675.378 | 0.028
675.64 | 0.013
675.901 | 0.008
676.162 | 0.024
676.424 | 0.017
676.685 | 0.028
676.946 | 0.025
677.207 | 0.033
677.468 | 0.031
677.729 | 0.012
677.99 | 0.01
678.251 | 0.028
678.512 | 0.025
678.772 | 0.019
679.033 | 0.033
679.294 | 0.018
679.554 | 0.022
679.815 | 0.026
680.075 | 0.014
680.335 | 0.025
680.596 | 0.015
680.856 | 0.025
681.116 | 0.042
681.376 | 0.028
681.636 | 0.02
681.896 | 0.033
682.156 | 0.031
682.416 | 0.019
682.676 | 0.019
682.936 | 0.031
683.196 | 0.023
683.455 | 0.021
683.715 | 0.02
683.974 | 0.027
684.234 | 0.017
684.493 | 0.008
684.753 | 0.026
685.012 | 0.035
685.271 | 0.012
685.53 | 0.021
685.789 | 0.034
686.049 | 0.026
686.308 | 0.019
686.567 | 0.018
686.825 | 0.033
687.084 | 0.025
687.343 | 0.023
687.602 | 0.028
687.86 | 0.031
688.119 | 0.027
688.378 | 0.032
688.636 | 0.028
688.894 | 0.022
689.153 | 0.026
689.411 | 0.033
689.669 | 0.025
689.928 | 0.027
690.186 | 0.008
690.444 | 0.041
690.702 | 0.043
690.96 | 0.041
691.218 | 0.019
691.476 | 0.028
691.733 | 0.03
691.991 | 0.025
692.249 | 0.016
692.506 | 0.024
692.764 | 0.031
693.021 | 0.033
693.279 | 0.026
693.536 | 0.032
693.794 | 0.013
694.051 | 0.02
694.308 | 0.026
694.565 | 0.03
694.822 | 0.039
695.079 | 0.013
695.336 | 0.024
695.593 | 0.017
695.85 | 0.015
696.107 | 0.017
696.364 | 0.021
696.62 | 0.029
696.877 | 0.029
697.133 | 0.014
697.39 | 0.028
697.646 | 0.029
697.903 | 0.017
698.159 | 0.006
698.415 | 0.016
698.672 | 0.031
698.928 | 0.018
699.184 | 0.031
699.44 | 0.02
699.696 | 0.03
699.952 | 0.008
700.208 | 0.017
700.464 | 0.016
700.719 | 0.039
700.975 | 0.02
701.231 | 0.025
701.486 | 0.038
701.742 | 0.027
701.997 | 0.028
702.253 | 0.025
702.508 | 0.003
702.763 | 0.029
703.019 | 0.025
703.274 | 0.028
703.529 | 0.027
703.784 | 0.023
704.039 | 0.046
704.294 | 0.018
704.549 | 0.017
704.804 | 0.016
705.058 | 0.026
705.313 | 0.009
705.568 | 0.038
705.822 | 0.031
706.077 | 0.018
706.331 | 0.013
706.586 | 0.022
706.84 | 0.038
707.095 | 0.026
707.349 | -0.004
707.603 | 0.019
707.857 | 0.025
708.111 | 0.015
708.365 | 0.023
708.619 | 0.019
708.873 | 0.039
709.127 | 0.017
709.381 | 0.011
709.635 | 0.016
709.888 | 0.034
710.142 | 0.021
710.396 | 0.022
710.649 | 0.022
710.903 | 0.012
711.156 | 0.022
711.409 | 0.032
711.663 | 0.034
711.916 | 0.021
712.169 | 0.02
712.422 | 0.025
712.675 | 0.026
712.928 | 0.042
713.181 | 0.037
713.434 | 0.02
713.687 | 0.038
713.94 | 0.022
714.192 | 0.027
714.445 | 0.005
714.698 | 0.032
714.95 | 0.032
715.203 | 0.051
715.455 | 0.025
715.708 | 0.033
715.96 | 0.046
716.212 | 0.026
716.464 | 0.038
716.717 | 0.032
716.969 | 0.034
717.221 | 0.029
717.473 | 0.006
717.725 | 0.028
717.977 | 0.027
718.228 | 0.041
718.48 | 0.017
718.732 | 0.029
718.984 | 0.039
719.235 | 0.042
719.487 | 0.035
719.738 | 0.02
719.99 | 0.044
720.241 | 0.019
720.492 | 0.034
720.744 | 0.018
720.995 | 0.02
721.246 | 0.025
721.497 | 0.04
721.748 | 0.029
721.999 | 0.022
722.25 | 0.05
722.501 | 0.046
722.752 | 0.028
723.002 | 0.023
723.253 | 0.017
723.504 | 0.045
723.754 | 0.029
724.005 | 0.041
724.255 | 0.037
724.506 | 0.039
724.756 | 0.042
725.006 | 0.018
725.257 | 0.028
725.507 | 0.033
725.757 | 0.028
726.007 | 0.009
726.257 | 0.046
726.507 | 0.011
726.757 | 0.059
727.007 | 0.045
727.257 | 0.019
727.506 | 0.038
727.756 | 0.026
728.006 | 0.044
728.255 | 0.041
728.505 | 0.017
728.754 | 0.029
729.004 | 0.042
729.253 | 0.056
729.502 | 0.062
729.751 | 0.034
730.001 | 0.011
730.25 | 0.025
730.499 | 0.043
730.748 | 0.032
730.997 | 0.026
731.246 | 0.044
731.495 | 0.034
731.743 | 0.04
731.992 | 0.029
732.241 | 0.018
732.489 | 0.034
732.738 | 0.052
732.987 | 0.029
733.235 | 0.031
733.483 | 0.045
733.732 | 0.044
733.98 | 0.016
734.228 | 0.019
734.477 | 0.018
734.725 | 0.044
734.973 | 0.023
735.221 | 0.04
735.469 | 0.043
735.717 | 0.048
735.965 | 0.026
736.212 | 0.029
736.46 | 0.039
736.708 | 0.038
736.955 | 0.055
737.203 | 0.022
737.451 | 0.048
737.698 | 0.037
737.945 | 0.03
738.193 | 0.051
738.44 | 0.044
738.687 | 0.041
738.935 | 0.033
739.182 | 0.029
739.429 | 0.041
739.676 | 0.022
739.923 | 0.037
740.17 | 0.044
740.417 | 0.032
740.663 | 0.026
740.91 | 0.034
741.157 | 0.045
741.403 | 0.03
741.65 | 0.008
741.897 | 0.04
742.143 | 0.02
742.389 | 0.037
742.636 | 0.019
742.882 | 0.026
743.128 | 0.032
743.375 | 0.011
743.621 | 0.055
743.867 | 0.043
744.113 | 0.032
744.359 | 0.028
744.605 | 0.047
744.851 | 0.067
745.097 | 0.052
745.342 | 0.032
745.588 | 0.039
745.834 | 0.029
746.079 | 0.051
746.325 | 0.041
746.57 | 0.047
746.816 | 0.008
747.061 | 0.04
747.307 | 0.036
747.552 | 0.032
747.797 | 0.052
748.042 | 0.027
748.287 | 0.029
748.532 | 0.017
748.777 | 0.039
749.022 | 0.038
749.267 | 0.046
749.512 | 0.028
749.757 | 0.029
750.002 | 0.036
750.246 | 0.044
750.491 | 0.048
750.736 | 0.03
750.98 | 0.052
751.225 | 0.038
751.469 | 0.055
751.713 | 0.051
751.958 | 0.037
752.202 | 0.03
752.446 | 0.032
752.69 | 0.017
752.934 | 0.022
753.178 | 0.033
753.422 | 0.03
753.666 | 0.053
753.91 | 0.02
754.154 | 0.04
754.398 | 0.067
754.641 | 0.017
754.885 | 0.021
755.129 | 0.034
755.372 | 0.016
755.616 | 0.033
755.859 | 0.045
756.103 | 0.055
756.346 | 0.009
756.589 | 0.039
756.832 | 0.033
757.076 | 0.046
757.319 | 0.04
757.562 | 0.04
757.805 | 0.056
758.048 | 0.041
758.291 | 0.034
758.533 | 0.029
758.776 | 0.037
759.019 | 0.048
759.262 | 0.039
759.504 | 0.038
759.747 | 0.033
759.989 | 0.04
760.232 | 0.051
760.474 | 0.03
760.717 | 0.068
760.959 | 0.029
761.201 | 0.029
761.444 | 0.037
761.686 | 0.044
761.928 | 0.035
762.17 | 0.036
762.412 | 0.049
762.654 | 0.044
762.896 | 0.027
763.137 | 0.044
763.379 | 0.061
763.621 | 0.088
763.863 | 0.054
764.104 | 0.036
764.346 | 0.061
764.587 | 0.033
764.829 | 0.035
765.07 | 0.025
765.312 | 0.057
765.553 | 0.054
765.794 | 0.049
766.035 | 0.053
766.276 | 0.061
766.518 | 0.045
766.759 | 0.06
767 | 0.032
767.241 | 0.04
767.481 | 0.038
767.722 | 0.041
767.963 | 0.052
768.204 | 0.047
768.444 | 0.059
768.685 | 0.055
768.926 | 0.026
769.166 | 0.028
769.407 | 0.034
769.647 | 0.068
769.887 | 0.023
770.128 | 0.049
770.368 | 0.043
770.608 | 0.051
770.848 | 0.052
771.088 | 0.031
771.328 | 0.048
771.568 | 0.025
771.808 | 0.047
772.048 | 0.037
772.288 | 0.043
772.528 | 0.045
772.767 | 0.047
773.007 | 0.048
773.247 | 0.048
773.486 | 0.042
773.726 | 0.027
773.965 | 0.054
774.205 | 0.059
774.444 | 0.057
774.683 | 0.067
774.922 | 0.063
775.162 | 0.042
775.401 | 0.068
775.64 | 0.044
775.879 | 0.05
776.118 | 0.045
776.357 | 0.042
776.596 | 0.056
776.834 | 0.061
777.073 | 0.053
777.312 | 0.055
777.551 | 0.041
777.789 | 0.021
778.028 | 0.044
778.266 | 0.054
778.505 | 0.042
778.743 | 0.032
778.982 | 0.027
779.22 | 0.04
779.458 | 0.052
779.696 | 0.044
779.934 | 0.056
780.173 | 0.049
780.411 | 0.048
780.649 | 0.045
780.887 | 0.048
781.124 | 0.057
781.362 | 0.028
781.6 | 0.057
781.838 | 0.034
782.075 | 0.043
782.313 | 0.039
782.551 | 0.016
782.788 | 0.021
783.026 | 0.037
783.263 | 0.045
783.5 | 0.035
783.738 | 0.049
783.975 | 0.027
784.212 | 0.006
784.449 | 0.042
784.686 | 0.031
784.924 | 0.042
785.161 | 0.032
785.397 | 0.048
785.634 | 0.053
785.871 | 0.073
786.108 | 0.038
786.345 | 0.078
786.581 | 0.032
786.818 | 0.061
787.055 | 0.039
787.291 | 0.059
787.528 | 0.075
787.764 | 0.045
788.001 | 0.028
788.237 | 0.069
788.473 | 0.029
788.709 | 0.042
788.946 | 0.041
789.182 | 0.015
789.418 | 0.027
789.654 | 0.055
789.89 | 0.033
790.126 | 0.064
790.362 | 0.067
790.597 | 0.031
790.833 | 0.051
791.069 | 0.026
791.305 | 0.025
791.54 | 0.038
791.776 | 0.059
792.011 | 0.059
792.247 | 0.019
792.482 | 0.04
792.717 | 0.063
792.953 | 0.036
793.188 | 0.062
793.423 | 0.068
793.658 | 0.07
793.894 | 0.074
794.129 | 0.043
794.364 | 0.053
794.599 | 0.058
794.833 | 0.055
795.068 | 0.064
795.303 | 0.058
795.538 | 0.023
795.772 | 0.05
796.007 | 0.05
796.242 | 0.069
796.476 | 0.063
796.711 | 0.03
796.945 | 0.04
797.18 | 0.034
797.414 | 0.017
797.648 | 0.068
797.882 | 0.035
798.117 | 0.041
798.351 | 0.03
798.585 | 0.102
798.819 | 0.041
799.053 | 0.065
799.287 | 0.053
799.521 | 0.072
799.755 | 0.06
799.988 | 0.067
800.222 | 0.04
800.456 | 0.036
800.689 | 0.023
800.923 | 0.036
801.156 | 0.061
801.39 | 0.045
801.623 | 0.065
801.857 | 0.071
802.09 | 0.071
802.323 | 0.062
802.557 | 0.061
802.79 | 0.046
803.023 | 0.027
803.256 | 0.036
803.489 | 0.061
803.722 | 0.023
803.955 | 0.027
804.188 | 0.076
804.421 | 0.043
804.653 | 0.056
804.886 | 0.042
805.119 | 0.02
805.351 | 0.044
805.584 | 0.054
805.816 | 0.029
806.049 | 0.034
806.281 | 0.045
806.514 | 0.075
806.746 | 0.083
806.978 | 0.023
807.211 | 0.017
807.443 | 0.021
807.675 | 0.038
254 3
AbsoluteIrradiance:211
195.562 | 0
195.934 | 0
196.307 | 0
196.679 | 0
197.051 | 0
197.423 | 0
197.795 | 0
198.167 | 0
198.539 | 0
198.911 | 0
199.283 | 0
199.654 | 0
200.026 | 0
200.397 | 0
200.769 | 0
201.14 | 0
201.511 | 0
201.882 | 0
202.253 | 0
202.624 | 0
202.995 | 0
203.366 | 0
203.736 | 0
204.107 | 0
204.477 | 0
204.848 | 3.275
205.218 | 7.195
205.588 | 2.141
205.959 | 2.532
206.329 | 2.764
206.699 | 1.813
207.068 | 2.919
207.438 | 3.119
207.808 | 1.62
208.178 | 2.293
208.547 | 1.957
208.917 | 1.579
209.286 | 3.308
209.655 | 2.442
210.025 | 1.754
210.394 | 1.779
210.763 | 2.038
211.132 | 1.047
211.501 | 0.914
211.869 | 1.114
212.238 | 1.206
212.607 | 1.84
212.975 | 0.893
213.344 | 0.694
213.712 | 0.718
214.08 | 1.096
214.448 | 1.061
214.817 | 0.85
215.185 | 1.415
215.553 | 1.189
215.92 | 0.979
216.288 | 0.991
216.656 | 1.142
217.023 | 0.766
217.391 | 0.661
217.758 | 1.178
218.126 | 0.793
218.493 | 0.796
218.86 | 0.697
219.227 | 0.577
219.594 | 0.575
219.961 | 0.92
220.328 | 0.584
220.695 | 0.714
221.061 | 0.869
221.428 | 0.715
221.795 | 0.789
222.161 | 0.949
222.527 | 0.486
222.893 | 0.848
223.26 | 0.305
223.626 | 0.88
223.992 | 0.59
224.358 | 0.539
224.723 | 0.577
225.089 | 0.27
225.455 | 0.548
225.82 | 0.375
226.186 | 0.31
226.551 | 0.16
226.917 | 0.339
227.282 | 0.33
227.647 | 0.487
228.012 | 0.366
228.377 | 0.3
228.742 | 0.338
229.107 | 0.203
229.471 | 0.322
229.836 | 0.235
230.201 | 0.286
230.565 | 0.278
230.929 | 0.212
231.294 | 0.144
231.658 | 0.201
232.022 | 0.161
232.386 | 0.144
232.75 | -0.005
233.114 | 0.161
233.478 | 0.161
233.842 | 0.183
234.205 | 0.189
234.569 | 0.137
234.932 | 0.104
235.296 | 0.112
235.659 | 0.133
236.022 | 0.122
236.385 | 0.097
236.748 | 0.15
237.111 | 0.084
237.474 | 0.163
237.837 | 0.157
238.2 | 0.124
238.562 | 0.121
238.925 | 0.183
239.287 | 0.139
239.65 | 0.11
240.012 | 0.161
240.374 | 0.14
240.736 | 0.149
241.098 | 0.125
241.46 | 0.153
241.822 | 0.172
242.184 | 0.083
242.546 | 0.146
242.907 | 0.122
243.269 | 0.093
243.63 | 0.098
243.992 | 0.153
244.353 | 0.146
244.714 | 0.124
245.075 | 0.164
245.436 | 0.175
245.797 | 0.147
246.158 | 0.108
246.519 | 0.188
246.879 | 0.205
247.24 | 0.2
247.601 | 0.22
247.961 | 0.252
248.321 | 0.235
248.682 | 0.256
249.042 | 0.234
249.402 | 0.277
249.762 | 0.353
250.122 | 0.359
250.482 | 0.47
250.842 | 0.51
251.201 | 0.84
251.561 | 2.823
251.921 | 8.569
252.28 | 11.224
252.639 | 13.037
252.999 | 18.001
253.358 | 33.415
253.717 | 61.318
254.076 | 28.012
254.435 | 3.684
254.794 | 1.216
255.153 | 0.856
255.511 | 0.713
255.87 | 0.589
256.228 | 0.541
256.587 | 0.485
256.945 | 0.404
257.304 | 0.377
257.662 | 0.389
258.02 | 0.372
258.378 | 0.355
258.736 | 0.301
259.094 | 0.369
259.452 | 0.336
259.809 | 0.292
260.167 | 0.284
260.524 | 0.276
260.882 | 0.244
261.239 | 0.266
261.597 | 0.251
261.954 | 0.248
262.311 | 0.267
262.668 | 0.277
263.025 | 0.258
263.382 | 0.278
263.739 | 0.262
264.095 | 0.296
264.452 | 0.254
264.809 | 0.272
265.165 | 0.291
265.521 | 0.285
265.878 | 0.193
266.234 | 0.164
266.59 | 0.139
266.946 | 0.184
267.302 | 0.159
267.658 | 0.161
268.014 | 0.13
268.37 | 0.145
268.725 | 0.131
269.081 | 0.131
269.436 | 0.135
269.792 | 0.115
270.147 | 0.137
270.502 | 0.148
270.857 | 0.158
271.212 | 0.091
271.567 | 0.058
271.922 | 0.116
272.277 | 0.088
272.632 | 0.105
272.987 | 0.103
273.341 | 0.092
273.696 | 0.083
274.05 | 0.046
274.404 | 0.064
274.759 | 0.1
275.113 | 0.137
275.467 | 0.065
275.821 | 0.068
276.175 | 0.082
276.529 | 0.098
276.882 | 0.12
277.236 | 0.079
277.59 | 0.095
277.943 | 0.051
278.297 | 0.115
278.65 | 0.052
279.003 | 0.038
279.356 | 0.069
279.709 | 0.082
280.062 | 0.078
280.415 | 0.105
280.768 | 0.074
281.121 | 0.108
281.474 | 0.07
281.826 | 0.07
282.179 | 0.063
282.531 | 0.043
282.884 | 0.063
283.236 | 0.073
283.588 | 0.1
283.94 | 0.044
284.292 | 0.067
284.644 | 0.037
284.996 | 0.056
285.348 | 0.061
285.7 | 0.084
286.051 | 0.058
286.403 | 0.055
286.754 | 0.058
287.106 | 0.083
287.457 | 0.072
287.808 | 0.07
288.159 | 0.089
288.51 | 0.091
288.861 | 0.09
289.212 | 0.088
289.563 | 0.07
289.914 | 0.074
290.264 | 0.037
290.615 | 0.052
290.966 | 0.066
291.316 | 0.036
291.666 | 0.071
292.017 | 0.061
292.367 | 0.043
292.717 | 0.039
293.067 | 0.042
293.417 | 0.049
293.767 | 0.046
294.116 | 0.063
294.466 | 0.046
294.816 | 0.06
295.165 | 0.051
295.515 | 0.066
295.864 | 0.109
296.213 | 0.15
296.562 | 0.248
296.911 | 0.162
297.261 | 0.055
297.609 | 0.039
297.958 | 0.015
298.307 | 0.038
298.656 | 0.034
299.004 | 0.032
299.353 | 0.003
299.701 | 0.031
300.05 | 0.025
300.398 | 0.04
300.746 | 0.047
301.095 | 0.066
301.443 | 0.081
301.791 | 0.133
302.138 | 0.164
302.486 | 0.066
302.834 | 0.031
303.182 | 0.016
303.529 | 0.036
303.877 | 0.032
304.224 | 0.035
304.572 | 0.034
304.919 | 0.014
305.266 | 0.036
305.613 | 0.024
305.96 | 0.043
306.307 | 0.026
306.654 | 0.049
307.001 | 0.043
307.347 | 0.033
307.694 | 0.017
308.04 | 0.022
308.387 | 0.04
308.733 | 0.031
309.08 | 0.029
309.426 | 0.046
309.772 | 0.033
310.118 | 0.037
310.464 | 0.027
310.81 | 0.04
311.156 | 0.054
311.501 | 0.142
311.847 | 0.288
312.193 | 0.601
312.538 | 0.797
312.883 | 0.784
313.229 | 0.525
313.574 | 0.107
313.919 | 0.041
314.264 | 0.039
314.609 | 0.038
314.954 | 0.029
315.299 | 0.028
315.644 | 0.021
315.988 | 0.034
316.333 | 0.036
316.678 | 0.027
317.022 | 0.026
317.366 | 0.026
317.711 | 0.027
318.055 | 0.021
318.399 | 0.028
318.743 | 0.029
319.087 | 0.025
319.431 | 0.033
319.775 | 0.025
320.118 | 0.04
320.462 | 0.026
320.806 | 0.026
321.149 | 0.031
321.493 | 0.027
321.836 | 0.021
322.179 | 0.011
322.522 | 0.026
322.865 | 0.031
323.208 | 0.024
323.551 | 0.037
323.894 | 0.021
324.237 | 0.019
324.58 | 0.023
324.922 | 0.026
325.265 | 0.032
325.607 | 0.027
325.95 | 0.03
326.292 | 0.013
326.634 | 0.02
326.976 | 0.019
327.318 | 0.016
327.66 | 0.028
328.002 | 0.013
328.344 | 0.026
328.686 | 0.028
329.027 | 0.013
329.369 | 0.013
329.711 | 0.026
330.052 | 0.025
330.393 | 0.015
330.735 | 0.018
331.076 | 0.023
331.417 | 0.017
331.758 | 0.023
332.099 | 0.009
332.44 | 0.025
332.78 | 0.017
333.121 | 0.033
333.462 | 0.051
333.802 | 0.086
334.143 | 0.071
334.483 | 0.03
334.823 | 0.016
335.164 | 0.025
335.504 | 0.013
335.844 | 0.027
336.184 | 0.015
336.524 | 0.014
336.864 | 0.017
337.203 | 0.018
337.543 | 0.033
337.883 | 0.022
338.222 | 0.034
338.562 | 0.016
338.901 | 0.027
339.24 | 0.01
339.58 | 0.019
339.919 | 0.015
340.258 | 0.025
340.597 | 0.018
340.936 | 0.014
341.274 | 0.022
341.613 | 0.026
341.952 | 0.023
342.29 | 0.021
342.629 | 0.027
342.967 | 0.018
343.306 | 0.02
343.644 | 0.021
343.982 | 0.019
344.32 | 0.021
344.658 | 0.013
344.996 | 0.019
345.334 | 0.017
345.672 | 0.027
346.01 | 0.018
346.347 | 0.037
346.685 | 0.022
347.022 | 0.027
347.36 | 0.028
347.697 | 0.024
348.034 | 0.022
348.371 | 0.023
348.708 | 0.018
349.045 | 0.03
349.382 | 0.024
349.719 | 0.016
350.056 | 0.02
350.393 | 0.016
350.729 | 0.031
351.066 | 0.02
351.402 | 0.018
351.739 | 0.033
352.075 | 0.026
352.411 | 0.018
352.747 | 0.015
353.083 | 0.02
353.419 | 0.019
353.755 | 0.018
354.091 | 0.027
354.427 | 0.024
354.763 | 0.023
355.098 | 0.029
355.434 | 0.017
355.769 | 0.035
356.104 | 0.011
356.44 | 0.016
356.775 | 0.026
357.11 | 0.023
357.445 | 0.02
357.78 | 0.015
358.115 | 0.027
358.45 | 0.023
358.785 | 0.026
359.119 | 0.016
359.454 | 0.019
359.788 | 0.023
360.123 | 0.023
360.457 | 0.014
360.791 | 0.024
361.126 | 0.012
361.46 | 0.017
361.794 | 0.023
362.128 | 0.034
362.462 | 0.035
362.795 | 0.029
363.129 | 0.027
363.463 | 0.021
363.796 | 0.054
364.13 | 0.191
364.463 | 0.527
364.797 | 0.741
365.13 | 0.595
365.463 | 0.253
365.796 | 0.164
366.129 | 0.148
366.462 | 0.09
366.795 | 0.036
367.128 | 0.037
367.461 | 0.02
367.793 | 0.032
368.126 | 0.029
368.458 | 0.026
368.791 | 0.028
369.123 | 0.025
369.455 | 0.02
369.787 | 0.023
370.12 | 0.022
370.452 | 0.013
370.784 | 0.019
371.115 | 0.023
371.447 | 0.013
371.779 | 0.019
372.111 | 0.025
372.442 | 0.02
372.774 | 0.03
373.105 | 0.023
373.436 | 0.023
373.768 | 0.022
374.099 | 0.015
374.43 | 0.029
374.761 | 0.019
375.092 | 0.021
375.423 | 0.02
375.754 | 0.021
376.084 | 0.019
376.415 | 0.018
376.745 | 0.02
377.076 | 0.027
377.406 | 0.024
377.737 | 0.037
378.067 | 0.026
378.397 | 0.019
378.727 | 0.019
379.057 | 0.018
379.387 | 0.018
379.717 | 0.02
380.047 | 0.024
380.377 | 0.024
380.706 | 0.024
381.036 | 0.025
381.365 | 0.033
381.695 | 0.019
382.024 | 0.016
382.353 | 0.009
382.682 | 0.02
383.012 | 0.017
383.341 | 0.029
383.67 | 0.023
383.998 | 0.011
384.327 | 0.023
384.656 | 0.014
384.985 | 0.014
385.313 | 0.01
385.642 | 0.017
385.97 | 0.017
386.298 | 0.016
386.627 | 0.015
386.955 | 0.027
387.283 | 0.014
387.611 | 0.014
387.939 | 0.019
388.267 | 0.015
388.595 | 0.026
388.922 | 0.008
389.25 | 0.033
389.578 | 0.019
389.905 | 0.024
390.233 | 0.022
390.56 | 0.027
390.887 | 0.027
391.214 | 0.023
391.542 | 0.014
391.869 | 0.016
392.196 | 0.015
392.522 | 0.01
392.849 | 0.017
393.176 | 0.015
393.503 | 0.018
393.829 | 0.016
394.156 | 0.018
394.482 | 0.021
394.809 | 0.024
395.135 | 0.023
395.461 | 0.013
395.787 | 0.024
396.113 | 0.011
396.439 | 0.014
396.765 | 0.015
397.091 | 0.014
397.417 | 0.013
397.742 | 0.02
398.068 | 0.025
398.394 | 0.021
398.719 | 0.018
399.044 | 0.017
399.37 | 0.027
399.695 | 0.018
400.02 | 0.023
400.345 | 0.027
400.67 | 0.022
400.995 | 0.029
401.32 | 0.02
401.645 | 0.022
401.969 | 0.014
402.294 | 0.019
402.618 | 0.038
402.943 | 0.021
403.267 | 0.031
403.592 | 0.114
403.916 | 0.379
404.24 | 0.787
404.564 | 1.04
404.888 | 0.608
405.212 | 0.117
405.536 | 0.023
405.86 | 0.026
406.183 | 0.023
406.507 | 0.027
406.831 | 0.038
407.154 | 0.058
407.478 | 0.092
407.801 | 0.094
408.124 | 0.042
408.447 | 0.022
408.77 | 0.024
409.093 | 0.024
409.416 | 0.017
409.739 | 0.026
410.062 | 0.011
410.385 | 0.028
410.708 | 0.023
411.03 | 0.026
411.353 | 0.013
411.675 | 0.025
411.997 | 0.022
412.32 | 0.014
412.642 | 0.024
412.964 | 0.019
413.286 | 0.017
413.608 | 0.02
413.93 | 0.021
414.252 | 0.011
414.574 | 0.015
414.895 | 0.015
415.217 | 0.025
415.538 | 0.015
415.86 | 0.016
416.181 | 0.027
416.503 | 0.017
416.824 | 0.017
417.145 | 0.015
417.466 | 0.019
417.787 | 0.023
418.108 | 0.012
418.429 | 0.015
418.75 | 0.019
419.07 | 0.018
419.391 | 0.019
419.712 | 0.018
420.032 | 0.017
420.353 | 0.018
420.673 | 0.014
420.993 | 0.021
421.313 | 0.022
421.633 | 0.02
421.953 | 0.013
422.273 | 0.017
422.593 | 0.021
422.913 | 0.015
423.233 | 0.014
423.553 | 0.013
423.872 | 0.014
424.192 | 0.015
424.511 | 0.012
424.83 | 0.021
425.15 | 0.021
425.469 | 0.019
425.788 | 0.024
426.107 | 0.018
426.426 | 0.024
426.745 | 0.027
427.064 | 0.019
427.383 | 0.023
427.702 | 0.009
428.02 | 0.024
428.339 | 0.018
428.657 | 0.019
428.976 | 0.022
429.294 | 0.017
429.612 | 0.013
429.93 | 0.022
430.248 | 0.027
430.567 | 0.02
430.884 | 0.02
431.202 | 0.014
431.52 | 0.015
431.838 | 0.015
432.156 | 0.018
432.473 | 0.022
432.791 | 0.027
433.108 | 0.021
433.426 | 0.038
433.743 | 0.04
434.06 | 0.045
434.377 | 0.074
434.694 | 0.196
435.011 | 0.636
435.328 | 1.247
435.645 | 2.16
435.962 | 2.355
436.279 | 0.753
436.595 | 0.092
436.912 | 0.032
437.228 | 0.025
437.545 | 0.02
437.861 | 0.02
438.177 | 0.018
438.493 | 0.017
438.81 | 0.027
439.126 | 0.025
439.442 | 0.029
439.757 | 0.014
440.073 | 0.022
440.389 | 0.022
440.705 | 0.032
441.02 | 0.025
441.336 | 0.022
441.651 | 0.024
441.967 | 0.017
442.282 | 0.014
442.597 | 0.023
442.912 | 0.03
443.227 | 0.026
443.543 | 0.026
443.857 | 0.03
444.172 | 0.024
444.487 | 0.026
444.802 | 0.032
445.116 | 0.028
445.431 | 0.015
445.746 | 0.015
446.06 | 0.018
446.374 | 0.016
446.689 | 0.023
447.003 | 0.018
447.317 | 0.015
447.631 | 0.008
447.945 | 0.019
448.259 | 0.021
448.573 | 0.022
448.887 | 0.01
449.2 | 0.027
449.514 | 0.021
449.827 | 0.018
450.141 | 0.018
450.454 | 0.013
450.768 | 0.019
451.081 | 0.019
451.394 | 0.014
451.707 | 0.014
452.02 | 0.022
452.333 | 0.015
452.646 | 0.013
452.959 | 0.012
453.272 | 0.015
453.584 | 0.013
453.897 | 0.013
454.21 | 0.019
454.522 | 0.013
454.834 | 0.011
455.147 | 0.014
455.459 | 0.005
455.771 | 0.006
456.083 | 0.011
456.395 | 0.012
456.707 | 0.015
457.019 | 0.016
457.331 | 0.017
457.643 | 0.012
457.954 | 0.017
458.266 | 0.014
458.577 | 0.014
458.889 | 0.011
459.2 | 0.016
459.512 | 0.009
459.823 | 0.013
460.134 | 0.015
460.445 | 0.02
460.756 | 0.015
461.067 | 0.016
461.378 | 0.011
461.689 | 0.011
461.999 | 0.025
462.31 | 0.021
462.621 | 0.008
462.931 | 0.015
463.242 | 0.016
463.552 | 0.019
463.862 | 0.01
464.172 | 0.013
464.483 | 0.024
464.793 | 0.012
465.103 | 0.018
465.413 | 0.02
465.722 | 0.014
466.032 | 0.018
466.342 | 0.013
466.652 | 0.016
466.961 | 0.018
467.271 | 0.02
467.58 | 0.015
467.889 | 0.019
468.199 | 0.017
468.508 | 0.015
468.817 | 0.019
469.126 | 0.013
469.435 | 0.023
469.744 | 0.015
470.053 | 0.014
470.362 | 0.013
470.67 | 0.011
470.979 | 0.019
471.287 | 0.009
471.596 | 0.013
471.904 | 0.021
472.213 | 0.008
472.521 | 0.02
472.829 | 0.017
473.137 | 0.017
473.445 | 0.011
473.753 | 0.026
474.061 | 0.01
474.369 | 0.016
474.677 | 0.02
474.985 | 0.01
475.292 | 0.007
475.6 | 0.021
475.907 | 0.017
476.215 | 0.014
476.522 | 0.024
476.829 | 0.019
477.137 | 0.021
477.444 | 0.012
477.751 | 0.019
478.058 | 0.01
478.365 | 0.015
478.671 | 0.023
478.978 | 0.02
479.285 | 0.012
479.592 | 0.011
479.898 | 0.01
480.205 | 0.015
480.511 | 0.004
480.817 | 0.022
481.124 | 0.015
481.43 | 0.017
481.736 | 0.025
482.042 | 0.015
482.348 | 0.012
482.654 | 0.022
482.96 | 0.019
483.266 | 0.016
483.571 | 0.025
483.877 | 0.007
484.182 | 0.008
484.488 | 0.017
484.793 | 0.006
485.099 | 0.015
485.404 | 0.016
485.709 | 0.019
486.014 | 0.017
486.319 | 0.02
486.624 | 0.011
486.929 | 0.01
487.234 | 0.004
487.539 | 0.019
487.844 | 0.02
488.148 | 0.017
488.453 | 0.023
488.757 | 0.018
489.062 | 0.012
489.366 | 0.019
489.67 | 0.014
489.975 | 0.014
490.279 | 0.014
490.583 | 0.017
490.887 | 0.016
491.191 | 0.02
491.495 | 0.026
491.798 | 0.032
492.102 | 0.022
492.406 | 0.013
492.709 | 0.018
493.013 | 0.023
493.316 | 0.017
493.62 | 0.023
493.923 | 0.014
494.226 | 0.022
494.529 | 0.016
494.832 | 0.018
495.135 | 0.015
495.438 | 0.022
495.741 | 0.028
496.044 | 0.011
496.347 | 0.011
496.649 | 0.023
496.952 | 0.015
497.255 | 0.016
497.557 | 0.017
497.859 | 0.016
498.162 | 0.005
498.464 | 0.023
498.766 | 0.016
499.068 | 0.015
499.37 | 0.016
499.672 | 0.02
499.974 | 0.014
500.276 | 0.017
500.578 | 0.017
500.879 | 0.008
501.181 | 0.022
501.483 | 0.015
501.784 | 0.022
502.085 | 0.012
502.387 | 0.01
502.688 | 0.01
502.989 | 0.02
503.29 | 0.008
503.591 | 0.009
503.892 | 0.007
504.193 | 0.024
504.494 | 0.016
504.795 | 0.008
505.096 | 0.024
505.396 | 0.012
505.697 | 0.004
505.997 | 0.017
506.298 | 0.017
506.598 | 0.015
506.898 | 0.023
507.199 | 0.012
507.499 | 0.017
507.799 | 0.012
508.099 | 0.012
508.399 | 0.016
508.699 | 0.014
508.998 | 0.014
509.298 | 0.02
509.598 | 0.012
509.897 | 0.019
510.197 | 0.011
510.496 | 0.015
510.796 | 0.012
511.095 | 0.013
511.394 | 0.015
511.693 | 0.025
511.992 | 0.016
512.291 | 0.018
512.59 | 0.019
512.889 | 0.018
513.188 | 0.021
513.487 | 0.016
513.786 | 0.022
514.084 | 0.019
514.383 | 0.022
514.681 | 0.026
514.98 | 0.015
515.278 | 0.024
515.576 | 0.024
515.874 | 0.02
516.173 | 0.012
516.471 | 0.023
516.769 | 0.019
517.067 | 0.016
517.364 | 0.012
517.662 | 0.006
517.96 | 0.016
518.257 | 0.015
518.555 | 0.013
518.853 | 0.011
519.15 | 0.011
519.447 | 0.015
519.745 | 0.016
520.042 | 0.024
520.339 | 0.012
520.636 | 0.021
520.933 | 0.023
521.23 | 0.009
521.527 | 0.013
521.824 | 0.019
522.121 | 0.014
522.417 | 0.01
522.714 | 0.008
523.01 | 0.019
523.307 | 0.013
523.603 | 0.021
523.9 | 0.018
524.196 | 0.02
524.492 | 0.006
524.788 | 0.023
525.084 | 0.017
525.38 | 0.024
525.676 | 0.01
525.972 | 0.017
526.268 | 0.018
526.564 | 0.019
526.859 | 0.015
527.155 | 0.014
527.45 | 0.019
527.746 | 0.02
528.041 | 0.023
528.336 | 0.022
528.632 | 0.009
528.927 | 0.012
529.222 | 0.02
529.517 | 0.023
529.812 | 0.015
530.107 | 0.011
530.402 | 0.018
530.696 | 0.025
530.991 | 0.024
531.286 | 0.017
531.58 | 0.019
531.875 | 0.016
532.169 | 0.019
532.463 | 0.013
532.758 | 0.009
533.052 | 0.015
533.346 | 0.01
533.64 | 0.018
533.934 | 0.021
534.228 | 0.023
534.522 | 0.021
534.816 | 0.019
535.11 | 0.012
535.403 | 0.024
535.697 | 0.021
535.99 | 0.02
536.284 | 0.015
536.577 | 0.026
536.871 | 0.018
537.164 | 0.015
537.457 | 0.018
537.75 | 0.013
538.043 | 0.018
538.336 | 0.021
538.629 | 0.015
538.922 | 0.018
539.215 | 0.008
539.507 | 0.019
539.8 | 0.007
540.093 | 0.023
540.385 | 0.014
540.678 | 0.014
540.97 | 0.006
541.262 | 0.024
541.554 | 0.015
541.847 | 0.016
542.139 | 0.022
542.431 | 0.021
542.723 | 0.015
543.015 | 0.024
543.307 | 0.029
543.598 | 0.025
543.89 | 0.025
544.182 | 0.02
544.473 | 0.031
544.765 | 0.046
545.056 | 0.081
545.347 | 0.344
545.639 | 0.704
545.93 | 0.959
546.221 | 1.471
546.512 | 0.756
546.803 | 0.119
547.094 | 0.041
547.385 | 0.029
547.676 | 0.026
547.967 | 0.036
548.257 | 0.021
548.548 | 0.023
548.839 | 0.011
549.129 | 0.018
549.419 | 0.025
549.71 | 0.021
550 | 0.026
550.29 | 0.021
550.58 | 0.013
550.871 | 0.017
551.161 | 0.018
551.45 | 0.015
551.74 | 0.018
552.03 | 0.015
552.32 | 0.022
552.61 | 0.024
552.899 | 0.029
553.189 | 0.016
553.478 | 0.018
553.768 | 0.016
554.057 | 0.032
554.346 | 0.027
554.635 | 0.023
554.925 | 0.022
555.214 | 0.026
555.503 | 0.014
555.792 | 0.021
556.081 | 0.011
556.369 | 0.018
556.658 | 0.015
556.947 | 0.022
557.235 | 0.012
557.524 | 0.032
557.812 | 0.017
558.101 | 0.022
558.389 | 0.024
558.677 | 0.018
558.966 | 0.027
559.254 | 0.03
559.542 | 0.019
559.83 | 0.018
560.118 | 0.024
560.406 | 0.017
560.693 | 0.032
560.981 | 0.028
561.269 | 0.018
561.556 | 0.019
561.844 | 0.021
562.131 | 0.02
562.419 | 0.027
562.706 | 0.021
562.994 | 0.017
563.281 | 0.028
563.568 | 0.022
563.855 | 0.03
564.142 | 0.015
564.429 | 0.024
564.716 | 0.024
565.003 | 0.027
565.289 | 0.023
565.576 | 0.027
565.863 | 0.026
566.149 | 0.025
566.436 | 0.029
566.722 | 0.015
567.008 | 0.018
567.295 | 0.027
567.581 | 0.016
567.867 | 0.018
568.153 | 0.012
568.439 | 0.023
568.725 | 0.018
569.011 | 0.025
569.297 | 0.021
569.583 | 0.022
569.868 | 0.021
570.154 | 0.019
570.439 | 0.014
570.725 | 0.021
571.01 | 0.019
571.296 | 0.029
571.581 | 0.018
571.866 | 0.029
572.151 | 0.024
572.436 | 0.016
572.721 | 0.022
573.006 | 0.021
573.291 | 0.029
573.576 | 0.024
573.861 | 0.027
574.146 | 0.016
574.43 | 0.025
574.715 | 0.03
574.999 | 0.031
575.284 | 0.032
575.568 | 0.022
575.852 | 0.024
576.137 | 0.054
576.421 | 0.13
576.705 | 0.168
576.989 | 0.249
577.273 | 0.207
577.557 | 0.063
577.841 | 0.032
578.124 | 0.046
578.408 | 0.09
578.692 | 0.135
578.975 | 0.199
579.259 | 0.245
579.542 | 0.1
579.826 | 0.028
580.109 | 0.023
580.392 | 0.028
580.676 | 0.019
580.959 | 0.023
581.242 | 0.028
581.525 | 0.024
581.808 | 0.023
582.091 | 0.018
582.373 | 0.013
582.656 | 0.024
582.939 | 0.02
583.221 | 0.015
583.504 | 0.024
583.786 | 0.023
584.069 | 0.017
584.351 | 0.021
584.633 | 0.016
584.916 | 0.014
585.198 | 0.02
585.48 | 0.024
585.762 | 0.012
586.044 | 0.017
586.326 | 0.025
586.607 | 0.019
586.889 | 0.036
587.171 | 0.028
587.453 | 0.018
587.734 | 0.025
588.016 | 0.027
588.297 | 0.013
588.578 | 0.021
588.86 | 0.025
589.141 | 0.016
589.422 | 0.024
589.703 | 0.029
589.984 | 0.013
590.265 | 0.015
590.546 | 0.021
590.827 | 0.012
591.108 | 0.013
591.389 | 0.026
591.669 | 0.016
591.95 | 0.024
592.23 | 0.027
592.511 | 0.03
592.791 | 0.022
593.071 | 0.03
593.352 | 0.012
593.632 | 0.026
593.912 | 0.024
594.192 | 0.031
594.472 | 0.026
594.752 | 0.025
595.032 | 0.023
595.312 | 0.014
595.592 | 0.025
595.871 | 0.023
596.151 | 0.022
596.43 | 0.004
596.71 | 0.014
596.989 | 0.034
597.269 | 0.021
597.548 | 0.024
597.827 | 0.031
598.106 | 0.026
598.385 | 0.021
598.665 | 0.007
598.944 | 0.01
599.222 | 0.017
599.501 | 0.021
599.78 | 0.021
600.059 | 0.022
600.337 | 0.02
600.616 | 0.016
600.895 | 0.021
601.173 | 0.045
601.451 | 0.018
601.73 | 0.019
602.008 | 0.029
602.286 | 0.006
602.564 | 0.023
602.842 | 0.02
603.12 | 0.015
603.398 | 0.016
603.676 | 0.023
603.954 | 0.015
604.232 | 0.02
604.51 | 0.014
604.787 | 0.009
605.065 | 0.022
605.342 | 0.013
605.62 | 0.011
605.897 | 0.019
606.174 | 0.024
606.452 | 0.014
606.729 | 0.026
607.006 | 0.026
607.283 | 0.022
607.56 | 0.012
607.837 | 0.008
608.114 | 0.021
608.391 | 0.046
608.667 | 0.005
608.944 | 0.02
609.221 | 0.012
609.497 | 0.012
609.774 | 0.006
610.05 | 0.024
610.326 | 0.02
610.603 | 0.024
610.879 | 0.032
611.155 | 0.023
611.431 | 0.013
611.707 | 0.024
611.983 | 0.013
612.259 | 0.026
612.535 | 0.015
612.811 | 0.018
613.087 | 0.02
613.362 | 0.014
613.638 | 0.017
613.913 | 0.015
614.189 | 0.016
614.464 | 0.025
614.74 | 0.026
615.015 | 0.005
615.29 | 0.021
615.565 | 0.023
615.84 | 0.003
616.115 | 0.025
616.39 | 0.022
616.665 | 0.024
616.94 | 0.019
617.215 | 0.027
617.49 | 0.022
617.764 | 0.029
618.039 | 0.023
618.313 | 0.024
618.588 | 0.03
618.862 | 0.02
619.137 | 0.029
619.411 | 0.019
619.685 | 0.023
619.959 | 0.016
620.233 | 0.032
620.507 | 0.016
620.781 | 0.021
621.055 | 0.015
621.329 | 0.021
621.603 | 0.021
621.876 | 0.026
622.15 | 0.02
622.424 | 0.018
622.697 | 0.015
622.971 | 0.015
623.244 | 0.014
623.517 | 0.012
623.791 | 0.023
624.064 | 0.026
624.337 | 0.029
624.61 | 0.017
624.883 | 0.023
625.156 | 0.011
625.429 | 0.021
625.702 | 0.012
625.974 | 0.03
626.247 | 0.006
626.52 | 0.023
626.792 | 0.02
627.065 | 0.019
627.337 | 0.023
627.61 | 0.022
627.882 | 0.015
628.154 | 0.021
628.427 | 0.033
628.699 | 0.027
628.971 | 0.021
629.243 | 0.009
629.515 | 0.029
629.787 | 0.021
630.058 | 0.029
630.33 | 0.03
630.602 | 0.006
630.874 | 0.027
631.145 | 0.031
631.417 | 0.015
631.688 | 0.025
631.959 | 0.018
632.231 | 0.018
632.502 | 0.027
632.773 | 0.027
633.044 | 0.037
633.316 | 0.027
633.587 | 0.01
633.857 | 0.014
634.128 | 0.033
634.399 | 0.037
634.67 | 0.013
634.941 | 0.021
635.211 | 0.028
635.482 | 0.014
635.752 | 0.027
636.023 | 0.008
636.293 | 0.017
636.564 | 0.035
636.834 | 0.026
637.104 | 0.031
637.374 | 0.021
637.644 | 0.024
637.914 | 0.027
638.184 | 0.027
638.454 | 0.019
638.724 | 0.035
638.994 | 0.025
639.264 | 0.033
639.533 | 0.02
639.803 | 0.018
640.072 | 0.023
640.342 | 0.021
640.611 | 0.029
640.881 | 0.02
641.15 | 0.046
641.419 | 0.013
641.688 | 0.025
641.957 | 0.016
642.227 | 0.027
642.496 | 0.033
642.764 | 0.025
643.033 | 0.035
643.302 | 0.011
643.571 | 0.016
643.84 | 0.016
644.108 | 0.038
644.377 | 0.004
644.645 | 0.013
644.914 | 0.021
645.182 | 0.031
645.45 | 0.027
645.719 | 0.036
645.987 | 0.027
646.255 | 0.019
646.523 | 0.021
646.791 | 0.019
647.059 | 0.04
647.327 | 0.021
647.595 | 0.022
647.862 | 0.007
648.13 | 0.018
648.398 | 0.029
648.665 | 0.017
648.933 | 0.039
649.2 | 0.027
649.468 | 0.023
649.735 | 0.027
650.002 | 0.024
650.269 | 0.023
650.536 | 0.013
650.804 | 0.017
651.071 | 0.041
651.338 | 0.019
651.604 | 0.026
651.871 | 0.017
652.138 | 0.024
652.405 | 0.007
652.671 | 0.028
652.938 | 0.026
653.205 | 0.025
653.471 | 0.011
653.737 | 0.027
654.004 | 0.013
654.27 | 0.018
654.536 | 0.022
654.803 | 0.017
655.069 | 0.021
655.335 | 0.018
655.601 | 0.027
655.867 | 0.024
656.132 | 0.027
656.398 | 0.022
656.664 | 0.023
656.93 | 0.016
657.195 | 0.014
657.461 | 0.021
657.726 | 0.021
657.992 | 0.006
658.257 | 0.024
658.523 | 0.024
658.788 | 0.014
659.053 | 0.02
659.318 | 0.023
659.583 | 0.019
659.848 | 0.012
660.113 | 0.027
660.378 | 0.026
660.643 | 0.021
660.908 | 0.011
661.172 | 0.002
661.437 | 0.016
661.702 | 0.021
661.966 | 0.026
662.231 | 0.022
662.495 | 0.023
662.76 | 0.04
663.024 | 0.02
663.288 | 0.013
663.552 | 0.013
663.816 | 0.023
664.08 | 0.033
664.344 | 0.018
664.608 | 0.026
664.872 | 0.015
665.136 | 0.014
665.4 | 0.024
665.664 | 0.028
665.927 | 0.026
666.191 | 0.023
666.454 | 0.013
666.718 | 0.021
666.981 | 0.009
667.244 | 0.013
667.508 | 0.031
667.771 | 0.019
668.034 | 0.024
668.297 | 0.015
668.56 | 0.022
668.823 | 0.007
669.086 | 0.033
669.349 | 0.015
669.612 | 0.02
669.875 | 0.015
670.137 | 0.026
670.4 | 0.026
670.662 | 0.019
670.925 | 0.024
671.187 | 0.034
671.45 | 0.001
671.712 | 0.037
671.974 | 0.027
672.237 | 0.01
672.499 | 0.023
672.761 | 0.019
673.023 | 0.02
673.285 | 0.025
673.547 | 0.026
673.808 | 0.025
674.07 | 0.016
674.332 | 0.023
674.594 | 0.019
674.855 | 0.023
675.117 | 0.015
675.378 | 0.028
675.64 | 0.018
675.901 | 0.003
676.162 | 0.026
676.424 | 0.02
676.685 | 0.027
676.946 | 0.027
677.207 | 0.038
677.468 | 0.032
677.729 | 0.014
677.99 | 0.01
678.251 | 0.029
678.512 | 0.02
678.772 | 0.03
679.033 | 0.037
679.294 | 0.02
679.554 | 0.026
679.815 | 0.023
680.075 | 0.017
680.335 | 0.025
680.596 | 0.012
680.856 | 0.027
681.116 | 0.042
681.376 | 0.03
681.636 | 0.021
681.896 | 0.037
682.156 | 0.032
682.416 | 0.018
682.676 | 0.02
682.936 | 0.031
683.196 | 0.022
683.455 | 0.023
683.715 | 0.019
683.974 | 0.024
684.234 | 0.02
684.493 | 0.013
684.753 | 0.026
685.012 | 0.031
685.271 | 0.015
685.53 | 0.02
685.789 | 0.037
686.049 | 0.027
686.308 | 0.017
686.567 | 0.023
686.825 | 0.034
687.084 | 0.026
687.343 | 0.026
687.602 | 0.035
687.86 | 0.03
688.119 | 0.024
688.378 | 0.033
688.636 | 0.029
688.894 | 0.025
689.153 | 0.025
689.411 | 0.036
689.669 | 0.024
689.928 | 0.038
690.186 | 0.012
690.444 | 0.043
690.702 | 0.045
690.96 | 0.041
691.218 | 0.017
691.476 | 0.032
691.733 | 0.031
691.991 | 0.028
692.249 | 0.02
692.506 | 0.021
692.764 | 0.025
693.021 | 0.038
693.279 | 0.028
693.536 | 0.027
693.794 | 0.008
694.051 | 0.024
694.308 | 0.033
694.565 | 0.035
694.822 | 0.044
695.079 | 0.017
695.336 | 0.026
695.593 | 0.016
695.85 | 0.021
696.107 | 0.022
696.364 | 0.023
696.62 | 0.03
696.877 | 0.031
697.133 | 0.011
697.39 | 0.035
697.646 | 0.031
697.903 | 0.019
698.159 | 0.011
698.415 | 0.02
698.672 | 0.03
698.928 | 0.023
699.184 | 0.035
699.44 | 0.025
699.696 | 0.037
699.952 | 0.006
700.208 | 0.027
700.464 | 0.022
700.719 | 0.043
700.975 | 0.028
701.231 | 0.026
701.486 | 0.034
701.742 | 0.028
701.997 | 0.027
702.253 | 0.034
702.508 | 0.012
702.763 | 0.031
703.019 | 0.03
703.274 | 0.027
703.529 | 0.028
703.784 | 0.024
704.039 | 0.055
704.294 | 0.021
704.549 | 0.019
704.804 | 0.017
705.058 | 0.037
705.313 | 0.02
705.568 | 0.043
705.822 | 0.038
706.077 | 0.028
706.331 | 0.016
706.586 | 0.023
706.84 | 0.039
707.095 | 0.031
707.349 | -0.003
707.603 | 0.02
707.857 | 0.027
708.111 | 0.017
708.365 | 0.024
708.619 | 0.02
708.873 | 0.036
709.127 | 0.02
709.381 | 0.019
709.635 | 0.023
709.888 | 0.039
710.142 | 0.017
710.396 | 0.024
710.649 | 0.029
710.903 | 0.016
711.156 | 0.025
711.409 | 0.033
711.663 | 0.036
711.916 | 0.02
712.169 | 0.027
712.422 | 0.026
712.675 | 0.032
712.928 | 0.045
713.181 | 0.035
713.434 | 0.031
713.687 | 0.043
713.94 | 0.025
714.192 | 0.031
714.445 | 0.004
714.698 | 0.034
714.95 | 0.038
715.203 | 0.055
715.455 | 0.027
715.708 | 0.038
715.96 | 0.047
716.212 | 0.025
716.464 | 0.045
716.717 | 0.033
716.969 | 0.033
717.221 | 0.042
717.473 | 0.005
717.725 | 0.032
717.977 | 0.028
718.228 | 0.04
718.48 | 0.02
718.732 | 0.038
718.984 | 0.039
719.235 | 0.045
719.487 | 0.038
719.738 | 0.028
719.99 | 0.047
720.241 | 0.023
720.492 | 0.032
720.744 | 0.026
720.995 | 0.02
721.246 | 0.029
721.497 | 0.048
721.748 | 0.028
721.999 | 0.02
722.25 | 0.055
722.501 | 0.047
722.752 | 0.03
723.002 | 0.023
723.253 | 0.015
723.504 | 0.045
723.754 | 0.036
724.005 | 0.045
724.255 | 0.035
724.506 | 0.041
724.756 | 0.045
725.006 | 0.024
725.257 | 0.033
725.507 | 0.038
725.757 | 0.028
726.007 | 0.016
726.257 | 0.043
726.507 | 0.021
726.757 | 0.064
727.007 | 0.042
727.257 | 0.02
727.506 | 0.043
727.756 | 0.032
728.006 | 0.037
728.255 | 0.046
728.505 | 0.017
728.754 | 0.037
729.004 | 0.053
729.253 | 0.06
729.502 | 0.061
729.751 | 0.036
730.001 | 0.019
730.25 | 0.029
730.499 | 0.045
730.748 | 0.041
730.997 | 0.02
731.246 | 0.048
731.495 | 0.037
731.743 | 0.041
731.992 | 0.036
732.241 | 0.024
732.489 | 0.043
732.738 | 0.06
732.987 | 0.025
733.235 | 0.037
733.483 | 0.047
733.732 | 0.043
733.98 | 0.018
734.228 | 0.02
734.477 | 0.025
734.725 | 0.056
734.973 | 0.024
735.221 | 0.045
735.469 | 0.047
735.717 | 0.045
735.965 | 0.025
736.212 | 0.026
736.46 | 0.043
736.708 | 0.041
736.955 | 0.06
737.203 | 0.022
737.451 | 0.054
737.698 | 0.039
737.945 | 0.034
738.193 | 0.061
738.44 | 0.039
738.687 | 0.036
738.935 | 0.039
739.182 | 0.034
739.429 | 0.043
739.676 | 0.026
739.923 | 0.039
740.17 | 0.047
740.417 | 0.03
740.663 | 0.029
740.91 | 0.039
741.157 | 0.049
741.403 | 0.032
741.65 | 0.009
741.897 | 0.039
742.143 | 0.026
742.389 | 0.039
742.636 | 0.019
742.882 | 0.037
743.128 | 0.032
743.375 | 0.013
743.621 | 0.054
743.867 | 0.041
744.113 | 0.03
744.359 | 0.029
744.605 | 0.041
744.851 | 0.062
745.097 | 0.057
745.342 | 0.033
745.588 | 0.049
745.834 | 0.036
746.079 | 0.055
746.325 | 0.048
746.57 | 0.046
746.816 | 0.018
747.061 | 0.045
747.307 | 0.045
747.552 | 0.037
747.797 | 0.053
748.042 | 0.037
748.287 | 0.038
748.532 | 0.032
748.777 | 0.04
749.022 | 0.041
749.267 | 0.049
749.512 | 0.028
749.757 | 0.028
750.002 | 0.042
750.246 | 0.054
750.491 | 0.05
750.736 | 0.033
750.98 | 0.043
751.225 | 0.05
751.469 | 0.058
751.713 | 0.05
751.958 | 0.041
752.202 | 0.029
752.446 | 0.043
752.69 | 0.025
752.934 | 0.016
753.178 | 0.047
753.422 | 0.033
753.666 | 0.057
753.91 | 0.017
754.154 | 0.053
754.398 | 0.072
754.641 | 0.025
754.885 | 0.024
755.129 | 0.048
755.372 | 0.017
755.616 | 0.038
755.859 | 0.045
756.103 | 0.061
756.346 | 0.014
756.589 | 0.032
756.832 | 0.037
757.076 | 0.041
757.319 | 0.038
757.562 | 0.055
757.805 | 0.045
758.048 | 0.046
758.291 | 0.033
758.533 | 0.038
758.776 | 0.034
759.019 | 0.052
759.262 | 0.044
759.504 | 0.041
759.747 | 0.042
759.989 | 0.052
760.232 | 0.055
760.474 | 0.031
760.717 | 0.076
760.959 | 0.031
761.201 | 0.034
761.444 | 0.037
761.686 | 0.056
761.928 | 0.043
762.17 | 0.034
762.412 | 0.05
762.654 | 0.05
762.896 | 0.04
763.137 | 0.049
763.379 | 0.064
763.621 | 0.099
763.863 | 0.059
764.104 | 0.042
764.346 | 0.068
764.587 | 0.041
764.829 | 0.045
765.07 | 0.034
765.312 | 0.057
765.553 | 0.063
765.794 | 0.058
766.035 | 0.044
766.276 | 0.065
766.518 | 0.049
766.759 | 0.063
767 | 0.043
767.241 | 0.053
767.481 | 0.044
767.722 | 0.051
767.963 | 0.059
768.204 | 0.052
768.444 | 0.064
768.685 | 0.055
768.926 | 0.035
769.166 | 0.04
769.407 | 0.044
769.647 | 0.074
769.887 | 0.029
770.128 | 0.064
770.368 | 0.048
770.608 | 0.062
770.848 | 0.052
771.088 | 0.044
771.328 | 0.047
771.568 | 0.04
771.808 | 0.053
772.048 | 0.043
772.288 | 0.062
772.528 | 0.049
772.767 | 0.049
773.007 | 0.052
773.247 | 0.057
773.486 | 0.042
773.726 | 0.031
773.965 | 0.066
774.205 | 0.054
774.444 | 0.061
774.683 | 0.062
774.922 | 0.062
775.162 | 0.052
775.401 | 0.065
775.64 | 0.054
775.879 | 0.048
776.118 | 0.044
776.357 | 0.04
776.596 | 0.056
776.834 | 0.059
777.073 | 0.059
777.312 | 0.064
777.551 | 0.036
777.789 | 0.039
778.028 | 0.052
778.266 | 0.057
778.505 | 0.032
778.743 | 0.032
778.982 | 0.016
779.22 | 0.044
779.458 | 0.051
779.696 | 0.041
779.934 | 0.065
780.173 | 0.055
780.411 | 0.043
780.649 | 0.046
780.887 | 0.047
781.124 | 0.069
781.362 | 0.035
781.6 | 0.053
781.838 | 0.052
782.075 | 0.031
782.313 | 0.043
782.551 | 0.025
782.788 | 0.023
783.026 | 0.048
783.263 | 0.048
783.5 | 0.048
783.738 | 0.051
783.975 | 0.029
784.212 | 0
784.449 | 0.044
784.686 | 0.034
784.924 | 0.051
785.161 | 0.03
785.397 | 0.045
785.634 | 0.065
785.871 | 0.059
786.108 | 0.042
786.345 | 0.073
786.581 | 0.042
786.818 | 0.057
787.055 | 0.051
787.291 | 0.068
787.528 | 0.078
787.764 | 0.047
788.001 | 0.035
788.237 | 0.065
788.473 | 0.032
788.709 | 0.04
788.946 | 0.058
789.182 | 0.025
789.418 | 0.019
789.654 | 0.073
789.89 | 0.033
790.126 | 0.08
790.362 | 0.07
790.597 | 0.032
790.833 | 0.053
791.069 | 0.04
791.305 | 0.028
791.54 | 0.056
791.776 | 0.071
792.011 | 0.051
792.247 | 0.021
792.482 | 0.041
792.717 | 0.052
792.953 | 0.046
793.188 | 0.066
793.423 | 0.071
793.658 | 0.071
793.894 | 0.075
794.129 | 0.052
794.364 | 0.063
794.599 | 0.052
794.833 | 0.052
795.068 | 0.075
795.303 | 0.06
795.538 | 0.027
795.772 | 0.056
796.007 | 0.056
796.242 | 0.071
796.476 | 0.056
796.711 | 0.044
796.945 | 0.05
797.18 | 0.037
797.414 | 0.025
797.648 | 0.07
797.882 | 0.034
798.117 | 0.053
798.351 | 0.025
798.585 | 0.093
798.819 | 0.035
799.053 | 0.063
799.287 | 0.046
799.521 | 0.083
799.755 | 0.061
799.988 | 0.067
800.222 | 0.04
800.456 | 0.044
800.689 | 0.017
800.923 | 0.047
801.156 | 0.078
801.39 | 0.042
801.623 | 0.065
801.857 | 0.07
802.09 | 0.074
802.323 | 0.06
802.557 | 0.054
802.79 | 0.049
803.023 | 0.034
803.256 | 0.043
803.489 | 0.05
803.722 | 0.042
803.955 | 0.033
804.188 | 0.072
804.421 | 0.046
804.653 | 0.058
804.886 | 0.042
805.119 | 0.021
805.351 | 0.047
805.584 | 0.052
805.816 | 0.023
806.049 | 0.048
806.281 | 0.052
806.514 | 0.077
806.746 | 0.089
806.978 | 0.021
807.211 | 0.017
807.443 | 0.033
807.675 | 0.054
background 4
AbsoluteIrradiance:211
195.562 | 0
195.934 | 0
196.307 | 0
196.679 | 0
197.051 | 0
197.423 | 0
197.795 | 0
198.167 | 0
198.539 | 0
198.911 | 0
199.283 | 0
199.654 | 0
200.026 | 0
200.397 | 0
200.769 | 0
201.14 | 0
201.511 | 0
201.882 | 0
202.253 | 0
202.624 | 0
202.995 | 0
203.366 | 0
203.736 | 0
204.107 | 0
204.477 | 0
204.848 | 3.686
205.218 | 7.954
205.588 | 1.764
205.959 | 2.663
206.329 | 2.827
206.699 | 2.094
207.068 | 2.368
207.438 | 3.475
207.808 | 1.935
208.178 | 2.45
208.547 | 1.921
208.917 | 1.822
209.286 | 2.977
209.655 | 2.637
210.025 | 1.507
210.394 | 1.671
210.763 | 2.137
211.132 | 0.92
211.501 | 1.173
211.869 | 1.272
212.238 | 1.657
212.607 | 1.712
212.975 | 0.912
213.344 | 0.893
213.712 | 0.823
214.08 | 1.156
214.448 | 1.1
214.817 | 0.802
215.185 | 1.447
215.553 | 1.173
215.92 | 1.32
216.288 | 0.962
216.656 | 1.076
217.023 | 0.717
217.391 | 0.557
217.758 | 1.21
218.126 | 0.921
218.493 | 0.844
218.86 | 0.768
219.227 | 0.543
219.594 | 0.458
219.961 | 0.892
220.328 | 0.479
220.695 | 0.675
221.061 | 0.839
221.428 | 0.717
221.795 | 0.735
222.161 | 0.756
222.527 | 0.547
222.893 | 0.886
223.26 | 0.287
223.626 | 0.859
223.992 | 0.547
224.358 | 0.534
224.723 | 0.648
225.089 | 0.219
225.455 | 0.485
225.82 | 0.359
226.186 | 0.234
226.551 | 0.143
226.917 | 0.327
227.282 | 0.343
227.647 | 0.483
228.012 | 0.385
228.377 | 0.296
228.742 | 0.364
229.107 | 0.203
229.471 | 0.28
229.836 | 0.174
230.201 | 0.304
230.565 | 0.302
230.929 | 0.244
231.294 | 0.139
231.658 | 0.154
232.022 | 0.168
232.386 | 0.12
232.75 | -0.028
233.114 | 0.141
233.478 | 0.143
233.842 | 0.145
234.205 | 0.162
234.569 | 0.125
234.932 | 0.098
235.296 | 0.085
235.659 | 0.117
236.022 | 0.077
236.385 | 0.083
236.748 | 0.136
237.111 | 0.074
237.474 | 0.152
237.837 | 0.124
238.2 | 0.096
238.562 | 0.103
238.925 | 0.155
239.287 | 0.131
239.65 | 0.082
240.012 | 0.119
240.374 | 0.111
240.736 | 0.12
241.098 | 0.107
241.46 | 0.115
241.822 | 0.133
242.184 | 0.04
242.546 | 0.103
242.907 | 0.083
243.269 | 0.059
243.63 | 0.047
243.992 | 0.112
244.353 | 0.087
244.714 | 0.072
245.075 | 0.087
245.436 | 0.096
245.797 | 0.055
246.158 | 0.011
246.519 | 0.09
246.879 | 0.103
247.24 | 0.088
247.601 | 0.086
247.961 | 0.092
248.321 | 0.045
248.682 | 0.075
249.042 | 0.05
249.402 | 0.063
249.762 | 0.107
250.122 | 0.061
250.482 | 0.112
250.842 | 0.058
251.201 | 0.071
251.561 | 0.078
251.921 | 0.079
252.28 | 0.08
252.639 | 0.082
252.999 | 0.076
253.358 | 0.068
253.717 | 0.067
254.076 | 0.071
254.435 | 0.045
254.794 | 0.045
255.153 | 0.112
255.511 | 0.072
255.87 | 0.035
256.228 | 0.073
256.587 | 0.075
256.945 | 0.057
257.304 | 0.024
257.662 | 0.054
258.02 | 0.055
258.378 | 0.072
258.736 | 0.036
259.094 | 0.103
259.452 | 0.085
259.809 | 0.07
260.167 | 0.08
260.524 | 0.049
260.882 | 0.04
261.239 | 0.052
261.597 | 0.048
261.954 | 0.05
262.311 | 0.056
262.668 | 0.061
263.025 | 0.059
263.382 | 0.097
263.739 | 0.051
264.095 | 0.085
264.452 | 0.033
264.809 | 0.062
265.165 | 0.046
265.521 | 0.087
265.878 | 0.071
266.234 | 0.067
266.59 | 0.049
266.946 | 0.099
267.302 | 0.068
267.658 | 0.075
268.014 | 0.049
268.37 | 0.063
268.725 | 0.056
269.081 | 0.053
269.436 | 0.06
269.792 | 0.042
270.147 | 0.068
270.502 | 0.087
270.857 | 0.067
271.212 | 0.062
271.567 | 0.027
271.922 | 0.096
272.277 | 0.058
272.632 | 0.066
272.987 | 0.081
273.341 | 0.078
273.696 | 0.066
274.05 | 0.033
274.404 | 0.053
274.759 | 0.083
275.113 | 0.083
275.467 | 0.034
275.821 | 0.049
276.175 | 0.063
276.529 | 0.084
276.882 | 0.11
277.236 | 0.062
277.59 | 0.085
277.943 | 0.022
278.297 | 0.106
278.65 | 0.049
279.003 | 0.037
279.356 | 0.047
279.709 | 0.061
280.062 | 0.046
280.415 | 0.059
280.768 | 0.063
281.121 | 0.084
281.474 | 0.047
281.826 | 0.048
282.179 | 0.052
282.531 | 0.035
282.884 | 0.07
283.236 | 0.067
283.588 | 0.092
283.94 | 0.042
284.292 | 0.05
284.644 | 0.043
284.996 | 0.05
285.348 | 0.05
285.7 | 0.067
286.051 | 0.057
286.403 | 0.053
286.754 | 0.061
287.106 | 0.087
287.457 | 0.074
287.808 | 0.066
288.159 | 0.071
288.51 | 0.072
288.861 | 0.054
289.212 | 0.025
289.563 | 0.04
289.914 | 0.069
290.264 | 0.037
290.615 | 0.059
290.966 | 0.064
291.316 | 0.032
291.666 | 0.066
292.017 | 0.051
292.367 | 0.029
292.717 | 0.038
293.067 | 0.043
293.417 | 0.044
293.767 | 0.043
294.116 | 0.064
294.466 | 0.04
294.816 | 0.061
295.165 | 0.033
295.515 | 0.036
295.864 | 0.048
296.213 | 0.03
296.562 | 0.029
296.911 | 0.054
297.261 | 0.039
297.609 | 0.042
297.958 | 0.014
298.307 | 0.035
298.656 | 0.04
299.004 | 0.032
299.353 | 0.003
299.701 | 0.027
300.05 | 0.025
300.398 | 0.04
300.746 | 0.036
301.095 | 0.029
301.443 | 0.031
301.791 | 0.028
302.138 | 0.034
302.486 | 0.021
302.834 | 0.018
303.182 | 0.017
303.529 | 0.035
303.877 | 0.033
304.224 | 0.034
304.572 | 0.035
304.919 | 0.018
305.266 | 0.038
305.613 | 0.028
305.96 | 0.047
306.307 | 0.024
306.654 | 0.046
307.001 | 0.039
307.347 | 0.031
307.694 | 0.02
308.04 | 0.024
308.387 | 0.036
308.733 | 0.032
309.08 | 0.029
309.426 | 0.041
309.772 | 0.033
310.118 | 0.042
310.464 | 0.023
310.81 | 0.031
311.156 | 0.031
311.501 | 0.034
311.847 | 0.032
312.193 | 0.024
312.538 | 0.024
312.883 | 0.024
313.229 | 0.027
313.574 | 0.024
313.919 | 0.028
314.264 | 0.038
314.609 | 0.036
314.954 | 0.024
315.299 | 0.019
315.644 | 0.019
315.988 | 0.033
316.333 | 0.036
316.678 | 0.021
317.022 | 0.025
317.366 | 0.023
317.711 | 0.03
318.055 | 0.016
318.399 | 0.02
318.743 | 0.028
319.087 | 0.025
319.431 | 0.031
319.775 | 0.023
320.118 | 0.039
320.462 | 0.027
320.806 | 0.019
321.149 | 0.031
321.493 | 0.026
321.836 | 0.022
322.179 | 0.009
322.522 | 0.031
322.865 | 0.026
323.208 | 0.026
323.551 | 0.035
323.894 | 0.023
324.237 | 0.019
324.58 | 0.022
324.922 | 0.025
325.265 | 0.03
325.607 | 0.03
325.95 | 0.036
326.292 | 0.01
326.634 | 0.026
326.976 | 0.017
327.318 | 0.015
327.66 | 0.026
328.002 | 0.01
328.344 | 0.025
328.686 | 0.028
329.027 | 0.015
329.369 | 0.016
329.711 | 0.025
330.052 | 0.023
330.393 | 0.014
330.735 | 0.018
331.076 | 0.024
331.417 | 0.014
331.758 | 0.027
332.099 | 0.011
332.44 | 0.028
332.78 | 0.015
333.121 | 0.022
333.462 | 0.018
333.802 | 0.023
334.143 | 0.019
334.483 | 0.021
334.823 | 0.012
335.164 | 0.024
335.504 | 0.019
335.844 | 0.029
336.184 | 0.015
336.524 | 0.014
336.864 | 0.02
337.203 | 0.02
337.543 | 0.032
337.883 | 0.02
338.222 | 0.034
338.562 | 0.016
338.901 | 0.027
339.24 | 0.007
339.58 | 0.019
339.919 | 0.015
340.258 | 0.027
340.597 | 0.019
340.936 | 0.011
341.274 | 0.023
341.613 | 0.024
341.952 | 0.023
342.29 | 0.021
342.629 | 0.031
342.967 | 0.017
343.306 | 0.021
343.644 | 0.024
343.982 | 0.018
344.32 | 0.026
344.658 | 0.017
344.996 | 0.019
345.334 | 0.017
345.672 | 0.027
346.01 | 0.014
346.347 | 0.036
346.685 | 0.024
347.022 | 0.028
347.36 | 0.028
347.697 | 0.024
348.034 | 0.026
348.371 | 0.019
348.708 | 0.021
349.045 | 0.026
349.382 | 0.026
349.719 | 0.015
350.056 | 0.018
350.393 | 0.017
350.729 | 0.031
351.066 | 0.016
351.402 | 0.023
351.739 | 0.029
352.075 | 0.025
352.411 | 0.019
352.747 | 0.019
353.083 | 0.022
353.419 | 0.022
353.755 | 0.019
354.091 | 0.023
354.427 | 0.026
354.763 | 0.024
355.098 | 0.025
355.434 | 0.014
355.769 | 0.032
356.104 | 0.012
356.44 | 0.019
356.775 | 0.022
357.11 | 0.021
357.445 | 0.019
357.78 | 0.02
358.115 | 0.026
358.45 | 0.021
358.785 | 0.023
359.119 | 0.014
359.454 | 0.015
359.788 | 0.019
360.123 | 0.021
360.457 | 0.011
360.791 | 0.03
361.126 | 0.015
361.46 | 0.015
361.794 | 0.026
362.128 | 0.029
362.462 | 0.036
362.795 | 0.032
363.129 | 0.023
363.463 | 0.012
363.796 | 0.024
364.13 | 0.034
364.463 | 0.026
364.797 | 0.028
365.13 | 0.029
365.463 | 0.017
365.796 | 0.018
366.129 | 0.022
366.462 | 0.012
366.795 | 0.023
367.128 | 0.03
367.461 | 0.019
367.793 | 0.026
368.126 | 0.021
368.458 | 0.03
368.791 | 0.026
369.123 | 0.026
369.455 | 0.019
369.787 | 0.023
370.12 | 0.017
370.452 | 0.012
370.784 | 0.014
371.115 | 0.02
371.447 | 0.011
371.779 | 0.021
372.111 | 0.022
372.442 | 0.015
372.774 | 0.031
373.105 | 0.022
373.436 | 0.017
373.768 | 0.018
374.099 | 0.012
374.43 | 0.026
374.761 | 0.021
375.092 | 0.024
375.423 | 0.019
375.754 | 0.019
376.084 | 0.02
376.415 | 0.016
376.745 | 0.017
377.076 | 0.021
377.406 | 0.019
377.737 | 0.036
378.067 | 0.026
378.397 | 0.019
378.727 | 0.017
379.057 | 0.015
379.387 | 0.017
379.717 | 0.017
380.047 | 0.023
380.377 | 0.025
380.706 | 0.024
381.036 | 0.018
381.365 | 0.032
381.695 | 0.017
382.024 | 0.011
382.353 | 0.006
382.682 | 0.021
383.012 | 0.015
383.341 | 0.026
383.67 | 0.023
383.998 | 0.012
384.327 | 0.024
384.656 | 0.013
384.985 | 0.011
385.313 | 0.01
385.642 | 0.015
385.97 | 0.014
386.298 | 0.016
386.627 | 0.012
386.955 | 0.025
387.283 | 0.012
387.611 | 0.013
387.939 | 0.016
388.267 | 0.017
388.595 | 0.027
388.922 | 0.003
389.25 | 0.029
389.578 | 0.015
389.905 | 0.021
390.233 | 0.013
390.56 | 0.02
390.887 | 0.029
391.214 | 0.021
391.542 | 0.01
391.869 | 0.017
392.196 | 0.01
392.522 | 0.011
392.849 | 0.016
393.176 | 0.011
393.503 | 0.017
393.829 | 0.015
394.156 | 0.013
394.482 | 0.019
394.809 | 0.017
395.135 | 0.021
395.461 | 0.011
395.787 | 0.023
396.113 | 0.01
396.439 | 0.014
396.765 | 0.012
397.091 | 0.014
397.417 | 0.01
397.742 | 0.015
398.068 | 0.024
398.394 | 0.019
398.719 | 0.014
399.044 | 0.016
399.37 | 0.019
399.695 | 0.016
400.02 | 0.019
400.345 | 0.023
400.67 | 0.019
400.995 | 0.025
401.32 | 0.017
401.645 | 0.016
401.969 | 0.011
402.294 | 0.012
402.618 | 0.029
402.943 | 0.01
403.267 | 0.012
403.592 | 0.014
403.916 | 0.026
404.24 | 0.014
404.564 | 0.016
404.888 | 0.03
405.212 | 0.023
405.536 | 0.011
405.86 | 0.02
406.183 | 0.016
406.507 | 0.017
406.831 | 0.022
407.154 | 0.014
407.478 | 0.015
407.801 | 0.017
408.124 | 0.011
408.447 | 0.013
408.77 | 0.019
409.093 | 0.019
409.416 | 0.013
409.739 | 0.02
410.062 | 0.007
410.385 | 0.02
410.708 | 0.021
411.03 | 0.023
411.353 | 0.012
411.675 | 0.022
411.997 | 0.016
412.32 | 0.012
412.642 | 0.02
412.964 | 0.015
413.286 | 0.016
413.608 | 0.016
413.93 | 0.019
414.252 | 0.009
414.574 | 0.009
414.895 | 0.012
415.217 | 0.023
415.538 | 0.011
415.86 | 0.016
416.181 | 0.022
416.503 | 0.013
416.824 | 0.015
417.145 | 0.012
417.466 | 0.014
417.787 | 0.024
418.108 | 0.012
418.429 | 0.014
418.75 | 0.015
419.07 | 0.013
419.391 | 0.018
419.712 | 0.018
420.032 | 0.014
420.353 | 0.015
420.673 | 0.013
420.993 | 0.018
421.313 | 0.021
421.633 | 0.02
421.953 | 0.007
422.273 | 0.016
422.593 | 0.015
422.913 | 0.013
423.233 | 0.014
423.553 | 0.01
423.872 | 0.011
424.192 | 0.013
424.511 | 0.009
424.83 | 0.02
425.15 | 0.019
425.469 | 0.018
425.788 | 0.023
426.107 | 0.016
426.426 | 0.022
426.745 | 0.023
427.064 | 0.018
427.383 | 0.02
427.702 | 0.01
428.02 | 0.019
428.339 | 0.015
428.657 | 0.015
428.976 | 0.02
429.294 | 0.014
429.612 | 0.011
429.93 | 0.019
430.248 | 0.023
430.567 | 0.016
430.884 | 0.015
431.202 | 0.011
431.52 | 0.01
431.838 | 0.012
432.156 | 0.016
432.473 | 0.016
432.791 | 0.019
433.108 | 0.015
433.426 | 0.022
433.743 | 0.016
434.06 | 0.015
434.377 | 0.019
434.694 | 0.02
435.011 | 0.013
435.328 | 0.016
435.645 | 0.014
435.962 | 0.016
436.279 | 0.016
436.595 | 0.013
436.912 | 0.015
437.228 | 0.015
437.545 | 0.011
437.861 | 0.013
438.177 | 0.011
438.493 | 0.009
438.81 | 0.016
439.126 | 0.014
439.442 | 0.018
439.757 | 0.003
440.073 | 0.011
440.389 | 0.014
440.705 | 0.023
441.02 | 0.013
441.336 | 0.013
441.651 | 0.017
441.967 | 0.011
442.282 | 0.007
442.597 | 0.018
442.912 | 0.02
443.227 | 0.017
443.543 | 0.016
443.857 | 0.019
444.172 | 0.012
444.487 | 0.017
444.802 | 0.025
445.116 | 0.021
445.431 | 0.01
445.746 | 0.01
446.06 | 0.017
446.374 | 0.013
446.689 | 0.022
447.003 | 0.013
447.317 | 0.016
447.631 | 0.009
447.945 | 0.018
448.259 | 0.018
448.573 | 0.025
448.887 | 0.012
449.2 | 0.023
449.514 | 0.021
449.827 | 0.015
450.141 | 0.016
450.454 | 0.014
450.768 | 0.015
451.081 | 0.021
451.394 | 0.013
451.707 | 0.011
452.02 | 0.021
452.333 | 0.016
452.646 | 0.012
452.959 | 0.015
453.272 | 0.016
453.584 | 0.013
453.897 | 0.016
454.21 | 0.017
454.522 | 0.012
454.834 | 0.012
455.147 | 0.016
455.459 | 0.005
455.771 | 0.003
456.083 | 0.012
456.395 | 0.014
456.707 | 0.015
457.019 | 0.017
457.331 | 0.016
457.643 | 0.011
457.954 | 0.017
458.266 | 0.012
458.577 | 0.012
458.889 | 0.008
459.2 | 0.018
459.512 | 0.011
459.823 | 0.016
460.134 | 0.016
460.445 | 0.017
460.756 | 0.013
461.067 | 0.016
461.378 | 0.011
461.689 | 0.011
461.999 | 0.028
462.31 | 0.021
462.621 | 0.011
462.931 | 0.018
463.242 | 0.017
463.552 | 0.019
463.862 | 0.01
464.172 | 0.015
464.483 | 0.025
464.793 | 0.013
465.103 | 0.018
465.413 | 0.018
465.722 | 0.015
466.032 | 0.019
466.342 | 0.013
466.652 | 0.015
466.961 | 0.016
467.271 | 0.019
467.58 | 0.016
467.889 | 0.019
468.199 | 0.018
468.508 | 0.014
468.817 | 0.02
469.126 | 0.017
469.435 | 0.021
469.744 | 0.016
470.053 | 0.013
470.362 | 0.015
470.67 | 0.012
470.979 | 0.019
471.287 | 0.011
471.596 | 0.013
471.904 | 0.024
472.213 | 0.009
472.521 | 0.02
472.829 | 0.017
473.137 | 0.016
473.445 | 0.01
473.753 | 0.026
474.061 | 0.012
474.369 | 0.016
474.677 | 0.019
474.985 | 0.01
475.292 | 0.009
475.6 | 0.02
475.907 | 0.019
476.215 | 0.013
476.522 | 0.023
476.829 | 0.017
477.137 | 0.019
477.444 | 0.014
477.751 | 0.025
478.058 | 0.011
478.365 | 0.018
478.671 | 0.02
478.978 | 0.018
479.285 | 0.01
479.592 | 0.014
479.898 | 0.011
480.205 | 0.018
480.511 | 0.005
480.817 | 0.019
481.124 | 0.016
481.43 | 0.021
481.736 | 0.028
482.042 | 0.014
482.348 | 0.013
482.654 | 0.022
482.96 | 0.02
483.266 | 0.022
483.571 | 0.024
483.877 | 0.006
484.182 | 0.009
484.488 | 0.015
484.793 | 0.012
485.099 | 0.017
485.404 | 0.019
485.709 | 0.021
486.014 | 0.017
486.319 | 0.024
486.624 | 0.015
486.929 | 0.009
487.234 | 0.006
487.539 | 0.019
487.844 | 0.019
488.148 | 0.015
488.453 | 0.024
488.757 | 0.021
489.062 | 0.015
489.366 | 0.021
489.67 | 0.013
489.975 | 0.012
490.279 | 0.016
490.583 | 0.017
490.887 | 0.014
491.191 | 0.017
491.495 | 0.02
491.798 | 0.022
492.102 | 0.019
492.406 | 0.011
492.709 | 0.018
493.013 | 0.022
493.316 | 0.02
493.62 | 0.023
493.923 | 0.016
494.226 | 0.02
494.529 | 0.016
494.832 | 0.019
495.135 | 0.015
495.438 | 0.022
495.741 | 0.028
496.044 | 0.014
496.347 | 0.012
496.649 | 0.026
496.952 | 0.013
497.255 | 0.013
497.557 | 0.015
497.859 | 0.014
498.162 | 0.006
498.464 | 0.024
498.766 | 0.016
499.068 | 0.015
499.37 | 0.017
499.672 | 0.024
499.974 | 0.017
500.276 | 0.022
500.578 | 0.02
500.879 | 0.008
501.181 | 0.025
501.483 | 0.013
501.784 | 0.022
502.085 | 0.011
502.387 | 0.01
502.688 | 0.012
502.989 | 0.018
503.29 | 0.009
503.591 | 0.01
503.892 | 0.01
504.193 | 0.024
504.494 | 0.015
504.795 | 0.009
505.096 | 0.025
505.396 | 0.011
505.697 | 0.007
505.997 | 0.018
506.298 | 0.016
506.598 | 0.012
506.898 | 0.021
507.199 | 0.015
507.499 | 0.019
507.799 | 0.013
508.099 | 0.015
508.399 | 0.015
508.699 | 0.016
508.998 | 0.012
509.298 | 0.026
509.598 | 0.012
509.897 | 0.018
510.197 | 0.014
510.496 | 0.015
510.796 | 0.008
511.095 | 0.016
511.394 | 0.017
511.693 | 0.026
511.992 | 0.021
512.291 | 0.017
512.59 | 0.021
512.889 | 0.021
513.188 | 0.02
513.487 | 0.02
513.786 | 0.022
514.084 | 0.02
514.383 | 0.022
514.681 | 0.024
514.98 | 0.019
515.278 | 0.023
515.576 | 0.027
515.874 | 0.021
516.173 | 0.012
516.471 | 0.026
516.769 | 0.022
517.067 | 0.016
517.364 | 0.013
517.662 | 0.01
517.96 | 0.014
518.257 | 0.014
518.555 | 0.017
518.853 | 0.013
519.15 | 0.011
519.447 | 0.019
519.745 | 0.017
520.042 | 0.023
520.339 | 0.009
520.636 | 0.021
520.933 | 0.027
521.23 | 0.01
521.527 | 0.014
521.824 | 0.021
522.121 | 0.011
522.417 | 0.014
522.714 | 0.011
523.01 | 0.017
523.307 | 0.016
523.603 | 0.021
523.9 | 0.02
524.196 | 0.024
524.492 | 0.01
524.788 | 0.027
525.084 | 0.02
525.38 | 0.028
525.676 | 0.012
525.972 | 0.018
526.268 | 0.018
526.564 | 0.016
526.859 | 0.02
527.155 | 0.014
527.45 | 0.023
527.746 | 0.019
528.041 | 0.026
528.336 | 0.019
528.632 | 0.009
528.927 | 0.01
529.222 | 0.02
529.517 | 0.019
529.812 | 0.016
530.107 | 0.013
530.402 | 0.02
530.696 | 0.023
530.991 | 0.023
531.286 | 0.022
531.58 | 0.021
531.875 | 0.016
532.169 | 0.02
532.463 | 0.016
532.758 | 0.007
533.052 | 0.019
533.346 | 0.012
533.64 | 0.021
533.934 | 0.02
534.228 | 0.023
534.522 | 0.023
534.816 | 0.022
535.11 | 0.012
535.403 | 0.025
535.697 | 0.024
535.99 | 0.018
536.284 | 0.018
536.577 | 0.028
536.871 | 0.019
537.164 | 0.017
537.457 | 0.024
537.75 | 0.019
538.043 | 0.019
538.336 | 0.021
538.629 | 0.019
538.922 | 0.018
539.215 | 0.009
539.507 | 0.012
539.8 | 0.006
540.093 | 0.026
540.385 | 0.013
540.678 | 0.015
540.97 | 0.005
541.262 | 0.017
541.554 | 0.015
541.847 | 0.022
542.139 | 0.021
542.431 | 0.024
542.723 | 0.018
543.015 | 0.022
543.307 | 0.021
543.598 | 0.02
543.89 | 0.022
544.182 | 0.017
544.473 | 0.024
544.765 | 0.028
545.056 | 0.008
545.347 | 0.015
545.639 | 0.013
545.93 | 0.01
546.221 | 0.018
546.512 | 0.018
546.803 | 0.011
547.094 | 0.015
547.385 | 0.017
547.676 | 0.02
547.967 | 0.034
548.257 | 0.017
548.548 | 0.021
548.839 | 0.01
549.129 | 0.016
549.419 | 0.026
549.71 | 0.02
550 | 0.024
550.29 | 0.015
550.58 | 0.011
550.871 | 0.017
551.161 | 0.013
551.45 | 0.01
551.74 | 0.015
552.03 | 0.012
552.32 | 0.017
552.61 | 0.02
552.899 | 0.027
553.189 | 0.014
553.478 | 0.014
553.768 | 0.012
554.057 | 0.03
554.346 | 0.024
554.635 | 0.02
554.925 | 0.017
555.214 | 0.023
555.503 | 0.01
555.792 | 0.02
556.081 | 0.012
556.369 | 0.017
556.658 | 0.019
556.947 | 0.021
557.235 | 0.01
557.524 | 0.03
557.812 | 0.017
558.101 | 0.021
558.389 | 0.023
558.677 | 0.019
558.966 | 0.024
559.254 | 0.03
559.542 | 0.018
559.83 | 0.022
560.118 | 0.024
560.406 | 0.017
560.693 | 0.033
560.981 | 0.027
561.269 | 0.017
561.556 | 0.018
561.844 | 0.023
562.131 | 0.022
562.419 | 0.025
562.706 | 0.023
562.994 | 0.017
563.281 | 0.029
563.568 | 0.021
563.855 | 0.037
564.142 | 0.019
564.429 | 0.027
564.716 | 0.025
565.003 | 0.03
565.289 | 0.024
565.576 | 0.026
565.863 | 0.023
566.149 | 0.029
566.436 | 0.026
566.722 | 0.022
567.008 | 0.021
567.295 | 0.031
567.581 | 0.017
567.867 | 0.021
568.153 | 0.016
568.439 | 0.024
568.725 | 0.019
569.011 | 0.024
569.297 | 0.022
569.583 | 0.024
569.868 | 0.021
570.154 | 0.02
570.439 | 0.015
570.725 | 0.025
571.01 | 0.02
571.296 | 0.027
571.581 | 0.017
571.866 | 0.029
572.151 | 0.027
572.436 | 0.02
572.721 | 0.021
573.006 | 0.02
573.291 | 0.032
573.576 | 0.02
573.861 | 0.028
574.146 | 0.018
574.43 | 0.025
574.715 | 0.03
574.999 | 0.032
575.284 | 0.03
575.568 | 0.021
575.852 | 0.021
576.137 | 0.031
576.421 | 0.033
576.705 | 0.021
576.989 | 0.025
577.273 | 0.012
577.557 | 0.026
577.841 | 0.022
578.124 | 0.031
578.408 | 0.026
578.692 | 0.018
578.975 | 0.025
579.259 | 0.018
579.542 | 0.019
579.826 | 0.014
580.109 | 0.025
580.392 | 0.029
580.676 | 0.016
580.959 | 0.021
581.242 | 0.019
581.525 | 0.021
581.808 | 0.027
582.091 | 0.022
582.373 | 0.016
582.656 | 0.025
582.939 | 0.022
583.221 | 0.015
583.504 | 0.026
583.786 | 0.024
584.069 | 0.018
584.351 | 0.017
584.633 | 0.017
584.916 | 0.007
585.198 | 0.022
585.48 | 0.024
585.762 | 0.006
586.044 | 0.015
586.326 | 0.027
586.607 | 0.022
586.889 | 0.037
587.171 | 0.028
587.453 | 0.018
587.734 | 0.02
588.016 | 0.024
588.297 | 0.009
588.578 | 0.03
588.86 | 0.027
589.141 | 0.023
589.422 | 0.021
589.703 | 0.033
589.984 | 0.015
590.265 | 0.02
590.546 | 0.023
590.827 | 0.016
591.108 | 0.019
591.389 | 0.028
591.669 | 0.017
591.95 | 0.024
592.23 | 0.025
592.511 | 0.034
592.791 | 0.024
593.071 | 0.03
593.352 | 0.012
593.632 | 0.028
593.912 | 0.028
594.192 | 0.036
594.472 | 0.027
594.752 | 0.024
595.032 | 0.023
595.312 | 0.013
595.592 | 0.027
595.871 | 0.024
596.151 | 0.022
596.43 | 0.008
596.71 | 0.016
596.989 | 0.03
597.269 | 0.026
597.548 | 0.022
597.827 | 0.033
598.106 | 0.03
598.385 | 0.028
598.665 | 0.002
598.944 | 0.013
599.222 | 0.017
599.501 | 0.021
599.78 | 0.023
600.059 | 0.025
600.337 | 0.022
600.616 | 0.016
600.895 | 0.028
601.173 | 0.051
601.451 | 0.023
601.73 | 0.015
602.008 | 0.029
602.286 | 0.012
602.564 | 0.026
602.842 | 0.015
603.12 | 0.014
603.398 | 0.018
603.676 | 0.023
603.954 | 0.019
604.232 | 0.02
604.51 | 0.013
604.787 | 0.01
605.065 | 0.022
605.342 | 0.015
605.62 | 0.011
605.897 | 0.019
606.174 | 0.031
606.452 | 0.015
606.729 | 0.03
607.006 | 0.024
607.283 | 0.019
607.56 | 0.015
607.837 | 0.011
608.114 | 0.019
608.391 | 0.045
608.667 | 0.006
608.944 | 0.024
609.221 | 0.012
609.497 | 0.019
609.774 | 0.012
610.05 | 0.03
610.326 | 0.023
610.603 | 0.025
610.879 | 0.032
611.155 | 0.026
611.431 | 0.013
611.707 | 0.022
611.983 | 0.014
612.259 | 0.025
612.535 | 0.017
612.811 | 0.02
613.087 | 0.023
613.362 | 0.019
613.638 | 0.014
613.913 | 0.02
614.189 | 0.018
614.464 | 0.022
614.74 | 0.032
615.015 | 0.005
615.29 | 0.024
615.565 | 0.022
615.84 | 0.01
616.115 | 0.03
616.39 | 0.019
616.665 | 0.027
616.94 | 0.02
617.215 | 0.031
617.49 | 0.023
617.764 | 0.027
618.039 | 0.026
618.313 | 0.026
618.588 | 0.031
618.862 | 0.024
619.137 | 0.031
619.411 | 0.016
619.685 | 0.024
619.959 | 0.021
620.233 | 0.032
620.507 | 0.017
620.781 | 0.017
621.055 | 0.017
621.329 | 0.024
621.603 | 0.018
621.876 | 0.029
622.15 | 0.022
622.424 | 0.021
622.697 | 0.024
622.971 | 0.02
623.244 | 0.013
623.517 | 0.015
623.791 | 0.021
624.064 | 0.026
624.337 | 0.029
624.61 | 0.017
624.883 | 0.025
625.156 | 0.014
625.429 | 0.026
625.702 | 0.014
625.974 | 0.031
626.247 | 0.012
626.52 | 0.026
626.792 | 0.026
627.065 | 0.026
627.337 | 0.022
627.61 | 0.022
627.882 | 0.021
628.154 | 0.024
628.427 | 0.031
628.699 | 0.021
628.971 | 0.021
629.243 | 0.012
629.515 | 0.035
629.787 | 0.022
630.058 | 0.034
630.33 | 0.029
630.602 | 0.007
630.874 | 0.027
631.145 | 0.035
631.417 | 0.014
631.688 | 0.026
631.959 | 0.019
632.231 | 0.022
632.502 | 0.031
632.773 | 0.03
633.044 | 0.038
633.316 | 0.03
633.587 | 0.009
633.857 | 0.013
634.128 | 0.034
634.399 | 0.041
634.67 | 0.019
634.941 | 0.026
635.211 | 0.03
635.482 | 0.02
635.752 | 0.028
636.023 | 0.01
636.293 | 0.022
636.564 | 0.038
636.834 | 0.022
637.104 | 0.033
637.374 | 0.021
637.644 | 0.031
637.914 | 0.029
638.184 | 0.028
638.454 | 0.021
638.724 | 0.039
638.994 | 0.026
639.264 | 0.033
639.533 | 0.024
639.803 | 0.024
640.072 | 0.024
640.342 | 0.023
640.611 | 0.032
640.881 | 0.025
641.15 | 0.044
641.419 | 0.015
641.688 | 0.025
641.957 | 0.022
642.227 | 0.027
642.496 | 0.032
642.764 | 0.03
643.033 | 0.035
643.302 | 0.022
643.571 | 0.022
643.84 | 0.017
644.108 | 0.033
644.377 | 0.006
644.645 | 0.017
644.914 | 0.021
645.182 | 0.034
645.45 | 0.032
645.719 | 0.041
645.987 | 0.036
646.255 | 0.024
646.523 | 0.02
646.791 | 0.022
647.059 | 0.04
647.327 | 0.021
647.595 | 0.024
647.862 | 0.014
648.13 | 0.02
648.398 | 0.029
648.665 | 0.02
648.933 | 0.04
649.2 | 0.03
649.468 | 0.024
649.735 | 0.032
650.002 | 0.026
650.269 | 0.025
650.536 | 0.02
650.804 | 0.02
651.071 | 0.039
651.338 | 0.021
651.604 | 0.024
651.871 | 0.019
652.138 | 0.021
652.405 | 0.011
652.671 | 0.025
652.938 | 0.022
653.205 | 0.028
653.471 | 0.008
653.737 | 0.023
654.004 | 0.012
654.27 | 0.018
654.536 | 0.022
654.803 | 0.02
655.069 | 0.02
655.335 | 0.023
655.601 | 0.025
655.867 | 0.02
656.132 | 0.032
656.398 | 0.023
656.664 | 0.024
656.93 | 0.019
657.195 | 0.018
657.461 | 0.023
657.726 | 0.024
657.992 | 0.006
658.257 | 0.022
658.523 | 0.029
658.788 | 0.015
659.053 | 0.022
659.318 | 0.021
659.583 | 0.018
659.848 | 0.015
660.113 | 0.024
660.378 | 0.026
660.643 | 0.019
660.908 | 0.01
661.172 | 0.012
661.437 | 0.017
661.702 | 0.02
661.966 | 0.032
662.231 | 0.023
662.495 | 0.025
662.76 | 0.043
663.024 | 0.018
663.288 | 0.014
663.552 | 0.015
663.816 | 0.02
664.08 | 0.039
664.344 | 0.023
664.608 | 0.025
664.872 | 0.016
665.136 | 0.017
665.4 | 0.023
665.664 | 0.028
665.927 | 0.025
666.191 | 0.024
666.454 | 0.01
666.718 | 0.023
666.981 | 0.01
667.244 | 0.015
667.508 | 0.032
667.771 | 0.017
668.034 | 0.03
668.297 | 0.018
668.56 | 0.027
668.823 | 0.008
669.086 | 0.033
669.349 | 0.017
669.612 | 0.018
669.875 | 0.011
670.137 | 0.027
670.4 | 0.025
670.662 | 0.017
670.925 | 0.025
671.187 | 0.032
671.45 | 0.005
671.712 | 0.043
671.974 | 0.033
672.237 | 0.012
672.499 | 0.02
672.761 | 0.025
673.023 | 0.022
673.285 | 0.025
673.547 | 0.028
673.808 | 0.024
674.07 | 0.017
674.332 | 0.025
674.594 | 0.018
674.855 | 0.024
675.117 | 0.014
675.378 | 0.028
675.64 | 0.017
675.901 | 0.008
676.162 | 0.027
676.424 | 0.023
676.685 | 0.029
676.946 | 0.026
677.207 | 0.037
677.468 | 0.038
677.729 | 0.017
677.99 | 0.012
678.251 | 0.028
678.512 | 0.029
678.772 | 0.027
679.033 | 0.038
679.294 | 0.022
679.554 | 0.027
679.815 | 0.025
680.075 | 0.018
680.335 | 0.028
680.596 | 0.016
680.856 | 0.026
681.116 | 0.039
681.376 | 0.029
681.636 | 0.022
681.896 | 0.039
682.156 | 0.037
682.416 | 0.024
682.676 | 0.023
682.936 | 0.035
683.196 | 0.023
683.455 | 0.023
683.715 | 0.025
683.974 | 0.028
684.234 | 0.023
684.493 | 0.011
684.753 | 0.024
685.012 | 0.034
685.271 | 0.017
685.53 | 0.025
685.789 | 0.035
686.049 | 0.029
686.308 | 0.023
686.567 | 0.024
686.825 | 0.038
687.084 | 0.028
687.343 | 0.032
687.602 | 0.033
687.86 | 0.033
688.119 | 0.026
688.378 | 0.036
688.636 | 0.039
688.894 | 0.028
689.153 | 0.024
689.411 | 0.038
689.669 | 0.027
689.928 | 0.033
690.186 | 0.008
690.444 | 0.036
690.702 | 0.029
690.96 | 0.04
691.218 | 0.02
691.476 | 0.031
691.733 | 0.031
691.991 | 0.029
692.249 | 0.022
692.506 | 0.023
692.764 | 0.032
693.021 | 0.041
693.279 | 0.036
693.536 | 0.034
693.794 | 0.009
694.051 | 0.023
694.308 | 0.031
694.565 | 0.038
694.822 | 0.039
695.079 | 0.015
695.336 | 0.03
695.593 | 0.019
695.85 | 0.029
696.107 | 0.022
696.364 | 0.027
696.62 | 0.034
696.877 | 0.031
697.133 | 0.023
697.39 | 0.037
697.646 | 0.035
697.903 | 0.022
698.159 | 0.015
698.415 | 0.02
698.672 | 0.035
698.928 | 0.028
699.184 | 0.03
699.44 | 0.028
699.696 | 0.038
699.952 | 0.011
700.208 | 0.025
700.464 | 0.024
700.719 | 0.044
700.975 | 0.037
701.231 | 0.036
701.486 | 0.041
701.742 | 0.027
701.997 | 0.034
702.253 | 0.03
702.508 | 0.01
702.763 | 0.032
703.019 | 0.036
703.274 | 0.035
703.529 | 0.036
703.784 | 0.026
704.039 | 0.048
704.294 | 0.024
704.549 | 0.021
704.804 | 0.03
705.058 | 0.036
705.313 | 0.02
705.568 | 0.043
705.822 | 0.04
706.077 | 0.021
706.331 | 0.021
706.586 | 0.024
706.84 | 0.042
707.095 | 0.027
707.349 | 0
707.603 | 0.022
707.857 | 0.026
708.111 | 0.019
708.365 | 0.023
708.619 | 0.024
708.873 | 0.043
709.127 | 0.024
709.381 | 0.02
709.635 | 0.026
709.888 | 0.035
710.142 | 0.025
710.396 | 0.034
710.649 | 0.024
710.903 | 0.019
711.156 | 0.027
711.409 | 0.035
711.663 | 0.042
711.916 | 0.023
712.169 | 0.025
712.422 | 0.03
712.675 | 0.026
712.928 | 0.047
713.181 | 0.034
713.434 | 0.034
713.687 | 0.045
713.94 | 0.028
714.192 | 0.035
714.445 | 0.013
714.698 | 0.03
714.95 | 0.037
715.203 | 0.056
715.455 | 0.032
715.708 | 0.037
715.96 | 0.05
716.212 | 0.025
716.464 | 0.044
716.717 | 0.038
716.969 | 0.037
717.221 | 0.04
717.473 | 0.015
717.725 | 0.033
717.977 | 0.027
718.228 | 0.045
718.48 | 0.024
718.732 | 0.036
718.984 | 0.041
719.235 | 0.04
719.487 | 0.034
719.738 | 0.025
719.99 | 0.047
720.241 | 0.03
720.492 | 0.041
720.744 | 0.023
720.995 | 0.019
721.246 | 0.027
721.497 | 0.046
721.748 | 0.03
721.999 | 0.025
722.25 | 0.052
722.501 | 0.046
722.752 | 0.039
723.002 | 0.032
723.253 | 0.024
723.504 | 0.051
723.754 | 0.041
724.005 | 0.051
724.255 | 0.041
724.506 | 0.042
724.756 | 0.041
725.006 | 0.029
725.257 | 0.035
725.507 | 0.043
725.757 | 0.033
726.007 | 0.022
726.257 | 0.047
726.507 | 0.025
726.757 | 0.071
727.007 | 0.045
727.257 | 0.024
727.506 | 0.046
727.756 | 0.022
728.006 | 0.043
728.255 | 0.044
728.505 | 0.021
728.754 | 0.042
729.004 | 0.051
729.253 | 0.064
729.502 | 0.064
729.751 | 0.042
730.001 | 0.018
730.25 | 0.028
730.499 | 0.047
730.748 | 0.036
730.997 | 0.034
731.246 | 0.05
731.495 | 0.036
731.743 | 0.046
731.992 | 0.04
732.241 | 0.025
732.489 | 0.044
732.738 | 0.065
732.987 | 0.033
733.235 | 0.04
733.483 | 0.051
733.732 | 0.044
733.98 | 0.023
734.228 | 0.028
734.477 | 0.03
734.725 | 0.048
734.973 | 0.031
735.221 | 0.046
735.469 | 0.042
735.717 | 0.05
735.965 | 0.03
736.212 | 0.027
736.46 | 0.043
736.708 | 0.041
736.955 | 0.06
737.203 | 0.026
737.451 | 0.057
737.698 | 0.039
737.945 | 0.036
738.193 | 0.046
738.44 | 0.032
738.687 | 0.038
738.935 | 0.034
739.182 | 0.038
739.429 | 0.046
739.676 | 0.031
739.923 | 0.043
740.17 | 0.05
740.417 | 0.034
740.663 | 0.027
740.91 | 0.045
741.157 | 0.051
741.403 | 0.036
741.65 | 0.022
741.897 | 0.043
742.143 | 0.03
742.389 | 0.038
742.636 | 0.027
742.882 | 0.029
743.128 | 0.041
743.375 | 0.013
743.621 | 0.063
743.867 | 0.049
744.113 | 0.037
744.359 | 0.027
744.605 | 0.05
744.851 | 0.072
745.097 | 0.056
745.342 | 0.036
745.588 | 0.043
745.834 | 0.038
746.079 | 0.063
746.325 | 0.045
746.57 | 0.057
746.816 | 0.021
747.061 | 0.053
747.307 | 0.05
747.552 | 0.034
747.797 | 0.054
748.042 | 0.043
748.287 | 0.037
748.532 | 0.031
748.777 | 0.046
749.022 | 0.049
749.267 | 0.05
749.512 | 0.028
749.757 | 0.04
750.002 | 0.042
750.246 | 0.037
750.491 | 0.044
750.736 | 0.034
750.98 | 0.042
751.225 | 0.042
751.469 | 0.048
751.713 | 0.049
751.958 | 0.052
752.202 | 0.036
752.446 | 0.035
752.69 | 0.026
752.934 | 0.024
753.178 | 0.046
753.422 | 0.028
753.666 | 0.068
753.91 | 0.022
754.154 | 0.051
754.398 | 0.08
754.641 | 0.027
754.885 | 0.027
755.129 | 0.048
755.372 | 0.027
755.616 | 0.043
755.859 | 0.039
756.103 | 0.068
756.346 | 0.013
756.589 | 0.049
756.832 | 0.036
757.076 | 0.044
757.319 | 0.049
757.562 | 0.057
757.805 | 0.056
758.048 | 0.054
758.291 | 0.043
758.533 | 0.038
758.776 | 0.037
759.019 | 0.052
759.262 | 0.049
759.504 | 0.05
759.747 | 0.037
759.989 | 0.057
760.232 | 0.063
760.474 | 0.031
760.717 | 0.066
760.959 | 0.031
761.201 | 0.034
761.444 | 0.049
761.686 | 0.059
761.928 | 0.037
762.17 | 0.046
762.412 | 0.058
762.654 | 0.052
762.896 | 0.037
763.137 | 0.056
763.379 | 0.029
763.621 | 0.052
763.863 | 0.041
764.104 | 0.051
764.346 | 0.066
764.587 | 0.037
764.829 | 0.045
765.07 | 0.028
765.312 | 0.067
765.553 | 0.064
765.794 | 0.066
766.035 | 0.059
766.276 | 0.061
766.518 | 0.058
766.759 | 0.077
767 | 0.04
767.241 | 0.065
767.481 | 0.049
767.722 | 0.051
767.963 | 0.05
768.204 | 0.063
768.444 | 0.079
768.685 | 0.057
768.926 | 0.036
769.166 | 0.044
769.407 | 0.044
769.647 | 0.077
769.887 | 0.029
770.128 | 0.057
770.368 | 0.049
770.608 | 0.066
770.848 | 0.057
771.088 | 0.051
771.328 | 0.069
771.568 | 0.041
771.808 | 0.056
772.048 | 0.043
772.288 | 0.056
772.528 | 0.04
772.767 | 0.052
773.007 | 0.056
773.247 | 0.064
773.486 | 0.047
773.726 | 0.039
773.965 | 0.069
774.205 | 0.069
774.444 | 0.057
774.683 | 0.063
774.922 | 0.066
775.162 | 0.066
775.401 | 0.077
775.64 | 0.053
775.879 | 0.064
776.118 | 0.049
776.357 | 0.045
776.596 | 0.064
776.834 | 0.07
777.073 | 0.063
777.312 | 0.061
777.551 | 0.037
777.789 | 0.033
778.028 | 0.049
778.266 | 0.06
778.505 | 0.043
778.743 | 0.035
778.982 | 0.035
779.22 | 0.052
779.458 | 0.055
779.696 | 0.057
779.934 | 0.061
780.173 | 0.065
780.411 | 0.061
780.649 | 0.052
780.887 | 0.049
781.124 | 0.064
781.362 | 0.049
781.6 | 0.068
781.838 | 0.041
782.075 | 0.042
782.313 | 0.044
782.551 | 0.031
782.788 | 0.041
783.026 | 0.048
783.263 | 0.048
783.5 | 0.04
783.738 | 0.055
783.975 | 0.032
784.212 | 0.018
784.449 | 0.045
784.686 | 0.046
784.924 | 0.044
785.161 | 0.027
785.397 | 0.057
785.634 | 0.057
785.871 | 0.071
786.108 | 0.055
786.345 | 0.087
786.581 | 0.049
786.818 | 0.073
787.055 | 0.059
787.291 | 0.076
787.528 | 0.085
787.764 | 0.06
788.001 | 0.031
788.237 | 0.075
788.473 | 0.033
788.709 | 0.057
788.946 | 0.041
789.182 | 0.032
789.418 | 0.026
789.654 | 0.075
789.89 | 0.036
790.126 | 0.073
790.362 | 0.067
790.597 | 0.034
790.833 | 0.064
791.069 | 0.041
791.305 | 0.034
791.54 | 0.052
791.776 | 0.081
792.011 | 0.071
792.247 | 0.023
792.482 | 0.054
792.717 | 0.069
792.953 | 0.058
793.188 | 0.071
793.423 | 0.084
793.658 | 0.08
793.894 | 0.078
794.129 | 0.055
794.364 | 0.048
794.599 | 0.066
794.833 | 0.053
795.068 | 0.065
795.303 | 0.054
795.538 | 0.029
795.772 | 0.047
796.007 | 0.056
796.242 | 0.086
796.476 | 0.07
796.711 | 0.043
796.945 | 0.046
797.18 | 0.033
797.414 | 0.028
797.648 | 0.077
797.882 | 0.042
798.117 | 0.046
798.351 | 0.035
798.585 | 0.087
798.819 | 0.034
799.053 | 0.066
799.287 | 0.053
799.521 | 0.084
799.755 | 0.073
799.988 | 0.059
800.222 | 0.056
800.456 | 0.046
800.689 | 0.018
800.923 | 0.049
801.156 | 0.077
801.39 | 0.044
801.623 | 0.056
801.857 | 0.05
802.09 | 0.061
802.323 | 0.06
802.557 | 0.061
802.79 | 0.043
803.023 | 0.042
803.256 | 0.048
803.489 | 0.054
803.722 | 0.029
803.955 | 0.04
804.188 | 0.082
804.421 | 0.045
804.653 | 0.06
804.886 | 0.043
805.119 | 0.026
805.351 | 0.039
805.584 | 0.072
805.816 | 0.028
806.049 | 0.051
806.281 | 0.051
806.514 | 0.085
806.746 | 0.099
806.978 | 0.027
807.211 | 0.013
807.443 | 0.028
807.675 | 0.044
background 5
AbsoluteIrradiance:211
195.562 | 0
195.934 | 0
196.307 | 0
196.679 | 0
197.051 | 0
197.423 | 0
197.795 | 0
198.167 | 0
198.539 | 0
198.911 | 0
199.283 | 0
199.654 | 0
200.026 | 0
200.397 | 0
200.769 | 0
201.14 | 0
201.511 | 0
201.882 | 0
202.253 | 0
202.624 | 0
202.995 | 0
203.366 | 0
203.736 | 0
204.107 | 0
204.477 | 0
204.848 | 3.677
205.218 | 7.649
205.588 | 2.727
205.959 | 2.498
206.329 | 2.171
206.699 | 1.998
207.068 | 3.08
207.438 | 3.175
207.808 | 2.132
208.178 | 2.647
208.547 | 1.969
208.917 | 1.957
209.286 | 2.961
209.655 | 2.551
210.025 | 1.557
210.394 | 1.746
210.763 | 1.995
211.132 | 1.219
211.501 | 1.109
211.869 | 1.262
212.238 | 1.602
212.607 | 1.977
212.975 | 1.175
213.344 | 1.133
213.712 | 0.953
214.08 | 1.063
214.448 | 0.932
214.817 | 0.912
215.185 | 1.557
215.553 | 1.058
215.92 | 1.102
216.288 | 0.974
216.656 | 1.08
217.023 | 0.722
217.391 | 0.702
217.758 | 1.265
218.126 | 0.958
218.493 | 0.842
218.86 | 0.694
219.227 | 0.618
219.594 | 0.581
219.961 | 0.953
220.328 | 0.529
220.695 | 0.658
221.061 | 0.843
221.428 | 0.807
221.795 | 0.897
222.161 | 1.032
222.527 | 0.561
222.893 | 0.93
223.26 | 0.473
223.626 | 0.872
223.992 | 0.615
224.358 | 0.524
224.723 | 0.685
225.089 | 0.256
225.455 | 0.568
225.82 | 0.423
226.186 | 0.345
226.551 | 0.239
226.917 | 0.338
227.282 | 0.301
227.647 | 0.557
228.012 | 0.425
228.377 | 0.314
228.742 | 0.348
229.107 | 0.248
229.471 | 0.309
229.836 | 0.217
230.201 | 0.292
230.565 | 0.291
230.929 | 0.257
231.294 | 0.181
231.658 | 0.181
232.022 | 0.188
232.386 | 0.116
232.75 | -0.016
233.114 | 0.128
233.478 | 0.161
233.842 | 0.175
234.205 | 0.174
234.569 | 0.135
234.932 | 0.09
235.296 | 0.095
235.659 | 0.14
236.022 | 0.106
236.385 | 0.089
236.748 | 0.137
237.111 | 0.08
237.474 | 0.165
237.837 | 0.156
238.2 | 0.107
238.562 | 0.101
238.925 | 0.157
239.287 | 0.15
239.65 | 0.084
240.012 | 0.135
240.374 | 0.104
240.736 | 0.134
241.098 | 0.108
241.46 | 0.125
241.822 | 0.139
242.184 | 0.045
242.546 | 0.121
242.907 | 0.084
243.269 | 0.057
243.63 | 0.058
243.992 | 0.119
244.353 | 0.097
244.714 | 0.078
245.075 | 0.084
245.436 | 0.095
245.797 | 0.076
246.158 | 0.03
246.519 | 0.105
246.879 | 0.095
247.24 | 0.082
247.601 | 0.098
247.961 | 0.082
248.321 | 0.046
248.682 | 0.081
249.042 | 0.062
249.402 | 0.061
249.762 | 0.113
250.122 | 0.073
250.482 | 0.105
250.842 | 0.057
251.201 | 0.082
251.561 | 0.073
251.921 | 0.081
252.28 | 0.084
252.639 | 0.078
252.999 | 0.087
253.358 | 0.077
253.717 | 0.082
254.076 | 0.077
254.435 | 0.052
254.794 | 0.031
255.153 | 0.114
255.511 | 0.087
255.87 | 0.04
256.228 | 0.078
256.587 | 0.069
256.945 | 0.066
257.304 | 0.04
257.662 | 0.057
258.02 | 0.064
258.378 | 0.078
258.736 | 0.041
259.094 | 0.106
259.452 | 0.099
259.809 | 0.07
260.167 | 0.069
260.524 | 0.06
260.882 | 0.045
261.239 | 0.065
261.597 | 0.055
261.954 | 0.069
262.311 | 0.072
262.668 | 0.062
263.025 | 0.061
263.382 | 0.089
263.739 | 0.059
264.095 | 0.083
264.452 | 0.038
264.809 | 0.062
265.165 | 0.062
265.521 | 0.094
265.878 | 0.064
266.234 | 0.064
266.59 | 0.058
266.946 | 0.096
267.302 | 0.073
267.658 | 0.076
268.014 | 0.062
268.37 | 0.066
268.725 | 0.066
269.081 | 0.054
269.436 | 0.07
269.792 | 0.028
270.147 | 0.072
270.502 | 0.084
270.857 | 0.072
271.212 | 0.068
271.567 | 0.023
271.922 | 0.093
272.277 | 0.062
272.632 | 0.08
272.987 | 0.08
273.341 | 0.073
273.696 | 0.072
274.05 | 0.036
274.404 | 0.062
274.759 | 0.09
275.113 | 0.108
275.467 | 0.045
275.821 | 0.046
276.175 | 0.066
276.529 | 0.093
276.882 | 0.111
277.236 | 0.078
277.59 | 0.088
277.943 | 0.038
278.297 | 0.122
278.65 | 0.051
279.003 | 0.049
279.356 | 0.053
279.709 | 0.073
280.062 | 0.05
280.415 | 0.072
280.768 | 0.067
281.121 | 0.083
281.474 | 0.074
281.826 | 0.064
282.179 | 0.067
282.531 | 0.043
282.884 | 0.077
283.236 | 0.061
283.588 | 0.096
283.94 | 0.049
284.292 | 0.063
284.644 | 0.049
284.996 | 0.056
285.348 | 0.052
285.7 | 0.083
286.051 | 0.062
286.403 | 0.055
286.754 | 0.069
287.106 | 0.085
287.457 | 0.072
287.808 | 0.069
288.159 | 0.072
288.51 | 0.066
288.861 | 0.057
289.212 | 0.026
289.563 | 0.038
289.914 | 0.083
290.264 | 0.039
290.615 | 0.056
290.966 | 0.072
291.316 | 0.036
291.666 | 0.072
292.017 | 0.056
292.367 | 0.028
292.717 | 0.042
293.067 | 0.043
293.417 | 0.056
293.767 | 0.045
294.116 | 0.064
294.466 | 0.052
294.816 | 0.065
295.165 | 0.034
295.515 | 0.029
295.864 | 0.055
296.213 | 0.036
296.562 | 0.028
296.911 | 0.059
297.261 | 0.04
297.609 | 0.043
297.958 | 0.016
298.307 | 0.038
298.656 | 0.038
299.004 | 0.038
299.353 | 0.003
299.701 | 0.035
300.05 | 0.028
300.398 | 0.045
300.746 | 0.04
301.095 | 0.039
301.443 | 0.032
301.791 | 0.029
302.138 | 0.037
302.486 | 0.024
302.834 | 0.025
303.182 | 0.016
303.529 | 0.036
303.877 | 0.032
304.224 | 0.035
304.572 | 0.039
304.919 | 0.019
305.266 | 0.041
305.613 | 0.029
305.96 | 0.045
306.307 | 0.03
306.654 | 0.05
307.001 | 0.042
307.347 | 0.034
307.694 | 0.018
308.04 | 0.024
308.387 | 0.036
308.733 | 0.029
309.08 | 0.03
309.426 | 0.046
309.772 | 0.035
310.118 | 0.046
310.464 | 0.022
310.81 | 0.036
311.156 | 0.031
311.501 | 0.034
311.847 | 0.03
312.193 | 0.024
312.538 | 0.029
312.883 | 0.029
313.229 | 0.025
313.574 | 0.024
313.919 | 0.029
314.264 | 0.036
314.609 | 0.038
314.954 | 0.027
315.299 | 0.025
315.644 | 0.018
315.988 | 0.034
316.333 | 0.036
316.678 | 0.028
317.022 | 0.028
317.366 | 0.025
317.711 | 0.026
318.055 | 0.02
318.399 | 0.027
318.743 | 0.028
319.087 | 0.029
319.431 | 0.036
319.775 | 0.023
320.118 | 0.041
320.462 | 0.026
320.806 | 0.024
321.149 | 0.033
321.493 | 0.026
321.836 | 0.025
322.179 | 0.01
322.522 | 0.024
322.865 | 0.025
323.208 | 0.025
323.551 | 0.034
323.894 | 0.023
324.237 | 0.019
324.58 | 0.026
324.922 | 0.025
325.265 | 0.03
325.607 | 0.031
325.95 | 0.035
326.292 | 0.015
326.634 | 0.023
326.976 | 0.016
327.318 | 0.02
327.66 | 0.032
328.002 | 0.008
328.344 | 0.024
328.686 | 0.027
329.027 | 0.015
329.369 | 0.013
329.711 | 0.03
330.052 | 0.022
330.393 | 0.015
330.735 | 0.021
331.076 | 0.023
331.417 | 0.016
331.758 | 0.024
332.099 | 0.011
332.44 | 0.027
332.78 | 0.014
333.121 | 0.023
333.462 | 0.019
333.802 | 0.025
334.143 | 0.019
334.483 | 0.019
334.823 | 0.016
335.164 | 0.026
335.504 | 0.017
335.844 | 0.028
336.184 | 0.016
336.524 | 0.014
336.864 | 0.018
337.203 | 0.018
337.543 | 0.032
337.883 | 0.025
338.222 | 0.038
338.562 | 0.014
338.901 | 0.026
339.24 | 0.011
339.58 | 0.024
339.919 | 0.014
340.258 | 0.027
340.597 | 0.019
340.936 | 0.014
341.274 | 0.022
341.613 | 0.028
341.952 | 0.021
342.29 | 0.019
342.629 | 0.028
342.967 | 0.018
343.306 | 0.022
343.644 | 0.02
343.982 | 0.019
344.32 | 0.026
344.658 | 0.017
344.996 | 0.019
345.334 | 0.022
345.672 | 0.028
346.01 | 0.014
346.347 | 0.037
346.685 | 0.026
347.022 | 0.036
347.36 | 0.032
347.697 | 0.02
348.034 | 0.029
348.371 | 0.022
348.708 | 0.021
349.045 | 0.029
349.382 | 0.026
349.719 | 0.011
350.056 | 0.023
350.393 | 0.017
350.729 | 0.034
351.066 | 0.02
351.402 | 0.024
351.739 | 0.032
352.075 | 0.028
352.411 | 0.022
352.747 | 0.02
353.083 | 0.024
353.419 | 0.023
353.755 | 0.021
354.091 | 0.027
354.427 | 0.023
354.763 | 0.024
355.098 | 0.026
355.434 | 0.017
355.769 | 0.037
356.104 | 0.012
356.44 | 0.016
356.775 | 0.026
357.11 | 0.02
357.445 | 0.02
357.78 | 0.019
358.115 | 0.024
358.45 | 0.026
358.785 | 0.028
359.119 | 0.018
359.454 | 0.02
359.788 | 0.025
360.123 | 0.024
360.457 | 0.014
360.791 | 0.031
361.126 | 0.017
361.46 | 0.014
361.794 | 0.028
362.128 | 0.029
362.462 | 0.036
362.795 | 0.031
363.129 | 0.027
363.463 | 0.013
363.796 | 0.028
364.13 | 0.033
364.463 | 0.029
364.797 | 0.028
365.13 | 0.025
365.463 | 0.015
365.796 | 0.022
366.129 | 0.023
366.462 | 0.017
366.795 | 0.022
367.128 | 0.031
367.461 | 0.017
367.793 | 0.032
368.126 | 0.022
368.458 | 0.027
368.791 | 0.027
369.123 | 0.029
369.455 | 0.025
369.787 | 0.025
370.12 | 0.021
370.452 | 0.013
370.784 | 0.02
371.115 | 0.023
371.447 | 0.014
371.779 | 0.024
372.111 | 0.024
372.442 | 0.019
372.774 | 0.034
373.105 | 0.026
373.436 | 0.017
373.768 | 0.019
374.099 | 0.012
374.43 | 0.028
374.761 | 0.016
375.092 | 0.027
375.423 | 0.02
375.754 | 0.017
376.084 | 0.018
376.415 | 0.017
376.745 | 0.019
377.076 | 0.024
377.406 | 0.023
377.737 | 0.035
378.067 | 0.029
378.397 | 0.019
378.727 | 0.018
379.057 | 0.017
379.387 | 0.018
379.717 | 0.024
380.047 | 0.024
380.377 | 0.023
380.706 | 0.028
381.036 | 0.019
381.365 | 0.033
381.695 | 0.017
382.024 | 0.016
382.353 | 0.007
382.682 | 0.023
383.012 | 0.019
383.341 | 0.028
383.67 | 0.022
383.998 | 0.017
384.327 | 0.025
384.656 | 0.014
384.985 | 0.008
385.313 | 0.014
385.642 | 0.017
385.97 | 0.016
386.298 | 0.017
386.627 | 0.016
386.955 | 0.022
387.283 | 0.016
387.611 | 0.013
387.939 | 0.018
388.267 | 0.019
388.595 | 0.026
388.922 | 0.006
389.25 | 0.031
389.578 | 0.014
389.905 | 0.02
390.233 | 0.015
390.56 | 0.021
390.887 | 0.026
391.214 | 0.021
391.542 | 0.01
391.869 | 0.014
392.196 | 0.011
392.522 | 0.012
392.849 | 0.017
393.176 | 0.011
393.503 | 0.02
393.829 | 0.016
394.156 | 0.015
394.482 | 0.019
394.809 | 0.018
395.135 | 0.021
395.461 | 0.01
395.787 | 0.022
396.113 | 0.009
396.439 | 0.016
396.765 | 0.014
397.091 | 0.01
397.417 | 0.012
397.742 | 0.014
398.068 | 0.021
398.394 | 0.016
398.719 | 0.013
399.044 | 0.013
399.37 | 0.023
399.695 | 0.013
400.02 | 0.022
400.345 | 0.027
400.67 | 0.021
400.995 | 0.027
401.32 | 0.017
401.645 | 0.017
401.969 | 0.014
402.294 | 0.014
402.618 | 0.029
402.943 | 0.009
403.267 | 0.011
403.592 | 0.015
403.916 | 0.027
404.24 | 0.016
404.564 | 0.018
404.888 | 0.031
405.212 | 0.022
405.536 | 0.011
405.86 | 0.023
406.183 | 0.017
406.507 | 0.019
406.831 | 0.023
407.154 | 0.017
407.478 | 0.015
407.801 | 0.018
408.124 | 0.01
408.447 | 0.016
408.77 | 0.021
409.093 | 0.019
409.416 | 0.013
409.739 | 0.024
410.062 | 0.007
410.385 | 0.021
410.708 | 0.021
411.03 | 0.022
411.353 | 0.013
411.675 | 0.022
411.997 | 0.019
412.32 | 0.013
412.642 | 0.019
412.964 | 0.015
413.286 | 0.015
413.608 | 0.021
413.93 | 0.022
414.252 | 0.009
414.574 | 0.011
414.895 | 0.013
415.217 | 0.023
415.538 | 0.013
415.86 | 0.017
416.181 | 0.022
416.503 | 0.018
416.824 | 0.017
417.145 | 0.014
417.466 | 0.016
417.787 | 0.024
418.108 | 0.009
418.429 | 0.015
418.75 | 0.016
419.07 | 0.015
419.391 | 0.018
419.712 | 0.019
420.032 | 0.017
420.353 | 0.015
420.673 | 0.015
420.993 | 0.016
421.313 | 0.02
421.633 | 0.021
421.953 | 0.012
422.273 | 0.017
422.593 | 0.018
422.913 | 0.017
423.233 | 0.014
423.553 | 0.016
423.872 | 0.013
424.192 | 0.015
424.511 | 0.011
424.83 | 0.021
425.15 | 0.023
425.469 | 0.019
425.788 | 0.021
426.107 | 0.02
426.426 | 0.026
426.745 | 0.027
427.064 | 0.019
427.383 | 0.021
427.702 | 0.008
428.02 | 0.021
428.339 | 0.016
428.657 | 0.018
428.976 | 0.022
429.294 | 0.018
429.612 | 0.011
429.93 | 0.021
430.248 | 0.021
430.567 | 0.016
430.884 | 0.016
431.202 | 0.013
431.52 | 0.01
431.838 | 0.013
432.156 | 0.019
432.473 | 0.019
432.791 | 0.022
433.108 | 0.016
433.426 | 0.024
433.743 | 0.019
434.06 | 0.016
434.377 | 0.024
434.694 | 0.023
435.011 | 0.012
435.328 | 0.017
435.645 | 0.014
435.962 | 0.02
436.279 | 0.017
436.595 | 0.013
436.912 | 0.012
437.228 | 0.017
437.545 | 0.014
437.861 | 0.014
438.177 | 0.012
438.493 | 0.011
438.81 | 0.019
439.126 | 0.015
439.442 | 0.02
439.757 | 0.007
440.073 | 0.012
440.389 | 0.014
440.705 | 0.024
441.02 | 0.016
441.336 | 0.013
441.651 | 0.018
441.967 | 0.011
442.282 | 0.008
442.597 | 0.018
442.912 | 0.022
443.227 | 0.017
443.543 | 0.018
443.857 | 0.019
444.172 | 0.015
444.487 | 0.018
444.802 | 0.026
445.116 | 0.024
445.431 | 0.011
445.746 | 0.011
446.06 | 0.018
446.374 | 0.014
446.689 | 0.023
447.003 | 0.018
447.317 | 0.015
447.631 | 0.008
447.945 | 0.017
448.259 | 0.021
448.573 | 0.025
448.887 | 0.013
449.2 | 0.028
449.514 | 0.022
449.827 | 0.016
450.141 | 0.018
450.454 | 0.013
450.768 | 0.017
451.081 | 0.023
451.394 | 0.016
451.707 | 0.011
452.02 | 0.023
452.333 | 0.014
452.646 | 0.013
452.959 | 0.016
453.272 | 0.017
453.584 | 0.015
453.897 | 0.017
454.21 | 0.019
454.522 | 0.014
454.834 | 0.011
455.147 | 0.015
455.459 | 0.008
455.771 | 0.004
456.083 | 0.012
456.395 | 0.017
456.707 | 0.013
457.019 | 0.018
457.331 | 0.018
457.643 | 0.015
457.954 | 0.017
458.266 | 0.012
458.577 | 0.015
458.889 | 0.013
459.2 | 0.017
459.512 | 0.012
459.823 | 0.016
460.134 | 0.016
460.445 | 0.019
460.756 | 0.016
461.067 | 0.017
461.378 | 0.014
461.689 | 0.014
461.999 | 0.029
462.31 | 0.022
462.621 | 0.012
462.931 | 0.017
463.242 | 0.02
463.552 | 0.021
463.862 | 0.012
464.172 | 0.02
464.483 | 0.025
464.793 | 0.012
465.103 | 0.017
465.413 | 0.018
465.722 | 0.015
466.032 | 0.022
466.342 | 0.016
466.652 | 0.015
466.961 | 0.017
467.271 | 0.021
467.58 | 0.017
467.889 | 0.019
468.199 | 0.02
468.508 | 0.015
468.817 | 0.022
469.126 | 0.016
469.435 | 0.023
469.744 | 0.016
470.053 | 0.015
470.362 | 0.015
470.67 | 0.015
470.979 | 0.019
471.287 | 0.014
471.596 | 0.015
471.904 | 0.026
472.213 | 0.01
472.521 | 0.022
472.829 | 0.018
473.137 | 0.02
473.445 | 0.011
473.753 | 0.026
474.061 | 0.012
474.369 | 0.017
474.677 | 0.02
474.985 | 0.012
475.292 | 0.008
475.6 | 0.022
475.907 | 0.019
476.215 | 0.016
476.522 | 0.024
476.829 | 0.016
477.137 | 0.021
477.444 | 0.013
477.751 | 0.026
478.058 | 0.008
478.365 | 0.017
478.671 | 0.024
478.978 | 0.019
479.285 | 0.012
479.592 | 0.016
479.898 | 0.014
480.205 | 0.017
480.511 | 0.005
480.817 | 0.025
481.124 | 0.017
481.43 | 0.018
481.736 | 0.029
482.042 | 0.017
482.348 | 0.013
482.654 | 0.022
482.96 | 0.022
483.266 | 0.019
483.571 | 0.02
483.877 | 0.008
484.182 | 0.009
484.488 | 0.016
484.793 | 0.011
485.099 | 0.02
485.404 | 0.016
485.709 | 0.022
486.014 | 0.017
486.319 | 0.023
486.624 | 0.015
486.929 | 0.012
487.234 | 0.008
487.539 | 0.022
487.844 | 0.02
488.148 | 0.016
488.453 | 0.029
488.757 | 0.019
489.062 | 0.018
489.366 | 0.023
489.67 | 0.016
489.975 | 0.013
490.279 | 0.014
490.583 | 0.018
490.887 | 0.013
491.191 | 0.016
491.495 | 0.02
491.798 | 0.018
492.102 | 0.021
492.406 | 0.017
492.709 | 0.019
493.013 | 0.024
493.316 | 0.016
493.62 | 0.024
493.923 | 0.017
494.226 | 0.023
494.529 | 0.018
494.832 | 0.018
495.135 | 0.016
495.438 | 0.022
495.741 | 0.033
496.044 | 0.014
496.347 | 0.013
496.649 | 0.026
496.952 | 0.017
497.255 | 0.012
497.557 | 0.016
497.859 | 0.014
498.162 | 0.006
498.464 | 0.027
498.766 | 0.02
499.068 | 0.018
499.37 | 0.017
499.672 | 0.021
499.974 | 0.015
500.276 | 0.02
500.578 | 0.021
500.879 | 0.01
501.181 | 0.023
501.483 | 0.016
501.784 | 0.024
502.085 | 0.013
502.387 | 0.012
502.688 | 0.015
502.989 | 0.022
503.29 | 0.012
503.591 | 0.012
503.892 | 0.013
504.193 | 0.027
504.494 | 0.018
504.795 | 0.007
505.096 | 0.026
505.396 | 0.012
505.697 | 0.003
505.997 | 0.017
506.298 | 0.019
506.598 | 0.012
506.898 | 0.024
507.199 | 0.015
507.499 | 0.018
507.799 | 0.014
508.099 | 0.016
508.399 | 0.02
508.699 | 0.018
508.998 | 0.015
509.298 | 0.024
509.598 | 0.013
509.897 | 0.021
510.197 | 0.013
510.496 | 0.017
510.796 | 0.009
511.095 | 0.016
511.394 | 0.02
511.693 | 0.031
511.992 | 0.018
512.291 | 0.02
512.59 | 0.02
512.889 | 0.023
513.188 | 0.022
513.487 | 0.019
513.786 | 0.023
514.084 | 0.022
514.383 | 0.023
514.681 | 0.027
514.98 | 0.017
515.278 | 0.024
515.576 | 0.029
515.874 | 0.022
516.173 | 0.013
516.471 | 0.025
516.769 | 0.019
517.067 | 0.015
517.364 | 0.014
517.662 | 0.012
517.96 | 0.014
518.257 | 0.016
518.555 | 0.016
518.853 | 0.013
519.15 | 0.015
519.447 | 0.017
519.745 | 0.014
520.042 | 0.023
520.339 | 0.011
520.636 | 0.02
520.933 | 0.028
521.23 | 0.018
521.527 | 0.018
521.824 | 0.021
522.121 | 0.012
522.417 | 0.015
522.714 | 0.009
523.01 | 0.017
523.307 | 0.016
523.603 | 0.021
523.9 | 0.018
524.196 | 0.022
524.492 | 0.009
524.788 | 0.028
525.084 | 0.02
525.38 | 0.027
525.676 | 0.012
525.972 | 0.019
526.268 | 0.021
526.564 | 0.024
526.859 | 0.022
527.155 | 0.018
527.45 | 0.02
527.746 | 0.021
528.041 | 0.025
528.336 | 0.021
528.632 | 0.013
528.927 | 0.014
529.222 | 0.016
529.517 | 0.025
529.812 | 0.018
530.107 | 0.013
530.402 | 0.018
530.696 | 0.025
530.991 | 0.026
531.286 | 0.024
531.58 | 0.023
531.875 | 0.016
532.169 | 0.023
532.463 | 0.017
532.758 | 0.012
533.052 | 0.021
533.346 | 0.014
533.64 | 0.018
533.934 | 0.026
534.228 | 0.024
534.522 | 0.024
534.816 | 0.023
535.11 | 0.012
535.403 | 0.028
535.697 | 0.028
535.99 | 0.025
536.284 | 0.02
536.577 | 0.029
536.871 | 0.02
537.164 | 0.021
537.457 | 0.024
537.75 | 0.018
538.043 | 0.023
538.336 | 0.02
538.629 | 0.014
538.922 | 0.021
539.215 | 0.008
539.507 | 0.017
539.8 | 0.009
540.093 | 0.028
540.385 | 0.014
540.678 | 0.016
540.97 | 0.006
541.262 | 0.019
541.554 | 0.017
541.847 | 0.022
542.139 | 0.022
542.431 | 0.027
542.723 | 0.022
543.015 | 0.023
543.307 | 0.026
543.598 | 0.022
543.89 | 0.027
544.182 | 0.016
544.473 | 0.023
544.765 | 0.029
545.056 | 0.011
545.347 | 0.017
545.639 | 0.012
545.93 | 0.016
546.221 | 0.019
546.512 | 0.021
546.803 | 0.014
547.094 | 0.018
547.385 | 0.021
547.676 | 0.021
547.967 | 0.031
548.257 | 0.017
548.548 | 0.026
548.839 | 0.013
549.129 | 0.015
549.419 | 0.026
549.71 | 0.02
550 | 0.026
550.29 | 0.015
550.58 | 0.01
550.871 | 0.013
551.161 | 0.014
551.45 | 0.013
551.74 | 0.017
552.03 | 0.011
552.32 | 0.019
552.61 | 0.024
552.899 | 0.029
553.189 | 0.017
553.478 | 0.018
553.768 | 0.011
554.057 | 0.029
554.346 | 0.025
554.635 | 0.022
554.925 | 0.019
555.214 | 0.028
555.503 | 0.014
555.792 | 0.026
556.081 | 0.01
556.369 | 0.02
556.658 | 0.018
556.947 | 0.029
557.235 | 0.017
557.524 | 0.032
557.812 | 0.019
558.101 | 0.019
558.389 | 0.028
558.677 | 0.024
558.966 | 0.026
559.254 | 0.03
559.542 | 0.021
559.83 | 0.029
560.118 | 0.022
560.406 | 0.018
560.693 | 0.036
560.981 | 0.029
561.269 | 0.02
561.556 | 0.018
561.844 | 0.026
562.131 | 0.023
562.419 | 0.026
562.706 | 0.024
562.994 | 0.019
563.281 | 0.03
563.568 | 0.025
563.855 | 0.032
564.142 | 0.023
564.429 | 0.029
564.716 | 0.027
565.003 | 0.032
565.289 | 0.024
565.576 | 0.028
565.863 | 0.022
566.149 | 0.027
566.436 | 0.027
566.722 | 0.02
567.008 | 0.021
567.295 | 0.031
567.581 | 0.022
567.867 | 0.022
568.153 | 0.018
568.439 | 0.027
568.725 | 0.02
569.011 | 0.029
569.297 | 0.026
569.583 | 0.024
569.868 | 0.023
570.154 | 0.02
570.439 | 0.015
570.725 | 0.024
571.01 | 0.023
571.296 | 0.027
571.581 | 0.018
571.866 | 0.029
572.151 | 0.027
572.436 | 0.018
572.721 | 0.02
573.006 | 0.021
573.291 | 0.033
573.576 | 0.025
573.861 | 0.029
574.146 | 0.021
574.43 | 0.026
574.715 | 0.029
574.999 | 0.03
575.284 | 0.028
575.568 | 0.021
575.852 | 0.025
576.137 | 0.028
576.421 | 0.035
576.705 | 0.017
576.989 | 0.026
577.273 | 0.014
577.557 | 0.024
577.841 | 0.022
578.124 | 0.032
578.408 | 0.033
578.692 | 0.016
578.975 | 0.026
579.259 | 0.019
579.542 | 0.022
579.826 | 0.023
580.109 | 0.027
580.392 | 0.031
580.676 | 0.017
580.959 | 0.028
581.242 | 0.025
581.525 | 0.026
581.808 | 0.026
582.091 | 0.02
582.373 | 0.019
582.656 | 0.031
582.939 | 0.026
583.221 | 0.01
583.504 | 0.028
583.786 | 0.023
584.069 | 0.02
584.351 | 0.024
584.633 | 0.018
584.916 | 0.013
585.198 | 0.023
585.48 | 0.022
585.762 | 0.006
586.044 | 0.017
586.326 | 0.027
586.607 | 0.021
586.889 | 0.036
587.171 | 0.035
587.453 | 0.022
587.734 | 0.028
588.016 | 0.028
588.297 | 0.014
588.578 | 0.029
588.86 | 0.027
589.141 | 0.019
589.422 | 0.025
589.703 | 0.04
589.984 | 0.018
590.265 | 0.019
590.546 | 0.025
590.827 | 0.017
591.108 | 0.015
591.389 | 0.028
591.669 | 0.021
591.95 | 0.025
592.23 | 0.027
592.511 | 0.033
592.791 | 0.024
593.071 | 0.032
593.352 | 0.018
593.632 | 0.025
593.912 | 0.028
594.192 | 0.036
594.472 | 0.028
594.752 | 0.028
595.032 | 0.022
595.312 | 0.017
595.592 | 0.025
595.871 | 0.022
596.151 | 0.024
596.43 | 0.011
596.71 | 0.023
596.989 | 0.037
597.269 | 0.021
597.548 | 0.022
597.827 | 0.036
598.106 | 0.033
598.385 | 0.029
598.665 | 0.009
598.944 | 0.01
599.222 | 0.02
599.501 | 0.025
599.78 | 0.029
600.059 | 0.026
600.337 | 0.022
600.616 | 0.022
600.895 | 0.024
601.173 | 0.052
601.451 | 0.023
601.73 | 0.024
602.008 | 0.03
602.286 | 0.014
602.564 | 0.028
602.842 | 0.023
603.12 | 0.016
603.398 | 0.021
603.676 | 0.023
603.954 | 0.018
604.232 | 0.024
604.51 | 0.02
604.787 | 0.014
605.065 | 0.024
605.342 | 0.019
605.62 | 0.012
605.897 | 0.019
606.174 | 0.033
606.452 | 0.017
606.729 | 0.032
607.006 | 0.027
607.283 | 0.023
607.56 | 0.018
607.837 | 0.011
608.114 | 0.023
608.391 | 0.055
608.667 | 0.016
608.944 | 0.028
609.221 | 0.015
609.497 | 0.022
609.774 | 0.011
610.05 | 0.031
610.326 | 0.027
610.603 | 0.026
610.879 | 0.034
611.155 | 0.025
611.431 | 0.015
611.707 | 0.025
611.983 | 0.019
612.259 | 0.033
612.535 | 0.017
612.811 | 0.023
613.087 | 0.021
613.362 | 0.02
613.638 | 0.019
613.913 | 0.021
614.189 | 0.019
614.464 | 0.023
614.74 | 0.034
615.015 | 0.011
615.29 | 0.024
615.565 | 0.028
615.84 | 0.013
616.115 | 0.031
616.39 | 0.026
616.665 | 0.026
616.94 | 0.019
617.215 | 0.031
617.49 | 0.024
617.764 | 0.03
618.039 | 0.031
618.313 | 0.031
618.588 | 0.036
618.862 | 0.023
619.137 | 0.036
619.411 | 0.021
619.685 | 0.024
619.959 | 0.02
620.233 | 0.031
620.507 | 0.022
620.781 | 0.029
621.055 | 0.019
621.329 | 0.023
621.603 | 0.027
621.876 | 0.032
622.15 | 0.026
622.424 | 0.021
622.697 | 0.022
622.971 | 0.015
623.244 | 0.015
623.517 | 0.017
623.791 | 0.027
624.064 | 0.027
624.337 | 0.029
624.61 | 0.018
624.883 | 0.028
625.156 | 0.017
625.429 | 0.022
625.702 | 0.019
625.974 | 0.033
626.247 | 0.024
626.52 | 0.024
626.792 | 0.024
627.065 | 0.026
627.337 | 0.026
627.61 | 0.024
627.882 | 0.022
628.154 | 0.026
628.427 | 0.035
628.699 | 0.028
628.971 | 0.023
629.243 | 0.014
629.515 | 0.034
629.787 | 0.022
630.058 | 0.03
630.33 | 0.034
630.602 | 0.01
630.874 | 0.032
631.145 | 0.038
631.417 | 0.017
631.688 | 0.033
631.959 | 0.022
632.231 | 0.022
632.502 | 0.033
632.773 | 0.03
633.044 | 0.043
633.316 | 0.035
633.587 | 0.013
633.857 | 0.019
634.128 | 0.041
634.399 | 0.042
634.67 | 0.016
634.941 | 0.024
635.211 | 0.03
635.482 | 0.02
635.752 | 0.025
636.023 | 0.014
636.293 | 0.027
636.564 | 0.037
636.834 | 0.023
637.104 | 0.033
637.374 | 0.021
637.644 | 0.032
637.914 | 0.028
638.184 | 0.026
638.454 | 0.024
638.724 | 0.04
638.994 | 0.027
639.264 | 0.034
639.533 | 0.023
639.803 | 0.023
640.072 | 0.026
640.342 | 0.023
640.611 | 0.035
640.881 | 0.026
641.15 | 0.046
641.419 | 0.017
641.688 | 0.026
641.957 | 0.024
642.227 | 0.032
642.496 | 0.034
642.764 | 0.033
643.033 | 0.038
643.302 | 0.024
643.571 | 0.021
643.84 | 0.021
644.108 | 0.033
644.377 | 0.012
644.645 | 0.021
644.914 | 0.022
645.182 | 0.036
645.45 | 0.032
645.719 | 0.041
645.987 | 0.034
646.255 | 0.025
646.523 | 0.024
646.791 | 0.025
647.059 | 0.041
647.327 | 0.03
647.595 | 0.029
647.862 | 0.014
648.13 | 0.021
648.398 | 0.037
648.665 | 0.026
648.933 | 0.043
649.2 | 0.031
649.468 | 0.026
649.735 | 0.036
650.002 | 0.03
650.269 | 0.029
650.536 | 0.022
650.804 | 0.022
651.071 | 0.044
651.338 | 0.024
651.604 | 0.027
651.871 | 0.02
652.138 | 0.028
652.405 | 0.013
652.671 | 0.031
652.938 | 0.028
653.205 | 0.032
653.471 | 0.015
653.737 | 0.03
654.004 | 0.014
654.27 | 0.021
654.536 | 0.027
654.803 | 0.023
655.069 | 0.026
655.335 | 0.023
655.601 | 0.034
655.867 | 0.028
656.132 | 0.034
656.398 | 0.027
656.664 | 0.03
656.93 | 0.021
657.195 | 0.02
657.461 | 0.025
657.726 | 0.025
657.992 | 0.01
658.257 | 0.026
658.523 | 0.029
658.788 | 0.017
659.053 | 0.026
659.318 | 0.023
659.583 | 0.02
659.848 | 0.02
660.113 | 0.026
660.378 | 0.025
660.643 | 0.024
660.908 | 0.012
661.172 | 0.011
661.437 | 0.021
661.702 | 0.027
661.966 | 0.033
662.231 | 0.022
662.495 | 0.033
662.76 | 0.044
663.024 | 0.022
663.288 | 0.017
663.552 | 0.015
663.816 | 0.025
664.08 | 0.034
664.344 | 0.025
664.608 | 0.03
664.872 | 0.021
665.136 | 0.016
665.4 | 0.029
665.664 | 0.031
665.927 | 0.028
666.191 | 0.024
666.454 | 0.016
666.718 | 0.022
666.981 | 0.013
667.244 | 0.021
667.508 | 0.039
667.771 | 0.024
668.034 | 0.023
668.297 | 0.019
668.56 | 0.028
668.823 | 0.013
669.086 | 0.036
669.349 | 0.019
669.612 | 0.022
669.875 | 0.016
670.137 | 0.031
670.4 | 0.023
670.662 | 0.02
670.925 | 0.026
671.187 | 0.037
671.45 | 0.007
671.712 | 0.041
671.974 | 0.032
672.237 | 0.019
672.499 | 0.028
672.761 | 0.028
673.023 | 0.024
673.285 | 0.027
673.547 | 0.03
673.808 | 0.024
674.07 | 0.026
674.332 | 0.026
674.594 | 0.018
674.855 | 0.027
675.117 | 0.018
675.378 | 0.034
675.64 | 0.021
675.901 | 0.004
676.162 | 0.028
676.424 | 0.02
676.685 | 0.031
676.946 | 0.029
677.207 | 0.038
677.468 | 0.039
677.729 | 0.015
677.99 | 0.018
678.251 | 0.03
678.512 | 0.027
678.772 | 0.029
679.033 | 0.041
679.294 | 0.024
679.554 | 0.028
679.815 | 0.028
680.075 | 0.02
680.335 | 0.03
680.596 | 0.022
680.856 | 0.026
681.116 | 0.046
681.376 | 0.032
681.636 | 0.025
681.896 | 0.04
682.156 | 0.034
682.416 | 0.025
682.676 | 0.023
682.936 | 0.033
683.196 | 0.027
683.455 | 0.021
683.715 | 0.024
683.974 | 0.027
684.234 | 0.021
684.493 | 0.014
684.753 | 0.031
685.012 | 0.035
685.271 | 0.025
685.53 | 0.026
685.789 | 0.035
686.049 | 0.031
686.308 | 0.023
686.567 | 0.027
686.825 | 0.041
687.084 | 0.035
687.343 | 0.036
687.602 | 0.036
687.86 | 0.029
688.119 | 0.026
688.378 | 0.038
688.636 | 0.029
688.894 | 0.028
689.153 | 0.028
689.411 | 0.035
689.669 | 0.029
689.928 | 0.037
690.186 | 0.013
690.444 | 0.039
690.702 | 0.035
690.96 | 0.04
691.218 | 0.02
691.476 | 0.037
691.733 | 0.03
691.991 | 0.031
692.249 | 0.021
692.506 | 0.026
692.764 | 0.032
693.021 | 0.039
693.279 | 0.04
693.536 | 0.034
693.794 | 0.014
694.051 | 0.024
694.308 | 0.034
694.565 | 0.042
694.822 | 0.045
695.079 | 0.018
695.336 | 0.034
695.593 | 0.025
695.85 | 0.029
696.107 | 0.022
696.364 | 0.028
696.62 | 0.027
696.877 | 0.035
697.133 | 0.023
697.39 | 0.042
697.646 | 0.036
697.903 | 0.025
698.159 | 0.012
698.415 | 0.025
698.672 | 0.036
698.928 | 0.025
699.184 | 0.04
699.44 | 0.024
699.696 | 0.043
699.952 | 0.015
700.208 | 0.025
700.464 | 0.023
700.719 | 0.042
700.975 | 0.04
701.231 | 0.037
701.486 | 0.04
701.742 | 0.034
701.997 | 0.037
702.253 | 0.039
702.508 | 0.013
702.763 | 0.033
703.019 | 0.034
703.274 | 0.04
703.529 | 0.032
703.784 | 0.026
704.039 | 0.052
704.294 | 0.028
704.549 | 0.02
704.804 | 0.024
705.058 | 0.033
705.313 | 0.025
705.568 | 0.047
705.822 | 0.039
706.077 | 0.027
706.331 | 0.025
706.586 | 0.029
706.84 | 0.041
707.095 | 0.028
707.349 | 0.004
707.603 | 0.023
707.857 | 0.037
708.111 | 0.023
708.365 | 0.024
708.619 | 0.025
708.873 | 0.042
709.127 | 0.026
709.381 | 0.018
709.635 | 0.031
709.888 | 0.044
710.142 | 0.026
710.396 | 0.033
710.649 | 0.032
710.903 | 0.022
711.156 | 0.034
711.409 | 0.046
711.663 | 0.043
711.916 | 0.025
712.169 | 0.023
712.422 | 0.032
712.675 | 0.033
712.928 | 0.052
713.181 | 0.039
713.434 | 0.034
713.687 | 0.048
713.94 | 0.029
714.192 | 0.034
714.445 | 0.002
714.698 | 0.041
714.95 | 0.041
715.203 | 0.059
715.455 | 0.037
715.708 | 0.044
715.96 | 0.051
716.212 | 0.03
716.464 | 0.047
716.717 | 0.044
716.969 | 0.039
717.221 | 0.045
717.473 | 0.018
717.725 | 0.036
717.977 | 0.039
718.228 | 0.047
718.48 | 0.032
718.732 | 0.037
718.984 | 0.05
719.235 | 0.05
719.487 | 0.039
719.738 | 0.026
719.99 | 0.051
720.241 | 0.033
720.492 | 0.042
720.744 | 0.031
720.995 | 0.029
721.246 | 0.034
721.497 | 0.049
721.748 | 0.033
721.999 | 0.032
722.25 | 0.06
722.501 | 0.053
722.752 | 0.045
723.002 | 0.025
723.253 | 0.024
723.504 | 0.052
723.754 | 0.039
724.005 | 0.051
724.255 | 0.045
724.506 | 0.045
724.756 | 0.05
725.006 | 0.032
725.257 | 0.033
725.507 | 0.039
725.757 | 0.029
726.007 | 0.031
726.257 | 0.047
726.507 | 0.027
726.757 | 0.076
727.007 | 0.051
727.257 | 0.03
727.506 | 0.045
727.756 | 0.035
728.006 | 0.046
728.255 | 0.047
728.505 | 0.023
728.754 | 0.04
729.004 | 0.053
729.253 | 0.077
729.502 | 0.061
729.751 | 0.044
730.001 | 0.024
730.25 | 0.036
730.499 | 0.062
730.748 | 0.039
730.997 | 0.034
731.246 | 0.05
731.495 | 0.038
731.743 | 0.045
731.992 | 0.044
732.241 | 0.025
732.489 | 0.047
732.738 | 0.069
732.987 | 0.041
733.235 | 0.038
733.483 | 0.048
733.732 | 0.056
733.98 | 0.024
734.228 | 0.032
734.477 | 0.033
734.725 | 0.056
734.973 | 0.037
735.221 | 0.053
735.469 | 0.053
735.717 | 0.049
735.965 | 0.04
736.212 | 0.039
736.46 | 0.055
736.708 | 0.047
736.955 | 0.066
737.203 | 0.035
737.451 | 0.065
737.698 | 0.051
737.945 | 0.042
738.193 | 0.048
738.44 | 0.04
738.687 | 0.043
738.935 | 0.044
739.182 | 0.037
739.429 | 0.047
739.676 | 0.032
739.923 | 0.051
740.17 | 0.048
740.417 | 0.039
740.663 | 0.029
740.91 | 0.048
741.157 | 0.053
741.403 | 0.042
741.65 | 0.014
741.897 | 0.047
742.143 | 0.031
742.389 | 0.054
742.636 | 0.028
742.882 | 0.036
743.128 | 0.038
743.375 | 0.026
743.621 | 0.052
743.867 | 0.046
744.113 | 0.047
744.359 | 0.032
744.605 | 0.054
744.851 | 0.072
745.097 | 0.061
745.342 | 0.038
745.588 | 0.047
745.834 | 0.043
746.079 | 0.061
746.325 | 0.05
746.57 | 0.058
746.816 | 0.023
747.061 | 0.044
747.307 | 0.052
747.552 | 0.043
747.797 | 0.058
748.042 | 0.046
748.287 | 0.04
748.532 | 0.033
748.777 | 0.045
749.022 | 0.056
749.267 | 0.055
749.512 | 0.036
749.757 | 0.041
750.002 | 0.045
750.246 | 0.045
750.491 | 0.052
750.736 | 0.035
750.98 | 0.058
751.225 | 0.044
751.469 | 0.042
751.713 | 0.054
751.958 | 0.054
752.202 | 0.034
752.446 | 0.039
752.69 | 0.024
752.934 | 0.032
753.178 | 0.052
753.422 | 0.042
753.666 | 0.059
753.91 | 0.025
754.154 | 0.061
754.398 | 0.078
754.641 | 0.03
754.885 | 0.019
755.129 | 0.048
755.372 | 0.026
755.616 | 0.051
755.859 | 0.048
756.103 | 0.067
756.346 | 0.014
756.589 | 0.043
756.832 | 0.043
757.076 | 0.055
757.319 | 0.054
757.562 | 0.057
757.805 | 0.055
758.048 | 0.058
758.291 | 0.045
758.533 | 0.046
758.776 | 0.044
759.019 | 0.059
759.262 | 0.058
759.504 | 0.048
759.747 | 0.042
759.989 | 0.057
760.232 | 0.067
760.474 | 0.038
760.717 | 0.084
760.959 | 0.032
761.201 | 0.034
761.444 | 0.042
761.686 | 0.062
761.928 | 0.043
762.17 | 0.051
762.412 | 0.062
762.654 | 0.054
762.896 | 0.042
763.137 | 0.049
763.379 | 0.035
763.621 | 0.063
763.863 | 0.044
764.104 | 0.052
764.346 | 0.069
764.587 | 0.037
764.829 | 0.05
765.07 | 0.038
765.312 | 0.066
765.553 | 0.072
765.794 | 0.057
766.035 | 0.057
766.276 | 0.083
766.518 | 0.052
766.759 | 0.071
767 | 0.049
767.241 | 0.053
767.481 | 0.049
767.722 | 0.055
767.963 | 0.059
768.204 | 0.058
768.444 | 0.072
768.685 | 0.057
768.926 | 0.035
769.166 | 0.042
769.407 | 0.046
769.647 | 0.078
769.887 | 0.037
770.128 | 0.057
770.368 | 0.054
770.608 | 0.073
770.848 | 0.057
771.088 | 0.052
771.328 | 0.06
771.568 | 0.039
771.808 | 0.062
772.048 | 0.053
772.288 | 0.06
772.528 | 0.042
772.767 | 0.057
773.007 | 0.061
773.247 | 0.066
773.486 | 0.054
773.726 | 0.039
773.965 | 0.072
774.205 | 0.061
774.444 | 0.063
774.683 | 0.075
774.922 | 0.074
775.162 | 0.056
775.401 | 0.076
775.64 | 0.055
775.879 | 0.066
776.118 | 0.054
776.357 | 0.055
776.596 | 0.075
776.834 | 0.069
777.073 | 0.064
777.312 | 0.068
777.551 | 0.044
777.789 | 0.043
778.028 | 0.056
778.266 | 0.064
778.505 | 0.04
778.743 | 0.045
778.982 | 0.033
779.22 | 0.048
779.458 | 0.068
779.696 | 0.069
779.934 | 0.068
780.173 | 0.059
780.411 | 0.056
780.649 | 0.055
780.887 | 0.065
781.124 | 0.071
781.362 | 0.05
781.6 | 0.066
781.838 | 0.048
782.075 | 0.045
782.313 | 0.049
782.551 | 0.03
782.788 | 0.048
783.026 | 0.058
783.263 | 0.052
783.5 | 0.048
783.738 | 0.062
783.975 | 0.043
784.212 | 0.009
784.449 | 0.052
784.686 | 0.052
784.924 | 0.06
785.161 | 0.042
785.397 | 0.055
785.634 | 0.063
785.871 | 0.092
786.108 | 0.058
786.345 | 0.088
786.581 | 0.046
786.818 | 0.07
787.055 | 0.066
787.291 | 0.078
787.528 | 0.09
787.764 | 0.059
788.001 | 0.044
788.237 | 0.072
788.473 | 0.041
788.709 | 0.055
788.946 | 0.065
789.182 | 0.04
789.418 | 0.032
789.654 | 0.075
789.89 | 0.042
790.126 | 0.082
790.362 | 0.095
790.597 | 0.041
790.833 | 0.062
791.069 | 0.043
791.305 | 0.038
791.54 | 0.064
791.776 | 0.075
792.011 | 0.073
792.247 | 0.026
792.482 | 0.061
792.717 | 0.067
792.953 | 0.058
793.188 | 0.075
793.423 | 0.074
793.658 | 0.077
793.894 | 0.085
794.129 | 0.049
794.364 | 0.06
794.599 | 0.07
794.833 | 0.045
795.068 | 0.063
795.303 | 0.067
795.538 | 0.037
795.772 | 0.068
796.007 | 0.066
796.242 | 0.085
796.476 | 0.068
796.711 | 0.054
796.945 | 0.053
797.18 | 0.042
797.414 | 0.035
797.648 | 0.09
797.882 | 0.053
798.117 | 0.048
798.351 | 0.045
798.585 | 0.108
798.819 | 0.041
799.053 | 0.073
799.287 | 0.053
799.521 | 0.08
799.755 | 0.081
799.988 | 0.064
800.222 | 0.068
800.456 | 0.056
800.689 | 0.029
800.923 | 0.04
801.156 | 0.078
801.39 | 0.054
801.623 | 0.058
801.857 | 0.055
802.09 | 0.068
802.323 | 0.071
802.557 | 0.055
802.79 | 0.049
803.023 | 0.053
803.256 | 0.042
803.489 | 0.055
803.722 | 0.042
803.955 | 0.047
804.188 | 0.085
804.421 | 0.054
804.653 | 0.064
804.886 | 0.06
805.119 | 0.03
805.351 | 0.049
805.584 | 0.07
805.816 | 0.032
806.049 | 0.057
806.281 | 0.068
806.514 | 0.102
806.746 | 0.086
806.978 | 0.033
807.211 | 0.015
807.443 | 0.039
807.675 | 0.061
background 6
AbsoluteIrradiance:211
195.562 | 0
195.934 | 0
196.307 | 0
196.679 | 0
197.051 | 0
197.423 | 0
197.795 | 0
198.167 | 0
198.539 | 0
198.911 | 0
199.283 | 0
199.654 | 0
200.026 | 0
200.397 | 0
200.769 | 0
201.14 | 0
201.511 | 0
201.882 | 0
202.253 | 0
202.624 | 0
202.995 | 0
203.366 | 0
203.736 | 0
204.107 | 0
204.477 | 0
204.848 | 4.275
205.218 | 7.649
205.588 | 2.423
205.959 | 2.76
206.329 | 3.275
206.699 | 2.503
207.068 | 3.408
207.438 | 3.699
207.808 | 2.068
208.178 | 2.643
208.547 | 1.852
208.917 | 1.762
209.286 | 3.236
209.655 | 2.724
210.025 | 1.69
210.394 | 1.789
210.763 | 2.201
211.132 | 1.354
211.501 | 1.189
211.869 | 1.239
212.238 | 1.558
212.607 | 1.984
212.975 | 1.255
213.344 | 0.932
213.712 | 0.872
214.08 | 1.282
214.448 | 1.069
214.817 | 0.974
215.185 | 1.527
215.553 | 1.189
215.92 | 1.229
216.288 | 1.198
216.656 | 1.363
217.023 | 0.811
217.391 | 0.745
217.758 | 1.185
218.126 | 0.995
218.493 | 0.873
218.86 | 0.715
219.227 | 0.755
219.594 | 0.49
219.961 | 0.961
220.328 | 0.579
220.695 | 0.808
221.061 | 0.9
221.428 | 0.91
221.795 | 0.83
222.161 | 0.895
222.527 | 0.612
222.893 | 0.889
223.26 | 0.422
223.626 | 0.91
223.992 | 0.599
224.358 | 0.452
224.723 | 0.653
225.089 | 0.205
225.455 | 0.568
225.82 | 0.443
226.186 | 0.371
226.551 | 0.176
226.917 | 0.413
227.282 | 0.392
227.647 | 0.525
228.012 | 0.391
228.377 | 0.358
228.742 | 0.319
229.107 | 0.22
229.471 | 0.311
229.836 | 0.215
230.201 | 0.34
230.565 | 0.319
230.929 | 0.211
231.294 | 0.186
231.658 | 0.196
232.022 | 0.175
232.386 | 0.129
232.75 | -0.018
233.114 | 0.161
233.478 | 0.159
233.842 | 0.186
234.205 | 0.178
234.569 | 0.122
234.932 | 0.113
235.296 | 0.093
235.659 | 0.126
236.022 | 0.106
236.385 | 0.099
236.748 | 0.161
237.111 | 0.086
237.474 | 0.173
237.837 | 0.165
238.2 | 0.114
238.562 | 0.103
238.925 | 0.158
239.287 | 0.156
239.65 | 0.077
240.012 | 0.129
240.374 | 0.12
240.736 | 0.139
241.098 | 0.12
241.46 | 0.119
241.822 | 0.162
242.184 | 0.068
242.546 | 0.125
242.907 | 0.09
243.269 | 0.078
243.63 | 0.054
243.992 | 0.129
244.353 | 0.091
244.714 | 0.081
245.075 | 0.114
245.436 | 0.098
245.797 | 0.079
246.158 | 0.023
246.519 | 0.095
246.879 | 0.109
247.24 | 0.101
247.601 | 0.091
247.961 | 0.106
248.321 | 0.057
248.682 | 0.076
249.042 | 0.063
249.402 | 0.079
249.762 | 0.111
250.122 | 0.068
250.482 | 0.124
250.842 | 0.064
251.201 | 0.082
251.561 | 0.065
251.921 | 0.091
252.28 | 0.086
252.639 | 0.072
252.999 | 0.08
253.358 | 0.077
253.717 | 0.07
254.076 | 0.08
254.435 | 0.06
254.794 | 0.034
255.153 | 0.108
255.511 | 0.086
255.87 | 0.037
256.228 | 0.075
256.587 | 0.078
256.945 | 0.061
257.304 | 0.032
257.662 | 0.067
258.02 | 0.062
258.378 | 0.08
258.736 | 0.048
259.094 | 0.112
259.452 | 0.1
259.809 | 0.066
260.167 | 0.081
260.524 | 0.071
260.882 | 0.041
261.239 | 0.057
261.597 | 0.061
261.954 | 0.062
262.311 | 0.083
262.668 | 0.08
263.025 | 0.069
263.382 | 0.088
263.739 | 0.063
264.095 | 0.085
264.452 | 0.039
264.809 | 0.078
265.165 | 0.068
265.521 | 0.094
265.878 | 0.078
266.234 | 0.071
266.59 | 0.057
266.946 | 0.111
267.302 | 0.06
267.658 | 0.073
268.014 | 0.065
268.37 | 0.074
268.725 | 0.063
269.081 | 0.075
269.436 | 0.077
269.792 | 0.039
270.147 | 0.076
270.502 | 0.076
270.857 | 0.078
271.212 | 0.052
271.567 | 0.03
271.922 | 0.109
272.277 | 0.073
272.632 | 0.086
272.987 | 0.083
273.341 | 0.084
273.696 | 0.074
274.05 | 0.043
274.404 | 0.067
274.759 | 0.078
275.113 | 0.101
275.467 | 0.049
275.821 | 0.045
276.175 | 0.055
276.529 | 0.082
276.882 | 0.103
277.236 | 0.072
277.59 | 0.097
277.943 | 0.036
278.297 | 0.127
278.65 | 0.051
279.003 | 0.045
279.356 | 0.059
279.709 | 0.072
280.062 | 0.051
280.415 | 0.079
280.768 | 0.063
281.121 | 0.094
281.474 | 0.075
281.826 | 0.07
282.179 | 0.065
282.531 | 0.051
282.884 | 0.078
283.236 | 0.07
283.588 | 0.093
283.94 | 0.049
284.292 | 0.067
284.644 | 0.052
284.996 | 0.063
285.348 | 0.062
285.7 | 0.082
286.051 | 0.058
286.403 | 0.058
286.754 | 0.058
287.106 | 0.084
287.457 | 0.082
287.808 | 0.065
288.159 | 0.074
288.51 | 0.066
288.861 | 0.052
289.212 | 0.035
289.563 | 0.046
289.914 | 0.072
290.264 | 0.039
290.615 | 0.056
290.966 | 0.072
291.316 | 0.037
291.666 | 0.064
292.017 | 0.055
292.367 | 0.031
292.717 | 0.037
293.067 | 0.051
293.417 | 0.051
293.767 | 0.046
294.116 | 0.073
294.466 | 0.05
294.816 | 0.058
295.165 | 0.046
295.515 | 0.039
295.864 | 0.048
296.213 | 0.035
296.562 | 0.032
296.911 | 0.06
297.261 | 0.044
297.609 | 0.045
297.958 | 0.021
298.307 | 0.039
298.656 | 0.04
299.004 | 0.033
299.353 | 0.007
299.701 | 0.04
300.05 | 0.031
300.398 | 0.048
300.746 | 0.038
301.095 | 0.036
301.443 | 0.035
301.791 | 0.025
302.138 | 0.035
302.486 | 0.022
302.834 | 0.023
303.182 | 0.02
303.529 | 0.034
303.877 | 0.028
304.224 | 0.039
304.572 | 0.035
304.919 | 0.017
305.266 | 0.043
305.613 | 0.026
305.96 | 0.049
306.307 | 0.031
306.654 | 0.045
307.001 | 0.042
307.347 | 0.031
307.694 | 0.019
308.04 | 0.023
308.387 | 0.039
308.733 | 0.028
309.08 | 0.025
309.426 | 0.048
309.772 | 0.038
310.118 | 0.038
310.464 | 0.025
310.81 | 0.032
311.156 | 0.025
311.501 | 0.034
311.847 | 0.031
312.193 | 0.026
312.538 | 0.038
312.883 | 0.03
313.229 | 0.026
313.574 | 0.024
313.919 | 0.03
314.264 | 0.032
314.609 | 0.036
314.954 | 0.025
315.299 | 0.028
315.644 | 0.022
315.988 | 0.036
316.333 | 0.038
316.678 | 0.028
317.022 | 0.027
317.366 | 0.023
317.711 | 0.028
318.055 | 0.02
318.399 | 0.023
318.743 | 0.03
319.087 | 0.03
319.431 | 0.033
319.775 | 0.027
320.118 | 0.04
320.462 | 0.027
320.806 | 0.025
321.149 | 0.036
321.493 | 0.027
321.836 | 0.025
322.179 | 0.012
322.522 | 0.028
322.865 | 0.03
323.208 | 0.027
323.551 | 0.036
323.894 | 0.023
324.237 | 0.022
324.58 | 0.026
324.922 | 0.03
325.265 | 0.037
325.607 | 0.03
325.95 | 0.033
326.292 | 0.017
326.634 | 0.023
326.976 | 0.021
327.318 | 0.022
327.66 | 0.03
328.002 | 0.012
328.344 | 0.03
328.686 | 0.031
329.027 | 0.016
329.369 | 0.016
329.711 | 0.028
330.052 | 0.029
330.393 | 0.015
330.735 | 0.02
331.076 | 0.023
331.417 | 0.019
331.758 | 0.025
332.099 | 0.014
332.44 | 0.031
332.78 | 0.017
333.121 | 0.029
333.462 | 0.024
333.802 | 0.031
334.143 | 0.026
334.483 | 0.018
334.823 | 0.019
335.164 | 0.031
335.504 | 0.018
335.844 | 0.026
336.184 | 0.014
336.524 | 0.023
336.864 | 0.02
337.203 | 0.021
337.543 | 0.033
337.883 | 0.026
338.222 | 0.034
338.562 | 0.019
338.901 | 0.026
339.24 | 0.012
339.58 | 0.028
339.919 | 0.018
340.258 | 0.032
340.597 | 0.02
340.936 | 0.014
341.274 | 0.024
341.613 | 0.03
341.952 | 0.022
342.29 | 0.022
342.629 | 0.029
342.967 | 0.021
343.306 | 0.025
343.644 | 0.024
343.982 | 0.018
344.32 | 0.028
344.658 | 0.018
344.996 | 0.024
345.334 | 0.022
345.672 | 0.033
346.01 | 0.015
346.347 | 0.038
346.685 | 0.024
347.022 | 0.032
347.36 | 0.031
347.697 | 0.025
348.034 | 0.027
348.371 | 0.025
348.708 | 0.022
349.045 | 0.029
349.382 | 0.028
349.719 | 0.019
350.056 | 0.024
350.393 | 0.02
350.729 | 0.034
351.066 | 0.024
351.402 | 0.026
351.739 | 0.034
352.075 | 0.029
352.411 | 0.022
352.747 | 0.022
353.083 | 0.024
353.419 | 0.023
353.755 | 0.025
354.091 | 0.029
354.427 | 0.025
354.763 | 0.025
355.098 | 0.028
355.434 | 0.017
355.769 | 0.039
356.104 | 0.015
356.44 | 0.02
356.775 | 0.026
357.11 | 0.024
357.445 | 0.025
357.78 | 0.018
358.115 | 0.027
358.45 | 0.028
358.785 | 0.027
359.119 | 0.016
359.454 | 0.023
359.788 | 0.023
360.123 | 0.024
360.457 | 0.016
360.791 | 0.035
361.126 | 0.017
361.46 | 0.016
361.794 | 0.029
362.128 | 0.031
362.462 | 0.037
362.795 | 0.035
363.129 | 0.025
363.463 | 0.014
363.796 | 0.028
364.13 | 0.035
364.463 | 0.033
364.797 | 0.028
365.13 | 0.03
365.463 | 0.016
365.796 | 0.019
366.129 | 0.026
366.462 | 0.015
366.795 | 0.024
367.128 | 0.029
367.461 | 0.023
367.793 | 0.032
368.126 | 0.024
368.458 | 0.032
368.791 | 0.026
369.123 | 0.027
369.455 | 0.026
369.787 | 0.023
370.12 | 0.017
370.452 | 0.016
370.784 | 0.02
371.115 | 0.024
371.447 | 0.015
371.779 | 0.024
372.111 | 0.023
372.442 | 0.019
372.774 | 0.031
373.105 | 0.026
373.436 | 0.021
373.768 | 0.02
374.099 | 0.012
374.43 | 0.028
374.761 | 0.019
375.092 | 0.024
375.423 | 0.021
375.754 | 0.02
376.084 | 0.019
376.415 | 0.018
376.745 | 0.022
377.076 | 0.027
377.406 | 0.022
377.737 | 0.036
378.067 | 0.025
378.397 | 0.02
378.727 | 0.02
379.057 | 0.016
379.387 | 0.019
379.717 | 0.02
380.047 | 0.024
380.377 | 0.021
380.706 | 0.027
381.036 | 0.018
381.365 | 0.032
381.695 | 0.018
382.024 | 0.018
382.353 | 0.008
382.682 | 0.021
383.012 | 0.018
383.341 | 0.026
383.67 | 0.025
383.998 | 0.015
384.327 | 0.028
384.656 | 0.016
384.985 | 0.013
385.313 | 0.011
385.642 | 0.016
385.97 | 0.016
386.298 | 0.016
386.627 | 0.015
386.955 | 0.029
387.283 | 0.015
387.611 | 0.013
387.939 | 0.017
388.267 | 0.019
388.595 | 0.026
388.922 | 0.009
389.25 | 0.028
389.578 | 0.019
389.905 | 0.021
390.233 | 0.015
390.56 | 0.022
390.887 | 0.032
391.214 | 0.021
391.542 | 0.012
391.869 | 0.018
392.196 | 0.012
392.522 | 0.011
392.849 | 0.015
393.176 | 0.012
393.503 | 0.017
393.829 | 0.016
394.156 | 0.014
394.482 | 0.021
394.809 | 0.021
395.135 | 0.022
395.461 | 0.013
395.787 | 0.023
396.113 | 0.012
396.439 | 0.016
396.765 | 0.015
397.091 | 0.016
397.417 | 0.014
397.742 | 0.017
398.068 | 0.022
398.394 | 0.019
398.719 | 0.018
399.044 | 0.016
399.37 | 0.023
399.695 | 0.017
400.02 | 0.022
400.345 | 0.024
400.67 | 0.022
400.995 | 0.028
401.32 | 0.021
401.645 | 0.019
401.969 | 0.015
402.294 | 0.016
402.618 | 0.029
402.943 | 0.008
403.267 | 0.012
403.592 | 0.016
403.916 | 0.027
404.24 | 0.02
404.564 | 0.016
404.888 | 0.031
405.212 | 0.027
405.536 | 0.011
405.86 | 0.02
406.183 | 0.017
406.507 | 0.022
406.831 | 0.023
407.154 | 0.019
407.478 | 0.017
407.801 | 0.018
408.124 | 0.012
408.447 | 0.01
408.77 | 0.021
409.093 | 0.021
409.416 | 0.015
409.739 | 0.022
410.062 | 0.01
410.385 | 0.02
410.708 | 0.022
411.03 | 0.024
411.353 | 0.011
411.675 | 0.021
411.997 | 0.018
412.32 | 0.011
412.642 | 0.02
412.964 | 0.016
413.286 | 0.016
413.608 | 0.02
413.93 | 0.021
414.252 | 0.01
414.574 | 0.01
414.895 | 0.015
415.217 | 0.023
415.538 | 0.013
415.86 | 0.017
416.181 | 0.024
416.503 | 0.016
416.824 | 0.017
417.145 | 0.014
417.466 | 0.017
417.787 | 0.028
418.108 | 0.012
418.429 | 0.016
418.75 | 0.015
419.07 | 0.013
419.391 | 0.018
419.712 | 0.018
420.032 | 0.016
420.353 | 0.017
420.673 | 0.015
420.993 | 0.018
421.313 | 0.021
421.633 | 0.021
421.953 | 0.015
422.273 | 0.02
422.593 | 0.019
422.913 | 0.016
423.233 | 0.016
423.553 | 0.015
423.872 | 0.012
424.192 | 0.016
424.511 | 0.012
424.83 | 0.02
425.15 | 0.028
425.469 | 0.021
425.788 | 0.024
426.107 | 0.019
426.426 | 0.025
426.745 | 0.029
427.064 | 0.022
427.383 | 0.022
427.702 | 0.014
428.02 | 0.024
428.339 | 0.016
428.657 | 0.018
428.976 | 0.022
429.294 | 0.017
429.612 | 0.011
429.93 | 0.02
430.248 | 0.024
430.567 | 0.014
430.884 | 0.016
431.202 | 0.015
431.52 | 0.016
431.838 | 0.013
432.156 | 0.018
432.473 | 0.019
432.791 | 0.024
433.108 | 0.017
433.426 | 0.028
433.743 | 0.02
434.06 | 0.016
434.377 | 0.023
434.694 | 0.026
435.011 | 0.016
435.328 | 0.017
435.645 | 0.016
435.962 | 0.017
436.279 | 0.016
436.595 | 0.013
436.912 | 0.016
437.228 | 0.016
437.545 | 0.013
437.861 | 0.011
438.177 | 0.013
438.493 | 0.009
438.81 | 0.02
439.126 | 0.016
439.442 | 0.021
439.757 | 0.004
440.073 | 0.015
440.389 | 0.014
440.705 | 0.026
441.02 | 0.015
441.336 | 0.013
441.651 | 0.019
441.967 | 0.013
442.282 | 0.006
442.597 | 0.015
442.912 | 0.022
443.227 | 0.019
443.543 | 0.019
443.857 | 0.021
444.172 | 0.014
444.487 | 0.016
444.802 | 0.023
445.116 | 0.024
445.431 | 0.009
445.746 | 0.01
446.06 | 0.015
446.374 | 0.014
446.689 | 0.024
447.003 | 0.015
447.317 | 0.017
447.631 | 0.006
447.945 | 0.019
448.259 | 0.022
448.573 | 0.026
448.887 | 0.014
449.2 | 0.024
449.514 | 0.021
449.827 | 0.018
450.141 | 0.018
450.454 | 0.016
450.768 | 0.017
451.081 | 0.024
451.394 | 0.015
451.707 | 0.014
452.02 | 0.024
452.333 | 0.014
452.646 | 0.015
452.959 | 0.015
453.272 | 0.014
453.584 | 0.014
453.897 | 0.015
454.21 | 0.021
454.522 | 0.016
454.834 | 0.016
455.147 | 0.016
455.459 | 0.009
455.771 | 0.006
456.083 | 0.012
456.395 | 0.013
456.707 | 0.017
457.019 | 0.021
457.331 | 0.018
457.643 | 0.014
457.954 | 0.018
458.266 | 0.016
458.577 | 0.015
458.889 | 0.013
459.2 | 0.018
459.512 | 0.014
459.823 | 0.017
460.134 | 0.018
460.445 | 0.02
460.756 | 0.014
461.067 | 0.014
461.378 | 0.01
461.689 | 0.014
461.999 | 0.031
462.31 | 0.019
462.621 | 0.013
462.931 | 0.018
463.242 | 0.018
463.552 | 0.021
463.862 | 0.012
464.172 | 0.02
464.483 | 0.023
464.793 | 0.015
465.103 | 0.017
465.413 | 0.024
465.722 | 0.014
466.032 | 0.022
466.342 | 0.014
466.652 | 0.019
466.961 | 0.02
467.271 | 0.022
467.58 | 0.018
467.889 | 0.019
468.199 | 0.02
468.508 | 0.017
468.817 | 0.022
469.126 | 0.019
469.435 | 0.025
469.744 | 0.018
470.053 | 0.017
470.362 | 0.016
470.67 | 0.013
470.979 | 0.021
471.287 | 0.011
471.596 | 0.015
471.904 | 0.026
472.213 | 0.013
472.521 | 0.023
472.829 | 0.02
473.137 | 0.02
473.445 | 0.013
473.753 | 0.026
474.061 | 0.013
474.369 | 0.021
474.677 | 0.02
474.985 | 0.012
475.292 | 0.012
475.6 | 0.021
475.907 | 0.02
476.215 | 0.017
476.522 | 0.026
476.829 | 0.018
477.137 | 0.023
477.444 | 0.013
477.751 | 0.025
478.058 | 0.013
478.365 | 0.018
478.671 | 0.026
478.978 | 0.021
479.285 | 0.013
479.592 | 0.014
479.898 | 0.011
480.205 | 0.018
480.511 | 0.005
480.817 | 0.025
481.124 | 0.015
481.43 | 0.023
481.736 | 0.03
482.042 | 0.017
482.348 | 0.017
482.654 | 0.022
482.96 | 0.02
483.266 | 0.023
483.571 | 0.023
483.877 | 0.007
484.182 | 0.013
484.488 | 0.018
484.793 | 0.013
485.099 | 0.017
485.404 | 0.02
485.709 | 0.023
486.014 | 0.017
486.319 | 0.027
486.624 | 0.015
486.929 | 0.014
487.234 | 0.01
487.539 | 0.021
487.844 | 0.023
488.148 | 0.019
488.453 | 0.029
488.757 | 0.021
489.062 | 0.016
489.366 | 0.022
489.67 | 0.015
489.975 | 0.013
490.279 | 0.02
490.583 | 0.021
490.887 | 0.015
491.191 | 0.021
491.495 | 0.021
491.798 | 0.026
492.102 | 0.023
492.406 | 0.013
492.709 | 0.02
493.013 | 0.026
493.316 | 0.019
493.62 | 0.026
493.923 | 0.017
494.226 | 0.023
494.529 | 0.019
494.832 | 0.02
495.135 | 0.018
495.438 | 0.024
495.741 | 0.03
496.044 | 0.018
496.347 | 0.015
496.649 | 0.026
496.952 | 0.016
497.255 | 0.016
497.557 | 0.019
497.859 | 0.021
498.162 | 0.007
498.464 | 0.033
498.766 | 0.023
499.068 | 0.016
499.37 | 0.016
499.672 | 0.022
499.974 | 0.017
500.276 | 0.022
500.578 | 0.023
500.879 | 0.011
501.181 | 0.022
501.483 | 0.017
501.784 | 0.024
502.085 | 0.013
502.387 | 0.014
502.688 | 0.014
502.989 | 0.025
503.29 | 0.013
503.591 | 0.016
503.892 | 0.011
504.193 | 0.026
504.494 | 0.016
504.795 | 0.012
505.096 | 0.031
505.396 | 0.014
505.697 | 0.007
505.997 | 0.016
506.298 | 0.021
506.598 | 0.017
506.898 | 0.021
507.199 | 0.016
507.499 | 0.021
507.799 | 0.016
508.099 | 0.015
508.399 | 0.019
508.699 | 0.02
508.998 | 0.019
509.298 | 0.029
509.598 | 0.015
509.897 | 0.024
510.197 | 0.014
510.496 | 0.017
510.796 | 0.012
511.095 | 0.016
511.394 | 0.019
511.693 | 0.031
511.992 | 0.022
512.291 | 0.022
512.59 | 0.021
512.889 | 0.022
513.188 | 0.025
513.487 | 0.02
513.786 | 0.026
514.084 | 0.023
514.383 | 0.025
514.681 | 0.03
514.98 | 0.022
515.278 | 0.024
515.576 | 0.026
515.874 | 0.023
516.173 | 0.014
516.471 | 0.026
516.769 | 0.022
517.067 | 0.018
517.364 | 0.019
517.662 | 0.01
517.96 | 0.015
518.257 | 0.017
518.555 | 0.018
518.853 | 0.015
519.15 | 0.016
519.447 | 0.019
519.745 | 0.017
520.042 | 0.023
520.339 | 0.013
520.636 | 0.023
520.933 | 0.034
521.23 | 0.018
521.527 | 0.014
521.824 | 0.025
522.121 | 0.015
522.417 | 0.017
522.714 | 0.014
523.01 | 0.022
523.307 | 0.018
523.603 | 0.026
523.9 | 0.023
524.196 | 0.023
524.492 | 0.011
524.788 | 0.029
525.084 | 0.025
525.38 | 0.026
525.676 | 0.015
525.972 | 0.019
526.268 | 0.021
526.564 | 0.024
526.859 | 0.024
527.155 | 0.018
527.45 | 0.024
527.746 | 0.024
528.041 | 0.029
528.336 | 0.021
528.632 | 0.011
528.927 | 0.014
529.222 | 0.02
529.517 | 0.026
529.812 | 0.017
530.107 | 0.018
530.402 | 0.018
530.696 | 0.027
530.991 | 0.026
531.286 | 0.021
531.58 | 0.023
531.875 | 0.015
532.169 | 0.02
532.463 | 0.019
532.758 | 0.011
533.052 | 0.022
533.346 | 0.013
533.64 | 0.019
533.934 | 0.022
534.228 | 0.024
534.522 | 0.022
534.816 | 0.025
535.11 | 0.012
535.403 | 0.028
535.697 | 0.025
535.99 | 0.025
536.284 | 0.021
536.577 | 0.031
536.871 | 0.023
537.164 | 0.021
537.457 | 0.022
537.75 | 0.019
538.043 | 0.022
538.336 | 0.021
538.629 | 0.022
538.922 | 0.02
539.215 | 0.011
539.507 | 0.016
539.8 | 0.008
540.093 | 0.03
540.385 | 0.012
540.678 | 0.016
540.97 | 0.009
541.262 | 0.023
541.554 | 0.018
541.847 | 0.022
542.139 | 0.022
542.431 | 0.024
542.723 | 0.021
543.015 | 0.025
543.307 | 0.022
543.598 | 0.025
543.89 | 0.026
544.182 | 0.02
544.473 | 0.025
544.765 | 0.03
545.056 | 0.008
545.347 | 0.017
545.639 | 0.015
545.93 | 0.014
546.221 | 0.023
546.512 | 0.023
546.803 | 0.016
547.094 | 0.015
547.385 | 0.022
547.676 | 0.02
547.967 | 0.032
548.257 | 0.018
548.548 | 0.026
548.839 | 0.007
549.129 | 0.021
549.419 | 0.025
549.71 | 0.022
550 | 0.027
550.29 | 0.017
550.58 | 0.009
550.871 | 0.019
551.161 | 0.014
551.45 | 0.013
551.74 | 0.017
552.03 | 0.009
552.32 | 0.02
552.61 | 0.021
552.899 | 0.028
553.189 | 0.013
553.478 | 0.017
553.768 | 0.013
554.057 | 0.031
554.346 | 0.027
554.635 | 0.022
554.925 | 0.021
555.214 | 0.024
555.503 | 0.013
555.792 | 0.026
556.081 | 0.01
556.369 | 0.021
556.658 | 0.017
556.947 | 0.023
557.235 | 0.017
557.524 | 0.029
557.812 | 0.019
558.101 | 0.02
558.389 | 0.026
558.677 | 0.02
558.966 | 0.032
559.254 | 0.029
559.542 | 0.019
559.83 | 0.025
560.118 | 0.024
560.406 | 0.015
560.693 | 0.038
560.981 | 0.029
561.269 | 0.023
561.556 | 0.02
561.844 | 0.024
562.131 | 0.025
562.419 | 0.025
562.706 | 0.024
562.994 | 0.017
563.281 | 0.031
563.568 | 0.023
563.855 | 0.032
564.142 | 0.018
564.429 | 0.029
564.716 | 0.027
565.003 | 0.03
565.289 | 0.025
565.576 | 0.029
565.863 | 0.025
566.149 | 0.026
566.436 | 0.029
566.722 | 0.02
567.008 | 0.024
567.295 | 0.029
567.581 | 0.02
567.867 | 0.025
568.153 | 0.022
568.439 | 0.022
568.725 | 0.023
569.011 | 0.028
569.297 | 0.025
569.583 | 0.024
569.868 | 0.025
570.154 | 0.023
570.439 | 0.018
570.725 | 0.024
571.01 | 0.023
571.296 | 0.031
571.581 | 0.022
571.866 | 0.033
572.151 | 0.032
572.436 | 0.022
572.721 | 0.025
573.006 | 0.026
573.291 | 0.031
573.576 | 0.023
573.861 | 0.03
574.146 | 0.022
574.43 | 0.031
574.715 | 0.036
574.999 | 0.031
575.284 | 0.033
575.568 | 0.022
575.852 | 0.025
576.137 | 0.029
576.421 | 0.038
576.705 | 0.022
576.989 | 0.028
577.273 | 0.015
577.557 | 0.027
577.841 | 0.026
578.124 | 0.032
578.408 | 0.034
578.692 | 0.018
578.975 | 0.03
579.259 | 0.022
579.542 | 0.026
579.826 | 0.02
580.109 | 0.026
580.392 | 0.03
580.676 | 0.021
580.959 | 0.027
581.242 | 0.027
581.525 | 0.027
581.808 | 0.022
582.091 | 0.025
582.373 | 0.021
582.656 | 0.028
582.939 | 0.025
583.221 | 0.011
583.504 | 0.026
583.786 | 0.024
584.069 | 0.024
584.351 | 0.022
584.633 | 0.016
584.916 | 0.013
585.198 | 0.027
585.48 | 0.021
585.762 | 0.012
586.044 | 0.019
586.326 | 0.027
586.607 | 0.024
586.889 | 0.041
587.171 | 0.03
587.453 | 0.017
587.734 | 0.026
588.016 | 0.028
588.297 | 0.014
588.578 | 0.027
588.86 | 0.029
589.141 | 0.024
589.422 | 0.024
589.703 | 0.033
589.984 | 0.013
590.265 | 0.022
590.546 | 0.028
590.827 | 0.017
591.108 | 0.02
591.389 | 0.026
591.669 | 0.02
591.95 | 0.027
592.23 | 0.028
592.511 | 0.033
592.791 | 0.027
593.071 | 0.028
593.352 | 0.011
593.632 | 0.028
593.912 | 0.03
594.192 | 0.034
594.472 | 0.027
594.752 | 0.03
595.032 | 0.027
595.312 | 0.015
595.592 | 0.026
595.871 | 0.02
596.151 | 0.024
596.43 | 0.012
596.71 | 0.023
596.989 | 0.038
597.269 | 0.023
597.548 | 0.029
597.827 | 0.034
598.106 | 0.026
598.385 | 0.028
598.665 | 0.008
598.944 | 0.017
599.222 | 0.022
599.501 | 0.026
599.78 | 0.026
600.059 | 0.025
600.337 | 0.025
600.616 | 0.022
600.895 | 0.026
601.173 | 0.05
601.451 | 0.023
601.73 | 0.02
602.008 | 0.03
602.286 | 0.007
602.564 | 0.03
602.842 | 0.022
603.12 | 0.013
603.398 | 0.022
603.676 | 0.026
603.954 | 0.019
604.232 | 0.019
604.51 | 0.018
604.787 | 0.014
605.065 | 0.023
605.342 | 0.02
605.62 | 0.012
605.897 | 0.021
606.174 | 0.029
606.452 | 0.018
606.729 | 0.033
607.006 | 0.03
607.283 | 0.025
607.56 | 0.019
607.837 | 0.014
608.114 | 0.022
608.391 | 0.048
608.667 | 0.01
608.944 | 0.025
609.221 | 0.014
609.497 | 0.019
609.774 | 0.013
610.05 | 0.028
610.326 | 0.022
610.603 | 0.033
610.879 | 0.035
611.155 | 0.025
611.431 | 0.019
611.707 | 0.027
611.983 | 0.016
612.259 | 0.027
612.535 | 0.02
612.811 | 0.02
613.087 | 0.024
613.362 | 0.021
613.638 | 0.023
613.913 | 0.023
614.189 | 0.026
614.464 | 0.026
614.74 | 0.036
615.015 | 0.01
615.29 | 0.024
615.565 | 0.023
615.84 | 0.01
616.115 | 0.03
616.39 | 0.023
616.665 | 0.03
616.94 | 0.02
617.215 | 0.03
617.49 | 0.026
617.764 | 0.026
618.039 | 0.028
618.313 | 0.03
618.588 | 0.036
618.862 | 0.024
619.137 | 0.035
619.411 | 0.023
619.685 | 0.027
619.959 | 0.019
620.233 | 0.033
620.507 | 0.02
620.781 | 0.024
621.055 | 0.016
621.329 | 0.025
621.603 | 0.026
621.876 | 0.033
622.15 | 0.028
622.424 | 0.023
622.697 | 0.024
622.971 | 0.018
623.244 | 0.021
623.517 | 0.017
623.791 | 0.025
624.064 | 0.034
624.337 | 0.032
624.61 | 0.025
624.883 | 0.028
625.156 | 0.019
625.429 | 0.024
625.702 | 0.018
625.974 | 0.034
626.247 | 0.026
626.52 | 0.027
626.792 | 0.026
627.065 | 0.029
627.337 | 0.029
627.61 | 0.029
627.882 | 0.024
628.154 | 0.027
628.427 | 0.04
628.699 | 0.027
628.971 | 0.026
629.243 | 0.016
629.515 | 0.033
629.787 | 0.025
630.058 | 0.038
630.33 | 0.032
630.602 | 0.012
630.874 | 0.032
631.145 | 0.037
631.417 | 0.017
631.688 | 0.035
631.959 | 0.023
632.231 | 0.025
632.502 | 0.034
632.773 | 0.035
633.044 | 0.043
633.316 | 0.035
633.587 | 0.012
633.857 | 0.026
634.128 | 0.033
634.399 | 0.042
634.67 | 0.019
634.941 | 0.025
635.211 | 0.034
635.482 | 0.018
635.752 | 0.034
636.023 | 0.019
636.293 | 0.028
636.564 | 0.041
636.834 | 0.025
637.104 | 0.039
637.374 | 0.026
637.644 | 0.031
637.914 | 0.036
638.184 | 0.03
638.454 | 0.019
638.724 | 0.04
638.994 | 0.027
639.264 | 0.033
639.533 | 0.026
639.803 | 0.022
640.072 | 0.028
640.342 | 0.022
640.611 | 0.033
640.881 | 0.029
641.15 | 0.047
641.419 | 0.015
641.688 | 0.03
641.957 | 0.027
642.227 | 0.034
642.496 | 0.04
642.764 | 0.033
643.033 | 0.039
643.302 | 0.025
643.571 | 0.024
643.84 | 0.025
644.108 | 0.039
644.377 | 0.01
644.645 | 0.017
644.914 | 0.026
645.182 | 0.034
645.45 | 0.033
645.719 | 0.04
645.987 | 0.035
646.255 | 0.03
646.523 | 0.026
646.791 | 0.023
647.059 | 0.04
647.327 | 0.028
647.595 | 0.026
647.862 | 0.015
648.13 | 0.021
648.398 | 0.039
648.665 | 0.025
648.933 | 0.046
649.2 | 0.032
649.468 | 0.031
649.735 | 0.035
650.002 | 0.033
650.269 | 0.03
650.536 | 0.023
650.804 | 0.019
651.071 | 0.042
651.338 | 0.024
651.604 | 0.031
651.871 | 0.021
652.138 | 0.026
652.405 | 0.014
652.671 | 0.034
652.938 | 0.03
653.205 | 0.027
653.471 | 0.016
653.737 | 0.028
654.004 | 0.016
654.27 | 0.023
654.536 | 0.03
654.803 | 0.024
655.069 | 0.026
655.335 | 0.023
655.601 | 0.03
655.867 | 0.028
656.132 | 0.034
656.398 | 0.029
656.664 | 0.024
656.93 | 0.02
657.195 | 0.021
657.461 | 0.027
657.726 | 0.032
657.992 | 0.014
658.257 | 0.022
658.523 | 0.032
658.788 | 0.017
659.053 | 0.028
659.318 | 0.025
659.583 | 0.023
659.848 | 0.021
660.113 | 0.027
660.378 | 0.029
660.643 | 0.031
660.908 | 0.017
661.172 | 0.008
661.437 | 0.024
661.702 | 0.025
661.966 | 0.034
662.231 | 0.024
662.495 | 0.027
662.76 | 0.052
663.024 | 0.022
663.288 | 0.015
663.552 | 0.018
663.816 | 0.027
664.08 | 0.03
664.344 | 0.028
664.608 | 0.027
664.872 | 0.022
665.136 | 0.018
665.4 | 0.029
665.664 | 0.034
665.927 | 0.028
666.191 | 0.027
666.454 | 0.017
666.718 | 0.027
666.981 | 0.011
667.244 | 0.02
667.508 | 0.038
667.771 | 0.023
668.034 | 0.032
668.297 | 0.025
668.56 | 0.029
668.823 | 0.01
669.086 | 0.031
669.349 | 0.024
669.612 | 0.022
669.875 | 0.017
670.137 | 0.031
670.4 | 0.027
670.662 | 0.027
670.925 | 0.029
671.187 | 0.038
671.45 | 0.008
671.712 | 0.046
671.974 | 0.035
672.237 | 0.022
672.499 | 0.024
672.761 | 0.025
673.023 | 0.026
673.285 | 0.031
673.547 | 0.033
673.808 | 0.025
674.07 | 0.023
674.332 | 0.029
674.594 | 0.024
674.855 | 0.022
675.117 | 0.018
675.378 | 0.037
675.64 | 0.024
675.901 | 0.007
676.162 | 0.03
676.424 | 0.023
676.685 | 0.035
676.946 | 0.031
677.207 | 0.039
677.468 | 0.042
677.729 | 0.016
677.99 | 0.016
678.251 | 0.036
678.512 | 0.029
678.772 | 0.035
679.033 | 0.042
679.294 | 0.025
679.554 | 0.027
679.815 | 0.032
680.075 | 0.021
680.335 | 0.034
680.596 | 0.014
680.856 | 0.035
681.116 | 0.045
681.376 | 0.034
681.636 | 0.027
681.896 | 0.039
682.156 | 0.036
682.416 | 0.03
682.676 | 0.026
682.936 | 0.034
683.196 | 0.027
683.455 | 0.029
683.715 | 0.029
683.974 | 0.028
684.234 | 0.025
684.493 | 0.017
684.753 | 0.031
685.012 | 0.039
685.271 | 0.022
685.53 | 0.029
685.789 | 0.036
686.049 | 0.039
686.308 | 0.022
686.567 | 0.026
686.825 | 0.04
687.084 | 0.031
687.343 | 0.037
687.602 | 0.037
687.86 | 0.029
688.119 | 0.035
688.378 | 0.04
688.636 | 0.038
688.894 | 0.031
689.153 | 0.032
689.411 | 0.039
689.669 | 0.034
689.928 | 0.035
690.186 | 0.015
690.444 | 0.037
690.702 | 0.031
690.96 | 0.04
691.218 | 0.029
691.476 | 0.035
691.733 | 0.034
691.991 | 0.03
692.249 | 0.026
692.506 | 0.032
692.764 | 0.035
693.021 | 0.045
693.279 | 0.035
693.536 | 0.035
693.794 | 0.014
694.051 | 0.025
694.308 | 0.028
694.565 | 0.044
694.822 | 0.046
695.079 | 0.017
695.336 | 0.032
695.593 | 0.024
695.85 | 0.03
696.107 | 0.026
696.364 | 0.03
696.62 | 0.031
696.877 | 0.04
697.133 | 0.025
697.39 | 0.039
697.646 | 0.031
697.903 | 0.024
698.159 | 0.014
698.415 | 0.022
698.672 | 0.043
698.928 | 0.029
699.184 | 0.036
699.44 | 0.026
699.696 | 0.045
699.952 | 0.012
700.208 | 0.031
700.464 | 0.024
700.719 | 0.048
700.975 | 0.032
701.231 | 0.034
701.486 | 0.046
701.742 | 0.035
701.997 | 0.042
702.253 | 0.039
702.508 | 0.016
702.763 | 0.036
703.019 | 0.04
703.274 | 0.034
703.529 | 0.035
703.784 | 0.028
704.039 | 0.057
704.294 | 0.028
704.549 | 0.025
704.804 | 0.025
705.058 | 0.037
705.313 | 0.025
705.568 | 0.044
705.822 | 0.043
706.077 | 0.027
706.331 | 0.024
706.586 | 0.025
706.84 | 0.039
707.095 | 0.028
707.349 | 0.01
707.603 | 0.028
707.857 | 0.028
708.111 | 0.022
708.365 | 0.026
708.619 | 0.026
708.873 | 0.043
709.127 | 0.026
709.381 | 0.02
709.635 | 0.025
709.888 | 0.033
710.142 | 0.027
710.396 | 0.035
710.649 | 0.03
710.903 | 0.024
711.156 | 0.024
711.409 | 0.038
711.663 | 0.04
711.916 | 0.026
712.169 | 0.025
712.422 | 0.036
712.675 | 0.036
712.928 | 0.047
713.181 | 0.04
713.434 | 0.027
713.687 | 0.048
713.94 | 0.027
714.192 | 0.039
714.445 | 0.007
714.698 | 0.043
714.95 | 0.037
715.203 | 0.057
715.455 | 0.038
715.708 | 0.04
715.96 | 0.056
716.212 | 0.034
716.464 | 0.043
716.717 | 0.047
716.969 | 0.041
717.221 | 0.044
717.473 | 0.017
717.725 | 0.036
717.977 | 0.036
718.228 | 0.049
718.48 | 0.028
718.732 | 0.041
718.984 | 0.04
719.235 | 0.053
719.487 | 0.043
719.738 | 0.033
719.99 | 0.053
720.241 | 0.032
720.492 | 0.044
720.744 | 0.03
720.995 | 0.022
721.246 | 0.037
721.497 | 0.047
721.748 | 0.039
721.999 | 0.032
722.25 | 0.06
722.501 | 0.053
722.752 | 0.041
723.002 | 0.031
723.253 | 0.032
723.504 | 0.055
723.754 | 0.044
724.005 | 0.05
724.255 | 0.045
724.506 | 0.047
724.756 | 0.045
725.006 | 0.033
725.257 | 0.042
725.507 | 0.045
725.757 | 0.037
726.007 | 0.028
726.257 | 0.046
726.507 | 0.024
726.757 | 0.07
727.007 | 0.052
727.257 | 0.025
727.506 | 0.045
727.756 | 0.035
728.006 | 0.054
728.255 | 0.051
728.505 | 0.018
728.754 | 0.041
729.004 | 0.061
729.253 | 0.073
729.502 | 0.071
729.751 | 0.039
730.001 | 0.024
730.25 | 0.037
730.499 | 0.054
730.748 | 0.047
730.997 | 0.037
731.246 | 0.05
731.495 | 0.041
731.743 | 0.052
731.992 | 0.039
732.241 | 0.027
732.489 | 0.046
732.738 | 0.075
732.987 | 0.04
733.235 | 0.041
733.483 | 0.059
733.732 | 0.048
733.98 | 0.024
734.228 | 0.029
734.477 | 0.037
734.725 | 0.049
734.973 | 0.038
735.221 | 0.049
735.469 | 0.054
735.717 | 0.055
735.965 | 0.035
736.212 | 0.033
736.46 | 0.052
736.708 | 0.056
736.955 | 0.065
737.203 | 0.033
737.451 | 0.057
737.698 | 0.045
737.945 | 0.035
738.193 | 0.052
738.44 | 0.036
738.687 | 0.043
738.935 | 0.043
739.182 | 0.039
739.429 | 0.054
739.676 | 0.035
739.923 | 0.048
740.17 | 0.06
740.417 | 0.043
740.663 | 0.032
740.91 | 0.046
741.157 | 0.055
741.403 | 0.036
741.65 | 0.024
741.897 | 0.044
742.143 | 0.034
742.389 | 0.044
742.636 | 0.032
742.882 | 0.034
743.128 | 0.043
743.375 | 0.025
743.621 | 0.066
743.867 | 0.049
744.113 | 0.048
744.359 | 0.038
744.605 | 0.053
744.851 | 0.074
745.097 | 0.066
745.342 | 0.046
745.588 | 0.053
745.834 | 0.035
746.079 | 0.06
746.325 | 0.059
746.57 | 0.056
746.816 | 0.018
747.061 | 0.044
747.307 | 0.051
747.552 | 0.037
747.797 | 0.053
748.042 | 0.051
748.287 | 0.05
748.532 | 0.038
748.777 | 0.051
749.022 | 0.053
749.267 | 0.06
749.512 | 0.04
749.757 | 0.039
750.002 | 0.041
750.246 | 0.047
750.491 | 0.053
750.736 | 0.039
750.98 | 0.057
751.225 | 0.055
751.469 | 0.052
751.713 | 0.052
751.958 | 0.056
752.202 | 0.036
752.446 | 0.042
752.69 | 0.035
752.934 | 0.028
753.178 | 0.048
753.422 | 0.05
753.666 | 0.065
753.91 | 0.031
754.154 | 0.057
754.398 | 0.082
754.641 | 0.027
754.885 | 0.029
755.129 | 0.047
755.372 | 0.027
755.616 | 0.043
755.859 | 0.051
756.103 | 0.074
756.346 | 0.021
756.589 | 0.051
756.832 | 0.054
757.076 | 0.056
757.319 | 0.051
757.562 | 0.061
757.805 | 0.06
758.048 | 0.057
758.291 | 0.049
758.533 | 0.044
758.776 | 0.042
759.019 | 0.061
759.262 | 0.055
759.504 | 0.053
759.747 | 0.052
759.989 | 0.055
760.232 | 0.067
760.474 | 0.035
760.717 | 0.084
760.959 | 0.042
761.201 | 0.041
761.444 | 0.05
761.686 | 0.063
761.928 | 0.047
762.17 | 0.051
762.412 | 0.066
762.654 | 0.058
762.896 | 0.047
763.137 | 0.065
763.379 | 0.035
763.621 | 0.072
763.863 | 0.049
764.104 | 0.05
764.346 | 0.083
764.587 | 0.044
764.829 | 0.05
765.07 | 0.041
765.312 | 0.074
765.553 | 0.088
765.794 | 0.056
766.035 | 0.057
766.276 | 0.073
766.518 | 0.061
766.759 | 0.07
767 | 0.05
767.241 | 0.059
767.481 | 0.048
767.722 | 0.05
767.963 | 0.061
768.204 | 0.064
768.444 | 0.084
768.685 | 0.064
768.926 | 0.035
769.166 | 0.053
769.407 | 0.048
769.647 | 0.087
769.887 | 0.03
770.128 | 0.064
770.368 | 0.061
770.608 | 0.072
770.848 | 0.069
771.088 | 0.051
771.328 | 0.068
771.568 | 0.045
771.808 | 0.062
772.048 | 0.053
772.288 | 0.059
772.528 | 0.039
772.767 | 0.049
773.007 | 0.063
773.247 | 0.064
773.486 | 0.05
773.726 | 0.044
773.965 | 0.076
774.205 | 0.065
774.444 | 0.07
774.683 | 0.076
774.922 | 0.078
775.162 | 0.067
775.401 | 0.075
775.64 | 0.061
775.879 | 0.064
776.118 | 0.053
776.357 | 0.057
776.596 | 0.075
776.834 | 0.079
777.073 | 0.065
777.312 | 0.069
777.551 | 0.047
777.789 | 0.036
778.028 | 0.055
778.266 | 0.067
778.505 | 0.04
778.743 | 0.036
778.982 | 0.037
779.22 | 0.055
779.458 | 0.064
779.696 | 0.058
779.934 | 0.075
780.173 | 0.059
780.411 | 0.053
780.649 | 0.054
780.887 | 0.059
781.124 | 0.075
781.362 | 0.053
781.6 | 0.062
781.838 | 0.055
782.075 | 0.058
782.313 | 0.06
782.551 | 0.034
782.788 | 0.052
783.026 | 0.055
783.263 | 0.062
783.5 | 0.051
783.738 | 0.067
783.975 | 0.048
784.212 | 0.02
784.449 | 0.052
784.686 | 0.059
784.924 | 0.062
785.161 | 0.043
785.397 | 0.061
785.634 | 0.07
785.871 | 0.083
786.108 | 0.053
786.345 | 0.086
786.581 | 0.062
786.818 | 0.073
787.055 | 0.061
787.291 | 0.087
787.528 | 0.08
787.764 | 0.066
788.001 | 0.046
788.237 | 0.067
788.473 | 0.04
788.709 | 0.056
788.946 | 0.071
789.182 | 0.036
789.418 | 0.042
789.654 | 0.073
789.89 | 0.051
790.126 | 0.074
790.362 | 0.088
790.597 | 0.051
790.833 | 0.067
791.069 | 0.039
791.305 | 0.05
791.54 | 0.063
791.776 | 0.086
792.011 | 0.066
792.247 | 0.03
792.482 | 0.062
792.717 | 0.068
792.953 | 0.064
793.188 | 0.074
793.423 | 0.082
793.658 | 0.079
793.894 | 0.087
794.129 | 0.071
794.364 | 0.066
794.599 | 0.074
794.833 | 0.051
795.068 | 0.066
795.303 | 0.067
795.538 | 0.038
795.772 | 0.056
796.007 | 0.06
796.242 | 0.085
796.476 | 0.066
796.711 | 0.044
796.945 | 0.054
797.18 | 0.053
797.414 | 0.03
797.648 | 0.083
797.882 | 0.062
798.117 | 0.047
798.351 | 0.039
798.585 | 0.104
798.819 | 0.045
799.053 | 0.074
799.287 | 0.075
799.521 | 0.083
799.755 | 0.077
799.988 | 0.073
800.222 | 0.057
800.456 | 0.059
800.689 | 0.03
800.923 | 0.043
801.156 | 0.075
801.39 | 0.048
801.623 | 0.071
801.857 | 0.076
802.09 | 0.061
802.323 | 0.066
802.557 | 0.066
802.79 | 0.052
803.023 | 0.056
803.256 | 0.052
803.489 | 0.065
803.722 | 0.04
803.955 | 0.036
804.188 | 0.084
804.421 | 0.054
804.653 | 0.066
804.886 | 0.045
805.119 | 0.033
805.351 | 0.054
805.584 | 0.07
805.816 | 0.035
806.049 | 0.05
806.281 | 0.069
806.514 | 0.09
806.746 | 0.097
806.978 | 0.031
807.211 | 0.019
807.443 | 0.037
807.675 | 0.039
integration time 5000 ms
1 scan average
1 scan average

## Interpolation
wavelength | raw blanked output (averaged), from Tab "Raw UV-VIS Measurements" | new wavelength | forecast over 1, 1/2 nm increments | new wavelength | forecast over 1, 0.2 nm increments | new wavelength | forecast over 1, 0.5 nm increments | new wavelength | forecast over 1, 1 nm increments | 29.1737 | RLE is shown in column k
195.562 | 0 | 199.8 | 0 | 199.8 | 0 | 199.5 | 0 | 200 | 0 | 0
195.934 | 0 | 199.9 | 0 | 200 | 0 | 200 | 0 | 201 | 0 | 0
196.307 | 0 | 200 | 0 | 200.2 | 0 | 200.5 | 0 | 202 | 0 | 0
196.679 | 0 | 200.1 | 0 | 200.4 | 0 | 201 | 0 | 203 | 0 | 0
197.051 | 0 | 200.2 | 0 | 200.6 | 0 | 201.5 | 0 | 204 | 0.019374 | 0.00066409
197.423 | 0 | 200.3 | 0 | 200.8 | 0 | 202 | 0 | 205 | 0.29837 | 0.010227
197.795 | 0 | 200.4 | 0 | 201 | 0 | 202.5 | 0 | 206 | 0.49142 | 0.016845
198.167 | 0 | 200.5 | 0 | 201.2 | 0 | 203 | 0 | 207 | 0.32894 | 0.011275
198.539 | 0 | 200.6 | 0 | 201.4 | 0 | 203.5 | 0 | 208 | 0.22508 | 0.0077152
198.911 | 0 | 200.7 | 0 | 201.6 | 0 | 204 | 0 | 209 | 0.2447 | 0.0083876
199.283 | 0 | 200.8 | 0 | 201.8 | 0 | 204.5 | 0.058121 | 210 | 0.30103 | 0.010318
199.654 | 0 | 200.9 | 0 | 202 | 0 | 205 | 0.25186 | 211 | 0.23096 | 0.0079166
200.026 | 0 | 201 | 0 | 202.2 | 0 | 205.5 | 0.58512 | 212 | 0.19082 | 0.006541
200.397 | 0 | 201.1 | 0 | 202.4 | 0 | 206 | 0.47932 | 213 | 0.13468 | 0.0046165
200.769 | 0 | 201.2 | 0 | 202.6 | 0 | 206.5 | 0.40982 | 214 | 0.090151 | 0.0030901
201.14 | 0 | 201.3 | 0 | 202.8 | 0 | 207 | 0.33692 | 215 | 0.072373 | 0.0024808
201.511 | 0 | 201.4 | 0 | 203 | 0 | 207.5 | 0.24007 | 216 | 0.075042 | 0.0025722
201.882 | 0 | 201.5 | 0 | 203.2 | 0 | 208 | 0.2521 | 217 | 0.12424 | 0.0042587
202.253 | 0 | 201.6 | 0 | 203.4 | 0 | 208.5 | 0.18308 | 218 | 0.10295 | 0.0035289
202.624 | 0 | 201.7 | 0 | 203.6 | 0 | 209 | 0.25126 | 219 | 0.11748 | 0.004027
202.995 | 0 | 201.8 | 0 | 203.8 | 0 | 209.5 | 0.29976 | 220 | 0.1167 | 0.0040002
203.366 | 0 | 201.9 | 0 | 204 | 0 | 210 | 0.32024 | 221 | 0.080292 | 0.0027522
203.736 | 0 | 202 | 0 | 204.2 | 0 | 210.5 | 0.28308 | 222 | 0.079772 | 0.0027344
204.107 | 0 | 202.1 | 0 | 204.4 | 0.02127 | 211 | 0.23637 | 223 | 0.06675 | 0.002288
204.477 | 0 | 202.2 | 0 | 204.6 | 0.087483 | 211.5 | 0.17342 | 224 | 0.071881 | 0.0024639
204.848 | 0.17533 | 202.3 | 0 | 204.8 | 0.14504 | 212 | 0.18501 | 225 | 0.082464 | 0.0028267
205.218 | 0.3235 | 202.4 | 0 | 205 | 0.23278 | 212.5 | 0.21404 | 226 | 0.076266 | 0.0026142
205.588 | 0.87783 | 202.5 | 0 | 205.2 | 0.37776 | 213 | 0.10412 | 227 | 0.057112 | 0.0019577
205.959 | 0.34917 | 202.6 | 0 | 205.4 | 0.63167 | 213.5 | 0.085881 | 228 | 0.050132 | 0.0017184
206.329 | 0.43033 | 202.7 | 0 | 205.6 | 0.58815 | 214 | 0.093751 | 229 | 0.051998 | 0.0017824
206.699 | 0.43333 | 202.8 | 0 | 205.8 | 0.52403 | 214.5 | 0.09082 | 230 | 0.049131 | 0.0016841
207.068 | 0.33383 | 202.9 | 0 | 206 | 0.50731 | 215 | 0.067483 | 231 | 0.044028 | 0.0015092
207.438 | 0.19317 | 203 | 0 | 206.2 | 0.40661 | 215.5 | 0.058815 | 232 | 0.036117 | 0.001238
207.808 | 0.244 | 203.1 | 0 | 206.4 | 0.39409 | 216 | 0.057031 | 233 | 0.030638 | 0.0010502
208.178 | 0.31517 | 203.2 | 0 | 206.6 | 0.41927 | 216.5 | 0.10928 | 234 | 0.040365 | 0.0013836
208.547 | 0.10833 | 203.3 | 0 | 206.8 | 0.38594 | 217 | 0.14987 | 235 | 0.038755 | 0.0013284
208.917 | 0.17933 | 203.4 | 0 | 207 | 0.34751 | 217.5 | 0.11358 | 236 | 0.033734 | 0.0011563
209.286 | 0.474 | 203.5 | 0 | 207.2 | 0.27732 | 218 | 0.084347 | 237 | 0.033261 | 0.0011401
209.655 | 0.19533 | 203.6 | 0 | 207.4 | 0.23553 | 218.5 | 0.11092 | 238 | 0.040978 | 0.0014046
210.025 | 0.34217 | 203.7 | 0 | 207.6 | 0.23733 | 219 | 0.10512 | 239 | 0.037961 | 0.0013012
210.394 | 0.38233 | 203.8 | 0 | 207.8 | 0.23473 | 219.5 | 0.1364 | 240 | 0.042826 | 0.001468
210.763 | 0.17617 | 203.9 | 0 | 208 | 0.28243 | 220 | 0.12861 | 241 | 0.046208 | 0.0015839
211.132 | 0.29067 | 204 | 0 | 208.2 | 0.23913 | 220.5 | 0.085091 | 242 | 0.045322 | 0.0015535
211.501 | 0.095833 | 204.1 | 0 | 208.4 | 0.18172 | 221 | 0.080053 | 243 | 0.050358 | 0.0017261
211.869 | 0.232 | 204.2 | 0 | 208.6 | 0.17585 | 221.5 | 0.07573 | 244 | 0.06327 | 0.0021687
212.238 | 0.16233 | 204.3 | 0 | 208.8 | 0.15455 | 222 | 0.097585 | 245 | 0.079841 | 0.0027368
212.607 | 0.26733 | 204.4 | 0 | 209 | 0.25862 | 222.5 | 0.066001 | 246 | 0.097696 | 0.0033488
212.975 | 0.065 | 204.5 | 0.06381 | 209.2 | 0.34061 | 223 | 0.0726 | 247 | 0.12029 | 0.0041234
213.344 | 0.049333 | 204.6 | 0.087483 | 209.4 | 0.28536 | 223.5 | 0.06165 | 248 | 0.16854 | 0.0057771
213.712 | 0.123 | 204.7 | 0.11115 | 209.6 | 0.30214 | 224 | 0.082353 | 249 | 0.20898 | 0.0071633
214.08 | 0.038 | 204.8 | 0.13483 | 209.8 | 0.3114 | 224.5 | 0.071638 | 250 | 0.30343 | 0.010401
214.448 | 0.17117 | 204.9 | 0.18913 | 210 | 0.29834 | 225 | 0.08445 | 251 | 1.4083 | 0.048274
214.817 | 0.024667 | 205 | 0.23278 | 210.2 | 0.35099 | 225.5 | 0.091304 | 252 | 8.0539 | 0.27607
215.185 | 0.081167 | 205.1 | 0.27644 | 210.4 | 0.31718 | 226 | 0.080611 | 253 | 26.4511 | 0.90668
215.553 | 0.055333 | 205.2 | 0.3201 | 210.6 | 0.25389 | 226.5 | 0.056884 | 254 | 29.1737 | 1
215.92 | 0.036833 | 205.3 | 0.53673 | 210.8 | 0.25863 | 227 | 0.05577 | 255 | 3.0845 | 0.10573
216.288 | 0.084167 | 205.4 | 0.63167 | 211 | 0.25362 | 227.5 | 0.058682 | 256 | 0.5481 | 0.018787
216.656 | 0.12817 | 205.5 | 0.7266 | 211.2 | 0.19687 | 228 | 0.045098 | 257 | 0.38515 | 0.013202
217.023 | 0.1715 | 205.6 | 0.51723 | 211.4 | 0.15838 | 228.5 | 0.046616 | 258 | 0.3139 | 0.01076
217.391 | 0.13617 | 205.7 | 0.52063 | 211.6 | 0.17867 | 229 | 0.053281 | 259 | 0.26973 | 0.0092456
217.758 | 0.074667 | 205.8 | 0.52403 | 211.8 | 0.17956 | 229.5 | 0.056098 | 260 | 0.23827 | 0.0081671
218.126 | 0.076667 | 205.9 | 0.52742 | 212 | 0.17516 | 230 | 0.045984 | 261 | 0.22261 | 0.0076305
218.493 | 0.12233 | 206 | 0.52747 | 212.2 | 0.20029 | 230.5 | 0.04531 | 262 | 0.21315 | 0.0073063
218.86 | 0.12767 | 206.1 | 0.46704 | 212.4 | 0.22831 | 231 | 0.04609 | 263 | 0.21236 | 0.0072792
219.227 | 0.0675 | 206.2 | 0.40661 | 212.6 | 0.20786 | 231.5 | 0.040683 | 264 | 0.22273 | 0.0076346
219.594 | 0.19167 | 206.3 | 0.34618 | 212.8 | 0.13939 | 232 | 0.036697 | 265 | 0.21739 | 0.0074517
219.961 | 0.11717 | 206.4 | 0.41235 | 213 | 0.11216 | 232.5 | 0.030971 | 266 | 0.14289 | 0.004898
220.328 | 0.10817 | 206.5 | 0.42373 | 213.2 | 0.060793 | 233 | 0.026066 | 267 | 0.099693 | 0.0034172
220.695 | 0.0565 | 206.6 | 0.4351 | 213.4 | 0.068724 | 233.5 | 0.034877 | 268 | 0.092734 | 0.0031787
221.061 | 0.0915 | 206.7 | 0.39899 | 213.6 | 0.09927 | 234 | 0.042959 | 269 | 0.087811 | 0.0030099
221.428 | 0.076833 | 206.8 | 0.38594 | 213.8 | 0.081035 | 234.5 | 0.043259 | 270 | 0.086273 | 0.0029572
221.795 | 0.066667 | 206.9 | 0.37289 | 214 | 0.081641 | 235 | 0.037679 | 271 | 0.066182 | 0.0022686
222.161 | 0.14767 | 207 | 0.35984 | 214.2 | 0.11858 | 235.5 | 0.035327 | 272 | 0.035915 | 0.0012311
222.527 | 0.0005 | 207.1 | 0.30982 | 214.4 | 0.11126 | 236 | 0.034322 | 273 | 0.031007 | 0.0010628
222.893 | 0.127 | 207.2 | 0.27732 | 214.6 | 0.075175 | 236.5 | 0.031553 | 274 | 0.030031 | 0.0010294
223.26 | 0.048833 | 207.3 | 0.24482 | 214.8 | 0.075672 | 237 | 0.030368 | 275 | 0.039775 | 0.0013634
223.626 | 0.059167 | 207.4 | 0.21231 | 215 | 0.069922 | 237.5 | 0.037861 | 276 | 0.031717 | 0.0010872
223.992 | 0.089333 | 207.5 | 0.24947 | 215.2 | 0.056853 | 238 | 0.044497 | 277 | 0.023565 | 0.00080775
224.358 | 0.092 | 207.6 | 0.23733 | 215.4 | 0.062681 | 238.5 | 0.040578 | 278 | 0.019389 | 0.00066459
224.723 | 0.050667 | 207.7 | 0.22519 | 215.6 | 0.056887 | 239 | 0.03845 | 279 | 0.0191 | 0.00065469
225.089 | 0.11317 | 207.8 | 0.21305 | 215.8 | 0.042858 | 239.5 | 0.034854 | 280 | 0.031542 | 0.0010812
225.455 | 0.080167 | 207.9 | 0.26595 | 216 | 0.054855 | 240 | 0.046203 | 281 | 0.027661 | 0.00094815
225.82 | 0.092833 | 208 | 0.28243 | 216.2 | 0.073379 | 240.5 | 0.04742 | 282 | 0.017849 | 0.00061182
226.186 | 0.0775 | 208.1 | 0.29892 | 216.4 | 0.096954 | 241 | 0.045838 | 283 | 0.013674 | 0.00046871
226.551 | 0.047167 | 208.2 | 0.2184 | 216.6 | 0.12145 | 241.5 | 0.045365 | 284 | 0.012374 | 0.00042416
226.917 | 0.047833 | 208.3 | 0.20006 | 216.8 | 0.14509 | 242 | 0.044872 | 285 | 0.012187 | 0.00041775
227.282 | 0.076333 | 208.4 | 0.18172 | 217 | 0.15731 | 242.5 | 0.045729 | 286 | 0.01289 | 0.00044183
227.647 | 0.0515 | 208.5 | 0.16338 | 217.2 | 0.14719 | 243 | 0.050029 | 287 | 0.010385 | 0.00035598
228.012 | 0.039833 | 208.6 | 0.19127 | 217.4 | 0.12918 | 243.5 | 0.055315 | 288 | 0.020678 | 0.0007088
228.377 | 0.046667 | 208.7 | 0.17291 | 217.6 | 0.099869 | 244 | 0.061919 | 289 | 0.044106 | 0.0015118
228.742 | 0.0485 | 208.8 | 0.15455 | 217.8 | 0.087843 | 244.5 | 0.072575 | 290 | 0.023145 | 0.00079336
229.107 | 0.059 | 208.9 | 0.13618 | 218 | 0.076279 | 245 | 0.079291 | 291 | 0.0099913 | 0.00034248
229.471 | 0.044333 | 209 | 0.29511 | 218.2 | 0.088917 | 245.5 | 0.087658 | 292 | 0.013978 | 0.00047913
229.836 | 0.0685 | 209.1 | 0.34458 | 218.4 | 0.10697 | 246 | 0.098376 | 293 | 0.012409 | 0.00042534
230.201 | 0.0245 | 209.2 | 0.39405 | 218.6 | 0.11632 | 246.5 | 0.10705 | 294 | 0.0089762 | 0.00030768
230.565 | 0.059333 | 209.3 | 0.28319 | 218.8 | 0.11878 | 247 | 0.1153 | 295 | 0.023171 | 0.00079426
230.929 | 0.039333 | 209.4 | 0.28536 | 219 | 0.095375 | 247.5 | 0.13853 | 296 | 0.10367 | 0.0035535
231.294 | 0.049333 | 209.5 | 0.28753 | 219.2 | 0.10122 | 248 | 0.17443 | 297 | 0.095586 | 0.0032764
231.658 | 0.035667 | 209.6 | 0.2897 | 219.4 | 0.14403 | 248.5 | 0.19266 | 298 | 0.010346 | 0.00035465
232.022 | 0.036167 | 209.7 | 0.32921 | 219.6 | 0.13707 | 249 | 0.20114 | 299 | 0.0051133 | 0.00017527
232.386 | 0.037167 | 209.8 | 0.3114 | 219.8 | 0.13938 | 249.5 | 0.23314 | 300 | 0.00682 | 0.00023377
232.75 | 0.024 | 209.9 | 0.29358 | 220 | 0.13463 | 250 | 0.29276 | 301 | 0.038239 | 0.0013107
233.114 | 0.0255 | 210 | 0.27577 | 220.2 | 0.11181 | 250.5 | 0.38439 | 302 | 0.082094 | 0.002814
233.478 | 0.029667 | 210.1 | 0.32568 | 220.4 | 0.089386 | 251 | 0.71343 | 303 | 0.02688 | 0.00092139
233.842 | 0.048833 | 210.2 | 0.35099 | 220.6 | 0.07882 | 251.5 | 3.1272 | 304 | 0.005715 | 0.00019589
234.205 | 0.040333 | 210.3 | 0.3763 | 220.8 | 0.082988 | 252 | 8.5308 | 305 | 0.0060263 | 0.00020657
234.569 | 0.047333 | 210.4 | 0.29887 | 221 | 0.078385 | 252.5 | 12.5036 | 306 | 0.0051642 | 0.00017702
234.932 | 0.035667 | 210.5 | 0.27638 | 221.2 | 0.078787 | 253 | 21.1764 | 307 | 0.0057136 | 0.00019585
235.296 | 0.035 | 210.6 | 0.25389 | 221.4 | 0.080595 | 253.5 | 45.6733 | 308 | 0.0056649 | 0.00019418
235.659 | 0.035167 | 210.7 | 0.23139 | 221.6 | 0.072514 | 254 | 34.339 | 309 | 0.0062198 | 0.0002132
236.022 | 0.037167 | 210.8 | 0.27846 | 221.8 | 0.09131 | 254.5 | 7.5089 | 310 | 0.0068666 | 0.00023537
236.385 | 0.027667 | 210.9 | 0.26604 | 222 | 0.11689 | 255 | 1.0863 | 311 | 0.054886 | 0.0018814
236.748 | 0.033167 | 211 | 0.25362 | 222.2 | 0.084559 | 255.5 | 0.6583 | 312 | 0.4163 | 0.01427
237.111 | 0.028 | 211.1 | 0.2412 | 222.4 | 0.050008 | 256 | 0.54583 | 313 | 0.52424 | 0.01797
237.474 | 0.034333 | 211.2 | 0.18015 | 222.6 | 0.072489 | 256.5 | 0.44018 | 314 | 0.086529 | 0.002966
237.837 | 0.049333 | 211.3 | 0.16927 | 222.8 | 0.07669 | 257 | 0.37124 | 315 | 0.0070132 | 0.00024039
238.2 | 0.043333 | 211.4 | 0.15838 | 223 | 0.065798 | 257.5 | 0.34404 | 316 | 0.0051444 | 0.00017634
238.562 | 0.036333 | 211.5 | 0.1475 | 223.2 | 0.075313 | 258 | 0.31375 | 317 | 0.0051326 | 0.00017593
238.925 | 0.046 | 211.6 | 0.19824 | 223.4 | 0.065339 | 258.5 | 0.28389 | 318 | 0.0059934 | 0.00020544
239.287 | 0.027667 | 211.7 | 0.19026 | 223.6 | 0.05759 | 259 | 0.2717 | 319 | 0.0050302 | 0.00017242
239.65 | 0.035833 | 211.8 | 0.18228 | 223.8 | 0.075405 | 259.5 | 0.25359 | 320 | 0.0047664 | 0.00016338
240.012 | 0.048667 | 211.9 | 0.16615 | 224 | 0.082158 | 260 | 0.2339 | 321 | 0.0051504 | 0.00017654
240.374 | 0.052833 | 212 | 0.17516 | 224.2 | 0.089496 | 260.5 | 0.22731 | 322 | 0.0051995 | 0.00017823
240.736 | 0.0425 | 212.1 | 0.18417 | 224.4 | 0.079628 | 261 | 0.22459 | 323 | 0.0049361 | 0.0001692
241.098 | 0.0485 | 212.2 | 0.19318 | 224.6 | 0.064522 | 261.5 | 0.21593 | 324 | 0.0030021 | 0.0001029
241.46 | 0.043333 | 212.3 | 0.22352 | 224.8 | 0.079047 | 262 | 0.20873 | 325 | 0.0032397 | 0.00011105
241.822 | 0.046 | 212.4 | 0.22831 | 225 | 0.088496 | 262.5 | 0.2148 | 326 | 0.0035609 | 0.00012206
242.184 | 0.044667 | 212.5 | 0.2331 | 225.2 | 0.085807 | 263 | 0.20857 | 327 | 0.0039871 | 0.00013667
242.546 | 0.0435 | 212.6 | 0.23789 | 225.4 | 0.09261 | 263.5 | 0.21371 | 328 | 0.0038766 | 0.00013288
242.907 | 0.0535 | 212.7 | 0.15258 | 225.6 | 0.091342 | 264 | 0.22787 | 329 | 0.0019734 | 6.7644e-05
243.269 | 0.046167 | 212.8 | 0.13939 | 225.8 | 0.085846 | 264.5 | 0.22661 | 330 | 0.0025877 | 8.87e-05
243.63 | 0.061333 | 212.9 | 0.1262 | 226 | 0.082841 | 265 | 0.22672 | 331 | 0.0030559 | 0.00010475
243.992 | 0.055 | 213 | 0.11993 | 226.2 | 0.073146 | 265.5 | 0.19885 | 332 | 0.002876 | 9.8582e-05
244.353 | 0.074333 | 213.1 | 0.090361 | 226.4 | 0.059112 | 266 | 0.13021 | 333 | 0.018603 | 0.00063765
244.714 | 0.074667 | 213.2 | 0.060793 | 226.6 | 0.053287 | 266.5 | 0.099616 | 334 | 0.038496 | 0.0013196
245.075 | 0.079667 | 213.3 | 0.031225 | 226.8 | 0.047413 | 267 | 0.1008 | 335 | 0.0085644 | 0.00029357
245.436 | 0.0855 | 213.4 | 0.083541 | 227 | 0.056071 | 267.5 | 0.098664 | 336 | 0.0018616 | 6.381e-05
245.797 | 0.094333 | 213.5 | 0.091406 | 227.2 | 0.063827 | 268 | 0.089938 | 337 | 0.0016754 | 5.743e-05
246.158 | 0.10217 | 213.6 | 0.09927 | 227.4 | 0.059148 | 268.5 | 0.0896 | 338 | 0.0025139 | 8.6171e-05
246.519 | 0.1075 | 213.7 | 0.10713 | 227.6 | 0.057681 | 269 | 0.086719 | 339 | 0.0028602 | 9.804e-05
246.879 | 0.11183 | 213.8 | 0.068756 | 227.8 | 0.048239 | 269.5 | 0.087115 | 340 | 0.0026444 | 9.0644e-05
247.24 | 0.11833 | 213.9 | 0.067216 | 228 | 0.042298 | 270 | 0.085181 | 341 | 0.0032978 | 0.00011304
247.601 | 0.14533 | 214 | 0.065676 | 228.2 | 0.044755 | 270.5 | 0.086522 | 342 | 0.0032307 | 0.00011074
247.961 | 0.17583 | 214.1 | 0.11203 | 228.4 | 0.045275 | 271 | 0.070409 | 343 | 0.0020364 | 6.9804e-05
248.321 | 0.19167 | 214.2 | 0.11858 | 228.6 | 0.047647 | 271.5 | 0.041616 | 344 | 0.0024347 | 8.3455e-05
248.682 | 0.19583 | 214.3 | 0.12512 | 228.8 | 0.051754 | 272 | 0.032428 | 345 | 0.0019443 | 6.6646e-05
249.042 | 0.194 | 214.4 | 0.13166 | 229 | 0.055748 | 272.5 | 0.033701 | 346 | 0.0025903 | 8.8789e-05
249.402 | 0.2235 | 214.5 | 0.077001 | 229.2 | 0.052342 | 273 | 0.031027 | 347 | 0.0019555 | 6.703e-05
249.762 | 0.25783 | 214.6 | 0.075175 | 229.4 | 0.052033 | 273.5 | 0.028291 | 348 | 0.001733 | 5.9403e-05
250.122 | 0.30667 | 214.7 | 0.073348 | 229.6 | 0.058958 | 274 | 0.02938 | 349 | 0.0025341 | 8.6863e-05
250.482 | 0.3725 | 214.8 | 0.071522 | 229.8 | 0.05529 | 274.5 | 0.032421 | 350 | 0.0023595 | 8.0879e-05
250.842 | 0.46383 | 214.9 | 0.082147 | 230 | 0.041322 | 275 | 0.044951 | 351 | 0.0019435 | 6.6617e-05
251.201 | 0.78083 | 215 | 0.069922 | 230.2 | 0.041339 | 275.5 | 0.041953 | 352 | 0.0027519 | 9.4329e-05
251.561 | 2.7695 | 215.1 | 0.057698 | 230.4 | 0.048261 | 276 | 0.028932 | 353 | 0.0018392 | 6.3042e-05
251.921 | 8.5195 | 215.2 | 0.054347 | 230.6 | 0.044191 | 276.5 | 0.024267 | 354 | 0.0026633 | 9.1292e-05
252.28 | 11.1588 | 215.3 | 0.058514 | 230.8 | 0.045844 | 277 | 0.02413 | 355 | 0.0029175 | 0.0001
252.639 | 13.0395 | 215.4 | 0.062681 | 231 | 0.047747 | 277.5 | 0.022298 | 356 | 0.0034128 | 0.00011698
252.999 | 17.9525 | 215.5 | 0.066847 | 231.2 | 0.044678 | 278 | 0.019869 | 357 | 0.0038506 | 0.00013199
253.358 | 33.4157 | 215.6 | 0.054923 | 231.4 | 0.040911 | 278.5 | 0.015999 | 358 | 0.0030131 | 0.00010328
253.717 | 61.4305 | 215.7 | 0.04889 | 231.6 | 0.039982 | 279 | 0.018148 | 359 | 0.0040255 | 0.00013799
254.076 | 28.0535 | 215.8 | 0.042858 | 231.8 | 0.037821 | 279.5 | 0.023152 | 360 | 0.0036533 | 0.00012523
254.435 | 3.6665 | 215.9 | 0.036826 | 232 | 0.03557 | 280 | 0.035763 | 361 | 0.003669 | 0.00012576
254.794 | 1.202 | 216 | 0.061906 | 232.2 | 0.0367 | 280.5 | 0.03571 | 362 | 0.0050713 | 0.00017383
255.153 | 0.77283 | 216.1 | 0.065833 | 232.4 | 0.033219 | 281 | 0.024596 | 363 | 0.0099071 | 0.00033959
255.511 | 0.6565 | 216.2 | 0.06976 | 232.6 | 0.028868 | 281.5 | 0.022677 | 364 | 0.21998 | 0.0075402
255.87 | 0.58467 | 216.3 | 0.084545 | 232.8 | 0.027256 | 282 | 0.01734 | 365 | 0.45551 | 0.015614
256.228 | 0.484 | 216.4 | 0.096954 | 233 | 0.024882 | 282.5 | 0.01353 | 366 | 0.16594 | 0.0056879
256.587 | 0.42267 | 216.5 | 0.10936 | 233.2 | 0.026058 | 283 | 0.013516 | 367 | 0.030436 | 0.0010433
256.945 | 0.36717 | 216.6 | 0.12177 | 233.4 | 0.030608 | 283.5 | 0.013976 | 368 | 0.0063648 | 0.00021817
257.304 | 0.349 | 216.7 | 0.13321 | 233.6 | 0.038577 | 284 | 0.012373 | 369 | 0.0053612 | 0.00018377
257.662 | 0.34417 | 216.8 | 0.14509 | 233.8 | 0.042412 | 284.5 | 0.010774 | 370 | 0.005996 | 0.00020553
258.02 | 0.3095 | 216.9 | 0.15698 | 234 | 0.041937 | 285 | 0.011517 | 371 | 0.0056668 | 0.00019424
258.378 | 0.28817 | 217 | 0.16886 | 234.2 | 0.044529 | 285.5 | 0.014271 | 372 | 0.0056387 | 0.00019328
258.736 | 0.27317 | 217.1 | 0.14611 | 234.4 | 0.0451 | 286 | 0.014138 | 373 | 0.0065886 | 0.00022584
259.094 | 0.27567 | 217.2 | 0.14719 | 234.6 | 0.042025 | 286.5 | 0.01026 | 374 | 0.0060741 | 0.0002082
259.452 | 0.25683 | 217.3 | 0.14828 | 234.8 | 0.039629 | 287 | 0.010318 | 375 | 0.0047315 | 0.00016218
259.809 | 0.23667 | 217.4 | 0.12621 | 235 | 0.037888 | 287.5 | 0.010578 | 376 | 0.0050975 | 0.00017473
260.167 | 0.22767 | 217.5 | 0.11304 | 235.2 | 0.03552 | 288 | 0.018978 | 377 | 0.0051414 | 0.00017623
260.524 | 0.2305 | 217.6 | 0.099869 | 235.4 | 0.035206 | 288.5 | 0.032478 | 378 | 0.0050327 | 0.00017251
260.882 | 0.21867 | 217.7 | 0.086696 | 235.6 | 0.035369 | 289 | 0.054879 | 379 | 0.0041187 | 0.00014118
261.239 | 0.22983 | 217.8 | 0.092462 | 235.8 | 0.036199 | 289.5 | 0.04496 | 380 | 0.004492 | 0.00015397
261.597 | 0.20967 | 217.9 | 0.084371 | 236 | 0.035273 | 290 | 0.014849 | 381 | 0.005237 | 0.00017951
261.954 | 0.205 | 218 | 0.076279 | 236.2 | 0.031494 | 290.5 | 0.0096272 | 382 | 0.0050454 | 0.00017294
262.311 | 0.21317 | 218.1 | 0.068188 | 236.4 | 0.031693 | 291 | 0.0096353 | 383 | 0.0047616 | 0.00016322
262.668 | 0.21967 | 218.2 | 0.096041 | 236.6 | 0.031482 | 291.5 | 0.010711 | 384 | 0.0035929 | 0.00012315
263.025 | 0.206 | 218.3 | 0.10252 | 236.8 | 0.030082 | 292 | 0.014434 | 385 | 0.0043851 | 0.00015031
263.382 | 0.202 | 218.4 | 0.10901 | 237 | 0.029727 | 292.5 | 0.016789 | 386 | 0.004839 | 0.00016587
263.739 | 0.22583 | 218.5 | 0.10938 | 237.2 | 0.031295 | 293 | 0.011017 | 387 | 0.0052901 | 0.00018133
264.095 | 0.234 | 218.6 | 0.11632 | 237.4 | 0.03414 | 293.5 | 0.0094207 | 388 | 0.0047512 | 0.00016286
264.452 | 0.22867 | 218.7 | 0.12327 | 237.6 | 0.040925 | 294 | 0.0089085 | 389 | 0.0053907 | 0.00018478
264.809 | 0.22 | 218.8 | 0.13022 | 237.8 | 0.044593 | 294.5 | 0.0085995 | 390 | 0.0075906 | 0.00026019
265.165 | 0.23833 | 218.9 | 0.10285 | 238 | 0.044354 | 295 | 0.017519 | 391 | 0.0068075 | 0.00023334
265.521 | 0.2055 | 219 | 0.095375 | 238.2 | 0.044543 | 295.5 | 0.043396 | 392 | 0.0054727 | 0.00018759
265.878 | 0.12767 | 219.1 | 0.087904 | 238.4 | 0.039408 | 296 | 0.097791 | 393 | 0.0054265 | 0.00018601
266.234 | 0.10867 | 219.2 | 0.080434 | 238.6 | 0.04068 | 296.5 | 0.16982 | 394 | 0.0052453 | 0.0001798
266.59 | 0.089667 | 219.3 | 0.13531 | 238.8 | 0.042766 | 297 | 0.1001 | 395 | 0.005505 | 0.0001887
266.946 | 0.10333 | 219.4 | 0.14403 | 239 | 0.037825 | 297.5 | 0.016835 | 396 | 0.0047281 | 0.00016207
267.302 | 0.10317 | 219.5 | 0.15275 | 239.2 | 0.034759 | 298 | 0.0089313 | 397 | 0.0053673 | 0.00018398
267.658 | 0.095833 | 219.6 | 0.12585 | 239.4 | 0.034922 | 298.5 | 0.0052725 | 398 | 0.0058529 | 0.00020062
268.014 | 0.083667 | 219.7 | 0.13262 | 239.6 | 0.03483 | 299 | 0.00422 | 399 | 0.0059907 | 0.00020534
268.37 | 0.095333 | 219.8 | 0.13938 | 239.8 | 0.041743 | 299.5 | 0.0058473 | 400 | 0.0056337 | 0.00019311
268.725 | 0.086333 | 219.9 | 0.14615 | 240 | 0.046673 | 300 | 0.0047737 | 401 | 0.0059061 | 0.00020245
269.081 | 0.0855 | 220 | 0.13456 | 240.2 | 0.050192 | 300.5 | 0.0098389 | 402 | 0.0074225 | 0.00025442
269.436 | 0.086333 | 220.1 | 0.12319 | 240.4 | 0.049082 | 301 | 0.031805 | 403 | 0.047451 | 0.0016265
269.792 | 0.088833 | 220.2 | 0.11181 | 240.6 | 0.046075 | 301.5 | 0.073072 | 404 | 0.47829 | 0.016395
270.147 | 0.082 | 220.3 | 0.10044 | 240.8 | 0.046583 | 302 | 0.11125 | 405 | 0.43622 | 0.014952
270.502 | 0.086167 | 220.4 | 0.087993 | 241 | 0.046036 | 302.5 | 0.06196 |  |  | 0
270.857 | 0.093333 | 220.5 | 0.079728 | 241.2 | 0.044895 | 303 | 0.013896 |  |  | 0
271.212 | 0.048 | 220.6 | 0.071463 | 241.4 | 0.045314 | 303.5 | 0.0047846 |  |  | 0
271.567 | 0.038833 | 220.7 | 0.085267 | 241.6 | 0.045461 | 304 | 0.0057367 |  |  | 0
271.922 | 0.029333 | 220.8 | 0.082988 | 241.8 | 0.044899 | 304.5 | 0.0066236 |  |  | 0
272.277 | 0.034 | 220.9 | 0.080709 | 242 | 0.044994 | 305 | 0.0062643 |  |  | 0
272.632 | 0.034333 | 221 | 0.07843 | 242.2 | 0.044722 | 305.5 | 0.0051911 |  |  | 0
272.987 | 0.031667 | 221.1 | 0.076016 | 242.4 | 0.043976 | 306 | 0.004586 |  |  | 0
273.341 | 0.0275 | 221.2 | 0.078787 | 242.6 | 0.046874 | 306.5 | 0.0057156 |  |  | 0
273.696 | 0.028167 | 221.3 | 0.081558 | 242.8 | 0.050328 | 307 | 0.005691 |  |  | 0
274.05 | 0.028667 | 221.4 | 0.084329 | 243 | 0.049347 | 307.5 | 0.0057341 |  |  | 0
274.404 | 0.032667 | 221.5 | 0.075897 | 243.2 | 0.050412 | 308 | 0.0057461 |  |  | 0
274.759 | 0.0335 | 221.6 | 0.072514 | 243.4 | 0.055088 | 308.5 | 0.0055146 |  |  | 0
275.113 | 0.053667 | 221.7 | 0.069131 | 243.6 | 0.056146 | 309 | 0.0062832 |  |  | 0
275.467 | 0.046667 | 221.8 | 0.097571 | 243.8 | 0.056237 | 309.5 | 0.0068617 |  |  | 0
275.821 | 0.028833 | 221.9 | 0.10723 | 244 | 0.062222 | 310 | 0.00512 |  |  | 0
276.175 | 0.026 | 222 | 0.11689 | 244.2 | 0.067298 | 310.5 | 0.0086182 |  |  | 0
276.529 | 0.0235 | 222.1 | 0.12654 | 244.4 | 0.070126 | 311 | 0.031039 |  |  | 0
276.882 | 0.024 | 222.2 | 0.068086 | 244.6 | 0.074728 | 311.5 | 0.125 |  |  | 0
277.236 | 0.025 | 222.3 | 0.059047 | 244.8 | 0.077302 | 312 | 0.411 |  |  | 0
277.59 | 0.019667 | 222.4 | 0.050008 | 245 | 0.07875 | 312.5 | 0.71288 |  |  | 0
277.943 | 0.024 | 222.5 | 0.040968 | 245.2 | 0.08182 | 313 | 0.62737 |  |  | 0
278.297 | 0.0135 | 222.6 | 0.089661 | 245.4 | 0.085314 | 313.5 | 0.23247 |  |  | 0
278.65 | 0.016167 | 222.7 | 0.086838 | 245.6 | 0.089831 | 314 | 0.019619 |  |  | 0
279.003 | 0.017333 | 222.8 | 0.084015 | 245.8 | 0.094103 | 314.5 | 0.0075001 |  |  | 0
279.356 | 0.021833 | 222.9 | 0.059217 | 246 | 0.098686 | 315 | 0.0073545 |  |  | 0
279.709 | 0.025333 | 223 | 0.065798 | 246.2 | 0.10234 | 315.5 | 0.0061849 |  |  | 0
280.062 | 0.0405 | 223.1 | 0.072379 | 246.4 | 0.10575 | 316 | 0.0043039 |  |  | 0
280.415 | 0.0445 | 223.2 | 0.07896 | 246.6 | 0.10848 | 316.5 | 0.0049445 |  |  | 0
280.768 | 0.023833 | 223.3 | 0.074599 | 246.8 | 0.11113 | 317 | 0.005174 |  |  | 0
281.121 | 0.0225 | 223.4 | 0.065339 | 247 | 0.11437 | 317.5 | 0.0052794 |  |  | 0
281.474 | 0.024333 | 223.5 | 0.056079 | 247.2 | 0.1204 | 318 | 0.0064847 |  |  | 0
281.826 | 0.021 | 223.6 | 0.04682 | 247.4 | 0.13259 | 318.5 | 0.0062162 |  |  | 0
282.179 | 0.012 | 223.7 | 0.069872 | 247.6 | 0.14451 | 319 | 0.0045233 |  |  | 0
282.531 | 0.016 | 223.8 | 0.075405 | 247.8 | 0.1624 | 319.5 | 0.004351 |  |  | 0
282.884 | 0.009 | 223.9 | 0.080938 | 248 | 0.17457 | 320 | 0.0051381 |  |  | 0
283.236 | 0.017 | 224 | 0.080526 | 248.2 | 0.18632 | 320.5 | 0.0048102 |  |  | 0
283.588 | 0.0135 | 224.1 | 0.085011 | 248.4 | 0.19182 | 321 | 0.0053679 |  |  | 0
283.94 | 0.011 | 224.2 | 0.089496 | 248.6 | 0.19404 | 321.5 | 0.0052732 |  |  | 0
284.292 | 0.0145 | 224.3 | 0.093982 | 248.8 | 0.19422 | 322 | 0.0048467 |  |  | 0
284.644 | 0.0081667 | 224.4 | 0.075095 | 249 | 0.19869 | 322.5 | 0.0054787 |  |  | 0
284.996 | 0.0115 | 224.5 | 0.069808 | 249.2 | 0.21052 | 323 | 0.0050775 |  |  | 0
285.348 | 0.013667 | 224.6 | 0.064522 | 249.4 | 0.22212 | 323.5 | 0.004252 |  |  | 0
285.7 | 0.015167 | 224.7 | 0.059235 | 249.6 | 0.24267 | 324 | 0.0026299 |  |  | 0
286.051 | 0.014833 | 224.8 | 0.087503 | 249.8 | 0.26573 | 324.5 | 0.0021244 |  |  | 0
286.403 | 0.011333 | 224.9 | 0.090405 | 250 | 0.29016 | 325 | 0.0040697 |  |  | 0
286.754 | 0.0075 | 225 | 0.093307 | 250.2 | 0.32238 | 325.5 | 0.0035249 |  |  | 0
287.106 | 0.0125 | 225.1 | 0.081777 | 250.4 | 0.36074 | 326 | 0.0035193 |  |  | 0
287.457 | 0.009 | 225.2 | 0.085807 | 250.6 | 0.40676 | 326.5 | 0.0036386 |  |  | 0
287.808 | 0.011333 | 225.3 | 0.089837 | 250.8 | 0.48373 | 327 | 0.0038416 |  |  | 0
288.159 | 0.026167 | 225.4 | 0.093867 | 251 | 0.62895 | 327.5 | 0.0044813 |  |  | 0
288.51 | 0.029 | 225.5 | 0.094127 | 251.2 | 1.0276 | 328 | 0.0043804 |  |  | 0
288.861 | 0.048167 | 225.6 | 0.091342 | 251.4 | 1.9753 | 328.5 | 0.002768 |  |  | 0
289.212 | 0.070833 | 225.7 | 0.088558 | 251.6 | 4.0853 | 329 | 0.0013197 |  |  | 0
289.563 | 0.040833 | 225.8 | 0.085773 | 251.8 | 6.5921 | 329.5 | 0.0018326 |  |  | 0
289.914 | 0.0105 | 225.9 | 0.083208 | 252 | 8.5503 | 330 | 0.0030872 |  |  | 0
290.264 | 0.010833 | 226 | 0.082841 | 252.2 | 10.4499 | 330.5 | 0.0028433 |  |  | 0
290.615 | 0.0085 | 226.1 | 0.082475 | 252.4 | 11.6614 | 331 | 0.0032094 |  |  | 0
290.966 | 0.011667 | 226.2 | 0.071605 | 252.6 | 13.2783 | 331.5 | 0.003115 |  |  | 0
291.316 | 0.0073333 | 226.3 | 0.065358 | 252.8 | 15.5687 | 332 | 0.0023558 |  |  | 0
291.666 | 0.013167 | 226.4 | 0.059112 | 253 | 20.787 | 332.5 | 0.0031573 |  |  | 0
292.017 | 0.012833 | 226.5 | 0.052866 | 253.2 | 27.1736 | 333 | 0.011873 |  |  | 0
292.367 | 0.019667 | 226.6 | 0.055526 | 253.4 | 38.7826 | 333.5 | 0.040777 |  |  | 0
292.717 | 0.015333 | 226.7 | 0.051469 | 253.6 | 52.2537 | 334 | 0.053359 |  |  | 0
293.067 | 0.0065 | 226.8 | 0.047413 | 253.8 | 46.0852 | 334.5 | 0.021353 |  |  | 0
293.417 | 0.010667 | 226.9 | 0.043357 | 254 | 35.8574 | 335 | 0.002892 |  |  | 0
293.767 | 0.0093333 | 227 | 0.060435 | 254.2 | 21.0742 | 335.5 | 0.0014486 |  |  | 0
294.116 | 0.0086667 | 227.1 | 0.064423 | 254.4 | 8.8521 | 336 | 0.0027744 |  |  | 0
294.466 | 0.0075 | 227.2 | 0.068411 | 254.6 | 4.8034 | 336.5 | 0.0013617 |  |  | 0
294.816 | 0.0098333 | 227.3 | 0.058646 | 254.8 | 1.4577 | 337 | 0.0014212 |  |  | 0
295.165 | 0.020667 | 227.4 | 0.059148 | 255 | 1.0502 | 337.5 | 0.0022434 |  |  | 0
295.515 | 0.044833 | 227.5 | 0.059651 | 255.2 | 0.7511 | 338 | 0.0027135 |  |  | 0
295.864 | 0.067833 | 227.6 | 0.060153 | 255.4 | 0.68889 | 338.5 | 0.0025849 |  |  | 0
296.213 | 0.12767 | 227.7 | 0.053239 | 255.6 | 0.62756 | 339 | 0.0034342 |  |  | 0
296.562 | 0.22167 | 227.8 | 0.048239 | 255.8 | 0.59506 | 339.5 | 0.0025615 |  |  | 0
296.911 | 0.1155 | 227.9 | 0.043239 | 256 | 0.5437 | 340 | 0.002626 |  |  | 0
297.261 | 0.0195 | 228 | 0.038239 | 256.2 | 0.49872 | 340.5 | 0.0027458 |  |  | 0
297.609 | 0.0096667 | 228.1 | 0.045417 | 256.4 | 0.45833 | 341 | 0.003543 |  |  | 0
297.958 | 0.0091667 | 228.2 | 0.044755 | 256.6 | 0.42144 | 341.5 | 0.0036045 |  |  | 0
298.307 | 0.0078333 | 228.3 | 0.044093 | 256.8 | 0.38985 | 342 | 0.0035352 |  |  | 0
298.656 | 0.0026667 | 228.4 | 0.045273 | 257 | 0.37042 | 342.5 | 0.0025523 |  |  | 0
299.004 | 0.0036667 | 228.5 | 0.04646 | 257.2 | 0.35345 | 343 | 0.00068736 |  |  | 0
299.353 | 0.0063333 | 228.6 | 0.047647 | 257.4 | 0.34689 | 343.5 | 0.0028697 |  |  | 0
299.701 | 0.0058333 | 228.7 | 0.048835 | 257.6 | 0.34107 | 344 | 0.0026372 |  |  | 0
300.05 | 0.0043333 | 228.8 | 0.052369 | 257.8 | 0.32661 | 344.5 | 0.0017971 |  |  | 0
300.398 | 0.0036667 | 228.9 | 0.054058 | 258 | 0.31479 | 345 | 0.0016883 |  |  | 0
300.746 | 0.017 | 229 | 0.055748 | 258.2 | 0.29987 | 345.5 | 0.0023476 |  |  | 0
301.095 | 0.037333 | 229.1 | 0.057437 | 258.4 | 0.28843 | 346 | 0.0026618 |  |  | 0
301.443 | 0.056167 | 229.2 | 0.050079 | 258.6 | 0.27901 | 346.5 | 0.0027616 |  |  | 0
301.791 | 0.11217 | 229.3 | 0.049509 | 258.8 | 0.27599 | 347 | 0.0015835 |  |  | 0
302.138 | 0.13483 | 229.4 | 0.048939 | 259 | 0.27298 | 347.5 | 0.0015214 |  |  | 0
302.486 | 0.051167 | 229.5 | 0.057652 | 259.2 | 0.26614 | 348 | 0.0016996 |  |  | 0
302.834 | 0.016333 | 229.6 | 0.058958 | 259.4 | 0.25973 | 348.5 | 0.001978 |  |  | 0
303.182 | 0.0046667 | 229.7 | 0.060263 | 259.6 | 0.2483 | 349 | 0.0023651 |  |  | 0
303.529 | 0.0051667 | 229.8 | 0.061569 | 259.8 | 0.23897 | 349.5 | 0.0032593 |  |  | 0
303.877 | 0.004 | 229.9 | 0.044039 | 260 | 0.23261 | 350 | 0.0023779 |  |  | 0
304.224 | 0.0075 | 230 | 0.041322 | 260.2 | 0.23011 | 350.5 | 0.0014414 |  |  | 0
304.572 | 0.0066667 | 230.1 | 0.038605 | 260.4 | 0.2296 | 351 | 0.0017103 |  |  | 0
304.919 | 0.0055 | 230.2 | 0.035888 | 260.6 | 0.2256 | 351.5 | 0.0026787 |  |  | 0
305.266 | 0.0076667 | 230.3 | 0.049524 | 260.8 | 0.22395 | 352 | 0.0034748 |  |  | 0
305.613 | 0.0038333 | 230.4 | 0.048261 | 261 | 0.22622 | 352.5 | 0.0021023 |  |  | 0
305.96 | 0.0041667 | 230.5 | 0.046999 | 261.2 | 0.22359 | 353 | 0.00090484 |  |  | 0
306.307 | 0.0055 | 230.6 | 0.041769 | 261.4 | 0.21736 | 353.5 | 0.0025104 |  |  | 0
306.654 | 0.0061667 | 230.7 | 0.043806 | 261.6 | 0.21402 | 354 | 0.0031429 |  |  | 0
307.001 | 0.0055 | 230.8 | 0.045844 | 261.8 | 0.20777 | 354.5 | 0.0023367 |  |  | 0
307.347 | 0.0056667 | 230.9 | 0.047881 | 262 | 0.20793 | 355 | 0.0028795 |  |  | 0
307.694 | 0.0058333 | 231 | 0.048365 | 262.2 | 0.21048 | 355.5 | 0.0035363 |  |  | 0
308.04 | 0.006 | 231.1 | 0.046995 | 262.4 | 0.21397 | 356 | 0.0027437 |  |  | 0
308.387 | 0.0051667 | 231.2 | 0.045626 | 262.6 | 0.21589 | 356.5 | 0.0039585 |  |  | 0
308.733 | 0.0056667 | 231.3 | 0.041413 | 262.8 | 0.21162 | 357 | 0.0047593 |  |  | 0
309.08 | 0.0058333 | 231.4 | 0.040911 | 263 | 0.2092 | 357.5 | 0.0028339 |  |  | 0
309.426 | 0.0085 | 231.5 | 0.04041 | 263.2 | 0.20489 | 358 | 0.0023176 |  |  | 0
309.772 | 0.0055 | 231.6 | 0.039908 | 263.4 | 0.20958 | 358.5 | 0.0038879 |  |  | 0
310.118 | 0.0038333 | 231.7 | 0.039629 | 263.6 | 0.21733 | 359 | 0.0045067 |  |  | 0
310.464 | 0.0073333 | 231.8 | 0.037821 | 263.8 | 0.22378 | 359.5 | 0.003682 |  |  | 0
310.81 | 0.013333 | 231.9 | 0.036012 | 264 | 0.2299 | 360 | 0.0038091 |  |  | 0
311.156 | 0.032333 | 232 | 0.034203 | 264.2 | 0.22991 | 360.5 | 0.0034688 |  |  | 0
311.501 | 0.1125 | 232.1 | 0.036494 | 264.4 | 0.22921 | 361 | 0.0029419 |  |  | 0
311.847 | 0.2595 | 232.2 | 0.0367 | 264.6 | 0.22465 | 361.5 | 0.0045963 |  |  | 0
312.193 | 0.58017 | 232.3 | 0.036906 | 264.8 | 0.22455 | 362 | 0.0056168 |  |  | 0
312.538 | 0.775 | 232.4 | 0.03221 | 265 | 0.23159 | 362.5 | 0.0050007 |  |  | 0
312.883 | 0.76667 | 232.5 | 0.030539 | 265.2 | 0.22401 | 363 | 0.0069397 |  |  | 0
313.229 | 0.50267 | 232.6 | 0.028868 | 265.4 | 0.21649 | 363.5 | 0.017781 |  |  | 0
313.574 | 0.090333 | 232.7 | 0.027197 | 265.6 | 0.18517 | 364 | 0.14419 |  |  | 0
313.919 | 0.016167 | 232.8 | 0.028088 | 265.8 | 0.15142 | 364.5 | 0.49796 |  |  | 0
314.264 | 0.0065 | 232.9 | 0.026485 | 266 | 0.13066 | 365 | 0.57797 |  |  | 0
314.609 | 0.0073333 | 233 | 0.024882 | 266.2 | 0.10857 | 365.5 | 0.29061 |  |  | 0
314.954 | 0.0083333 | 233.1 | 0.02328 | 266.4 | 0.099807 | 366 | 0.1393 |  |  | 0
315.299 | 0.0058333 | 233.2 | 0.027058 | 266.6 | 0.098227 | 366.5 | 0.067907 |  |  | 0
315.644 | 0.0066667 | 233.3 | 0.027837 | 266.8 | 0.098983 | 367 | 0.017103 |  |  | 0
315.988 | 0.0025 | 233.4 | 0.028615 | 267 | 0.099874 | 367.5 | 0.0062981 |  |  | 0
316.333 | 0.0051667 | 233.5 | 0.035372 | 267.2 | 0.10354 | 368 | 0.0067522 |  |  | 0
316.678 | 0.0055 | 233.6 | 0.038577 | 267.4 | 0.10129 | 368.5 | 0.0060441 |  |  | 0
317.022 | 0.0043333 | 233.7 | 0.041782 | 267.6 | 0.096468 | 369 | 0.0050861 |  |  | 0
317.366 | 0.0065 | 233.8 | 0.044987 | 267.8 | 0.090333 | 369.5 | 0.0049534 |  |  | 0
317.711 | 0.0045 | 233.9 | 0.040468 | 268 | 0.088 | 370 | 0.0069779 |  |  | 0
318.055 | 0.0086667 | 234 | 0.041937 | 268.2 | 0.09148 | 370.5 | 0.0060567 |  |  | 0
318.399 | 0.005 | 234.1 | 0.043406 | 268.4 | 0.089635 | 371 | 0.0061652 |  |  | 0
318.743 | 0.0065 | 234.2 | 0.044875 | 268.6 | 0.089311 | 371.5 | 0.0047784 |  |  | 0
319.087 | 0.0021667 | 234.3 | 0.045306 | 268.8 | 0.088117 | 372 | 0.0053705 |  |  | 0
319.431 | 0.006 | 234.4 | 0.0451 | 269 | 0.085985 | 372.5 | 0.0067671 |  |  | 0
319.775 | 0.0035 | 234.5 | 0.044895 | 269.2 | 0.086055 | 373 | 0.006082 |  |  | 0
320.118 | 0.0065 | 234.6 | 0.04091 | 269.4 | 0.086433 | 373.5 | 0.0069168 |  |  | 0
320.462 | 0.0035 | 234.7 | 0.04027 | 269.6 | 0.087656 | 374 | 0.006432 |  |  | 0
320.806 | 0.0053333 | 234.8 | 0.039629 | 269.8 | 0.086286 | 374.5 | 0.0048735 |  |  | 0
321.149 | 0.0058333 | 234.9 | 0.038988 | 270 | 0.084454 | 375 | 0.0040196 |  |  | 0
321.493 | 0.0056667 | 235 | 0.038186 | 270.2 | 0.084802 | 375.5 | 0.0053013 |  |  | 0
321.836 | 0.0043333 | 235.1 | 0.03649 | 270.4 | 0.084716 | 376 | 0.0053569 |  |  | 0
322.179 | 0.005 | 235.2 | 0.034794 | 270.6 | 0.0878 | 376.5 | 0.0046344 |  |  | 0
322.522 | 0.0053333 | 235.3 | 0.035275 | 270.8 | 0.085258 | 377 | 0.004984 |  |  | 0
322.865 | 0.0065 | 235.4 | 0.035206 | 271 | 0.068146 | 377.5 | 0.0058059 |  |  | 0
323.208 | 0.0036667 | 235.5 | 0.035137 | 271.2 | 0.057822 | 378 | 0.0051441 |  |  | 0
323.551 | 0.0051667 | 235.6 | 0.035068 | 271.4 | 0.045625 | 378.5 | 0.0041481 |  |  | 0
323.894 | 0.0025 | 235.7 | 0.0359 | 271.6 | 0.03701 | 379 | 0.0042508 |  |  | 0
324.237 | 0.0018333 | 235.8 | 0.036199 | 271.8 | 0.032596 | 379.5 | 0.0039573 |  |  | 0
324.58 | 0.0015 | 235.9 | 0.036497 | 272 | 0.032112 | 380 | 0.0047029 |  |  | 0
324.922 | 0.0043333 | 236 | 0.036795 | 272.2 | 0.032575 | 380.5 | 0.0048158 |  |  | 0
325.265 | 0.0053333 | 236.1 | 0.032528 | 272.4 | 0.033422 | 381 | 0.005182 |  |  | 0
325.607 | 0.0025 | 236.2 | 0.031494 | 272.6 | 0.034022 | 381.5 | 0.0057132 |  |  | 0
325.95 | 0.0028333 | 236.3 | 0.030461 | 272.8 | 0.032781 | 382 | 0.004766 |  |  | 0
326.292 | 0.0055 | 236.4 | 0.032584 | 273 | 0.031188 | 382.5 | 0.0046572 |  |  | 0
326.634 | 0.0023333 | 236.5 | 0.032033 | 273.2 | 0.029111 | 383 | 0.0052541 |  |  | 0
326.976 | 0.005 | 236.6 | 0.031482 | 273.4 | 0.028432 | 383.5 | 0.0043736 |  |  | 0
327.318 | 0.003 | 236.7 | 0.030931 | 273.6 | 0.028094 | 384 | 0.0027242 |  |  | 0
327.66 | 0.0056667 | 236.8 | 0.029635 | 273.8 | 0.028283 | 384.5 | 0.0036808 |  |  | 0
328.002 | 0.0043333 | 236.9 | 0.029681 | 274 | 0.02907 | 385 | 0.0048971 |  |  | 0
328.344 | 0.0035 | 237 | 0.029727 | 274.2 | 0.030787 | 385.5 | 0.0045774 |  |  | 0
328.686 | 0.0021667 | 237.1 | 0.029773 | 274.4 | 0.031914 | 386 | 0.0043588 |  |  | 0
329.027 | -0.00016667 | 237.2 | 0.031976 | 274.6 | 0.032945 | 386.5 | 0.0055808 |  |  | 0
329.369 | 0.003 | 237.3 | 0.032137 | 274.8 | 0.039641 | 387 | 0.0053696 |  |  | 0
329.711 | 0.0013333 | 237.4 | 0.032298 | 275 | 0.047089 | 387.5 | 0.0049199 |  |  | 0
330.052 | 0.0036667 | 237.5 | 0.037986 | 275.2 | 0.048123 | 388 | 0.0046508 |  |  | 0
330.393 | 0.0041667 | 237.6 | 0.040925 | 275.4 | 0.046645 | 388.5 | 0.004683 |  |  | 0
330.735 | 0.0015 | 237.7 | 0.043863 | 275.6 | 0.038391 | 389 | 0.0052662 |  |  | 0
331.076 | 0.0045 | 237.8 | 0.046802 | 275.8 | 0.032595 | 389.5 | 0.0062229 |  |  | 0
331.417 | 0.003 | 237.9 | 0.043114 | 276 | 0.028608 | 390 | 0.0078033 |  |  | 0
331.758 | 0.0026667 | 238 | 0.044354 | 276.2 | 0.025594 | 390.5 | 0.0087456 |  |  | 0
332.099 | 0.002 | 238.1 | 0.045594 | 276.4 | 0.024416 | 391 | 0.0059392 |  |  | 0
332.44 | 0.0023333 | 238.2 | 0.046833 | 276.6 | 0.023992 | 391.5 | 0.0057377 |  |  | 0
332.78 | 0.0046667 | 238.3 | 0.041201 | 276.8 | 0.023984 | 392 | 0.005322 |  |  | 0
333.121 | 0.012667 | 238.4 | 0.039408 | 277 | 0.024416 | 392.5 | 0.0053585 |  |  | 0
333.462 | 0.035167 | 238.5 | 0.037615 | 277.2 | 0.023989 | 393 | 0.0059401 |  |  | 0
333.802 | 0.066 | 238.6 | 0.042028 | 277.4 | 0.021885 | 393.5 | 0.0049809 |  |  | 0
334.143 | 0.054 | 238.7 | 0.042397 | 277.6 | 0.022293 | 394 | 0.005191 |  |  | 0
334.483 | 0.014333 | 238.8 | 0.042766 | 277.8 | 0.02259 | 394.5 | 0.0055641 |  |  | 0
334.823 | 0.0028333 | 238.9 | 0.043134 | 278 | 0.019565 | 395 | 0.0059122 |  |  | 0
335.164 | 0.002 | 239 | 0.035767 | 278.2 | 0.017451 | 395.5 | 0.0050387 |  |  | 0
335.504 | 0.00033333 | 239.1 | 0.034574 | 278.4 | 0.016743 | 396 | 0.004442 |  |  | 0
335.844 | 0.0038333 | 239.2 | 0.03338 | 278.6 | 0.015366 | 396.5 | 0.0047037 |  |  | 0
336.184 | 0.003 | 239.3 | 0.036323 | 278.8 | 0.016481 | 397 | 0.0056202 |  |  | 0
336.524 | 0.00083333 | 239.4 | 0.034922 | 279 | 0.017938 | 397.5 | 0.0057779 |  |  | 0
336.864 | 0.0005 | 239.5 | 0.033521 | 279.2 | 0.020026 | 398 | 0.0058764 |  |  | 0
337.203 | 0.0023333 | 239.6 | 0.032121 | 279.4 | 0.021986 | 398.5 | 0.0059045 |  |  | 0
337.543 | 0.0023333 | 239.7 | 0.038847 | 279.6 | 0.024265 | 399 | 0.0060387 |  |  | 0
337.883 | 0.002 | 239.8 | 0.041743 | 279.8 | 0.030433 | 399.5 | 0.0060288 |  |  | 0
338.222 | 0.0036667 | 239.9 | 0.044639 | 280 | 0.036333 | 400 | 0.0055052 |  |  | 0
338.562 | 0.0016667 | 240 | 0.047535 | 280.2 | 0.040524 | 400.5 | 0.0053671 |  |  | 0
338.901 | 0.0033333 | 240.1 | 0.047844 | 280.4 | 0.041155 | 401 | 0.0062106 |  |  | 0
339.24 | 0.005 | 240.2 | 0.050192 | 280.6 | 0.03191 | 401.5 | 0.0061405 |  |  | 0
339.58 | 0.00033333 | 240.3 | 0.05254 | 280.8 | 0.028332 | 402 | 0.0067098 |  |  | 0
339.919 | 0.004 | 240.4 | 0.047779 | 281 | 0.023048 | 402.5 | 0.0094172 |  |  | 0
340.258 | 0.0018333 | 240.5 | 0.046927 | 281.2 | 0.022409 | 403 | 0.019784 |  |  | 0
340.597 | 0.0028333 | 240.6 | 0.046075 | 281.4 | 0.02333 | 403.5 | 0.11315 |  |  | 0
340.936 | 0.0038333 | 240.7 | 0.045223 | 281.6 | 0.022343 | 404 | 0.4786 |  |  | 0
341.274 | 0.0035 | 240.8 | 0.047561 | 281.8 | 0.020623 | 404.5 | 0.84313 |  |  | 0
341.613 | 0.0036667 | 240.9 | 0.046963 | 282 | 0.016072 | 405 | 0.42189 |  |  | 0
341.952 | 0.0035 | 241 | 0.046364 | 282.2 | 0.015325 | 405.5 | 0.043629 |  |  | 0
342.29 | 0.0036667 | 241.1 | 0.04478 | 282.4 | 0.014762 |  |  |  |  | 0
342.629 | 0.0023333 | 241.2 | 0.044895 | 282.6 | 0.012569 |  |  |  |  | 0
342.967 | -0.0011667 | 241.3 | 0.04501 | 282.8 | 0.012275 |  |  |  |  | 0
343.306 | 0.0021667 | 241.4 | 0.045125 | 283 | 0.014164 |  |  |  |  | 0
343.644 | 0.0043333 | 241.5 | 0.045806 | 283.2 | 0.014109 |  |  |  |  | 0
343.982 | 0.002 | 241.6 | 0.045461 | 283.4 | 0.014215 |  |  |  |  | 0
344.32 | 0.0025 | 241.7 | 0.045116 | 283.6 | 0.013821 |  |  |  |  | 0
344.658 | 0.0011667 | 241.8 | 0.04477 | 283.8 | 0.012027 |  |  |  |  | 0
344.996 | 0.002 | 241.9 | 0.04481 | 284 | 0.012496 |  |  |  |  | 0
345.334 | 0.0015 | 242 | 0.044994 | 284.2 | 0.012596 |  |  |  |  | 0
345.672 | 0.0031667 | 242.1 | 0.045179 | 284.4 | 0.010788 |  |  |  |  | 0
346.01 | 0.0025 | 242.2 | 0.044667 | 284.6 | 0.010506 |  |  |  |  | 0
346.347 | 0.0026667 | 242.3 | 0.044322 | 284.8 | 0.010724 |  |  |  |  | 0
346.685 | 0.0033333 | 242.4 | 0.043976 | 285 | 0.011121 |  |  |  |  | 0
347.022 | 0.00033333 | 242.5 | 0.043631 | 285.2 | 0.012705 |  |  |  |  | 0
347.36 | 0.0021667 | 242.6 | 0.047886 | 285.4 | 0.013812 |  |  |  |  | 0
347.697 | 0.0013333 | 242.7 | 0.049107 | 285.6 | 0.014757 |  |  |  |  | 0
348.034 | 0.0011667 | 242.8 | 0.050328 | 285.8 | 0.014963 |  |  |  |  | 0
348.371 | 0.003 | 242.9 | 0.051549 | 286 | 0.014485 |  |  |  |  | 0
348.708 | 0.0013333 | 243 | 0.048063 | 286.2 | 0.012967 |  |  |  |  | 0
349.045 | 0.0031667 | 243.1 | 0.048431 | 286.4 | 0.011501 |  |  |  |  | 0
349.382 | 0.002 | 243.2 | 0.048798 | 286.6 | 0.0091638 |  |  |  |  | 0
349.719 | 0.0043333 | 243.3 | 0.054006 | 286.8 | 0.0097761 |  |  |  |  | 0
350.056 | 0.0015 | 243.4 | 0.055088 | 287 | 0.010854 |  |  |  |  | 0
350.393 | 0.0015 | 243.5 | 0.05617 | 287.2 | 0.010324 |  |  |  |  | 0
350.729 | 0.0013333 | 243.6 | 0.057252 | 287.4 | 0.010417 |  |  |  |  | 0
351.066 | 0.0025 | 243.7 | 0.055017 | 287.6 | 0.010707 |  |  |  |  | 0
351.402 | 0.00083333 | 243.8 | 0.056237 | 287.8 | 0.012888 |  |  |  |  | 0
351.739 | 0.0046667 | 243.9 | 0.057458 | 288 | 0.020195 |  |  |  |  | 0
352.075 | 0.0028333 | 244 | 0.063705 | 288.2 | 0.023851 |  |  |  |  | 0
352.411 | 0.0038333 | 244.1 | 0.065502 | 288.4 | 0.028232 |  |  |  |  | 0
352.747 | 0 | 244.2 | 0.067298 | 288.6 | 0.036137 |  |  |  |  | 0
353.083 | 0.00066667 | 244.3 | 0.069095 | 288.8 | 0.045196 |  |  |  |  | 0
353.419 | 0.0016667 | 244.4 | 0.06928 | 289 | 0.057617 |  |  |  |  | 0
353.755 | 0.0041667 | 244.5 | 0.072004 | 289.2 | 0.061823 |  |  |  |  | 0
354.091 | 0.0031667 | 244.6 | 0.074728 | 289.4 | 0.051314 |  |  |  |  | 0
354.427 | 0.002 | 244.7 | 0.077452 | 289.6 | 0.03892 |  |  |  |  | 0
354.763 | 0.0023333 | 244.8 | 0.076857 | 289.8 | 0.020353 |  |  |  |  | 0
355.098 | 0.0035 | 244.9 | 0.077596 | 290 | 0.013843 |  |  |  |  | 0
355.434 | 0.003 | 245 | 0.078335 | 290.2 | 0.01035 |  |  |  |  | 0
355.769 | 0.0041667 | 245.1 | 0.08032 | 290.4 | 0.0095571 |  |  |  |  | 0
356.104 | 0.0016667 | 245.2 | 0.08182 | 290.6 | 0.009564 |  |  |  |  | 0
356.44 | 0.0031667 | 245.3 | 0.08332 | 290.8 | 0.010553 |  |  |  |  | 0
356.775 | 0.006 | 245.4 | 0.084821 | 291 | 0.009575 |  |  |  |  | 0
357.11 | 0.0043333 | 245.5 | 0.0878 | 291.2 | 0.0087779 |  |  |  |  | 0
357.445 | 0.0038333 | 245.6 | 0.089831 | 291.4 | 0.01021 |  |  |  |  | 0
357.78 | 0.00083333 | 245.7 | 0.091863 | 291.6 | 0.011274 |  |  |  |  | 0
358.115 | 0.003 | 245.8 | 0.094069 | 291.8 | 0.012159 |  |  |  |  | 0
358.45 | 0.0028333 | 245.9 | 0.096378 | 292 | 0.014222 |  |  |  |  | 0
358.785 | 0.0056667 | 246 | 0.098686 | 292.2 | 0.016921 |  |  |  |  | 0
359.119 | 0.0041667 | 246.1 | 0.10099 | 292.4 | 0.016777 |  |  |  |  | 0
359.454 | 0.0036667 | 246.2 | 0.1021 | 292.6 | 0.016777 |  |  |  |  | 0
359.788 | 0.0033333 | 246.3 | 0.10392 | 292.8 | 0.013266 |  |  |  |  | 0
360.123 | 0.004 | 246.4 | 0.10575 | 293 | 0.0098383 |  |  |  |  | 0
360.457 | 0.0045 | 246.5 | 0.10757 | 293.2 | 0.0099467 |  |  |  |  | 0
360.791 | 0.002 | 246.6 | 0.10826 | 293.4 | 0.0090209 |  |  |  |  | 0
361.126 | 0.0026667 | 246.7 | 0.1096 | 293.6 | 0.009574 |  |  |  |  | 0
361.46 | 0.0056667 | 246.8 | 0.11094 | 293.8 | 0.0095377 |  |  |  |  | 0
361.794 | 0.0046667 | 246.9 | 0.11287 | 294 | 0.0088878 |  |  |  |  | 0
362.128 | 0.0063333 | 247 | 0.11437 | 294.2 | 0.0083001 |  |  |  |  | 0
362.462 | 0.0053333 | 247.1 | 0.11587 | 294.4 | 0.0081658 |  |  |  |  | 0
362.795 | 0.0038333 | 247.2 | 0.11737 | 294.6 | 0.00889 |  |  |  |  | 0
363.129 | 0.0083333 | 247.3 | 0.12795 | 294.8 | 0.010845 |  |  |  |  | 0
363.463 | 0.0115 | 247.4 | 0.13259 | 295 | 0.016138 |  |  |  |  | 0
363.796 | 0.031333 | 247.5 | 0.13723 | 295.2 | 0.025574 |  |  |  |  | 0
364.13 | 0.1615 | 247.6 | 0.14187 | 295.4 | 0.036863 |  |  |  |  | 0
364.463 | 0.50283 | 247.7 | 0.15442 | 295.6 | 0.049675 |  |  |  |  | 0
364.797 | 0.7195 | 247.8 | 0.1624 | 295.8 | 0.068344 |  |  |  |  | 0
365.13 | 0.57083 | 247.9 | 0.17037 | 296 | 0.096251 |  |  |  |  | 0
365.463 | 0.24433 | 248 | 0.17345 | 296.2 | 0.12878 |  |  |  |  | 0
365.796 | 0.151 | 248.1 | 0.17989 | 296.4 | 0.18027 |  |  |  |  | 0
366.129 | 0.12883 | 248.2 | 0.18632 | 296.6 | 0.16971 |  |  |  |  | 0
366.462 | 0.076333 | 248.3 | 0.19276 | 296.8 | 0.1508 |  |  |  |  | 0
366.795 | 0.018667 | 248.4 | 0.18996 | 297 | 0.10221 |  |  |  |  | 0
367.128 | 0.0083333 | 248.5 | 0.19273 | 297.2 | 0.04731 |  |  |  |  | 0
367.461 | 0.0058333 | 248.6 | 0.19551 | 297.4 | 0.027029 |  |  |  |  | 0
367.793 | 0.0056667 | 248.7 | 0.19389 | 297.6 | 0.0066566 |  |  |  |  | 0
368.126 | 0.0081667 | 248.8 | 0.19422 | 297.8 | 0.0099523 |  |  |  |  | 0
368.458 | 0.0048333 | 248.9 | 0.19454 | 298 | 0.0085883 |  |  |  |  | 0
368.791 | 0.0061667 | 249 | 0.19486 | 298.2 | 0.0082533 |  |  |  |  | 0
369.123 | 0.0043333 | 249.1 | 0.20667 | 298.4 | 0.0061461 |  |  |  |  | 0
369.455 | 0.0053333 | 249.2 | 0.21052 | 298.6 | 0.0043474 |  |  |  |  | 0
369.787 | 0.0048333 | 249.3 | 0.21436 | 298.8 | 0.0038588 |  |  |  |  | 0
370.12 | 0.0091667 | 249.4 | 0.2182 | 299 | 0.0035495 |  |  |  |  | 0
370.452 | 0.0051667 | 249.5 | 0.2338 | 299.2 | 0.0052518 |  |  |  |  | 0
370.784 | 0.0053333 | 249.6 | 0.24267 | 299.4 | 0.0056464 |  |  |  |  | 0
371.115 | 0.0075 | 249.7 | 0.25153 | 299.6 | 0.0060472 |  |  |  |  | 0
371.447 | 0.0038333 | 249.8 | 0.26706 | 299.8 | 0.0055016 |  |  |  |  | 0
371.779 | 0.0043333 | 249.9 | 0.27861 | 300 | 0.0046757 |  |  |  |  | 0
372.111 | 0.0058333 | 250 | 0.29016 | 300.2 | 0.0041437 |  |  |  |  | 0
372.442 | 0.0073333 | 250.1 | 0.30171 | 300.4 | 0.0074641 |  |  |  |  | 0
372.774 | 0.0066667 | 250.2 | 0.32476 | 300.6 | 0.01201 |  |  |  |  | 0
373.105 | 0.0051667 | 250.3 | 0.34068 | 300.8 | 0.020837 |  |  |  |  | 0
373.436 | 0.0075 | 250.4 | 0.35661 | 301 | 0.031826 |  |  |  |  | 0
373.768 | 0.0071667 | 250.5 | 0.38493 | 301.2 | 0.042752 |  |  |  |  | 0
374.099 | 0.006 | 250.6 | 0.40676 | 301.4 | 0.059016 |  |  |  |  | 0
374.43 | 0.0055 | 250.7 | 0.42859 | 301.6 | 0.085436 |  |  |  |  | 0
374.761 | 0.0036667 | 250.8 | 0.45042 | 301.8 | 0.10391 |  |  |  |  | 0
375.092 | 0.0036667 | 250.9 | 0.57218 | 302 | 0.12475 |  |  |  |  | 0
375.423 | 0.0046667 | 251 | 0.62895 | 302.2 | 0.10508 |  |  |  |  | 0
375.754 | 0.0066667 | 251.1 | 0.68573 | 302.4 | 0.075551 |  |  |  |  | 0
376.084 | 0.0043333 | 251.2 | 0.74251 | 302.6 | 0.048035 |  |  |  |  | 0
376.415 | 0.0055 | 251.3 | 1.6546 | 302.8 | 0.021546 |  |  |  |  | 0
376.745 | 0.0038333 | 251.4 | 1.9753 | 303 | 0.012965 |  |  |  |  | 0
377.076 | 0.0056667 | 251.5 | 2.2961 | 303.2 | 0.0071772 |  |  |  |  | 0
377.406 | 0.0051667 | 251.6 | 4.4425 | 303.4 | 0.0052124 |  |  |  |  | 0
377.737 | 0.0065 | 251.7 | 5.5173 | 303.6 | 0.0041984 |  |  |  |  | 0
378.067 | 0.0048333 | 251.8 | 6.5921 | 303.8 | 0.0048107 |  |  |  |  | 0
378.397 | 0.0038333 | 251.9 | 7.6669 | 304 | 0.0059692 |  |  |  |  | 0
378.727 | 0.0041667 | 252 | 8.4084 | 304.2 | 0.0064301 |  |  |  |  | 0
379.057 | 0.005 | 252.1 | 9.5754 | 304.4 | 0.0067291 |  |  |  |  | 0
379.387 | 0.0031667 | 252.2 | 10.7424 | 304.6 | 0.0065909 |  |  |  |  | 0
379.717 | 0.0043333 | 252.3 | 11.0318 | 304.8 | 0.0058985 |  |  |  |  | 0
380.047 | 0.0046667 | 252.4 | 11.6614 | 305 | 0.0064035 |  |  |  |  | 0
380.377 | 0.0056667 | 252.5 | 12.2909 | 305.2 | 0.006491 |  |  |  |  | 0
380.706 | 0.0041667 | 252.6 | 12.9204 | 305.4 | 0.0053449 |  |  |  |  | 0
381.036 | 0.0053333 | 252.7 | 14.6236 | 305.6 | 0.0049176 |  |  |  |  | 0
381.365 | 0.006 | 252.8 | 15.5687 | 305.8 | 0.0042791 |  |  |  |  | 0
381.695 | 0.0058333 | 252.9 | 16.5138 | 306 | 0.0044024 |  |  |  |  | 0
382.024 | 0.004 | 253 | 21.507 | 306.2 | 0.0050764 |  |  |  |  | 0
382.353 | 0.005 | 253.1 | 24.3403 | 306.4 | 0.0055654 |  |  |  |  | 0
382.682 | 0.0043333 | 253.2 | 27.1736 | 306.6 | 0.0058928 |  |  |  |  | 0
383.012 | 0.0058333 | 253.3 | 30.0068 | 306.8 | 0.0057222 |  |  |  |  | 0
383.341 | 0.005 | 253.4 | 40.1428 | 307 | 0.0057168 |  |  |  |  | 0
383.67 | 0.0038333 | 253.5 | 46.1983 | 307.2 | 0.0056338 |  |  |  |  | 0
383.998 | 0.002 | 253.6 | 52.2537 | 307.4 | 0.0056646 |  |  |  |  | 0
384.327 | 0.0026667 | 253.7 | 58.3091 | 307.6 | 0.0057882 |  |  |  |  | 0
384.656 | 0.0046667 | 253.8 | 40.3467 | 307.8 | 0.0058845 |  |  |  |  | 0
384.985 | 0.0053333 | 253.9 | 39.5999 | 308 | 0.0058408 |  |  |  |  | 0
385.313 | 0.0046667 | 254 | 38.8531 | 308.2 | 0.005513 |  |  |  |  | 0
385.642 | 0.0045 | 254.1 | 29.1193 | 308.4 | 0.005526 |  |  |  |  | 0
385.97 | 0.0038333 | 254.2 | 21.0742 | 308.6 | 0.0055083 |  |  |  |  | 0
386.298 | 0.0048333 | 254.3 | 13.0291 | 308.8 | 0.0055985 |  |  |  |  | 0
386.627 | 0.0063333 | 254.4 | 4.984 | 309 | 0.0060925 |  |  |  |  | 0
386.955 | 0.0053333 | 254.5 | 8.5432 | 309.2 | 0.0071584 |  |  |  |  | 0
387.283 | 0.0048333 | 254.6 | 4.8034 | 309.4 | 0.0073728 |  |  |  |  | 0
387.611 | 0.005 | 254.7 | 1.0636 | 309.6 | 0.0065273 |  |  |  |  | 0
387.939 | 0.0045 | 254.8 | 1.8563 | 309.8 | 0.005772 |  |  |  |  | 0
388.267 | 0.0046667 | 254.9 | 1.4532 | 310 | 0.0044069 |  |  |  |  | 0
388.595 | 0.0046667 | 255 | 1.0502 | 310.2 | 0.005181 |  |  |  |  | 0
388.922 | 0.0048333 | 255.1 | 0.64721 | 310.4 | 0.0070004 |  |  |  |  | 0
389.25 | 0.0061667 | 255.2 | 0.84109 | 310.6 | 0.010034 |  |  |  |  | 0
389.578 | 0.006 | 255.3 | 0.76499 | 310.8 | 0.015035 |  |  |  |  | 0
389.905 | 0.0073333 | 255.4 | 0.68889 | 311 | 0.024531 |  |  |  |  | 0
390.233 | 0.009 | 255.5 | 0.61279 | 311.2 | 0.053552 |  |  |  |  | 0
390.56 | 0.010167 | 255.6 | 0.64807 | 311.4 | 0.087776 |  |  |  |  | 0
390.887 | 0.0048333 | 255.7 | 0.62183 | 311.6 | 0.15648 |  |  |  |  | 0
391.214 | 0.0053333 | 255.8 | 0.59559 | 311.8 | 0.26209 |  |  |  |  | 0
391.542 | 0.0063333 | 255.9 | 0.56776 | 312 | 0.42079 |  |  |  |  | 0
391.869 | 0.005 | 256 | 0.5437 | 312.2 | 0.55012 |  |  |  |  | 0
392.196 | 0.0051667 | 256.1 | 0.51964 | 312.4 | 0.69292 |  |  |  |  | 0
392.522 | 0.0056667 | 256.2 | 0.49559 | 312.6 | 0.74754 |  |  |  |  | 0
392.849 | 0.0048333 | 256.3 | 0.48092 | 312.8 | 0.73468 |  |  |  |  | 0
393.176 | 0.0076667 | 256.4 | 0.45833 | 313 | 0.63544 |  |  |  |  | 0
393.503 | 0.0038333 | 256.5 | 0.43574 | 313.2 | 0.512 |  |  |  |  | 0
393.829 | 0.0046667 | 256.6 | 0.42244 | 313.4 | 0.28554 |  |  |  |  | 0
394.156 | 0.0066667 | 256.7 | 0.40614 | 313.6 | 0.16221 |  |  |  |  | 0
394.482 | 0.0031667 | 256.8 | 0.38985 | 313.8 | 0.043709 |  |  |  |  | 0
394.809 | 0.0076667 | 256.9 | 0.37355 | 314 | 0.0055676 |  |  |  |  | 0
395.135 | 0.0051667 | 257 | 0.374 | 314.2 | 0.0095802 |  |  |  |  | 0
395.461 | 0.0051667 | 257.1 | 0.36372 | 314.4 | 0.0082589 |  |  |  |  | 0
395.787 | 0.0046667 | 257.2 | 0.35345 | 314.6 | 0.0067693 |  |  |  |  | 0
396.113 | 0.0043333 | 257.3 | 0.34318 | 314.8 | 0.0078964 |  |  |  |  | 0
396.439 | 0.0038333 | 257.4 | 0.35035 | 315 | 0.0073593 |  |  |  |  | 0
396.765 | 0.0058333 | 257.5 | 0.34714 | 315.2 | 0.0068077 |  |  |  |  | 0
397.091 | 0.006 | 257.6 | 0.34394 | 315.4 | 0.0067005 |  |  |  |  | 0
397.417 | 0.0053333 | 257.7 | 0.33213 | 315.6 | 0.0058013 |  |  |  |  | 0
397.742 | 0.0061667 | 257.8 | 0.32661 | 315.8 | 0.0042442 |  |  |  |  | 0
398.068 | 0.0063333 | 257.9 | 0.32109 | 316 | 0.0043495 |  |  |  |  | 0
398.394 | 0.0048333 | 258 | 0.31558 | 316.2 | 0.004318 |  |  |  |  | 0
398.719 | 0.0066667 | 258.1 | 0.30769 | 316.4 | 0.004632 |  |  |  |  | 0
399.044 | 0.0053333 | 258.2 | 0.29987 | 316.6 | 0.0052126 |  |  |  |  | 0
399.37 | 0.0068333 | 258.3 | 0.29205 | 316.8 | 0.0048522 |  |  |  |  | 0
399.695 | 0.0055 | 258.4 | 0.28916 | 317 | 0.0049665 |  |  |  |  | 0
400.02 | 0.0058333 | 258.5 | 0.28409 | 317.2 | 0.0057032 |  |  |  |  | 0
400.345 | 0.0046667 | 258.6 | 0.27901 | 317.4 | 0.0053702 |  |  |  |  | 0
400.67 | 0.0055 | 258.7 | 0.27394 | 317.6 | 0.005167 |  |  |  |  | 0
400.995 | 0.0076667 | 258.8 | 0.27788 | 317.8 | 0.0063924 |  |  |  |  | 0
401.32 | 0.0048333 | 258.9 | 0.27614 | 318 | 0.0069008 |  |  |  |  | 0
401.645 | 0.007 | 259 | 0.27439 | 318.2 | 0.0061609 |  |  |  |  | 0
401.969 | 0.006 | 259.1 | 0.26842 | 318.4 | 0.0064523 |  |  |  |  | 0
402.294 | 0.0073333 | 259.2 | 0.26614 | 318.6 | 0.0060892 |  |  |  |  | 0
402.618 | 0.010333 | 259.3 | 0.26386 | 318.8 | 0.004668 |  |  |  |  | 0
402.943 | 0.014 | 259.4 | 0.26158 | 319 | 0.0040952 |  |  |  |  | 0
403.267 | 0.0235 | 259.5 | 0.25375 | 319.2 | 0.0048068 |  |  |  |  | 0
403.592 | 0.10167 | 259.6 | 0.2483 | 319.4 | 0.0044727 |  |  |  |  | 0
403.916 | 0.35633 | 259.7 | 0.24284 | 319.6 | 0.0042164 |  |  |  |  | 0
404.24 | 0.77417 | 259.8 | 0.23739 | 319.8 | 0.005062 |  |  |  |  | 0
404.564 | 1.0243 | 259.9 | 0.23669 | 320 | 0.0054965 |  |  |  |  | 0
404.888 | 0.58367 | 260 | 0.23261 | 320.2 | 0.0048559 |  |  |  |  | 0
405.212 | 0.097333 | 260.1 | 0.22853 | 320.4 | 0.0046816 |  |  |  |  | 0
405.536 | 0.0155 | 260.2 | 0.23132 | 320.6 | 0.0048771 |  |  |  |  | 0
405.86 | 0.010667 | 260.3 | 0.23046 | 320.8 | 0.0048183 |  |  |  |  | 0
406.183 | 0.0083333 | 260.4 | 0.2296 | 321 | 0.0055491 |  |  |  |  | 0
406.507 | 0.011833 | 260.5 | 0.22873 | 321.2 | 0.0057362 |  |  |  |  | 0
406.831 | 0.018667 | 260.6 | 0.22466 | 321.4 | 0.0055595 |  |  |  |  | 0
407.154 | 0.044333 | 260.7 | 0.2234 | 321.6 | 0.0050435 |  |  |  |  | 0
407.478 | 0.080167 | 260.8 | 0.22214 | 321.8 | 0.00479 |  |  |  |  | 0
407.801 | 0.083 | 260.9 | 0.22632 | 322 | 0.0048406 |  |  |  |  | 0
408.124 | 0.037167 | 261 | 0.22622 | 322.2 | 0.0049094 |  |  |  |  | 0
408.447 | 0.010333 | 261.1 | 0.22613 | 322.4 | 0.005211 |  |  |  |  | 0
408.77 | 0.0063333 | 261.2 | 0.22603 | 322.6 | 0.0057129 |  |  |  |  | 0
409.093 | 0.0073333 | 261.3 | 0.21862 | 322.8 | 0.005767 |  |  |  |  | 0
409.416 | 0.0055 | 261.4 | 0.21736 | 323 | 0.0048387 |  |  |  |  | 0
409.739 | 0.0068333 | 261.5 | 0.2161 | 323.2 | 0.0046269 |  |  |  |  | 0
410.062 | 0.0075 | 261.6 | 0.21472 | 323.4 | 0.0047379 |  |  |  |  | 0
410.385 | 0.0083333 | 261.7 | 0.21124 | 323.6 | 0.0039208 |  |  |  |  | 0
410.708 | 0.0053333 | 261.8 | 0.20777 | 323.8 | 0.0033387 |  |  |  |  | 0
411.03 | 0.0068333 | 261.9 | 0.20429 | 324 | 0.0026516 |  |  |  |  | 0
411.353 | 0.0048333 | 262 | 0.2095 | 324.2 | 0.0018994 |  |  |  |  | 0
411.675 | 0.0071667 | 262.1 | 0.20999 | 324.4 | 0.0017068 |  |  |  |  | 0
411.997 | 0.0058333 | 262.2 | 0.21048 | 324.6 | 0.0023951 |  |  |  |  | 0
412.32 | 0.0065 | 262.3 | 0.21097 | 324.8 | 0.0033592 |  |  |  |  | 0
412.642 | 0.0081667 | 262.4 | 0.21444 | 325 | 0.004199 |  |  |  |  | 0
412.964 | 0.0073333 | 262.5 | 0.21649 | 325.2 | 0.004651 |  |  |  |  | 0
413.286 | 0.006 | 262.6 | 0.21855 | 325.4 | 0.0036937 |  |  |  |  | 0
413.608 | 0.0051667 | 262.7 | 0.21262 | 325.6 | 0.0032676 |  |  |  |  | 0
413.93 | 0.0058333 | 262.8 | 0.21162 | 325.8 | 0.0028528 |  |  |  |  | 0
414.252 | 0.0061667 | 262.9 | 0.21062 | 326 | 0.0035296 |  |  |  |  | 0
414.574 | 0.0071667 | 263 | 0.20961 | 326.2 | 0.0041753 |  |  |  |  | 0
414.895 | 0.0053333 | 263.1 | 0.20737 | 326.4 | 0.0034766 |  |  |  |  | 0
415.217 | 0.006 | 263.2 | 0.20489 | 326.6 | 0.0036545 |  |  |  |  | 0
415.538 | 0.0043333 | 263.3 | 0.20242 | 326.8 | 0.0041564 |  |  |  |  | 0
415.86 | 0.0058333 | 263.4 | 0.21178 | 327 | 0.0037055 |  |  |  |  | 0
416.181 | 0.0078333 | 263.5 | 0.21456 | 327.2 | 0.0036628 |  |  |  |  | 0
416.503 | 0.0036667 | 263.6 | 0.21733 | 327.4 | 0.0043762 |  |  |  |  | 0
416.824 | 0.0043333 | 263.7 | 0.22011 | 327.6 | 0.0046582 |  |  |  |  | 0
417.145 | 0.0051667 | 263.8 | 0.22336 | 327.8 | 0.0046062 |  |  |  |  | 0
417.466 | 0.0065 | 263.9 | 0.22785 | 328 | 0.0046623 |  |  |  |  | 0
417.787 | 0.0043333 | 264 | 0.23234 | 328.2 | 0.0038728 |  |  |  |  | 0
418.108 | 0.0026667 | 264.1 | 0.22952 | 328.4 | 0.0031837 |  |  |  |  | 0
418.429 | 0.0051667 | 264.2 | 0.22991 | 328.6 | 0.0023727 |  |  |  |  | 0
418.75 | 0.0053333 | 264.3 | 0.23031 | 328.8 | 0.0012196 |  |  |  |  | 0
419.07 | 0.0065 | 264.4 | 0.23071 | 329 | 0.00086151 |  |  |  |  | 0
419.391 | 0.004 | 264.5 | 0.22661 | 329.2 | 0.001878 |  |  |  |  | 0
419.712 | 0.0045 | 264.6 | 0.22465 | 329.4 | 0.0017112 |  |  |  |  | 0
420.032 | 0.005 | 264.7 | 0.22269 | 329.6 | 0.0018955 |  |  |  |  | 0
420.353 | 0.0063333 | 264.8 | 0.22073 | 329.8 | 0.0025731 |  |  |  |  | 0
420.673 | 0.0031667 | 264.9 | 0.23024 | 330 | 0.0030181 |  |  |  |  | 0
420.993 | 0.0046667 | 265 | 0.23159 | 330.2 | 0.0036704 |  |  |  |  | 0
421.313 | 0.0046667 | 265.1 | 0.23294 | 330.4 | 0.0033161 |  |  |  |  | 0
421.633 | 0.0041667 | 265.2 | 0.22056 | 330.6 | 0.002455 |  |  |  |  | 0
421.953 | 0.0038333 | 265.3 | 0.21853 | 330.8 | 0.003009 |  |  |  |  | 0
422.273 | 0.0026667 | 265.4 | 0.21649 | 331 | 0.0033463 |  |  |  |  | 0
422.593 | 0.0051667 | 265.5 | 0.21446 | 331.2 | 0.0032727 |  |  |  |  | 0
422.913 | 0.0046667 | 265.6 | 0.17829 | 331.4 | 0.003457 |  |  |  |  | 0
423.233 | 0.0021667 | 265.7 | 0.16276 | 331.6 | 0.002897 |  |  |  |  | 0
423.553 | 0.0033333 | 265.8 | 0.14724 | 331.8 | 0.0024898 |  |  |  |  | 0
423.872 | 0.004 | 265.9 | 0.14424 | 332 | 0.0022936 |  |  |  |  | 0
424.192 | 0.0028333 | 266 | 0.13066 | 332.2 | 0.002284 |  |  |  |  | 0
424.511 | 0.0025 | 266.1 | 0.11707 | 332.4 | 0.0025525 |  |  |  |  | 0
424.83 | 0.0038333 | 266.2 | 0.10349 | 332.6 | 0.0036276 |  |  |  |  | 0
425.15 | 0.0016667 | 266.3 | 0.10514 | 332.8 | 0.006415 |  |  |  |  | 0
425.469 | 0.003 | 266.4 | 0.099807 | 333 | 0.0098896 |  |  |  |  | 0
425.788 | 0.0035 | 266.5 | 0.09447 | 333.2 | 0.019315 |  |  |  |  | 0
426.107 | 0.005 | 266.6 | 0.10048 | 333.4 | 0.032143 |  |  |  |  | 0
426.426 | 0.0031667 | 266.7 | 0.099732 | 333.6 | 0.048777 |  |  |  |  | 0
426.745 | 0.0038333 | 266.8 | 0.098983 | 333.8 | 0.058489 |  |  |  |  | 0
427.064 | 0.0043333 | 266.9 | 0.098233 | 334 | 0.057183 |  |  |  |  | 0
427.383 | 0.006 | 267 | 0.099746 | 334.2 | 0.044406 |  |  |  |  | 0
427.702 | 0.0028333 | 267.1 | 0.10164 | 334.4 | 0.026849 |  |  |  |  | 0
428.02 | 0.006 | 267.2 | 0.10354 | 334.6 | 0.014919 |  |  |  |  | 0
428.339 | 0.0045 | 267.3 | 0.10543 | 334.8 | 0.0040882 |  |  |  |  | 0
428.657 | 0.0058333 | 267.4 | 0.099745 | 335 | 0.0031907 |  |  |  |  | 0
428.976 | 0.0036667 | 267.5 | 0.098692 | 335.2 | 0.001397 |  |  |  |  | 0
429.294 | 0.0023333 | 267.6 | 0.097639 | 335.4 | 0.00085477 |  |  |  |  | 0
429.612 | 0.0045 | 267.7 | 0.093072 | 335.6 | 0.0017954 |  |  |  |  | 0
429.93 | 0.0038333 | 267.8 | 0.090333 | 335.8 | 0.002682 |  |  |  |  | 0
430.248 | 0.0048333 | 267.9 | 0.087594 | 336 | 0.0030007 |  |  |  |  | 0
430.567 | 0.006 | 268 | 0.084856 | 336.2 | 0.0026405 |  |  |  |  | 0
430.884 | 0.0056667 | 268.1 | 0.091551 | 336.4 | 0.0016026 |  |  |  |  | 0
431.202 | 0.003 | 268.2 | 0.09148 | 336.6 | 0.0010413 |  |  |  |  | 0
431.52 | 0.0048333 | 268.3 | 0.09141 | 336.8 | 0.00084318 |  |  |  |  | 0
431.838 | 0.0056667 | 268.4 | 0.088559 | 337 | 0.0015232 |  |  |  |  | 0
432.156 | 0.0056667 | 268.5 | 0.088935 | 337.2 | 0.0018972 |  |  |  |  | 0
432.473 | 0.0058333 | 268.6 | 0.089311 | 337.4 | 0.002253 |  |  |  |  | 0
432.791 | 0.0085 | 268.7 | 0.089688 | 337.6 | 0.0022875 |  |  |  |  | 0
433.108 | 0.0076667 | 268.8 | 0.088023 | 337.8 | 0.0023141 |  |  |  |  | 0
433.426 | 0.014333 | 268.9 | 0.086641 | 338 | 0.0028969 |  |  |  |  | 0
433.743 | 0.021167 | 269 | 0.085258 | 338.2 | 0.0029295 |  |  |  |  | 0
434.06 | 0.034 | 269.1 | 0.086056 | 338.4 | 0.0023568 |  |  |  |  | 0
434.377 | 0.056333 | 269.2 | 0.086055 | 338.6 | 0.002666 |  |  |  |  | 0
434.694 | 0.17883 | 269.3 | 0.086055 | 338.8 | 0.0027713 |  |  |  |  | 0
435.011 | 0.63 | 269.4 | 0.086055 | 339 | 0.0036179 |  |  |  |  | 0
435.328 | 1.2393 | 269.5 | 0.087187 | 339.2 | 0.0039133 |  |  |  |  | 0
435.645 | 2.1515 | 269.6 | 0.087656 | 339.4 | 0.0021827 |  |  |  |  | 0
435.962 | 2.3445 | 269.7 | 0.088125 | 339.6 | 0.0025849 |  |  |  |  | 0
436.279 | 0.74117 | 269.8 | 0.085671 | 339.8 | 0.0027853 |  |  |  |  | 0
436.595 | 0.080833 | 269.9 | 0.085063 | 340 | 0.0024427 |  |  |  |  | 0
436.912 | 0.020333 | 270 | 0.084454 | 340.2 | 0.00265 |  |  |  |  | 0
437.228 | 0.012167 | 270.1 | 0.083845 | 340.4 | 0.0026445 |  |  |  |  | 0
437.545 | 0.010667 | 270.2 | 0.085468 | 340.6 | 0.0028173 |  |  |  |  | 0
437.861 | 0.010167 | 270.3 | 0.085092 | 340.8 | 0.0034322 |  |  |  |  | 0
438.177 | 0.010833 | 270.4 | 0.084716 | 341 | 0.0035768 |  |  |  |  | 0
438.493 | 0.010167 | 270.5 | 0.084341 | 341.2 | 0.0036202 |  |  |  |  | 0
438.81 | 0.010667 | 270.6 | 0.088731 | 341.4 | 0.0036358 |  |  |  |  | 0
439.126 | 0.012833 | 270.7 | 0.090327 | 341.6 | 0.0035845 |  |  |  |  | 0
439.442 | 0.012833 | 270.8 | 0.091923 | 341.8 | 0.0035556 |  |  |  |  | 0
439.757 | 0.012833 | 270.9 | 0.073522 | 342 | 0.0035926 |  |  |  |  | 0
440.073 | 0.012333 | 271 | 0.068146 | 342.2 | 0.0034574 |  |  |  |  | 0
440.389 | 0.011833 | 271.1 | 0.062771 | 342.4 | 0.0029776 |  |  |  |  | 0
440.705 | 0.012667 | 271.2 | 0.057395 | 342.6 | 0.0021799 |  |  |  |  | 0
441.02 | 0.0115 | 271.3 | 0.053301 | 342.8 | 0.00038818 |  |  |  |  | 0
441.336 | 0.011 | 271.4 | 0.045625 | 343 | 0.00061891 |  |  |  |  | 0
441.651 | 0.0091667 | 271.5 | 0.037949 | 343.2 | 0.001055 |  |  |  |  | 0
441.967 | 0.0096667 | 271.6 | 0.037855 | 343.4 | 0.0023106 |  |  |  |  | 0
442.282 | 0.010167 | 271.7 | 0.035226 | 343.6 | 0.0034485 |  |  |  |  | 0
442.597 | 0.0078333 | 271.8 | 0.032596 | 343.8 | 0.0027949 |  |  |  |  | 0
442.912 | 0.011833 | 271.9 | 0.029967 | 344 | 0.0027634 |  |  |  |  | 0
443.227 | 0.011 | 272 | 0.033525 | 344.2 | 0.0023532 |  |  |  |  | 0
443.543 | 0.010167 | 272.1 | 0.032844 | 344.4 | 0.0018464 |  |  |  |  | 0
443.857 | 0.012333 | 272.2 | 0.032163 | 344.6 | 0.0016895 |  |  |  |  | 0
444.172 | 0.012667 | 272.3 | 0.032718 | 344.8 | 0.0017839 |  |  |  |  | 0
444.487 | 0.011 | 272.4 | 0.033422 | 345 | 0.0016248 |  |  |  |  | 0
444.802 | 0.010667 | 272.5 | 0.034126 | 345.2 | 0.0016561 |  |  |  |  | 0
445.116 | 0.010167 | 272.6 | 0.03483 | 345.4 | 0.0021834 |  |  |  |  | 0
445.431 | 0.009 | 272.7 | 0.03311 | 345.6 | 0.0025401 |  |  |  |  | 0
445.746 | 0.0076667 | 272.8 | 0.032781 | 345.8 | 0.0025782 |  |  |  |  | 0
446.06 | 0.0055 | 272.9 | 0.032453 | 346 | 0.0027704 |  |  |  |  | 0
446.374 | 0.0063333 | 273 | 0.031038 | 346.2 | 0.0026367 |  |  |  |  | 0
446.689 | 0.0038333 | 273.1 | 0.030074 | 346.4 | 0.0028276 |  |  |  |  | 0
447.003 | 0.0053333 | 273.2 | 0.029111 | 346.6 | 0.0027418 |  |  |  |  | 0
447.317 | 0.0025 | 273.3 | 0.028147 | 346.8 | 0.0017127 |  |  |  |  | 0
447.631 | 0.0045 | 273.4 | 0.028822 | 347 | 0.0013999 |  |  |  |  | 0
447.945 | 0.0033333 | 273.5 | 0.028328 | 347.2 | 0.001638 |  |  |  |  | 0
448.259 | 0.004 | 273.6 | 0.027835 | 347.4 | 0.0014297 |  |  |  |  | 0
448.573 | 0.001 | 273.7 | 0.028118 | 347.6 | 0.0015571 |  |  |  |  | 0
448.887 | 0.0025 | 273.8 | 0.028283 | 347.8 | 0.0014027 |  |  |  |  | 0
449.2 | 0.0036667 | 273.9 | 0.028447 | 348 | 0.0014523 |  |  |  |  | 0
449.514 | 0.0031667 | 274 | 0.028612 | 348.2 | 0.0022438 |  |  |  |  | 0
449.827 | 0.0038333 | 274.1 | 0.030151 | 348.4 | 0.0020656 |  |  |  |  | 0
450.141 | 0.0045 | 274.2 | 0.030787 | 348.6 | 0.00189 |  |  |  |  | 0
450.454 | 0.0031667 | 274.3 | 0.031422 | 348.8 | 0.0023283 |  |  |  |  | 0
450.768 | 0.0038333 | 274.4 | 0.032058 | 349 | 0.0024469 |  |  |  |  | 0
451.081 | 0.0033333 | 274.5 | 0.032263 | 349.2 | 0.00232 |  |  |  |  | 0
451.394 | 0.0026667 | 274.6 | 0.032945 | 349.4 | 0.0029959 |  |  |  |  | 0
451.707 | 0.0035 | 274.7 | 0.033626 | 349.6 | 0.003544 |  |  |  |  | 0
452.02 | 0.0031667 | 274.8 | 0.041168 | 349.8 | 0.002915 |  |  |  |  | 0
452.333 | 0.0038333 | 274.9 | 0.044129 | 350 | 0.0023797 |  |  |  |  | 0
452.646 | 0.0028333 | 275 | 0.047089 | 350.2 | 0.0018391 |  |  |  |  | 0
452.959 | 0.0021667 | 275.1 | 0.05005 | 350.4 | 0.0014264 |  |  |  |  | 0
453.272 | 0.0028333 | 275.2 | 0.046229 | 350.6 | 0.0013931 |  |  |  |  | 0
453.584 | 0.0028333 | 275.3 | 0.048089 | 350.8 | 0.0017609 |  |  |  |  | 0
453.897 | 0.0035 | 275.4 | 0.049948 | 351 | 0.001914 |  |  |  |  | 0
454.21 | 0.0021667 | 275.5 | 0.041898 | 351.2 | 0.001456 |  |  |  |  | 0
454.522 | 0.0026667 | 275.6 | 0.038391 | 351.4 | 0.0018904 |  |  |  |  | 0
454.834 | 0.001 | 275.7 | 0.034883 | 351.6 | 0.0033038 |  |  |  |  | 0
455.147 | 0.00066667 | 275.8 | 0.031375 | 351.8 | 0.0032815 |  |  |  |  | 0
455.459 | 0.0013333 | 275.9 | 0.031527 | 352 | 0.0035201 |  |  |  |  | 0
455.771 | 0.0035 | 276 | 0.028608 | 352.2 | 0.0036228 |  |  |  |  | 0
456.083 | 0.0013333 | 276.1 | 0.025689 | 352.4 | 0.0029068 |  |  |  |  | 0
456.395 | 0.0026667 | 276.2 | 0.025923 | 352.6 | 0.0014253 |  |  |  |  | 0
456.707 | 0.002 | 276.3 | 0.025169 | 352.8 | 0.001011 |  |  |  |  | 0
457.019 | 0.0011667 | 276.4 | 0.024416 | 353 | 0.00063558 |  |  |  |  | 0
457.331 | 0.0016667 | 276.5 | 0.023663 | 353.2 | 0.001068 |  |  |  |  | 0
457.643 | 0.0021667 | 276.6 | 0.024298 | 353.4 | 0.0018228 |  |  |  |  | 0
457.954 | 0.003 | 276.7 | 0.024015 | 353.6 | 0.0031094 |  |  |  |  | 0
458.266 | 0.0026667 | 276.8 | 0.023732 | 353.8 | 0.0033514 |  |  |  |  | 0
458.577 | 0.003 | 276.9 | 0.024204 | 354 | 0.0033175 |  |  |  |  | 0
458.889 | 0.0023333 | 277 | 0.024416 | 354.2 | 0.0027597 |  |  |  |  | 0
459.2 | 0.0018333 | 277.1 | 0.024629 | 354.4 | 0.0023205 |  |  |  |  | 0
459.512 | 0.0018333 | 277.2 | 0.024841 | 354.6 | 0.0022855 |  |  |  |  | 0
459.823 | 0.0031667 | 277.3 | 0.022497 | 354.8 | 0.0025913 |  |  |  |  | 0
460.134 | 0.0018333 | 277.4 | 0.021885 | 355 | 0.0030019 |  |  |  |  | 0
460.445 | 0.0031667 | 277.5 | 0.021273 | 355.2 | 0.0030453 |  |  |  |  | 0
460.756 | 0.0043333 | 277.6 | 0.022874 | 355.4 | 0.0033366 |  |  |  |  | 0
461.067 | 0.0023333 | 277.7 | 0.022732 | 355.6 | 0.0037206 |  |  |  |  | 0
461.378 | 0.0021667 | 277.8 | 0.02259 | 355.8 | 0.0031288 |  |  |  |  | 0
461.689 | 0.00083333 | 277.9 | 0.022448 | 356 | 0.0024847 |  |  |  |  | 0
461.999 | 0.0025 | 278 | 0.018561 | 356.2 | 0.0026175 |  |  |  |  | 0
462.31 | 0.0028333 | 278.1 | 0.017688 | 356.4 | 0.00309 |  |  |  |  | 0
462.621 | 0.0023333 | 278.2 | 0.016814 | 356.6 | 0.0046464 |  |  |  |  | 0
462.931 | 0.0011667 | 278.3 | 0.017852 | 356.8 | 0.0048511 |  |  |  |  | 0
463.242 | 0.0018333 | 278.4 | 0.016743 | 357 | 0.0048918 |  |  |  |  | 0
463.552 | 0.0035 | 278.5 | 0.015634 | 357.2 | 0.004535 |  |  |  |  | 0
463.862 | 0.0025 | 278.6 | 0.014525 | 357.4 | 0.003535 |  |  |  |  | 0
464.172 | -0.00033333 | 278.7 | 0.015938 | 357.6 | 0.0021903 |  |  |  |  | 0
464.483 | 0.0023333 | 278.8 | 0.016481 | 357.8 | 0.0022016 |  |  |  |  | 0
464.793 | 0.001 | 278.9 | 0.017024 | 358 | 0.0022819 |  |  |  |  | 0
465.103 | 0.0026667 | 279 | 0.017567 | 358.2 | 0.0024693 |  |  |  |  | 0
465.413 | 0.0013333 | 279.1 | 0.019223 | 358.4 | 0.0032933 |  |  |  |  | 0
465.722 | 0.0018333 | 279.2 | 0.020026 | 358.6 | 0.0044303 |  |  |  |  | 0
466.032 | 0.0031667 | 279.3 | 0.020828 | 358.8 | 0.0045112 |  |  |  |  | 0
466.342 | 0.0023333 | 279.4 | 0.021999 | 359 | 0.0046521 |  |  |  |  | 0
466.652 | 0.0026667 | 279.5 | 0.023132 | 359.2 | 0.0043569 |  |  |  |  | 0
466.961 | 0.0015 | 279.6 | 0.024265 | 359.4 | 0.0037619 |  |  |  |  | 0
467.271 | 0.0015 | 279.7 | 0.025398 | 359.6 | 0.0035399 |  |  |  |  | 0
467.58 | 0.002 | 279.8 | 0.031628 | 359.8 | 0.0036034 |  |  |  |  | 0
467.889 | 0.0026667 | 279.9 | 0.034272 | 360 | 0.0037723 |  |  |  |  | 0
468.199 | 0.0015 | 280 | 0.036916 | 360.2 | 0.0040517 |  |  |  |  | 0
468.508 | 0.002 | 280.1 | 0.037809 | 360.4 | 0.0040177 |  |  |  |  | 0
468.817 | 0.0015 | 280.2 | 0.040524 | 360.6 | 0.0030719 |  |  |  |  | 0
469.126 | 0.0018333 | 280.3 | 0.043239 | 360.8 | 0.0028541 |  |  |  |  | 0
469.435 | 0.0023333 | 280.4 | 0.045954 | 361 | 0.0024842 |  |  |  |  | 0
469.744 | 0.0015 | 280.5 | 0.034271 | 361.2 | 0.0034873 |  |  |  |  | 0
470.053 | 0.00066667 | 280.6 | 0.03191 | 361.4 | 0.0046001 |  |  |  |  | 0
470.362 | 0.0018333 | 280.7 | 0.02955 | 361.6 | 0.0047525 |  |  |  |  | 0
470.67 | 0.0021667 | 280.8 | 0.029281 | 361.8 | 0.0054249 |  |  |  |  | 0
470.979 | 0.0013333 | 280.9 | 0.026164 | 362 | 0.0057611 |  |  |  |  | 0
471.287 | 0.00066667 | 281 | 0.023048 | 362.2 | 0.0056644 |  |  |  |  | 0
471.596 | 0.0013333 | 281.1 | 0.019932 | 362.4 | 0.0054517 |  |  |  |  | 0
471.904 | 0.00033333 | 281.2 | 0.023612 | 362.6 | 0.0046482 |  |  |  |  | 0
472.213 | 0.0026667 | 281.3 | 0.023682 | 362.8 | 0.0054775 |  |  |  |  | 0
472.521 | 0.0018333 | 281.4 | 0.023753 | 363 | 0.0067548 |  |  |  |  | 0
472.829 | 0.003 | 281.5 | 0.022555 | 363.2 | 0.0085867 |  |  |  |  | 0
473.137 | 0.0031667 | 281.6 | 0.022343 | 363.4 | 0.013064 |  |  |  |  | 0
473.445 | 0.0036667 | 281.7 | 0.02213 | 363.6 | 0.021789 |  |  |  |  | 0
473.753 | 0.0013333 | 281.8 | 0.021918 | 363.8 | 0.061869 |  |  |  |  | 0
474.061 | 0.0005 | 281.9 | 0.017822 | 364 | 0.11393 |  |  |  |  | 0
474.369 | 0.0018333 | 282 | 0.016072 | 364.2 | 0.25676 |  |  |  |  | 0
474.677 | 0.0021667 | 282.1 | 0.014323 | 364.4 | 0.42239 |  |  |  |  | 0
474.985 | 0.0015 | 282.2 | 0.016182 | 364.6 | 0.5756 |  |  |  |  | 0
475.292 | 0.00066667 | 282.3 | 0.015472 | 364.8 | 0.62186 |  |  |  |  | 0
475.6 | 0.0026667 | 282.4 | 0.014762 | 365 | 0.61851 |  |  |  |  | 0
475.907 | 0.0018333 | 282.5 | 0.014052 | 365.2 | 0.49354 |  |  |  |  | 0
476.215 | 0 | 282.6 | 0.012041 | 365.4 | 0.33597 |  |  |  |  | 0
476.522 | 0.0023333 | 282.7 | 0.011614 | 365.6 | 0.23569 |  |  |  |  | 0
476.829 | 0.003 | 282.8 | 0.011188 | 365.8 | 0.16779 |  |  |  |  | 0
477.137 | 0.0015 | 282.9 | 0.014023 | 366 | 0.13934 |  |  |  |  | 0
477.444 | 0.0011667 | 283 | 0.014164 | 366.2 | 0.11077 |  |  |  |  | 0
477.751 | -0.0005 | 283.1 | 0.014305 | 366.4 | 0.085405 |  |  |  |  | 0
478.058 | 0.0023333 | 283.2 | 0.014446 | 366.6 | 0.051784 |  |  |  |  | 0
478.365 | 0.001 | 283.3 | 0.013576 | 366.8 | 0.030967 |  |  |  |  | 0
478.671 | 0.0025 | 283.4 | 0.014215 | 367 | 0.013514 |  |  |  |  | 0
478.978 | 0.0025 | 283.5 | 0.014854 | 367.2 | 0.0068302 |  |  |  |  | 0
479.285 | 0.0031667 | 283.6 | 0.013731 | 367.4 | 0.0065955 |  |  |  |  | 0
479.592 | 0.0011667 | 283.7 | 0.012879 | 367.6 | 0.0060521 |  |  |  |  | 0
479.898 | 0.00083333 | 283.8 | 0.012027 | 367.8 | 0.0063867 |  |  |  |  | 0
480.205 | 0.0016667 | 283.9 | 0.011174 | 368 | 0.0072811 |  |  |  |  | 0
480.511 | 0.0011667 | 284 | 0.013085 | 368.2 | 0.0065887 |  |  |  |  | 0
480.817 | 0.0018333 | 284.1 | 0.013227 | 368.4 | 0.0060493 |  |  |  |  | 0
481.124 | 0.002 | 284.2 | 0.013369 | 368.6 | 0.0059633 |  |  |  |  | 0
481.43 | 0.002 | 284.3 | 0.01119 | 368.8 | 0.0052654 |  |  |  |  | 0
481.736 | 0.00066667 | 284.4 | 0.010788 | 369 | 0.0049542 |  |  |  |  | 0
482.042 | 0.0016667 | 284.5 | 0.010385 | 369.2 | 0.0050387 |  |  |  |  | 0
482.348 | 0.0016667 | 284.6 | 0.0099826 | 369.4 | 0.004951 |  |  |  |  | 0
482.654 | 0.0025 | 284.7 | 0.01115 | 369.6 | 0.0049425 |  |  |  |  | 0
482.96 | 0.002 | 284.8 | 0.010724 | 369.8 | 0.0062099 |  |  |  |  | 0
483.266 | 0.0015 | 284.9 | 0.010298 | 370 | 0.0076711 |  |  |  |  | 0
483.571 | 0.0043333 | 285 | 0.011142 | 370.2 | 0.0070527 |  |  |  |  | 0
483.877 | 0.0036667 | 285.1 | 0.011924 | 370.4 | 0.00643 |  |  |  |  | 0
484.182 | 0.00016667 | 285.2 | 0.012705 | 370.6 | 0.0057011 |  |  |  |  | 0
484.488 | 0.0035 | 285.3 | 0.013486 | 370.8 | 0.0058635 |  |  |  |  | 0
484.793 | 0.00033333 | 285.4 | 0.013715 | 371 | 0.006761 |  |  |  |  | 0
485.099 | 0.0011667 | 285.5 | 0.014236 | 371.2 | 0.0058711 |  |  |  |  | 0
485.404 | 0.00033333 | 285.6 | 0.014757 | 371.4 | 0.0050056 |  |  |  |  | 0
485.709 | 0.00066667 | 285.7 | 0.015278 | 371.6 | 0.0044926 |  |  |  |  | 0
486.014 | 0.002 | 285.8 | 0.014722 | 371.8 | 0.0045922 |  |  |  |  | 0
486.319 | 0.0015 | 285.9 | 0.014888 | 372 | 0.0053323 |  |  |  |  | 0
486.624 | 0.001 | 286 | 0.015054 | 372.2 | 0.006187 |  |  |  |  | 0
486.929 | 0.00066667 | 286.1 | 0.013512 | 372.4 | 0.0068387 |  |  |  |  | 0
487.234 | 0.001 | 286.2 | 0.012967 | 372.6 | 0.006809 |  |  |  |  | 0
487.539 | 0.0028333 | 286.3 | 0.012421 | 372.8 | 0.0064045 |  |  |  |  | 0
487.844 | 0.0015 | 286.4 | 0.011876 | 373 | 0.0056494 |  |  |  |  | 0
488.148 | 0.002 | 286.5 | 0.010207 | 373.2 | 0.0061922 |  |  |  |  | 0
488.453 | 0.00066667 | 286.6 | 0.0091638 | 373.4 | 0.0067696 |  |  |  |  | 0
488.757 | 0.0011667 | 286.7 | 0.0081207 | 373.6 | 0.0071045 |  |  |  |  | 0
489.062 | 0.0016667 | 286.8 | 0.010521 | 373.8 | 0.0069371 |  |  |  |  | 0
489.366 | 0 | 286.9 | 0.010687 | 374 | 0.0063909 |  |  |  |  | 0
489.67 | 0.0031667 | 287 | 0.010854 | 374.2 | 0.0059679 |  |  |  |  | 0
489.975 | 0.0038333 | 287.1 | 0.01102 | 374.4 | 0.0053298 |  |  |  |  | 0
490.279 | 0.0011667 | 287.2 | 0.0098685 | 374.6 | 0.0044564 |  |  |  |  | 0
490.583 | 0.0033333 | 287.3 | 0.010082 | 374.8 | 0.0040555 |  |  |  |  | 0
490.887 | 0.005 | 287.4 | 0.010296 | 375 | 0.0038403 |  |  |  |  | 0
491.191 | 0.0036667 | 287.5 | 0.010873 | 375.2 | 0.0041631 |  |  |  |  | 0
491.495 | 0.0095 | 287.6 | 0.010707 | 375.4 | 0.0047095 |  |  |  |  | 0
491.798 | 0.011167 | 287.7 | 0.010541 | 375.6 | 0.0058021 |  |  |  |  | 0
492.102 | 0.0056667 | 287.8 | 0.010374 | 375.8 | 0.0055344 |  |  |  |  | 0
492.406 | 0.0018333 | 287.9 | 0.01775 | 376 | 0.00524 |  |  |  |  | 0
492.709 | 0.001 | 288 | 0.020195 | 376.2 | 0.0052962 |  |  |  |  | 0
493.013 | 0.00033333 | 288.1 | 0.022641 | 376.4 | 0.0048516 |  |  |  |  | 0
493.316 | 0.0013333 | 288.2 | 0.023198 | 376.6 | 0.0044158 |  |  |  |  | 0
493.62 | 0.0018333 | 288.3 | 0.025715 | 376.8 | 0.0047979 |  |  |  |  | 0
493.923 | 0.00033333 | 288.4 | 0.028232 | 377 | 0.0050141 |  |  |  |  | 0
494.226 | 0.0026667 | 288.5 | 0.030748 | 377.2 | 0.0051399 |  |  |  |  | 0
494.529 | 0.00033333 | 288.6 | 0.037265 | 377.4 | 0.0055938 |  |  |  |  | 0
494.832 | 0.0011667 | 288.7 | 0.040399 | 377.6 | 0.0060222 |  |  |  |  | 0
495.135 | 0.0016667 | 288.8 | 0.043533 | 377.8 | 0.0056782 |  |  |  |  | 0
495.438 | 0.0028333 | 288.9 | 0.051657 | 378 | 0.005236 |  |  |  |  | 0
495.741 | 0.002 | 289 | 0.057617 | 378.2 | 0.0045182 |  |  |  |  | 0
496.044 | 0.00016667 | 289.1 | 0.063576 | 378.4 | 0.0041875 |  |  |  |  | 0
496.347 | 0.0018333 | 289.2 | 0.069535 | 378.6 | 0.0040727 |  |  |  |  | 0
496.649 | 0.0013333 | 289.3 | 0.052358 | 378.8 | 0.0043577 |  |  |  |  | 0
496.952 | 0.002 | 289.4 | 0.051314 | 379 | 0.0045003 |  |  |  |  | 0
497.255 | 0.003 | 289.5 | 0.050269 | 379.2 | 0.0038944 |  |  |  |  | 0
497.557 | 0.0023333 | 289.6 | 0.037542 | 379.4 | 0.003983 |  |  |  |  | 0
497.859 | 0.0018333 | 289.7 | 0.028948 | 379.6 | 0.0039515 |  |  |  |  | 0
498.162 | 0.0013333 | 289.8 | 0.020353 | 379.8 | 0.0041887 |  |  |  |  | 0
498.464 | 0.0013333 | 289.9 | 0.011759 | 380 | 0.0047221 |  |  |  |  | 0
498.766 | 0.001 | 290 | 0.017026 | 380.2 | 0.005198 |  |  |  |  | 0
499.068 | 0.0016667 | 290.1 | 0.012744 | 380.4 | 0.0049852 |  |  |  |  | 0
499.37 | 0.00083333 | 290.2 | 0.0084624 | 380.6 | 0.0046643 |  |  |  |  | 0
499.672 | 0.0005 | 290.3 | 0.0098426 | 380.8 | 0.0048517 |  |  |  |  | 0
499.974 | 0.00066667 | 290.4 | 0.0095571 | 381 | 0.0050704 |  |  |  |  | 0
500.276 | 0.00033333 | 290.5 | 0.0092716 | 381.2 | 0.0056239 |  |  |  |  | 0
500.578 | 0.00066667 | 290.6 | 0.0089862 | 381.4 | 0.005825 |  |  |  |  | 0
500.879 | 0.0011667 | 290.7 | 0.010434 | 381.6 | 0.005662 |  |  |  |  | 0
501.181 | 0.003 | 290.8 | 0.010553 | 381.8 | 0.0049582 |  |  |  |  | 0
501.483 | 0.0011667 | 290.9 | 0.010672 | 382 | 0.0046182 |  |  |  |  | 0
501.784 | 0.002 | 291 | 0.0091097 | 382.2 | 0.0047215 |  |  |  |  | 0
502.085 | 0.0005 | 291.1 | 0.0089438 | 382.4 | 0.0045274 |  |  |  |  | 0
502.387 | 0.00083333 | 291.2 | 0.0087779 | 382.6 | 0.0047221 |  |  |  |  | 0
502.688 | 0.0018333 | 291.3 | 0.0086119 | 382.8 | 0.0052045 |  |  |  |  | 0
502.989 | 0.003 | 291.4 | 0.010902 | 383 | 0.0053113 |  |  |  |  | 0
503.29 | 0.0016667 | 291.5 | 0.011117 | 383.2 | 0.0052464 |  |  |  |  | 0
503.591 | 0.0021667 | 291.6 | 0.011331 | 383.4 | 0.004821 |  |  |  |  | 0
503.892 | 0.00083333 | 291.7 | 0.011375 | 383.6 | 0.0039933 |  |  |  |  | 0
504.193 | -0.00016667 | 291.8 | 0.012159 | 383.8 | 0.003016 |  |  |  |  | 0
504.494 | 0.0025 | 291.9 | 0.012943 | 384 | 0.002681 |  |  |  |  | 0
504.795 | 0.0018333 | 292 | 0.013728 | 384.2 | 0.0024756 |  |  |  |  | 0
505.096 | -0.00016667 | 292.1 | 0.015995 | 384.4 | 0.0031725 |  |  |  |  | 0
505.396 | 0.0026667 | 292.2 | 0.016921 | 384.6 | 0.0041434 |  |  |  |  | 0
505.697 | 0.0023333 | 292.3 | 0.017848 | 384.8 | 0.0048058 |  |  |  |  | 0
505.997 | 0.0015 | 292.4 | 0.016062 | 385 | 0.0049963 |  |  |  |  | 0
506.298 | 0.002 | 292.5 | 0.016419 | 385.2 | 0.0048891 |  |  |  |  | 0
506.598 | 0.0035 | 292.6 | 0.016777 | 385.4 | 0.0047364 |  |  |  |  | 0
506.898 | 0.0046667 | 292.7 | 0.017134 | 385.6 | 0.0044419 |  |  |  |  | 0
507.199 | 0.0013333 | 292.8 | 0.012272 | 385.8 | 0.0041326 |  |  |  |  | 0
507.499 | 0.0013333 | 292.9 | 0.010391 | 386 | 0.0042883 |  |  |  |  | 0
507.799 | 0.00016667 | 293 | 0.0085102 | 386.2 | 0.0046557 |  |  |  |  | 0
508.099 | 0.0018333 | 293.1 | 0.010613 | 386.4 | 0.0053869 |  |  |  |  | 0
508.399 | 0.0016667 | 293.2 | 0.0099467 | 386.6 | 0.0058238 |  |  |  |  | 0
508.699 | 0.00083333 | 293.3 | 0.00928 | 386.8 | 0.0056322 |  |  |  |  | 0
508.998 | 0.0011667 | 293.4 | 0.0086133 | 387 | 0.0054247 |  |  |  |  | 0
509.298 | 0.00083333 | 293.5 | 0.0091693 | 387.2 | 0.0050517 |  |  |  |  | 0
509.598 | 0.0015 | 293.6 | 0.009574 | 387.4 | 0.0049961 |  |  |  |  | 0
509.897 | 0.00083333 | 293.7 | 0.0099788 | 387.6 | 0.0048574 |  |  |  |  | 0
510.197 | 0.0011667 | 293.8 | 0.0094602 | 387.8 | 0.0046817 |  |  |  |  | 0
510.496 | 0.0015 | 293.9 | 0.009174 | 388 | 0.0046542 |  |  |  |  | 0
510.796 | 0.0041667 | 294 | 0.0088878 | 388.2 | 0.0046165 |  |  |  |  | 0
511.095 | 0.0021667 | 294.1 | 0.0086017 | 388.4 | 0.0046449 |  |  |  |  | 0
511.394 | 0.00016667 | 294.2 | 0.0082805 | 388.6 | 0.0047143 |  |  |  |  | 0
511.693 | 0.00066667 | 294.3 | 0.0080182 | 388.8 | 0.0047744 |  |  |  |  | 0
511.992 | 0.00016667 | 294.4 | 0.0077559 | 389 | 0.0052764 |  |  |  |  | 0
512.291 | 0.0025 | 294.5 | 0.0087233 | 389.2 | 0.0057477 |  |  |  |  | 0
512.59 | 0.0021667 | 294.6 | 0.00889 | 389.4 | 0.0059334 |  |  |  |  | 0
512.889 | 0.0011667 | 294.7 | 0.0090567 | 389.6 | 0.0064563 |  |  |  |  | 0
513.188 | 0.0028333 | 294.8 | 0.0092233 | 389.8 | 0.0068958 |  |  |  |  | 0
513.487 | 0.0021667 | 294.9 | 0.014255 | 390 | 0.0077626 |  |  |  |  | 0
513.786 | 0.00016667 | 295 | 0.016138 | 390.2 | 0.0087516 |  |  |  |  | 0
514.084 | 0.00066667 | 295.1 | 0.018021 | 390.4 | 0.0095572 |  |  |  |  | 0
514.383 | 0.0015 | 295.2 | 0.026847 | 390.6 | 0.008281 |  |  |  |  | 0
514.681 | 0.0011667 | 295.3 | 0.031855 | 390.8 | 0.0067536 |  |  |  |  | 0
514.98 | 0.00066667 | 295.4 | 0.036863 | 391 | 0.0059427 |  |  |  |  | 0
515.278 | 0.003 | 295.5 | 0.041871 | 391.2 | 0.0051215 |  |  |  |  | 0
515.576 | -0.00016667 | 295.6 | 0.050203 | 391.4 | 0.0059253 |  |  |  |  | 0
515.874 | 0.0013333 | 295.7 | 0.05695 | 391.6 | 0.0057185 |  |  |  |  | 0
516.173 | 0.0041667 | 295.8 | 0.063698 | 391.8 | 0.0054482 |  |  |  |  | 0
516.471 | -0.0005 | 295.9 | 0.084383 | 392 | 0.0052663 |  |  |  |  | 0
516.769 | 0.00033333 | 296 | 0.096251 | 392.2 | 0.0052515 |  |  |  |  | 0
517.067 | 0.0021667 | 296.1 | 0.10812 | 392.4 | 0.0054863 |  |  |  |  | 0
517.364 | 0.0013333 | 296.2 | 0.11999 | 392.6 | 0.0053007 |  |  |  |  | 0
517.662 | 0.0016667 | 296.3 | 0.15823 | 392.8 | 0.0054744 |  |  |  |  | 0
517.96 | 0.0028333 | 296.4 | 0.18027 | 393 | 0.0065173 |  |  |  |  | 0
518.257 | 0.0025 | 296.5 | 0.20231 | 393.2 | 0.0058286 |  |  |  |  | 0
518.555 | 0.003 | 296.6 | 0.15428 | 393.4 | 0.0051019 |  |  |  |  | 0
518.853 | 0 | 296.7 | 0.15254 | 393.6 | 0.0047907 |  |  |  |  | 0
519.15 | 0.0016667 | 296.8 | 0.1508 | 393.8 | 0.0046219 |  |  |  |  | 0
519.447 | 0.00083333 | 296.9 | 0.14905 | 394 | 0.0057962 |  |  |  |  | 0
519.745 | 0.0016667 | 297 | 0.093245 | 394.2 | 0.0051548 |  |  |  |  | 0
520.042 | 0.0026667 | 297.1 | 0.064324 | 394.4 | 0.0048787 |  |  |  |  | 0
520.339 | 0.0028333 | 297.2 | 0.035402 | 394.6 | 0.0060143 |  |  |  |  | 0
520.636 | 0.0011667 | 297.3 | 0.042203 | 394.8 | 0.0060345 |  |  |  |  | 0
520.933 | -0.00083333 | 297.4 | 0.027029 | 395 | 0.0059204 |  |  |  |  | 0
521.23 | 0.00033333 | 297.5 | 0.011855 | 395.2 | 0.0057818 |  |  |  |  | 0
521.527 | 0.0038333 | 297.6 | -0.0033192 | 395.4 | 0.0051071 |  |  |  |  | 0
521.824 | 0.0016667 | 297.7 | 0.011434 | 395.6 | 0.0048934 |  |  |  |  | 0
522.121 | 0.0018333 | 297.8 | 0.0099523 | 395.8 | 0.0047 |  |  |  |  | 0
522.417 | -0.00016667 | 297.9 | 0.0084704 | 396 | 0.00445 |  |  |  |  | 0
522.714 | 0.00016667 | 298 | 0.0087786 | 396.2 | 0.0041758 |  |  |  |  | 0
523.01 | 0.0021667 | 298.1 | 0.0085159 | 396.4 | 0.0042522 |  |  |  |  | 0
523.307 | -0.00066667 | 298.2 | 0.0082533 | 396.6 | 0.0050371 |  |  |  |  | 0
523.603 | 0.0011667 | 298.3 | 0.0079906 | 396.8 | 0.0054255 |  |  |  |  | 0
523.9 | 0.00066667 | 298.4 | 0.0056895 | 397 | 0.0057964 |  |  |  |  | 0
524.196 | 0.0021667 | 298.5 | 0.0047583 | 397.2 | 0.0056386 |  |  |  |  | 0
524.492 | 0.002 | 298.6 | 0.003827 | 397.4 | 0.0056339 |  |  |  |  | 0
524.788 | 0 | 298.7 | 0.004457 | 397.6 | 0.0058801 |  |  |  |  | 0
525.084 | 0.00066667 | 298.8 | 0.0038588 | 397.8 | 0.0060417 |  |  |  |  | 0
525.38 | 0.0015 | 298.9 | 0.0032606 | 398 | 0.0060797 |  |  |  |  | 0
525.676 | 0.0021667 | 299 | 0.0026623 | 398.2 | 0.0055078 |  |  |  |  | 0
525.972 | 0.0016667 | 299.1 | 0.0047256 | 398.4 | 0.0057499 |  |  |  |  | 0
526.268 | 0.0013333 | 299.2 | 0.0052518 | 398.6 | 0.0060496 |  |  |  |  | 0
526.564 | 0.003 | 299.3 | 0.005778 | 398.8 | 0.0058414 |  |  |  |  | 0
526.859 | -0.00016667 | 299.4 | 0.005425 | 399 | 0.0059566 |  |  |  |  | 0
527.155 | 0.0011667 | 299.5 | 0.0057361 | 399.2 | 0.006318 |  |  |  |  | 0
527.45 | 0.0011667 | 299.6 | 0.0060472 | 399.4 | 0.0060544 |  |  |  |  | 0
527.746 | 0.0015 | 299.7 | 0.0063582 | 399.6 | 0.0059729 |  |  |  |  | 0
528.041 | 0.0011667 | 299.8 | 0.0052168 | 399.8 | 0.005894 |  |  |  |  | 0
528.336 | 0.0021667 | 299.9 | 0.0049298 | 400 | 0.0055191 |  |  |  |  | 0
528.632 | 0.0013333 | 300 | 0.0046428 | 400.2 | 0.0051026 |  |  |  |  | 0
528.927 | 0.0016667 | 300.1 | 0.0044546 | 400.4 | 0.0051778 |  |  |  |  | 0
529.222 | 0.0035 | 300.2 | 0.0041437 | 400.6 | 0.0055131 |  |  |  |  | 0
529.517 | 0.001 | 300.3 | 0.0038328 | 400.8 | 0.0065444 |  |  |  |  | 0
529.812 | 0.0018333 | 300.4 | 0.0083697 | 401 | 0.0062977 |  |  |  |  | 0
530.107 | 0.0015 | 300.5 | 0.01019 | 401.2 | 0.0057897 |  |  |  |  | 0
530.402 | 0.0025 | 300.6 | 0.01201 | 401.4 | 0.0061402 |  |  |  |  | 0
530.696 | 0.002 | 300.7 | 0.01383 | 401.6 | 0.0061908 |  |  |  |  | 0
530.991 | 0.0025 | 300.8 | 0.021926 | 401.8 | 0.0062241 |  |  |  |  | 0
531.286 | -0.00016667 | 300.9 | 0.026757 | 402 | 0.0066809 |  |  |  |  | 0
531.58 | -0.00083333 | 301 | 0.031587 | 402.2 | 0.0072243 |  |  |  |  | 0
531.875 | 0.0021667 | 301.1 | 0.037133 | 402.4 | 0.0085987 |  |  |  |  | 0
532.169 | 0.001 | 301.2 | 0.042752 | 402.6 | 0.010198 |  |  |  |  | 0
532.463 | 0.0005 | 301.3 | 0.048372 | 402.8 | 0.012422 |  |  |  |  | 0
532.758 | 0.004 | 301.4 | 0.053991 | 403 | 0.016564 |  |  |  |  | 0
533.052 | 0.00066667 | 301.5 | 0.074684 | 403.2 | 0.030368 |  |  |  |  | 0
533.346 | 0.0016667 | 301.6 | 0.085436 | 403.4 | 0.064317 |  |  |  |  | 0
533.64 | 0.0018333 | 301.7 | 0.096188 | 403.6 | 0.15288 |  |  |  |  | 0
533.934 | 0.0031667 | 301.8 | 0.10211 | 403.8 | 0.26731 |  |  |  |  | 0
534.228 | 0.0011667 | 301.9 | 0.11343 | 404 | 0.47272 |  |  |  |  | 0
534.522 | 0.0011667 | 302 | 0.12475 | 404.2 | 0.69576 |  |  |  |  | 0
534.816 | 0.002 | 302.1 | 0.13608 | 404.4 | 0.88322 |  |  |  |  | 0
535.11 | 0.0011667 | 302.2 | 0.093972 | 404.6 | 0.84128 |  |  |  |  | 0
535.403 | -0.0013333 | 302.3 | 0.085188 | 404.8 | 0.67668 |  |  |  |  | 0
535.697 | 0.0025 | 302.4 | 0.076403 | 405 | 0.40822 |  |  |  |  | 0
535.99 | 0.00066667 | 302.5 | 0.065061 | 405.2 | 0.18076 |  |  |  |  | 0
536.284 | 0.0016667 | 302.6 | 0.048035 | 405.4 | 0.067328 |  |  |  |  | 0
536.577 | 0.0015 | 302.7 | 0.031009 | 405.6 | 0.010496 |  |  |  |  | 0
536.871 | 0.0016667 | 302.8 | 0.013983 | 405.8 | 0.012048 |  |  |  |  | 0
537.164 | 0.0005 | 302.9 | 0.019646 | 406 | 0.0099453 |  |  |  |  | 0
537.457 | -0.0005 | 303 | 0.012965 | [6 empty] | 0
537.75 | 0.00083333 | 303.1 | 0.006284 | [6 empty] | 0
538.043 | 0.0026667 | 303.2 | 0.0084275 | [6 empty] | 0
538.336 | 0.0016667 | 303.3 | 0.00682 | [6 empty] | 0
538.629 | 0.002 | 303.4 | 0.0052124 | [6 empty] | 0
538.922 | 0.0015 | 303.5 | 0.0036048 | [6 empty] | 0
539.215 | -0.0005 | 303.6 | 0.0045432 | [6 empty] | 0
539.507 | 0.0035 | 303.7 | 0.0044472 | [6 empty] | 0
539.8 | 0.0021667 | 303.8 | 0.0043512 | [6 empty] | 0
540.093 | 0.0021667 | 303.9 | 0.0056338 | [6 empty] | 0
540.385 | 0.003 | 304 | 0.0059692 | [6 empty] | 0
540.678 | 0.0045 | 304.1 | 0.0063046 | [6 empty] | 0
540.97 | 0.0025 | 304.2 | 0.00664 | [6 empty] | 0
541.262 | 0.0045 | 304.3 | 0.0063457 | [6 empty] | 0
541.554 | 0.0011667 | 304.4 | 0.0067291 | [6 empty] | 0
541.847 | 0.001 | 304.5 | 0.0071124 | [6 empty] | 0
542.139 | 0.0046667 | 304.6 | 0.006474 | [6 empty] | 0
542.431 | 0.0011667 | 304.7 | 0.0061863 | [6 empty] | 0
542.723 | -0.00033333 | 304.8 | 0.0058985 | [6 empty] | 0
543.015 | 0.003 | 304.9 | 0.0056108 | [6 empty] | 0
543.307 | 0.0061667 | 305 | 0.0067278 | [6 empty] | 0
543.598 | 0.007 | 305.1 | 0.0068719 | [6 empty] | 0
543.89 | 0.0035 | 305.2 | 0.007016 | [6 empty] | 0
544.182 | 0.0055 | 305.3 | 0.005585 | [6 empty] | 0
544.473 | 0.009 | 305.4 | 0.0053449 | [6 empty] | 0
544.765 | 0.020667 | 305.5 | 0.0051047 | [6 empty] | 0
545.056 | 0.076 | 305.6 | 0.0048646 | [6 empty] | 0
545.347 | 0.32817 | 305.7 | 0.0047835 | [6 empty] | 0
545.639 | 0.69467 | 305.8 | 0.0042791 | [6 empty] | 0
545.93 | 0.948 | 305.9 | 0.0037748 | [6 empty] | 0
546.221 | 1.4593 | 306 | 0.0045961 | [6 empty] | 0
546.512 | 0.743 | 306.1 | 0.0048362 | [6 empty] | 0
546.803 | 0.1125 | 306.2 | 0.0050764 | [6 empty] | 0
547.094 | 0.0285 | 306.3 | 0.0053165 | [6 empty] | 0
547.385 | 0.012167 | 306.4 | 0.0055458 | [6 empty] | 0
547.676 | 0.0093333 | 306.5 | 0.005834 | [6 empty] | 0
547.967 | 0.0073333 | 306.6 | 0.0061222 | [6 empty] | 0
548.257 | 0.0065 | 306.7 | 0.0057222 | [6 empty] | 0
548.548 | 0.004 | 306.8 | 0.0057222 | [6 empty] | 0
548.839 | 0.0026667 | 306.9 | 0.0057222 | [6 empty] | 0
549.129 | 0.0038333 | 307 | 0.0057222 | [6 empty] | 0
549.419 | 0.0048333 | 307.1 | 0.0057061 | [6 empty] | 0
549.71 | 0.0055 | 307.2 | 0.0056338 | [6 empty] | 0
550 | 0.0073333 | 307.3 | 0.0055616 | [6 empty] | 0
550.29 | 0.0078333 | 307.4 | 0.005692 | [6 empty] | 0
550.58 | 0.0066667 | 307.5 | 0.0057401 | [6 empty] | 0
550.871 | 0.0075 | 307.6 | 0.0057882 | [6 empty] | 0
551.161 | 0.007 | 307.7 | 0.0058364 | [6 empty] | 0
551.45 | 0.0053333 | 307.8 | 0.0058845 | [6 empty] | 0
551.74 | 0.005 | 307.9 | 0.0059326 | [6 empty] | 0
552.03 | 0.0061667 | 308 | 0.0059807 | [6 empty] | 0
552.32 | 0.0075 | 308.1 | 0.0056092 | [6 empty] | 0
552.61 | 0.0055 | 308.2 | 0.005513 | [6 empty] | 0
552.899 | 0.0046667 | 308.3 | 0.0054167 | [6 empty] | 0
553.189 | 0.0058333 | 308.4 | 0.0056047 | [6 empty] | 0
553.478 | 0.0061667 | 308.5 | 0.0055565 | [6 empty] | 0
553.768 | 0.0081667 | 308.6 | 0.0055083 | [6 empty] | 0
554.057 | 0.0063333 | 308.7 | 0.0054601 | [6 empty] | 0
554.346 | 0.006 | 308.8 | 0.0056197 | [6 empty] | 0
554.635 | 0.007 | 308.9 | 0.0057159 | [6 empty] | 0
554.925 | 0.0083333 | 309 | 0.005812 | [6 empty] | 0
555.214 | 0.0045 | 309.1 | 0.0067498 | [6 empty] | 0
555.503 | 0.0063333 | 309.2 | 0.0071584 | [6 empty] | 0
555.792 | 0.0035 | 309.3 | 0.0075671 | [6 empty] | 0
556.081 | 0.0051667 | 309.4 | 0.0079758 | [6 empty] | 0
556.369 | 0.0016667 | 309.5 | 0.0065755 | [6 empty] | 0
556.658 | 0.0015 | 309.6 | 0.0065273 | [6 empty] | 0
556.947 | 0.0013333 | 309.7 | 0.0064791 | [6 empty] | 0
557.235 | 0.0045 | 309.8 | 0.0057556 | [6 empty] | 0
557.524 | 0.0035 | 309.9 | 0.0050812 | [6 empty] | 0
557.812 | 0.0043333 | 310 | 0.0044069 | [6 empty] | 0
558.101 | 0.0033333 | 310.1 | 0.0037325 | [6 empty] | 0
558.389 | 0.0018333 | 310.2 | 0.0057728 | [6 empty] | 0
558.677 | -0.00016667 | 310.3 | 0.0060377 | [6 empty] | 0
558.966 | 0.0031667 | 310.4 | 0.0063027 | [6 empty] | 0
559.254 | 0.0036667 | 310.5 | 0.0086609 | [6 empty] | 0
559.542 | 0.0046667 | 310.6 | 0.010034 | [6 empty] | 0
559.83 | 0 | 310.7 | 0.011407 | [6 empty] | 0
560.118 | 0.0025 | 310.8 | 0.012779 | [6 empty] | 0
560.406 | 0.0033333 | 310.9 | 0.020918 | [6 empty] | 0
560.693 | 0.0011667 | 311 | 0.024531 | [6 empty] | 0
560.981 | 0.0055 | 311.1 | 0.028144 | [6 empty] | 0
561.269 | 0.0021667 | 311.2 | 0.059083 | [6 empty] | 0
561.556 | 0.0026667 | 311.3 | 0.07343 | [6 empty] | 0
561.844 | 0.0018333 | 311.4 | 0.087776 | [6 empty] | 0
562.131 | 0.0038333 | 311.5 | 0.10212 | [6 empty] | 0
562.419 | 0.004 | 311.6 | 0.16722 | [6 empty] | 0
562.706 | 0.0011667 | 311.7 | 0.2001 | [6 empty] | 0
562.994 | 0.0026667 | 311.8 | 0.23298 | [6 empty] | 0
563.281 | 0.0025 | 311.9 | 0.35321 | [6 empty] | 0
563.568 | 0.0046667 | 312 | 0.42079 | [6 empty] | 0
563.855 | 0.00066667 | 312.1 | 0.48837 | [6 empty] | 0
564.142 | -0.0021667 | 312.2 | 0.54369 | [6 empty] | 0
564.429 | 0.0021667 | 312.3 | 0.6183 | [6 empty] | 0
564.716 | 0.0023333 | 312.4 | 0.69292 | [6 empty] | 0
565.003 | 0.0016667 | 312.5 | 0.76753 | [6 empty] | 0
565.289 | 0.0031667 | 312.6 | 0.72404 | [6 empty] | 0
565.576 | 0.00016667 | 312.7 | 0.75106 | [6 empty] | 0
565.863 | 0.0028333 | 312.8 | 0.77809 | [6 empty] | 0
566.149 | 0.0023333 | 312.9 | 0.67487 | [6 empty] | 0
566.436 | 0.0023333 | 313 | 0.63544 | [6 empty] | 0
566.722 | -0.00033333 | 313.1 | 0.59601 | [6 empty] | 0
567.008 | 0.0016667 | 313.2 | 0.55658 | [6 empty] | 0
567.295 | 0.00083333 | 313.3 | 0.38341 | [6 empty] | 0
567.581 | 0.00083333 | 313.4 | 0.28554 | [6 empty] | 0
567.867 | 0.00066667 | 313.5 | 0.18768 | [6 empty] | 0
568.153 | -0.00016667 | 313.6 | 0.18472 | [6 empty] | 0
568.439 | 0.00066667 | 313.7 | 0.11422 | [6 empty] | 0
568.725 | 0.0021667 | 313.8 | 0.043709 | [6 empty] | 0
569.011 | 0.0016667 | 313.9 | -0.026798 | [6 empty] | 0
569.297 | 0.0016667 | 314 | 0.027825 | [6 empty] | 0
569.583 | 0.0021667 | 314.1 | 0.015676 | [6 empty] | 0
569.868 | -0.00016667 | 314.2 | 0.0035258 | [6 empty] | 0
570.154 | 0.005 | 314.3 | 0.0095391 | [6 empty] | 0
570.439 | 0.001 | 314.4 | 0.0082589 | [6 empty] | 0
570.725 | 0.002 | 314.5 | 0.0069787 | [6 empty] | 0
571.01 | 0.0015 | 314.6 | 0.0056986 | [6 empty] | 0
571.296 | 0.0035 | 314.7 | 0.0076307 | [6 empty] | 0
571.581 | 0.0011667 | 314.8 | 0.0078964 | [6 empty] | 0
571.866 | 0.0005 | 314.9 | 0.0081621 | [6 empty] | 0
572.151 | 0.00066667 | 315 | 0.0070667 | [6 empty] | 0
572.436 | 0.002 | 315.1 | 0.0068493 | [6 empty] | 0
572.721 | 0.0018333 | 315.2 | 0.0066319 | [6 empty] | 0
573.006 | 0.00083333 | 315.3 | 0.006942 | [6 empty] | 0
573.291 | 0.0015 | 315.4 | 0.0067005 | [6 empty] | 0
573.576 | 0.0033333 | 315.5 | 0.0064589 | [6 empty] | 0
573.861 | 0.0021667 | 315.6 | 0.0062174 | [6 empty] | 0
574.146 | -0.00083333 | 315.7 | 0.0047277 | [6 empty] | 0
574.43 | 0.0033333 | 315.8 | 0.0042442 | [6 empty] | 0
574.715 | 0.0031667 | 315.9 | 0.0037608 | [6 empty] | 0
574.999 | 0.0023333 | 316 | 0.0047524 | [6 empty] | 0
575.284 | 0.0048333 | 316.1 | 0.0045352 | [6 empty] | 0
575.568 | 0.003 | 316.2 | 0.004318 | [6 empty] | 0
575.852 | 0.0061667 | 316.3 | 0.0041008 | [6 empty] | 0
576.137 | 0.028 | 316.4 | 0.0046802 | [6 empty] | 0
576.421 | 0.097833 | 316.5 | 0.005115 | [6 empty] | 0
576.705 | 0.14883 | 316.6 | 0.0055498 | [6 empty] | 0
576.989 | 0.22783 | 316.7 | 0.004973 | [6 empty] | 0
577.273 | 0.19867 | 316.8 | 0.0048522 | [6 empty] | 0
577.557 | 0.0435 | 316.9 | 0.0047313 | [6 empty] | 0
577.841 | 0.012667 | 317 | 0.0046105 | [6 empty] | 0
578.124 | 0.016 | 317.1 | 0.0055578 | [6 empty] | 0
578.408 | 0.065 | 317.2 | 0.0057032 | [6 empty] | 0
578.692 | 0.12717 | 317.3 | 0.0058485 | [6 empty] | 0
578.975 | 0.17833 | 317.4 | 0.0051192 | [6 empty] | 0
579.259 | 0.231 | 317.5 | 0.0051431 | [6 empty] | 0
579.542 | 0.082 | 317.6 | 0.005167 | [6 empty] | 0
579.826 | 0.015333 | 317.7 | 0.0051908 | [6 empty] | 0
580.109 | 0.0053333 | 317.8 | 0.0068361 | [6 empty] | 0
580.392 | 0.0026667 | 317.9 | 0.0071501 | [6 empty] | 0
580.676 | 0.0058333 | 318 | 0.0074642 | [6 empty] | 0
580.959 | 0.0036667 | 318.1 | 0.0060883 | [6 empty] | 0
581.242 | 0.0071667 | 318.2 | 0.0061609 | [6 empty] | 0
581.525 | 0.0025 | 318.3 | 0.0062336 | [6 empty] | 0
581.808 | 0.0028333 | 318.4 | 0.0067191 | [6 empty] | 0
582.091 | 0.0021667 | 318.5 | 0.0064042 | [6 empty] | 0
582.373 | 0.0013333 | 318.6 | 0.0060892 | [6 empty] | 0
582.656 | 0.0023333 | 318.7 | 0.0057743 | [6 empty] | 0
582.939 | 0.0023333 | 318.8 | 0.0043208 | [6 empty] | 0
583.221 | 0.0043333 | 318.9 | 0.003909 | [6 empty] | 0
583.504 | 0.0011667 | 319 | 0.0034972 | [6 empty] | 0
583.786 | 0.0026667 | 319.1 | 0.0048794 | [6 empty] | 0
584.069 | 0.0011667 | 319.2 | 0.0048068 | [6 empty] | 0
584.351 | 0.0033333 | 319.3 | 0.0047341 | [6 empty] | 0
584.633 | 0.0031667 | 319.4 | 0.0046614 | [6 empty] | 0
584.916 | 0.006 | 319.5 | 0.0040226 | [6 empty] | 0
585.198 | 0.002 | 319.6 | 0.0042164 | [6 empty] | 0
585.48 | 0.0048333 | 319.7 | 0.0044102 | [6 empty] | 0
585.762 | 0.0048333 | 319.8 | 0.0053517 | [6 empty] | 0
586.044 | 0.0048333 | 319.9 | 0.0054241 | [6 empty] | 0
586.326 | 0.00033333 | 320 | 0.0054965 | [6 empty] | 0
586.607 | 0.0028333 | 320.1 | 0.0055688 | [6 empty] | 0
586.889 | 0.0045 | 320.2 | 0.0044997 | [6 empty] | 0
587.171 | 0.0005 | 320.3 | 0.0044992 | [6 empty] | 0
587.453 | 0.0033333 | 320.4 | 0.0044988 | [6 empty] | 0
587.734 | 0.004 | 320.5 | 0.0050467 | [6 empty] | 0
588.016 | 0.0053333 | 320.6 | 0.0048771 | [6 empty] | 0
588.297 | 0.0041667 | 320.7 | 0.0047075 | [6 empty] | 0
588.578 | 0.0005 | 320.8 | 0.004538 | [6 empty] | 0
588.86 | 0.0031667 | 320.9 | 0.0052094 | [6 empty] | 0
589.141 | 0.001 | 321 | 0.0055491 | [6 empty] | 0
589.422 | 0.0025 | 321.1 | 0.0058888 | [6 empty] | 0
589.703 | 0.0005 | 321.2 | 0.0056357 | [6 empty] | 0
589.984 | 0.0018333 | 321.3 | 0.0056841 | [6 empty] | 0
590.265 | 0.0033333 | 321.4 | 0.0057326 | [6 empty] | 0
590.546 | 0.0015 | 321.5 | 0.0052618 | [6 empty] | 0
590.827 | 0.00083333 | 321.6 | 0.0050435 | [6 empty] | 0
591.108 | -0.00016667 | 321.7 | 0.0048253 | [6 empty] | 0
591.389 | 0.0028333 | 321.8 | 0.004607 | [6 empty] | 0
591.669 | 0.0016667 | 321.9 | 0.0049378 | [6 empty] | 0
591.95 | 0.004 | 322 | 0.0048406 | [6 empty] | 0
592.23 | 0.003 | 322.1 | 0.0047434 | [6 empty] | 0
592.511 | 0.0031667 | 322.2 | 0.0049195 | [6 empty] | 0
592.791 | 0.0015 | 322.3 | 0.0050653 | [6 empty] | 0
593.071 | 0.0018333 | 322.4 | 0.005211 | [6 empty] | 0
593.352 | 0.0048333 | 322.5 | 0.0053568 | [6 empty] | 0
593.632 | 0.00066667 | 322.6 | 0.0057817 | [6 empty] | 0
593.912 | 0.001 | 322.7 | 0.0060003 | [6 empty] | 0
594.192 | 0 | 322.8 | 0.006219 | [6 empty] | 0
594.472 | 0.0031667 | 322.9 | 0.0050816 | [6 empty] | 0
594.752 | 0.0021667 | 323 | 0.0048387 | [6 empty] | 0
595.032 | 0.004 | 323.1 | 0.0045957 | [6 empty] | 0
595.312 | 0.0011667 | 323.2 | 0.0043528 | [6 empty] | 0
595.592 | 0.0016667 | 323.3 | 0.0049323 | [6 empty] | 0
595.871 | 0.0033333 | 323.4 | 0.0047379 | [6 empty] | 0
596.151 | 0.0023333 | 323.5 | 0.0045436 | [6 empty] | 0
596.43 | 0.00016667 | 323.6 | 0.0036944 | [6 empty] | 0
596.71 | 0.0013333 | 323.7 | 0.0035244 | [6 empty] | 0
596.989 | -0.0005 | 323.8 | 0.0033543 | [6 empty] | 0
597.269 | 0.0031667 | 323.9 | 0.0031375 | [6 empty] | 0
597.548 | 0.0036667 | 324 | 0.0026516 | [6 empty] | 0
597.827 | 0.0025 | 324.1 | 0.0021657 | [6 empty] | 0
598.106 | 0.0035 | 324.2 | 0.0016798 | [6 empty] | 0
598.385 | -0.0018333 | 324.3 | 0.0018526 | [6 empty] | 0
598.665 | 0.0025 | 324.4 | 0.0017068 | [6 empty] | 0
598.944 | 0.001 | 324.5 | 0.0015611 | [6 empty] | 0
599.222 | -0.0005 | 324.6 | 0.0026297 | [6 empty] | 0
599.501 | 0.0031667 | 324.7 | 0.0029945 | [6 empty] | 0
599.78 | 0 | 324.8 | 0.0033592 | [6 empty] | 0
600.059 | 0.00066667 | 324.9 | 0.0037239 | [6 empty] | 0
600.337 | -0.00033333 | 325 | 0.0041568 | [6 empty] | 0
600.616 | -0.001 | 325.1 | 0.0047162 | [6 empty] | 0
600.895 | 0.0023333 | 325.2 | 0.0052757 | [6 empty] | 0
601.173 | 0.005 | 325.3 | 0.0039611 | [6 empty] | 0
601.451 | 0.0015 | 325.4 | 0.0036937 | [6 empty] | 0
601.73 | 0.0011667 | 325.5 | 0.0034264 | [6 empty] | 0
602.008 | 0.0011667 | 325.6 | 0.003159 | [6 empty] | 0
602.286 | 0.00083333 | 325.7 | 0.0032176 | [6 empty] | 0
602.564 | 0.0018333 | 325.8 | 0.0028528 | [6 empty] | 0
602.842 | 0.0026667 | 325.9 | 0.0024881 | [6 empty] | 0
603.12 | 0.0015 | 326 | 0.0038315 | [6 empty] | 0
603.398 | 0.0033333 | 326.1 | 0.0042693 | [6 empty] | 0
603.676 | 0 | 326.2 | 0.004707 | [6 empty] | 0
603.954 | 0.00033333 | 326.3 | 0.0035497 | [6 empty] | 0
604.232 | 0.0026667 | 326.4 | 0.0034766 | [6 empty] | 0
604.51 | 0.0021667 | 326.5 | 0.0034035 | [6 empty] | 0
604.787 | 0.001 | 326.6 | 0.0033304 | [6 empty] | 0
605.065 | -0.0011667 | 326.7 | 0.0042295 | [6 empty] | 0
605.342 | 0.0011667 | 326.8 | 0.0041564 | [6 empty] | 0
605.62 | 0.0066667 | 326.9 | 0.0040833 | [6 empty] | 0
605.897 | 0.0025 | 327 | 0.0034678 | [6 empty] | 0
606.174 | -0.00016667 | 327.1 | 0.0035653 | [6 empty] | 0
606.452 | 0.00033333 | 327.2 | 0.0036628 | [6 empty] | 0
606.729 | 0.0013333 | 327.3 | 0.0037602 | [6 empty] | 0
607.006 | 0.0023333 | 327.4 | 0.0046355 | [6 empty] | 0
607.283 | 0.001 | 327.5 | 0.0047329 | [6 empty] | 0
607.56 | 0.0013333 | 327.6 | 0.0048304 | [6 empty] | 0
607.837 | -0.001 | 327.7 | 0.0044113 | [6 empty] | 0
608.114 | 0.0028333 | 327.8 | 0.0046062 | [6 empty] | 0
608.391 | 0.002 | 327.9 | 0.0048012 | [6 empty] | 0
608.667 | 0.0018333 | 328 | 0.0049961 | [6 empty] | 0
608.944 | 0 | 328.1 | 0.0041896 | [6 empty] | 0
609.221 | 0.0015 | 328.2 | 0.0038728 | [6 empty] | 0
609.497 | -0.002 | 328.3 | 0.003556 | [6 empty] | 0
609.774 | 0.00066667 | 328.4 | 0.0031559 | [6 empty] | 0
610.05 | -0.0028333 | 328.5 | 0.0028392 | [6 empty] | 0
610.326 | 0.0013333 | 328.6 | 0.0025224 | [6 empty] | 0
610.603 | -0.00083333 | 328.7 | 0.0017564 | [6 empty] | 0
610.879 | 0.0013333 | 328.8 | 0.0012196 | [6 empty] | 0
611.155 | 0.0026667 | 328.9 | 0.00068284 | [6 empty] | 0
611.431 | 0.0025 | 329 | 0.00014607 | [6 empty] | 0
611.707 | 0.0013333 | 329.1 | 0.0017556 | [6 empty] | 0
611.983 | 0.001 | 329.2 | 0.001878 | [6 empty] | 0
612.259 | 0.0031667 | 329.3 | 0.0020004 | [6 empty] | 0
612.535 | 0.0028333 | 329.4 | 0.0014569 | [6 empty] | 0
612.811 | -0.00083333 | 329.5 | 0.0016762 | [6 empty] | 0
613.087 | 0.00066667 | 329.6 | 0.0018955 | [6 empty] | 0
613.362 | -0.0011667 | 329.7 | 0.0021148 | [6 empty] | 0
613.638 | 0.003 | 329.8 | 0.0027536 | [6 empty] | 0
613.913 | 0.0005 | 329.9 | 0.0028509 | [6 empty] | 0
614.189 | 0.00066667 | 330 | 0.0029483 | [6 empty] | 0
614.464 | 0.001 | 330.1 | 0.003255 | [6 empty] | 0
614.74 | -0.00016667 | 330.2 | 0.0036704 | [6 empty] | 0
615.015 | -0.00033333 | 330.3 | 0.0040859 | [6 empty] | 0
615.29 | 0.00033333 | 330.4 | 0.0030899 | [6 empty] | 0
615.565 | 0.0026667 | 330.5 | 0.0027725 | [6 empty] | 0
615.84 | 0.0011667 | 330.6 | 0.002455 | [6 empty] | 0
616.115 | -0.00066667 | 330.7 | 0.0021376 | [6 empty] | 0
616.39 | 0.004 | 330.8 | 0.0034205 | [6 empty] | 0
616.665 | 0.0005 | 330.9 | 0.0034689 | [6 empty] | 0
616.94 | 0.0055 | 331 | 0.0035173 | [6 empty] | 0
617.215 | -0.00016667 | 331.1 | 0.0030528 | [6 empty] | 0
617.49 | 0.0013333 | 331.2 | 0.0032727 | [6 empty] | 0
617.764 | 0.0045 | 331.3 | 0.0034927 | [6 empty] | 0
618.039 | 0.001 | 331.4 | 0.0037126 | [6 empty] | 0
618.313 | 0.001 | 331.5 | 0.0031658 | [6 empty] | 0
618.588 | 0.002 | 331.6 | 0.002897 | [6 empty] | 0
618.862 | 0.00066667 | 331.7 | 0.0026281 | [6 empty] | 0
619.137 | 0.0016667 | 331.8 | 0.002494 | [6 empty] | 0
619.411 | 0.0015 | 331.9 | 0.0023473 | [6 empty] | 0
619.685 | 0.00066667 | 332 | 0.0022007 | [6 empty] | 0
619.959 | 0.0018333 | 332.1 | 0.0023328 | [6 empty] | 0
620.233 | 0.0031667 | 332.2 | 0.002284 | [6 empty] | 0
620.507 | 0.0015 | 332.3 | 0.0022351 | [6 empty] | 0
620.781 | 0.003 | 332.4 | 0.0021862 | [6 empty] | 0
621.055 | 0.0031667 | 332.5 | 0.0032362 | [6 empty] | 0
621.329 | 0.00033333 | 332.6 | 0.0036276 | [6 empty] | 0
621.603 | 0.0031667 | 332.7 | 0.004019 | [6 empty] | 0
621.876 | 0.0013333 | 332.8 | 0.0068541 | [6 empty] | 0
622.15 | 0.00066667 | 332.9 | 0.0083718 | [6 empty] | 0
622.424 | 0 | 333 | 0.0098896 | [6 empty] | 0
622.697 | 0.0011667 | 333.1 | 0.011407 | [6 empty] | 0
622.971 | 0.0015 | 333.2 | 0.021033 | [6 empty] | 0
623.244 | 0.0031667 | 333.3 | 0.025505 | [6 empty] | 0
623.517 | 0.001 | 333.4 | 0.029977 | [6 empty] | 0
623.791 | 0.001 | 333.5 | 0.040946 | [6 empty] | 0
624.064 | 0.0033333 | 333.6 | 0.048777 | [6 empty] | 0
624.337 | 0.002 | 333.7 | 0.056608 | [6 empty] | 0
624.61 | 0.0005 | 333.8 | 0.064439 | [6 empty] | 0
624.883 | 0.00033333 | 333.9 | 0.05442 | [6 empty] | 0
625.156 | 0 | 334 | 0.057183 | [6 empty] | 0
625.429 | 0.002 | 334.1 | 0.059945 | [6 empty] | 0
625.702 | 0.0015 | 334.2 | 0.040429 | [6 empty] | 0
625.974 | 0.00016667 | 334.3 | 0.032844 | [6 empty] | 0
626.247 | -0.014167 | 334.4 | 0.025259 | [6 empty] | 0
626.52 | -0.0005 | 334.5 | 0.022443 | [6 empty] | 0
626.792 | 0.00083333 | 334.6 | 0.014919 | [6 empty] | 0
627.065 | -0.00066667 | 334.7 | 0.007394 | [6 empty] | 0
627.337 | 0.0035 | 334.8 | -0.00013047 | [6 empty] | 0
627.61 | 0.002 | 334.9 | 0.005001 | [6 empty] | 0
627.882 | 0.00083333 | 335 | 0.0031907 | [6 empty] | 0
628.154 | -0.00033333 | 335.1 | 0.0013804 | [6 empty] | 0
628.427 | 0.0026667 | 335.2 | 0.0015889 | [6 empty] | 0
628.699 | 0.0013333 | 335.3 | 0.0012218 | [6 empty] | 0
628.971 | 0.0036667 | 335.4 | 0.00085477 | [6 empty] | 0
629.243 | 0.0023333 | 335.5 | 0.00048772 | [6 empty] | 0
629.515 | 0.002 | 335.6 | 0.0023144 | [6 empty] | 0
629.787 | 0.0043333 | 335.7 | 0.002584 | [6 empty] | 0
630.058 | 0.0018333 | 335.8 | 0.0028536 | [6 empty] | 0
630.33 | 0.0018333 | 335.9 | 0.0026085 | [6 empty] | 0
630.602 | 0.0028333 | 336 | 0.0030007 | [6 empty] | 0
630.874 | 0.00066667 | 336.1 | 0.0033928 | [6 empty] | 0
631.145 | 0.0025 | 336.2 | 0.002485 | [6 empty] | 0
631.417 | 0.0026667 | 336.3 | 0.0020438 | [6 empty] | 0
631.688 | -0.00033333 | 336.4 | 0.0016026 | [6 empty] | 0
631.959 | 0.0036667 | 336.5 | 0.0011614 | [6 empty] | 0
632.231 | 0.0026667 | 336.6 | 0.001165 | [6 empty] | 0
632.502 | -0.0018333 | 336.7 | 0.00079739 | [6 empty] | 0
632.773 | 0.0011667 | 336.8 | 0.00042974 | [6 empty] | 0
633.044 | 0.0013333 | 336.9 | 0.0013024 | [6 empty] | 0
633.316 | -0.00016667 | 337 | 0.0015232 | [6 empty] | 0
633.587 | 0.0025 | 337.1 | 0.0017439 | [6 empty] | 0
633.857 | 0.0011667 | 337.2 | 0.0019647 | [6 empty] | 0
634.128 | 0.0036667 | 337.3 | 0.0019831 | [6 empty] | 0
634.399 | -0.001 | 337.4 | 0.002253 | [6 empty] | 0
634.67 | 0.002 | 337.5 | 0.0025228 | [6 empty] | 0
634.941 | 0.0018333 | 337.6 | 0.0021943 | [6 empty] | 0
635.211 | 0.00016667 | 337.7 | 0.0021453 | [6 empty] | 0
635.482 | 0.0011667 | 337.8 | 0.0020962 | [6 empty] | 0
635.752 | 0.0026667 | 337.9 | 0.0027007 | [6 empty] | 0
636.023 | 0.0026667 | 338 | 0.0028969 | [6 empty] | 0
636.293 | -0.0021667 | 338.1 | 0.0030931 | [6 empty] | 0
636.564 | 0.0013333 | 338.2 | 0.0032893 | [6 empty] | 0
636.834 | 0.004 | 338.3 | 0.0024061 | [6 empty] | 0
637.104 | 0.0023333 | 338.4 | 0.0023568 | [6 empty] | 0
637.374 | 0 | 338.5 | 0.0023074 | [6 empty] | 0
637.644 | -0.0021667 | 338.6 | 0.00287 | [6 empty] | 0
637.914 | -0.0016667 | 338.7 | 0.0028206 | [6 empty] | 0
638.184 | 0.0045 | 338.8 | 0.0027713 | [6 empty] | 0
638.454 | 0.0026667 | 338.9 | 0.0027219 | [6 empty] | 0
638.724 | -0.002 | 339 | 0.0038201 | [6 empty] | 0
638.994 | 0.003 | 339.1 | 0.0043117 | [6 empty] | 0
639.264 | 0.0028333 | 339.2 | 0.0048033 | [6 empty] | 0
639.533 | -0.00066667 | 339.3 | 0.002625 | [6 empty] | 0
639.803 | -0.001 | 339.4 | 0.0021827 | [6 empty] | 0
640.072 | 0.0011667 | 339.5 | 0.0017404 | [6 empty] | 0
640.342 | 0.0021667 | 339.6 | 0.003081 | [6 empty] | 0
640.611 | -0.0021667 | 339.7 | 0.0029332 | [6 empty] | 0
640.881 | 0.0013333 | 339.8 | 0.0027853 | [6 empty] | 0
641.15 | 0.0023333 | 339.9 | 0.0026374 | [6 empty] | 0
641.419 | 0.00033333 | 340 | 0.0022348 | [6 empty] | 0
641.688 | 0.003 | 340.1 | 0.002456 | [6 empty] | 0
641.957 | 0.00033333 | 340.2 | 0.0026772 | [6 empty] | 0
642.227 | 0.00083333 | 340.3 | 0.0028166 | [6 empty] | 0
642.496 | 0.0016667 | 340.4 | 0.0026445 | [6 empty] | 0
642.764 | 0.0026667 | 340.5 | 0.0024725 | [6 empty] | 0
643.033 | 0.0023333 | 340.6 | 0.0028422 | [6 empty] | 0
643.302 | 0.00066667 | 340.7 | 0.0031372 | [6 empty] | 0
643.571 | 0.0025 | 340.8 | 0.0034322 | [6 empty] | 0
643.84 | 0.0011667 | 340.9 | 0.0037271 | [6 empty] | 0
644.108 | 0.0058333 | 341 | 0.0034523 | [6 empty] | 0
644.377 | 0.0011667 | 341.1 | 0.0035509 | [6 empty] | 0
644.645 | -0.0005 | 341.2 | 0.0036494 | [6 empty] | 0
644.914 | 0.00016667 | 341.3 | 0.0036604 | [6 empty] | 0
645.182 | 0.0011667 | 341.4 | 0.0036358 | [6 empty] | 0
645.45 | -0.00016667 | 341.5 | 0.0036112 | [6 empty] | 0
645.719 | -0.00083333 | 341.6 | 0.0035866 | [6 empty] | 0
645.987 | 0.0045 | 341.7 | 0.0035556 | [6 empty] | 0
646.255 | -0.003 | 341.8 | 0.0035556 | [6 empty] | 0
646.523 | 0.00033333 | 341.9 | 0.0035556 | [6 empty] | 0
646.791 | 0.00033333 | 342 | 0.0036111 | [6 empty] | 0
647.059 | 0.0026667 | 342.1 | 0.0036111 | [6 empty] | 0
647.327 | 0.00033333 | 342.2 | 0.0036111 | [6 empty] | 0
647.595 | 0.0013333 | 342.3 | 0.00315 | [6 empty] | 0
647.862 | -0.00016667 | 342.4 | 0.0029776 | [6 empty] | 0
648.13 | 0.0013333 | 342.5 | 0.0028051 | [6 empty] | 0
648.398 | 0.0033333 | 342.6 | 0.0026327 | [6 empty] | 0
648.665 | 0.00016667 | 342.7 | 0.001102 | [6 empty] | 0
648.933 | 0.0016667 | 342.8 | 0.00038818 | [6 empty] | 0
649.2 | 0.0005 | 342.9 | -0.0003256 | [6 empty] | 0
649.468 | 0.00066667 | 343 | 0.0011032 | [6 empty] | 0
649.735 | -0.0005 | 343.1 | 0.0010791 | [6 empty] | 0
650.002 | -0.0005 | 343.2 | 0.001055 | [6 empty] | 0
650.269 | 0.0015 | 343.3 | 0.0010309 | [6 empty] | 0
650.536 | -0.0005 | 343.4 | 0.0025442 | [6 empty] | 0
650.804 | 0.0018333 | 343.5 | 0.0033567 | [6 empty] | 0
651.071 | 0.0033333 | 343.6 | 0.0041692 | [6 empty] | 0
651.338 | 0.00066667 | 343.7 | 0.0028195 | [6 empty] | 0
651.604 | 0.0033333 | 343.8 | 0.0027949 | [6 empty] | 0
651.871 | 0.0023333 | 343.9 | 0.0027702 | [6 empty] | 0
652.138 | 0.0026667 | 344 | 0.0028956 | [6 empty] | 0
652.405 | 0.0013333 | 344.1 | 0.0026244 | [6 empty] | 0
652.671 | 0.0025 | 344.2 | 0.0023532 | [6 empty] | 0
652.938 | 0.0015 | 344.3 | 0.002082 | [6 empty] | 0
653.205 | -0.00016667 | 344.4 | 0.0017903 | [6 empty] | 0
653.471 | 0.0015 | 344.5 | 0.001667 | [6 empty] | 0
653.737 | 0.001 | 344.6 | 0.0015437 | [6 empty] | 0
654.004 | 0.0025 | 344.7 | 0.0018578 | [6 empty] | 0
654.27 | -0.001 | 344.8 | 0.0017839 | [6 empty] | 0
654.536 | 0.0016667 | 344.9 | 0.0017099 | [6 empty] | 0
654.803 | 0.00016667 | 345 | 0.0015575 | [6 empty] | 0
655.069 | 0.002 | 345.1 | 0.0016068 | [6 empty] | 0
655.335 | 0.0005 | 345.2 | 0.0016561 | [6 empty] | 0
655.601 | 0.0013333 | 345.3 | 0.0017055 | [6 empty] | 0
655.867 | 0.0021667 | 345.4 | 0.0023361 | [6 empty] | 0
656.132 | 0.0021667 | 345.5 | 0.0025087 | [6 empty] | 0
656.398 | 0 | 345.6 | 0.0026813 | [6 empty] | 0
656.664 | 0.002 | 345.7 | 0.0024303 | [6 empty] | 0
656.93 | 0.0026667 | 345.8 | 0.0025782 | [6 empty] | 0
657.195 | 0.00033333 | 345.9 | 0.0027262 | [6 empty] | 0
657.461 | -0.001 | 346 | 0.0028741 | [6 empty] | 0
657.726 | 0.00016667 | 346.1 | 0.0027108 | [6 empty] | 0
657.992 | 0.00016667 | 346.2 | 0.0026367 | [6 empty] | 0
658.257 | 0.0025 | 346.3 | 0.0025625 | [6 empty] | 0
658.523 | -0.001 | 346.4 | 0.0028984 | [6 empty] | 0
658.788 | 0.0013333 | 346.5 | 0.0030219 | [6 empty] | 0
659.053 | -0.00083333 | 346.6 | 0.0031454 | [6 empty] | 0
659.318 | 0.003 | 346.7 | 0.0020581 | [6 empty] | 0
659.583 | 0.0018333 | 346.8 | 0.0017127 | [6 empty] | 0
659.848 | 0.00066667 | 346.9 | 0.0013673 | [6 empty] | 0
660.113 | 0.0033333 | 347 | 0.0010219 | [6 empty] | 0
660.378 | 0.0025 | 347.1 | 0.0018105 | [6 empty] | 0
660.643 | -0.00033333 | 347.2 | 0.001638 | [6 empty] | 0
660.908 | 0.0025 | 347.3 | 0.0014655 | [6 empty] | 0
661.172 | -0.00083333 | 347.4 | 0.0013376 | [6 empty] | 0
661.437 | 0.001 | 347.5 | 0.001486 | [6 empty] | 0
661.702 | 0.00083333 | 347.6 | 0.0016343 | [6 empty] | 0
661.966 | -0.0005 | 347.7 | 0.0015511 | [6 empty] | 0
662.231 | 0.0013333 | 347.8 | 0.0014027 | [6 empty] | 0
662.495 | 0.0011667 | 347.9 | 0.0012544 | [6 empty] | 0
662.76 | -0.0023333 | 348 | 0.001106 | [6 empty] | 0
663.024 | 0.0028333 | 348.1 | 0.0019965 | [6 empty] | 0
663.288 | 0.0023333 | 348.2 | 0.0022438 | [6 empty] | 0
663.552 | 0.001 | 348.3 | 0.0024911 | [6 empty] | 0
663.816 | 0.0023333 | 348.4 | 0.0018405 | [6 empty] | 0
664.08 | 0.002 | 348.5 | 0.0018652 | [6 empty] | 0
664.344 | -0.00066667 | 348.6 | 0.00189 | [6 empty] | 0
664.608 | 0.0023333 | 348.7 | 0.0019147 | [6 empty] | 0
664.872 | 0.0028333 | 348.8 | 0.0025227 | [6 empty] | 0
665.136 | 0.0028333 | 348.9 | 0.0025475 | [6 empty] | 0
665.4 | 0.0005 | 349 | 0.0025722 | [6 empty] | 0
665.664 | 0.0015 | 349.1 | 0.0022211 | [6 empty] | 0
665.927 | 0.001 | 349.2 | 0.00232 | [6 empty] | 0
666.191 | 0.004 | 349.3 | 0.0024189 | [6 empty] | 0
666.454 | 0.00083333 | 349.4 | 0.0031978 | [6 empty] | 0
666.718 | 0.0011667 | 349.5 | 0.0033709 | [6 empty] | 0
666.981 | 0.001 | 349.6 | 0.003544 | [6 empty] | 0
667.244 | 0.0018333 | 349.7 | 0.0037171 | [6 empty] | 0
667.508 | 0.0018333 | 349.8 | 0.002551 | [6 empty] | 0
667.771 | 0.002 | 349.9 | 0.0024768 | [6 empty] | 0
668.034 | 0.0021667 | 350 | 0.0024027 | [6 empty] | 0
668.297 | 0.00016667 | 350.1 | 0.0022595 | [6 empty] | 0
668.56 | -0.0005 | 350.2 | 0.0018391 | [6 empty] | 0
668.823 | 0.0015 | 350.3 | 0.0014187 | [6 empty] | 0
669.086 | 0.004 | 350.4 | 0.0014426 | [6 empty] | 0
669.349 | 0.002 | 350.5 | 0.0014179 | [6 empty] | 0
669.612 | -0.00033333 | 350.6 | 0.0013931 | [6 empty] | 0
669.875 | 0.0036667 | 350.7 | 0.0013684 | [6 empty] | 0
670.137 | 0.0018333 | 350.8 | 0.0018828 | [6 empty] | 0
670.4 | 0.0046667 | 350.9 | 0.0020315 | [6 empty] | 0
670.662 | 0.0016667 | 351 | 0.0021802 | [6 empty] | 0
670.925 | -0.00016667 | 351.1 | 0.0015301 | [6 empty] | 0
671.187 | 0.0011667 | 351.2 | 0.001456 | [6 empty] | 0
671.45 | 0.001 | 351.3 | 0.0013819 | [6 empty] | 0
671.712 | 0.0033333 | 351.4 | 0.0013079 | [6 empty] | 0
671.974 | -0.0018333 | 351.5 | 0.0029815 | [6 empty] | 0
672.237 | 0.00016667 | 351.6 | 0.0033038 | [6 empty] | 0
672.499 | 0.0031667 | 351.7 | 0.0036262 | [6 empty] | 0
672.761 | 0.00083333 | 351.8 | 0.0029603 | [6 empty] | 0
673.023 | -0.00066667 | 351.9 | 0.0032579 | [6 empty] | 0
673.285 | 0.0011667 | 352 | 0.0035555 | [6 empty] | 0
673.547 | 0.00033333 | 352.1 | 0.0037468 | [6 empty] | 0
673.808 | 0.006 | 352.2 | 0.0036228 | [6 empty] | 0
674.07 | -0.0005 | 352.3 | 0.0034988 | [6 empty] | 0
674.332 | 0.00083333 | 352.4 | 0.0033748 | [6 empty] | 0
674.594 | 0.0023333 | 352.5 | 0.001847 | [6 empty] | 0
674.855 | 0.0018333 | 352.6 | 0.0014253 | [6 empty] | 0
675.117 | 0.0018333 | 352.7 | 0.0010037 | [6 empty] | 0
675.378 | 0.00083333 | 352.8 | 0.0012502 | [6 empty] | 0
675.64 | 0.0005 | 352.9 | 0.00077902 | [6 empty] | 0
675.901 | 0.0028333 | 353 | 0.00030779 | [6 empty] | 0
676.162 | 0.0016667 | 353.1 | 0.00081994 | [6 empty] | 0
676.424 | 0.0035 | 353.2 | 0.001068 | [6 empty] | 0
676.685 | 0.0016667 | 353.3 | 0.001316 | [6 empty] | 0
676.946 | 0.00066667 | 353.4 | 0.001564 | [6 empty] | 0
677.207 | 0.0021667 | 353.5 | 0.0025885 | [6 empty] | 0
677.468 | 0.0015 | 353.6 | 0.0031094 | [6 empty] | 0
677.729 | 0.0011667 | 353.7 | 0.0036302 | [6 empty] | 0
677.99 | 0.00083333 | 353.8 | 0.0031004 | [6 empty] | 0
678.251 | 0.002 | 353.9 | 0.0033237 | [6 empty] | 0
678.512 | 0.00016667 | 354 | 0.0035469 | [6 empty] | 0
678.772 | 0.0013333 | 354.1 | 0.0030821 | [6 empty] | 0
679.033 | 0.0011667 | 354.2 | 0.0027597 | [6 empty] | 0
679.294 | 0.0025 | 354.3 | 0.0024373 | [6 empty] | 0
679.554 | 0.0023333 | 354.4 | 0.0021148 | [6 empty] | 0
679.815 | 0.0016667 | 354.5 | 0.0024095 | [6 empty] | 0
680.075 | -0.00033333 | 354.6 | 0.0022855 | [6 empty] | 0
680.335 | -0.00083333 | 354.7 | 0.0021615 | [6 empty] | 0
680.596 | 0.0015 | 354.8 | 0.0026945 | [6 empty] | 0
680.856 | 0.0026667 | 354.9 | 0.002918 | [6 empty] | 0
681.116 | 0.0018333 | 355 | 0.0031415 | [6 empty] | 0
681.376 | 0.0021667 | 355.1 | 0.0029461 | [6 empty] | 0
681.636 | 0.0026667 | 355.2 | 0.0030453 | [6 empty] | 0
681.896 | 0.0018333 | 355.3 | 0.0031446 | [6 empty] | 0
682.156 | 0.0025 | 355.4 | 0.0032438 | [6 empty] | 0
682.416 | -0.0011667 | 355.5 | 0.0036214 | [6 empty] | 0
682.676 | 0.0023333 | 355.6 | 0.0037206 | [6 empty] | 0
682.936 | 0.0018333 | 355.7 | 0.0038198 | [6 empty] | 0
683.196 | 0.001 | 355.8 | 0.0028828 | [6 empty] | 0
683.455 | 0.0023333 | 355.9 | 0.0026837 | [6 empty] | 0
683.715 | 0 | 356 | 0.0024847 | [6 empty] | 0
683.974 | 0.0025 | 356.1 | 0.0022857 | [6 empty] | 0
684.234 | 0.0015 | 356.2 | 0.0028577 | [6 empty] | 0
684.493 | 0.0028333 | 356.3 | 0.002709 | [6 empty] | 0
684.753 | 0.0045 | 356.4 | 0.0025602 | [6 empty] | 0
685.012 | 0.0021667 | 356.5 | 0.0040007 | [6 empty] | 0
685.271 | -0.00016667 | 356.6 | 0.0046464 | [6 empty] | 0
685.53 | 0.002 | 356.7 | 0.0052921 | [6 empty] | 0
685.789 | 0.003 | 356.8 | 0.0045435 | [6 empty] | 0
686.049 | -0.00016667 | 356.9 | 0.0047177 | [6 empty] | 0
686.308 | 0.0025 | 357 | 0.0048918 | [6 empty] | 0
686.567 | -0.0005 | 357.1 | 0.0050659 | [6 empty] | 0
686.825 | 0.0005 | 357.2 | 0.0044312 | [6 empty] | 0
687.084 | 0.0023333 | 357.3 | 0.0041078 | [6 empty] | 0
687.343 | -0.00083333 | 357.4 | 0.0037844 | [6 empty] | 0
687.602 | 0.0013333 | 357.5 | 0.0027127 | [6 empty] | 0
687.86 | 0.0045 | 357.6 | 0.0021903 | [6 empty] | 0
688.119 | 0.0023333 | 357.7 | 0.0016679 | [6 empty] | 0
688.378 | 0.0011667 | 357.8 | 0.0025307 | [6 empty] | 0
688.636 | 0.0016667 | 357.9 | 0.0024063 | [6 empty] | 0
688.894 | 0.001 | 358 | 0.0022819 | [6 empty] | 0
689.153 | 0.0035 | 358.1 | 0.0021575 | [6 empty] | 0
689.411 | 0.0038333 | 358.2 | 0.002476 | [6 empty] | 0
689.669 | 0.0016667 | 358.3 | 0.0027745 | [6 empty] | 0
689.928 | 0.0011667 | 358.4 | 0.003073 | [6 empty] | 0
690.186 | 0.0041667 | 358.5 | 0.0040323 | [6 empty] | 0
690.444 | 0.012667 | 358.6 | 0.0044303 | [6 empty] | 0
690.702 | 0.017167 | 358.7 | 0.0048284 | [6 empty] | 0
690.96 | 0.0053333 | 358.8 | 0.0042528 | [6 empty] | 0
691.218 | 0.0028333 | 358.9 | 0.0044525 | [6 empty] | 0
691.476 | 0.00083333 | 359 | 0.0046521 | [6 empty] | 0
691.733 | 0.0035 | 359.1 | 0.0048517 | [6 empty] | 0
691.991 | 0.0045 | 359.2 | 0.0042589 | [6 empty] | 0
692.249 | 0.0016667 | 359.3 | 0.00396 | [6 empty] | 0
692.506 | 0.0018333 | 359.4 | 0.0036611 | [6 empty] | 0
692.764 | -0.001 | 359.5 | 0.0036645 | [6 empty] | 0
693.021 | 0.003 | 359.6 | 0.0035399 | [6 empty] | 0
693.279 | -0.0013333 | 359.7 | 0.0034153 | [6 empty] | 0
693.536 | 0.0021667 | 359.8 | 0.0036725 | [6 empty] | 0
693.794 | 0.0023333 | 359.9 | 0.0037224 | [6 empty] | 0
694.051 | 0.0036667 | 360 | 0.0037723 | [6 empty] | 0
694.308 | 0.0033333 | 360.1 | 0.0038222 | [6 empty] | 0
694.565 | -0.001 | 360.2 | 0.0040793 | [6 empty] | 0
694.822 | 0.0035 | 360.3 | 0.0042537 | [6 empty] | 0
695.079 | 0.0015 | 360.4 | 0.0044281 | [6 empty] | 0
695.336 | -0.00016667 | 360.5 | 0.0033713 | [6 empty] | 0
695.593 | 0.002 | 360.6 | 0.0030719 | [6 empty] | 0
695.85 | -0.0016667 | 360.7 | 0.0027725 | [6 empty] | 0
696.107 | 0.001 | 360.8 | 0.0030318 | [6 empty] | 0
696.364 | 0.001 | 360.9 | 0.002758 | [6 empty] | 0
696.62 | 0.0083333 | 361 | 0.0024842 | [6 empty] | 0
696.877 | 0.0025 | 361.1 | 0.0022104 | [6 empty] | 0
697.133 | -0.0038333 | 361.2 | 0.0038517 | [6 empty] | 0
697.39 | 0.00083333 | 361.3 | 0.0043996 | [6 empty] | 0
697.646 | 0.001 | 361.4 | 0.0049475 | [6 empty] | 0
697.903 | 0.00066667 | 361.5 | 0.0044531 | [6 empty] | 0
698.159 | 0.0011667 | 361.6 | 0.0047525 | [6 empty] | 0
698.415 | 0.0031667 | 361.7 | 0.0050519 | [6 empty] | 0
698.672 | -0.00066667 | 361.8 | 0.0055615 | [6 empty] | 0
698.928 | 0 | 361.9 | 0.0056613 | [6 empty] | 0
699.184 | 0.0036667 | 362 | 0.0057611 | [6 empty] | 0
699.44 | 0.0023333 | 362.1 | 0.0058609 | [6 empty] | 0
699.696 | 0.0021667 | 362.2 | 0.0055163 | [6 empty] | 0
699.952 | -0.00083333 | 362.3 | 0.0056161 | [6 empty] | 0
700.208 | 0.00016667 | 362.4 | 0.0057159 | [6 empty] | 0
700.464 | 0.0013333 | 362.5 | 0.005023 | [6 empty] | 0
700.719 | 0.0046667 | 362.6 | 0.0046482 | [6 empty] | 0
700.975 | -0.0031667 | 362.7 | 0.0042735 | [6 empty] | 0
701.231 | 0.001 | 362.8 | 0.0058543 | [6 empty] | 0
701.486 | 0.0011667 | 362.9 | 0.0063046 | [6 empty] | 0
701.742 | 0.0031667 | 363 | 0.0067548 | [6 empty] | 0
701.997 | -0.0015 | 363.1 | 0.007205 | [6 empty] | 0
702.253 | 0.00083333 | 363.2 | 0.0087038 | [6 empty] | 0
702.508 | -0.0005 | 363.3 | 0.0098515 | [6 empty] | 0
702.763 | 0.0026667 | 363.4 | 0.010999 | [6 empty] | 0
703.019 | -0.00066667 | 363.5 | 0.018342 | [6 empty] | 0
703.274 | -0.0018333 | 363.6 | 0.021789 | [6 empty] | 0
703.529 | 0.0023333 | 363.7 | 0.025236 | [6 empty] | 0
703.784 | 0.0046667 | 363.8 | 0.068936 | [6 empty] | 0
704.039 | 0.0041667 | 363.9 | 0.091433 | [6 empty] | 0
704.294 | -0.0005 | 364 | 0.11393 | [6 empty] | 0
704.549 | 0.0011667 | 364.1 | 0.13643 | [6 empty] | 0
704.804 | -0.0011667 | 364.2 | 0.2816 | [6 empty] | 0
705.058 | 0.0046667 | 364.3 | 0.35227 | [6 empty] | 0
705.313 | -0.001 | 364.4 | 0.42294 | [6 empty] | 0
705.568 | 0.0011667 | 364.5 | 0.49195 | [6 empty] | 0
705.822 | 0.00066667 | 364.6 | 0.5756 | [6 empty] | 0
706.077 | 0.0016667 | 364.7 | 0.65925 | [6 empty] | 0
706.331 | -0.0016667 | 364.8 | 0.59806 | [6 empty] | 0
706.586 | 0.005 | 364.9 | 0.60829 | [6 empty] | 0
706.84 | 0.0058333 | 365 | 0.61851 | [6 empty] | 0
707.095 | 0.005 | 365.1 | 0.62873 | [6 empty] | 0
707.349 | -0.0005 | 365.2 | 0.46161 | [6 empty] | 0
707.603 | 0.003 | 365.3 | 0.39027 | [6 empty] | 0
707.857 | 0.0021667 | 365.4 | 0.31892 | [6 empty] | 0
708.111 | 0.0026667 | 365.5 | 0.29873 | [6 empty] | 0
708.365 | 0.0026667 | 365.6 | 0.23569 | [6 empty] | 0
708.619 | 0.0031667 | 365.7 | 0.17266 | [6 empty] | 0
708.873 | 0.002 | 365.8 | 0.17403 | [6 empty] | 0
709.127 | -0.001 | 365.9 | 0.15669 | [6 empty] | 0
709.381 | 0.00066667 | 366 | 0.13934 | [6 empty] | 0
709.635 | -0.00033333 | 366.1 | 0.122 | [6 empty] | 0
709.888 | 0.003 | 366.2 | 0.11076 | [6 empty] | 0
710.142 | 0.00066667 | 366.3 | 0.099551 | [6 empty] | 0
710.396 | -0.002 | 366.4 | 0.08834 | [6 empty] | 0
710.649 | 0.0036667 | 366.5 | 0.068325 | [6 empty] | 0
710.903 | -0.0005 | 366.6 | 0.051784 | [6 empty] | 0
711.156 | 0.0028333 | 366.7 | 0.035242 | [6 empty] | 0
711.409 | 0.001 | 366.8 | 0.033934 | [6 empty] | 0
711.663 | 0 | 366.9 | 0.023724 | [6 empty] | 0
711.916 | 0.00033333 | 367 | 0.013514 | [6 empty] | 0
712.169 | 0.003 | 367.1 | 0.0033033 | [6 empty] | 0
712.422 | 0.0016667 | 367.2 | 0.0095571 | [6 empty] | 0
712.675 | 0.00083333 | 367.3 | 0.0076301 | [6 empty] | 0
712.928 | 0.0036667 | 367.4 | 0.0057032 | [6 empty] | 0
713.181 | 0.0045 | 367.5 | 0.0064533 | [6 empty] | 0
713.434 | 0.0021667 | 367.6 | 0.0060521 | [6 empty] | 0
713.687 | -0.00016667 | 367.7 | 0.005651 | [6 empty] | 0
713.94 | 0.0038333 | 367.8 | 0.006579 | [6 empty] | 0
714.192 | 0.0005 | 367.9 | 0.00693 | [6 empty] | 0
714.445 | 0.001 | 368 | 0.0072811 | [6 empty] | 0
714.698 | 0.003 | 368.1 | 0.0076322 | [6 empty] | 0
714.95 | 0.001 | 368.2 | 0.0061294 | [6 empty] | 0
715.203 | 0.004 | 368.3 | 0.0060045 | [6 empty] | 0
715.455 | -0.0005 | 368.4 | 0.0058797 | [6 empty] | 0
715.708 | 0.007 | 368.5 | 0.0062637 | [6 empty] | 0
715.96 | 0.0023333 | 368.6 | 0.0059633 | [6 empty] | 0
716.212 | 0.001 | 368.7 | 0.0056629 | [6 empty] | 0
716.464 | 0.0021667 | 368.8 | 0.0051041 | [6 empty] | 0
716.717 | -0.0015 | 368.9 | 0.0050292 | [6 empty] | 0
716.969 | 0 | 369 | 0.0049542 | [6 empty] | 0
717.221 | 0.0005 | 369.1 | 0.0048793 | [6 empty] | 0
717.473 | 0.0021667 | 369.2 | 0.0051811 | [6 empty] | 0
717.725 | 0.0015 | 369.3 | 0.0050556 | [6 empty] | 0
717.977 | 0.002 | 369.4 | 0.0049301 | [6 empty] | 0
718.228 | 0.0018333 | 369.5 | 0.0048672 | [6 empty] | 0
718.48 | 0.00016667 | 369.6 | 0.0049425 | [6 empty] | 0
718.732 | 0.0026667 | 369.7 | 0.0050178 | [6 empty] | 0
718.984 | 0.0018333 | 369.8 | 0.0065175 | [6 empty] | 0
719.235 | 0.0026667 | 369.9 | 0.0070943 | [6 empty] | 0
719.487 | 0.0026667 | 370 | 0.0076711 | [6 empty] | 0
719.738 | 0.0023333 | 370.1 | 0.0082479 | [6 empty] | 0
719.99 | 0.0011667 | 370.2 | 0.0064297 | [6 empty] | 0
720.241 | -0.0013333 | 370.3 | 0.0064804 | [6 empty] | 0
720.492 | -0.002 | 370.4 | 0.0065312 | [6 empty] | 0
720.744 | 0.00016667 | 370.5 | 0.0062784 | [6 empty] | 0
720.995 | 0.0028333 | 370.6 | 0.0057011 | [6 empty] | 0
721.246 | 0.00066667 | 370.7 | 0.0051238 | [6 empty] | 0
721.497 | 0.0036667 | 370.8 | 0.0060575 | [6 empty] | 0
721.748 | 0 | 370.9 | 0.0064092 | [6 empty] | 0
721.999 | 0.001 | 371 | 0.006761 | [6 empty] | 0
722.25 | 0.0045 | 371.1 | 0.0071128 | [6 empty] | 0
722.501 | 0.007 | 371.2 | 0.0053636 | [6 empty] | 0
722.752 | -0.0043333 | 371.3 | 0.0051369 | [6 empty] | 0
723.002 | 0.0018333 | 371.4 | 0.0049103 | [6 empty] | 0
723.253 | -0.0018333 | 371.5 | 0.0049695 | [6 empty] | 0
723.504 | 0.0015 | 371.6 | 0.0044926 | [6 empty] | 0
723.754 | -0.00016667 | 371.7 | 0.0040156 | [6 empty] | 0
724.005 | -0.00083333 | 371.8 | 0.0047299 | [6 empty] | 0
724.255 | -0.00033333 | 371.9 | 0.0050311 | [6 empty] | 0
724.506 | 0.0041667 | 372 | 0.0053323 | [6 empty] | 0
724.756 | 0.0036667 | 372.1 | 0.0056335 | [6 empty] | 0
725.006 | 0.0005 | 372.2 | 0.0062376 | [6 empty] | 0
725.257 | 0.0045 | 372.3 | 0.00669 | [6 empty] | 0
725.507 | 0.0015 | 372.4 | 0.0071425 | [6 empty] | 0
725.757 | 0.0028333 | 372.5 | 0.0066835 | [6 empty] | 0
726.007 | -0.0031667 | 372.6 | 0.006809 | [6 empty] | 0
726.257 | 0.0028333 | 372.7 | 0.0069346 | [6 empty] | 0
726.507 | -0.001 | 372.8 | 0.0063028 | [6 empty] | 0
726.757 | 0.0005 | 372.9 | 0.0059761 | [6 empty] | 0
727.007 | 0.003 | 373 | 0.0056494 | [6 empty] | 0
727.257 | 0.0013333 | 373.1 | 0.0053226 | [6 empty] | 0
727.506 | 0.0018333 | 373.2 | 0.006564 | [6 empty] | 0
727.756 | 0.0033333 | 373.3 | 0.0066899 | [6 empty] | 0
728.006 | 0.002 | 373.4 | 0.0068158 | [6 empty] | 0
728.255 | 0.0043333 | 373.5 | 0.006803 | [6 empty] | 0
728.505 | 0.0036667 | 373.6 | 0.0071045 | [6 empty] | 0
728.754 | 0.0011667 | 373.7 | 0.007406 | [6 empty] | 0
729.004 | 0.0041667 | 373.8 | 0.0068158 | [6 empty] | 0
729.253 | -0.00033333 | 373.9 | 0.0065896 | [6 empty] | 0
729.502 | 0.0041667 | 374 | 0.0063634 | [6 empty] | 0
729.751 | 0.001 | 374.1 | 0.0062197 | [6 empty] | 0
730.001 | 0.0028333 | 374.2 | 0.0059679 | [6 empty] | 0
730.25 | 0.0028333 | 374.3 | 0.0057162 | [6 empty] | 0
730.499 | -0.00033333 | 374.4 | 0.0054644 | [6 empty] | 0
730.748 | 0.004 | 374.5 | 0.0048088 | [6 empty] | 0
730.997 | -0.0011667 | 374.6 | 0.0044564 | [6 empty] | 0
731.246 | 0.0016667 | 374.7 | 0.0041039 | [6 empty] | 0
731.495 | 0.0046667 | 374.8 | 0.0041698 | [6 empty] | 0
731.743 | 0 | 374.9 | 0.0038928 | [6 empty] | 0
731.992 | 0 | 375 | 0.0036159 | [6 empty] | 0
732.241 | 0.003 | 375.1 | 0.0040121 | [6 empty] | 0
732.489 | 0.00066667 | 375.2 | 0.0041631 | [6 empty] | 0
732.738 | 0.00016667 | 375.3 | 0.0043142 | [6 empty] | 0
732.987 | -0.0011667 | 375.4 | 0.0044653 | [6 empty] | 0
733.235 | 0.0035 | 375.5 | 0.0053489 | [6 empty] | 0
733.483 | 0.0015 | 375.6 | 0.0058021 | [6 empty] | 0
733.732 | 0.0028333 | 375.7 | 0.0062553 | [6 empty] | 0
733.98 | 0.0036667 | 375.8 | 0.005199 | [6 empty] | 0
734.228 | 0.0011667 | 375.9 | 0.0051489 | [6 empty] | 0
734.477 | -0.0013333 | 376 | 0.0050988 | [6 empty] | 0
734.725 | 0.0056667 | 376.1 | 0.0054724 | [6 empty] | 0
734.973 | -0.0026667 | 376.2 | 0.0052962 | [6 empty] | 0
735.221 | 0.0026667 | 376.3 | 0.0051199 | [6 empty] | 0
735.469 | 0.0041667 | 376.4 | 0.0049437 | [6 empty] | 0
735.717 | 0.0025 | 376.5 | 0.0044912 | [6 empty] | 0
735.965 | 0.0031667 | 376.6 | 0.0044158 | [6 empty] | 0
736.212 | 0.0023333 | 376.7 | 0.0043403 | [6 empty] | 0
736.46 | -0.001 | 376.8 | 0.0050139 | [6 empty] | 0
736.708 | 0.00083333 | 376.9 | 0.0050394 | [6 empty] | 0
736.955 | 0.0018333 | 377 | 0.0050649 | [6 empty] | 0
737.203 | 0.0013333 | 377.1 | 0.004938 | [6 empty] | 0
737.451 | -0.001 | 377.2 | 0.0051399 | [6 empty] | 0
737.698 | -0.00016667 | 377.3 | 0.0053418 | [6 empty] | 0
737.945 | 0.0028333 | 377.4 | 0.0055437 | [6 empty] | 0
738.193 | 0.015667 | 377.5 | 0.005896 | [6 empty] | 0
738.44 | 0.013833 | 377.6 | 0.0060222 | [6 empty] | 0
738.687 | 0.0038333 | 377.7 | 0.0061484 | [6 empty] | 0
738.935 | 0.0026667 | 377.8 | 0.0054682 | [6 empty] | 0
739.182 | 0.0048333 | 377.9 | 0.005418 | [6 empty] | 0
739.429 | -0.001 | 378 | 0.0053678 | [6 empty] | 0
739.676 | 0.0015 | 378.1 | 0.0049222 | [6 empty] | 0
739.923 | 0.00066667 | 378.2 | 0.0045182 | [6 empty] | 0
740.17 | 0.00066667 | 378.3 | 0.0041141 | [6 empty] | 0
740.417 | 0.0018333 | 378.4 | 0.0042747 | [6 empty] | 0
740.663 | 0.0065 | 378.5 | 0.0041737 | [6 empty] | 0
740.91 | -0.0025 | 378.6 | 0.0040727 | [6 empty] | 0
741.157 | 0.0015 | 378.7 | 0.0039717 | [6 empty] | 0
741.403 | 0.0023333 | 378.8 | 0.0044624 | [6 empty] | 0
741.65 | 0 | 378.9 | 0.0046391 | [6 empty] | 0
741.897 | 0.0033333 | 379 | 0.0048159 | [6 empty] | 0
742.143 | -0.00083333 | 379.1 | 0.004046 | [6 empty] | 0
742.389 | 0.0015 | 379.2 | 0.0038944 | [6 empty] | 0
742.636 | 0.0018333 | 379.3 | 0.0037429 | [6 empty] | 0
742.882 | 0.0046667 | 379.4 | 0.0041535 | [6 empty] | 0
743.128 | -0.0005 | 379.5 | 0.0040525 | [6 empty] | 0
743.375 | -0.0005 | 379.6 | 0.0039515 | [6 empty] | 0
743.621 | 0.0025 | 379.7 | 0.0038505 | [6 empty] | 0
743.867 | 0.003 | 379.8 | 0.0042442 | [6 empty] | 0
744.113 | -0.0045 | 379.9 | 0.0044715 | [6 empty] | 0
744.359 | 0.0035 | 380 | 0.0046987 | [6 empty] | 0
744.605 | 0.0031667 | 380.1 | 0.004996 | [6 empty] | 0
744.851 | 0.0021667 | 380.2 | 0.005198 | [6 empty] | 0
745.097 | -0.0011667 | 380.3 | 0.0054 | [6 empty] | 0
745.342 | 0.0015 | 380.4 | 0.0048157 | [6 empty] | 0
745.588 | 0.0051667 | 380.5 | 0.00474 | [6 empty] | 0
745.834 | 0.0026667 | 380.6 | 0.0046643 | [6 empty] | 0
746.079 | 0.0015 | 380.7 | 0.0045886 | [6 empty] | 0
746.325 | 0.0046667 | 380.8 | 0.0050084 | [6 empty] | 0
746.57 | 0.0011667 | 380.9 | 0.004958 | [6 empty] | 0
746.816 | -0.0021667 | 381 | 0.0049076 | [6 empty] | 0
747.061 | 0.0056667 | 381.1 | 0.0053457 | [6 empty] | 0
747.307 | 0.001 | 381.2 | 0.0056239 | [6 empty] | 0
747.552 | 0.0045 | 381.3 | 0.0059021 | [6 empty] | 0
747.797 | 0.0031667 | 381.4 | 0.0057485 | [6 empty] | 0
748.042 | -0.00016667 | 381.5 | 0.0058243 | [6 empty] | 0
748.287 | -0.00066667 | 381.6 | 0.0059001 | [6 empty] | 0
748.532 | -0.00066667 | 381.7 | 0.0052616 | [6 empty] | 0
748.777 | 0.00083333 | 381.8 | 0.0049582 | [6 empty] | 0
749.022 | -0.0023333 | 381.9 | 0.0046549 | [6 empty] | 0
749.267 | 0.0015 | 382 | 0.0043515 | [6 empty] | 0
749.512 | 0.0021667 | 382.1 | 0.0048482 | [6 empty] | 0
749.757 | 0.0011667 | 382.2 | 0.0047215 | [6 empty] | 0
750.002 | 0.003 | 382.3 | 0.0045949 | [6 empty] | 0
750.246 | 0.0105 | 382.4 | 0.0044683 | [6 empty] | 0
750.491 | 0.0115 | 382.5 | 0.0045189 | [6 empty] | 0
750.736 | 0.0051667 | 382.6 | 0.0045696 | [6 empty] | 0
750.98 | 0.0058333 | 382.7 | 0.0050779 | [6 empty] | 0
751.225 | 0.0086667 | 382.8 | 0.0052045 | [6 empty] | 0
751.469 | 0.0195 | 382.9 | 0.0053312 | [6 empty] | 0
751.713 | 0.0085 | 383 | 0.0054578 | [6 empty] | 0
751.958 | -0.0036667 | 383.1 | 0.0051451 | [6 empty] | 0
752.202 | 0.0043333 | 383.2 | 0.0052464 | [6 empty] | 0
752.446 | 0.0075 | 383.3 | 0.0053478 | [6 empty] | 0
752.69 | -0.00016667 | 383.4 | 0.0047096 | [6 empty] | 0
752.934 | 0.001 | 383.5 | 0.0044056 | [6 empty] | 0
753.178 | 0.00066667 | 383.6 | 0.0041017 | [6 empty] | 0
753.422 | 0.0016667 | 383.7 | 0.0034726 | [6 empty] | 0
753.666 | 0.0018333 | 383.8 | 0.003016 | [6 empty] | 0
753.91 | -0.00016667 | 383.9 | 0.0025595 | [6 empty] | 0
754.154 | 0.0023333 | 384 | 0.0028304 | [6 empty] | 0
754.398 | 0.0018333 | 384.1 | 0.002653 | [6 empty] | 0
754.641 | 0.0011667 | 384.2 | 0.0024756 | [6 empty] | 0
754.885 | 0.0033333 | 384.3 | 0.0022982 | [6 empty] | 0
755.129 | 0.0028333 | 384.4 | 0.003407 | [6 empty] | 0
755.372 | -0.00016667 | 384.5 | 0.0038122 | [6 empty] | 0
755.616 | -0.0023333 | 384.6 | 0.0042175 | [6 empty] | 0
755.859 | 0.0045 | 384.7 | 0.0044005 | [6 empty] | 0
756.103 | -0.0033333 | 384.8 | 0.0048058 | [6 empty] | 0
756.346 | 0.0033333 | 384.9 | 0.0052111 | [6 empty] | 0
756.589 | -0.00033333 | 385 | 0.0048889 | [6 empty] | 0
756.832 | 0.00033333 | 385.1 | 0.004889 | [6 empty] | 0
757.076 | 0.0015 | 385.2 | 0.0048891 | [6 empty] | 0
757.319 | -0.0033333 | 385.3 | 0.0048892 | [6 empty] | 0
757.562 | 0.0033333 | 385.4 | 0.0047234 | [6 empty] | 0
757.805 | 0.003 | 385.5 | 0.0045966 | [6 empty] | 0
758.048 | -0.0026667 | 385.6 | 0.0044698 | [6 empty] | 0
758.291 | -0.0031667 | 385.7 | 0.0042594 | [6 empty] | 0
758.533 | -0.0005 | 385.8 | 0.0041326 | [6 empty] | 0
758.776 | 0.0076667 | 385.9 | 0.0040058 | [6 empty] | 0
759.019 | 0.0036667 | 386 | 0.0044041 | [6 empty] | 0
759.262 | -0.0031667 | 386.1 | 0.0044549 | [6 empty] | 0
759.504 | -0.00066667 | 386.2 | 0.0045058 | [6 empty] | 0
759.747 | 0.0033333 | 386.3 | 0.0050063 | [6 empty] | 0
759.989 | 0.0013333 | 386.4 | 0.0053869 | [6 empty] | 0
760.232 | -0.001 | 386.5 | 0.0057675 | [6 empty] | 0
760.474 | 0.0056667 | 386.6 | 0.006148 | [6 empty] | 0
760.717 | 0.004 | 386.7 | 0.005556 | [6 empty] | 0
760.959 | 0.0021667 | 386.8 | 0.0056322 | [6 empty] | 0
761.201 | 0.0043333 | 386.9 | 0.0057085 | [6 empty] | 0
761.444 | -0.00016667 | 387 | 0.0053971 | [6 empty] | 0
761.686 | -0.001 | 387.1 | 0.0051684 | [6 empty] | 0
761.928 | 0.0096667 | 387.2 | 0.0049398 | [6 empty] | 0
762.17 | -0.0028333 | 387.3 | 0.0050469 | [6 empty] | 0
762.412 | 0.00016667 | 387.4 | 0.0049961 | [6 empty] | 0
762.654 | -0.00033333 | 387.5 | 0.0049453 | [6 empty] | 0
762.896 | 0.004 | 387.6 | 0.0048945 | [6 empty] | 0
763.137 | 0.0046667 | 387.7 | 0.0047326 | [6 empty] | 0
763.379 | 0.037 | 387.8 | 0.0046817 | [6 empty] | 0
763.621 | 0.045333 | 387.9 | 0.0046309 | [6 empty] | 0
763.863 | 0.017667 | 388 | 0.0046912 | [6 empty] | 0
764.104 | 0.0021667 | 388.1 | 0.0046404 | [6 empty] | 0
764.346 | 0.0076667 | 388.2 | 0.0045896 | [6 empty] | 0
764.587 | 0.0031667 | 388.3 | 0.0046195 | [6 empty] | 0
764.829 | 0.001 | 388.4 | 0.0046449 | [6 empty] | 0
765.07 | 0.0025 | 388.5 | 0.0046703 | [6 empty] | 0
765.312 | 0.00016667 | 388.6 | 0.0047236 | [6 empty] | 0
765.553 | -0.0023333 | 388.7 | 0.004749 | [6 empty] | 0
765.794 | 0.0016667 | 388.8 | 0.0047744 | [6 empty] | 0
766.035 | -0.001 | 388.9 | 0.0047999 | [6 empty] | 0
766.276 | 0.0028333 | 389 | 0.0054002 | [6 empty] | 0
766.518 | -0.00083333 | 389.1 | 0.0056293 | [6 empty] | 0
766.759 | 0.00066667 | 389.2 | 0.0058584 | [6 empty] | 0
767 | 0.0018333 | 389.3 | 0.0057556 | [6 empty] | 0
767.241 | -0.00083333 | 389.4 | 0.0059334 | [6 empty] | 0
767.481 | 0.0031667 | 389.5 | 0.0061113 | [6 empty] | 0
767.722 | 0.005 | 389.6 | 0.0065398 | [6 empty] | 0
767.963 | 0.0045 | 389.7 | 0.0067178 | [6 empty] | 0
768.204 | 0.0021667 | 389.8 | 0.0068958 | [6 empty] | 0
768.444 | -0.0028333 | 389.9 | 0.0070738 | [6 empty] | 0
768.685 | 0.0013333 | 390 | 0.0078781 | [6 empty] | 0
768.926 | 0.0036667 | 390.1 | 0.0083361 | [6 empty] | 0
769.166 | -0.0013333 | 390.2 | 0.0087941 | [6 empty] | 0
769.407 | 0.005 | 390.3 | 0.0091246 | [6 empty] | 0
769.647 | 0.0046667 | 390.4 | 0.0095572 | [6 empty] | 0
769.887 | 0.001 | 390.5 | 0.0099898 | [6 empty] | 0
770.128 | 0.0048333 | 390.6 | 0.0077452 | [6 empty] | 0
770.368 | -0.004 | 390.7 | 0.0071081 | [6 empty] | 0
770.608 | -0.00066667 | 390.8 | 0.0064709 | [6 empty] | 0
770.848 | -0.00066667 | 390.9 | 0.0066817 | [6 empty] | 0
771.088 | -0.0018333 | 391 | 0.0059427 | [6 empty] | 0
771.328 | -0.0056667 | 391.1 | 0.0052036 | [6 empty] | 0
771.568 | 0.00033333 | 391.2 | 0.0044646 | [6 empty] | 0
771.808 | 0.003 | 391.3 | 0.0056962 | [6 empty] | 0
772.048 | 0.0016667 | 391.4 | 0.0059253 | [6 empty] | 0
772.288 | 0.0045 | 391.5 | 0.0061543 | [6 empty] | 0
772.528 | 0.018667 | 391.6 | 0.005526 | [6 empty] | 0
772.767 | 0.0076667 | 391.7 | 0.0054753 | [6 empty] | 0
773.007 | 0.0045 | 391.8 | 0.0054246 | [6 empty] | 0
773.247 | -0.001 | 391.9 | 0.0054447 | [6 empty] | 0
773.486 | 0.00016667 | 392 | 0.0052663 | [6 empty] | 0
773.726 | 0.0015 | 392.1 | 0.0050879 | [6 empty] | 0
773.965 | 0.0021667 | 392.2 | 0.0052822 | [6 empty] | 0
774.205 | 0.0056667 | 392.3 | 0.0053843 | [6 empty] | 0
774.444 | 0.0045 | 392.4 | 0.0054863 | [6 empty] | 0
774.683 | 0.0061667 | 392.5 | 0.0055884 | [6 empty] | 0
774.922 | 0.002 | 392.6 | 0.0051825 | [6 empty] | 0
775.162 | 0.00083333 | 392.7 | 0.0051313 | [6 empty] | 0
775.401 | 0.0036667 | 392.8 | 0.0050802 | [6 empty] | 0
775.64 | 0.0053333 | 392.9 | 0.0062115 | [6 empty] | 0
775.879 | -0.0023333 | 393 | 0.0065173 | [6 empty] | 0
776.118 | 0.0045 | 393.1 | 0.0068231 | [6 empty] | 0
776.357 | 0.0015 | 393.2 | 0.0054077 | [6 empty] | 0
776.596 | 0.0016667 | 393.3 | 0.0052548 | [6 empty] | 0
776.834 | -0.0051667 | 393.4 | 0.0051019 | [6 empty] | 0
777.073 | 0.0036667 | 393.5 | 0.004949 | [6 empty] | 0
777.312 | 0.0081667 | 393.6 | 0.0049414 | [6 empty] | 0
777.551 | 0.0058333 | 393.7 | 0.0044816 | [6 empty] | 0
777.789 | 0.0046667 | 393.8 | 0.0040218 | [6 empty] | 0
778.028 | 0.0026667 | 393.9 | 0.0053622 | [6 empty] | 0
778.266 | 0.003 | 394 | 0.0057962 | [6 empty] | 0
778.505 | 0.0025 | 394.1 | 0.0062302 | [6 empty] | 0
778.743 | 0.002 | 394.2 | 0.0047317 | [6 empty] | 0
778.982 | 0.001 | 394.3 | 0.0045024 | [6 empty] | 0
779.22 | 0.0036667 | 394.4 | 0.0042731 | [6 empty] | 0
779.458 | 0.0023333 | 394.5 | 0.0058605 | [6 empty] | 0
779.696 | 0.001 | 394.6 | 0.0060143 | [6 empty] | 0
779.934 | 0.0045 | 394.7 | 0.006168 | [6 empty] | 0
780.173 | 0.0026667 | 394.8 | 0.0063218 | [6 empty] | 0
780.411 | 0.0053333 | 394.9 | 0.0056136 | [6 empty] | 0
780.649 | 0.0075 | 395 | 0.0059204 | [6 empty] | 0
780.887 | 0.0016667 | 395.1 | 0.0062272 | [6 empty] | 0
781.124 | 0.0068333 | 395.2 | 0.0057508 | [6 empty] | 0
781.362 | 0.00066667 | 395.3 | 0.0053673 | [6 empty] | 0
781.6 | 0.0013333 | 395.4 | 0.0049839 | [6 empty] | 0
781.838 | 0.0046667 | 395.5 | 0.0049701 | [6 empty] | 0
782.075 | 0.0021667 | 395.6 | 0.0048934 | [6 empty] | 0
782.313 | 0.0031667 | 395.7 | 0.0048167 | [6 empty] | 0
782.551 | 0.0045 | 395.8 | 0.0047056 | [6 empty] | 0
782.788 | -0.010833 | 395.9 | 0.0045778 | [6 empty] | 0
783.026 | 0.0031667 | 396 | 0.00445 | [6 empty] | 0
783.263 | 0.0033333 | 396.1 | 0.0043222 | [6 empty] | 0
783.5 | 0.0075 | 396.2 | 0.0041666 | [6 empty] | 0
783.738 | 0.0018333 | 396.3 | 0.0040388 | [6 empty] | 0
783.975 | 0.0018333 | 396.4 | 0.003911 | [6 empty] | 0
784.212 | 0.00066667 | 396.5 | 0.004807 | [6 empty] | 0
784.449 | 0.0073333 | 396.6 | 0.0050371 | [6 empty] | 0
784.686 | -0.0031667 | 396.7 | 0.0052671 | [6 empty] | 0
784.924 | 0.0026667 | 396.8 | 0.0053385 | [6 empty] | 0
785.161 | 0.004 | 396.9 | 0.0056708 | [6 empty] | 0
785.397 | 0.00033333 | 397 | 0.0060032 | [6 empty] | 0
785.634 | 0.0038333 | 397.1 | 0.0057153 | [6 empty] | 0
785.871 | -0.0021667 | 397.2 | 0.0056386 | [6 empty] | 0
786.108 | 0.0048333 | 397.3 | 0.0055619 | [6 empty] | 0
786.345 | 0.0055 | 397.4 | 0.0054853 | [6 empty] | 0
786.581 | -0.0013333 | 397.5 | 0.0058546 | [6 empty] | 0
786.818 | 0.0038333 | 397.6 | 0.0058801 | [6 empty] | 0
787.055 | 0.001 | 397.7 | 0.0059055 | [6 empty] | 0
787.291 | -0.0015 | 397.8 | 0.006033 | [6 empty] | 0
787.528 | 0.0061667 | 397.9 | 0.0061866 | [6 empty] | 0
787.764 | 0.0028333 | 398 | 0.0063401 | [6 empty] | 0
788.001 | 0.001 | 398.1 | 0.0057123 | [6 empty] | 0
788.237 | 0.009 | 398.2 | 0.0055078 | [6 empty] | 0
788.473 | 0.004 | 398.3 | 0.0053033 | [6 empty] | 0
788.709 | 0.0016667 | 398.4 | 0.0059477 | [6 empty] | 0
788.946 | 0.0066667 | 398.5 | 0.0059986 | [6 empty] | 0
789.182 | 0.00033333 | 398.6 | 0.0060496 | [6 empty] | 0
789.418 | 0.0031667 | 398.7 | 0.0061005 | [6 empty] | 0
789.654 | 0.0041667 | 398.8 | 0.0056734 | [6 empty] | 0
789.89 | 0.0016667 | 398.9 | 0.0057503 | [6 empty] | 0
790.126 | 0.008 | 399 | 0.0058273 | [6 empty] | 0
790.362 | 0.0033333 | 399.1 | 0.0062922 | [6 empty] | 0
790.597 | 0.0031667 | 399.2 | 0.006318 | [6 empty] | 0
790.833 | 0.0033333 | 399.3 | 0.0063438 | [6 empty] | 0
791.069 | 0.0025 | 399.4 | 0.0058967 | [6 empty] | 0
791.305 | 0.0021667 | 399.5 | 0.0059225 | [6 empty] | 0
791.54 | 0.00083333 | 399.6 | 0.0059484 | [6 empty] | 0
791.776 | -0.0021667 | 399.7 | 0.0060479 | [6 empty] | 0
792.011 | 0.00033333 | 399.8 | 0.005894 | [6 empty] | 0
792.247 | 0.0026667 | 399.9 | 0.0057402 | [6 empty] | 0
792.482 | -0.0025 | 400 | 0.0055863 | [6 empty] | 0
792.717 | -0.00066667 | 400.1 | 0.0052308 | [6 empty] | 0
792.953 | -0.0028333 | 400.2 | 0.0051026 | [6 empty] | 0
793.188 | 0.0088333 | 400.3 | 0.0049744 | [6 empty] | 0
793.423 | 0.0018333 | 400.4 | 0.0053051 | [6 empty] | 0
793.658 | 0.0036667 | 400.5 | 0.0052538 | [6 empty] | 0
793.894 | 0.0031667 | 400.6 | 0.0052026 | [6 empty] | 0
794.129 | 0.00016667 | 400.7 | 0.0060829 | [6 empty] | 0
794.364 | 0.0055 | 400.8 | 0.0065444 | [6 empty] | 0
794.599 | -0.0046667 | 400.9 | 0.007006 | [6 empty] | 0
794.833 | 0.014167 | 401 | 0.0059949 | [6 empty] | 0
795.068 | 0.020667 | 401.1 | 0.0058923 | [6 empty] | 0
795.303 | 0.0071667 | 401.2 | 0.0057897 | [6 empty] | 0
795.538 | 0.0048333 | 401.3 | 0.0056872 | [6 empty] | 0
795.772 | 0.0085 | 401.4 | 0.0064179 | [6 empty] | 0
796.007 | 0.0071667 | 401.5 | 0.0063154 | [6 empty] | 0
796.242 | -0.0041667 | 401.6 | 0.0062128 | [6 empty] | 0
796.476 | 0.0026667 | 401.7 | 0.0060441 | [6 empty] | 0
796.711 | 0.0051667 | 401.8 | 0.0062241 | [6 empty] | 0
796.945 | 0.0096667 | 401.9 | 0.0064041 | [6 empty] | 0
797.18 | 0.0031667 | 402 | 0.0067936 | [6 empty] | 0
797.414 | 0.003 | 402.1 | 0.0068451 | [6 empty] | 0
797.648 | 0.0038333 | 402.2 | 0.0068967 | [6 empty] | 0
797.882 | -0.0051667 | 402.3 | 0.0079312 | [6 empty] | 0
798.117 | 0.010833 | 402.4 | 0.0085987 | [6 empty] | 0
798.351 | 0.0033333 | 402.5 | 0.0092663 | [6 empty] | 0
798.585 | 0.004 | 402.6 | 0.0099339 | [6 empty] | 0
798.819 | 0.0071667 | 402.7 | 0.011394 | [6 empty] | 0
799.053 | 0.0068333 | 402.8 | 0.012422 | [6 empty] | 0
799.287 | 0.0026667 | 402.9 | 0.013449 | [6 empty] | 0
799.521 | 0.0048333 | 403 | 0.017107 | [6 empty] | 0
799.755 | -0.0011667 | 403.1 | 0.019136 | [6 empty] | 0
799.988 | 0.011 | 403.2 | 0.021164 | [6 empty] | 0
800.222 | -0.0018333 | 403.3 | 0.050803 | [6 empty] | 0
800.456 | 0.0038333 | 403.4 | 0.064317 | [6 empty] | 0
800.689 | 0.0105 | 403.5 | 0.07783 | [6 empty] | 0
800.923 | 0.0061667 | 403.6 | 0.16477 | [6 empty] | 0
801.156 | 0.0115 | 403.7 | 0.21604 | [6 empty] | 0
801.39 | 0.0061667 | 403.8 | 0.26731 | [6 empty] | 0
801.623 | 0.013667 | 403.9 | 0.31858 | [6 empty] | 0
801.857 | 0.025667 | 404 | 0.4979 | [6 empty] | 0
802.09 | 0.016833 | 404.1 | 0.60168 | [6 empty] | 0
802.323 | 0.0035 | 404.2 | 0.70546 | [6 empty] | 0
802.557 | 0.0095 | 404.3 | 0.78013 | [6 empty] | 0
802.79 | 0.0066667 | 404.4 | 0.88322 | [6 empty] | 0
803.023 | -0.0025 | 404.5 | 0.9863 | [6 empty] | 0
803.256 | 0.0038333 | 404.6 | 0.78347 | [6 empty] | 0
803.489 | 0.0083333 | 404.7 | 0.75407 | [6 empty] | 0
803.722 | 0.0056667 | 404.8 | 0.72468 | [6 empty] | 0
803.955 | 0.0028333 | 404.9 | 0.55128 | [6 empty] | 0
804.188 | 0.0045 | 405 | 0.40822 | [6 empty] | 0
804.421 | 0.0093333 | 405.1 | 0.26517 | [6 empty] | 0
804.653 | 0.0045 | 405.2 | 0.12211 | [6 empty] | 0
804.886 | -0.0005 | 405.3 | 0.15501 | [6 empty] | 0
805.119 | 0.00066667 | 405.4 | 0.067328 | [6 empty] | 0
805.351 | 0.0096667 | 405.5 | -0.020352 | [6 empty] | 0
805.584 | -0.0018333 | 405.6 | 0.032607 | [6 empty] | 0
805.816 | 0.0055 | 405.7 | 0.019233 | [6 empty] | 0
806.049 | -0.0036667 | 405.8 | 0.005858 | [6 empty] | 0
806.281 | 0.0033333 | 405.9 | 0.011053 | [6 empty] | 0
806.514 | 0.012667 | 406 | 0.0099453 | [6 empty] | 0
806.746 | 0.0125 | 406.1 | 0.0088374 | [6 empty] | 0
806.978 | -0.0031667 | [8 empty] | 0
807.211 | 0.0065 | [8 empty] | 0
807.443 | 0.00016667 | [8 empty] | 0
807.675 | 0.0036667 | [8 empty] | 0
807.443 | 0
807.675 | 0

## Reflection Factor
Recommended set of data from Jim Bolton, personal communication, June 2016
Wavelength
| range | [11 empty] | Recommended absorption coefficients
| 170-197 nm | a1 = | 7.2264e-05 | b1 = | -0.0399 | c1 = | 7.1409 | d1 = | -412.8021 | [5 empty] | 1/10
| 198-206 nm |  |  | b2 = | 0.0071158 | c2 = | -2.9461 | d2 = | 302.1054 |  |  |  | wavelength |  | absorption
| 207-320 nm | a3 = | -2.921e-07 | b3 = | 0.00028708 | c3 = | -0.099558 | d3 = | 8.0526 |  |  |  | / nm | a / cm-1 | length / cm
| 320-400 nm | a4 = | 9.325e-07 | b4 = | -0.00088013 | c4 = | 0.26227 | d4 = | -28.3699 |  |  |  | 172 | 575.44 | 0.0017378
[5 empty] | 0.00063175 |  |  | fraction | 1/10 |  |  |  | 184.9 | 1.8017 | 0.55502
see hidden columns for calcuations
| wavelength | [6 empty] | absorbed in | absorption | Refractive | Reflection | Reflection | 253.7 | 0.00031814 | 3143.2527
| / nm | a / cm-1 | log(a) | log(a) fit | deviation | devation2 | UVT | 10 cm | length / cm | Index | coeficient | Factor
| 170 | 1174.9 | 3.07 | 3.0739 | -0.0039311 | 1.5454e-05 | 0% | 1 | 0.00085114
| 170.5 | 977.24 | 2.99 | 2.9933 | -0.0032729 | 1.0712e-05 | 0% | 1 | 0.0010233
| 171 | 831.76 | 2.92 | 2.9111 | 0.0088503 | 7.8328e-05 | 0% | 1 | 0.0012023 |  |  |  | For absorption coefficients from 170 to 190 nm:
| 171.5 | 691.83 | 2.84 | 2.8276 | 0.012392 | 0.00015357 | 0% | 1 | 0.0014454 |  |  |  | J. L. Weeks, G. M. A. C. Meaburn and S. Gordon, Absorption Coefficients of Liquid Water
| 172 | 575.44 | 2.76 | 2.7427 | 0.017294 | 0.00029907 | 0% | 1 | 0.0017378 |  |  |  | and aqueous solutions in the far ultraviolet, Radiation Research, 19(3), 559-567 (1963)
| 172.5 | 467.74 | 2.67 | 2.6565 | 0.013505 | 0.00018238 | 0% | 1 | 0.0021379
| 173 | 380.19 | 2.58 | 2.569 | 0.010959 | 0.00012011 | 0% | 1 | 0.0026303 |  |  |  | For absorption coefficients from 196 to 320 nm:
| 173.5 | 309.03 | 2.49 | 2.4804 | 0.0096152 | 9.2452e-05 | 0% | 1 | 0.0032359 |  |  |  | T. I. Quickenden and J. A. Irvin, "The ultraviolet absorption spectrum
| 174 | 251.19 | 2.4 | 2.3906 | 0.009416 | 8.8661e-05 | 6.457e-250% | 1 | 0.0039811 |  |  |  | of liquid water," J. Chem Phys., 72, 4416--4428, (1980).
| 174.5 | 204.17 | 2.31 | 2.2997 | 0.010294 | 0.00010596 | 6.761e-203% | 1 | 0.0048979
| 175 | 165.96 | 2.22 | 2.2078 | 0.012228 | 0.00014953 | 1.096e-164% | 1 | 0.0060255 |  |  |  | For absorption coefficients from 320 to 400 nm:
| 175.5 | 134.9 | 2.13 | 2.1149 | 0.015141 | 0.00022924 | 1.259e-133% | 1 | 0.0074129 |  |  |  | H. Buiteveld, J. M. H. Hakvoort and M. Donze, The optical properties of pure water,
| 176 | 107.15 | 2.03 | 2.021 | 0.0089513 | 8.0126e-05 | 7.079e-106% | 1 | 0.0093327 |  |  |  | in SPIE Proceedings on Ocean Optics XII, edited by J. S. Jaffe, 2258, 174--183, (1994).
| 176.5 | 87.1 | 1.94 | 1.9263 | 0.01368 | 0.00018714 | 7.943e-86% | 1 | 0.011481
| 177 | 67.61 | 1.83 | 1.8308 | -0.00080658 | 6.5056e-07 | 2.455e-66% | 1 | 0.014791
| 177.5 | 54.95 | 1.74 | 1.7345 | 0.0054348 | 2.9537e-05 | 1.122e-53% | 1 | 0.018198
| 178 | 44.67 | 1.65 | 1.6375 | 0.012477 | 0.00015569 | 2.138e-43% | 1 | 0.022386
| 178.5 | 34.67 | 1.54 | 1.5399 | 6.5173e-05 | 4.2475e-09 | 2.138e-33% | 1 | 0.028843
| 179 | 28.84 | 1.46 | 1.4416 | 0.018358 | 0.00033701 | 1.445e-27% | 1 | 0.034674
| 179.5 | 22.387 | 1.35 | 1.3428 | 0.0071567 | 5.1218e-05 | 4.102e-21% | 1 | 0.044669 |  |  |  | For refractive indices from 182 to 400 nm:
| 180 | 17.783 | 1.25 | 1.2435 | 0.006457 | 4.1693e-05 | 1.648e-16% | 1 | 0.056233 |  |  |  | M. Daimon and A. Masumura, "Measurement of the refractive index of distilled water from
| 180.5 | 14.125 | 1.15 | 1.1438 | 0.0061704 | 3.8074e-05 | 7.499e-13% | 1 | 0.070796 |  |  |  | the near-infrared region to the ultraviolet region", Applied Optics, 46(18), 3811-3820 (2007)
| 181 | 11.382 | 1.0562 | 1.0437 | 0.012515 | 0.00015662 | 4.15e-10% | 1 | 0.087858
| 181.5 | 8.913 | 0.95002 | 0.94326 | 0.0067651 | 4.5766e-05 | 1.222e-07% | 1 | 0.1122
| 182 | 6.761 | 0.83001 | 0.84254 | -0.012527 | 0.00015693 | 1.734e-05% | 1 | 0.14791 | 1.4677 | 0.035922 | 0.96408
| 182.5 | 5.37 | 0.72997 | 0.7416 | -0.011621 | 0.00013504 | 0.0004266% | 1 | 0.18622 | 1.4659 | 0.035699 | 0.9643
| 183 | 4.074 | 0.61002 | 0.64048 | -0.030464 | 0.00092803 | 0.008433% | 1 | 0.24546 | 1.4642 | 0.035482 | 0.96452
| 183.5 | 3.311 | 0.51996 | 0.53926 | -0.019301 | 0.00037254 | 0.04887% | 1 | 0.30202 | 1.4625 | 0.035271 | 0.96473
| 184 | 2.692 | 0.43008 | 0.43798 | -0.007902 | 6.2442e-05 | 0.2032% | 1 | 0.37147 | 1.4608 | 0.035067 | 0.96493
| 184.5 | 2.138 | 0.33001 | 0.33669 | -0.0066808 | 4.4633e-05 | 0.7278% | 1 | 0.46773 | 1.4592 | 0.034869 | 0.96513
| 184.9 | 1.8017 | 0.25569 | 0.25569 | 0 | 3.0815e-33 | 1.579% | 1 | 0.55502 | 1.458 | 0.034714 | 0.96529
| 185 | 1.66 | 0.22011 | 0.23545 | -0.015341 | 0.00023534 | 2.188% | 1 | 0.60241 | 1.4576 | 0.034676 | 0.96532
| 185.5 | 1.259 | 0.10003 | 0.13431 | -0.034287 | 0.0011756 | 5.508% | 1 | 0.79428 | 1.4561 | 0.034488 | 0.96551
| 186 | 1.023 | 0.0098756 | 0.033334 | -0.023458 | 0.00055029 | 9.484% | 1 | 0.97752 | 1.4546 | 0.034306 | 0.96569
| 186.5 | 0.813 | -0.089909 | -0.067433 | -0.022476 | 0.00050518 | 15.38% | 1 | 1.23 | 1.4532 | 0.034128 | 0.96587
| 187 | 0.661 | -0.1798 | -0.16793 | -0.011864 | 0.00014076 | 21.83% | 1 | 1.5129 | 1.4518 | 0.033954 | 0.96605
| 187.5 | 0.525 | -0.27984 | -0.26812 | -0.011725 | 0.00013748 | 29.85% | 0.99999 | 1.9048 | 1.4504 | 0.033786 | 0.96621
| 188 | 0.417 | -0.37986 | -0.36792 | -0.011941 | 0.0001426 | 38.28% | 0.99993 | 2.3981 | 1.4491 | 0.033621 | 0.96638
| 188.5 | 0.331 | -0.48017 | -0.4673 | -0.012871 | 0.00016566 | 46.67% | 0.99951 | 3.0211 | 1.4477 | 0.03346 | 0.96654
| 189 | 0.282 | -0.54975 | -0.5662 | 0.016446 | 0.00027047 | 52.24% | 0.99849 | 3.5461 | 1.4465 | 0.033303 | 0.9667
| 189.5 | 0.224 | -0.64975 | -0.66456 | 0.014804 | 0.00021916 | 59.7% | 0.99425 | 4.4643 | 1.4452 | 0.03315 | 0.96685
| 190 | 0.2 | -0.69897 | -0.76232 | 0.063354 | 0.0040137 | 63.1% | 0.99 | 5 | 1.444 | 0.033 | 0.967
| 196 | 0.0126 | -1.8996 | -1.8697 | -0.029919 | 0.00089514 | 97.14% | 0.25183 | 79.3651 | 1.431 | 0.031435 | 0.96857
| 197 | 0.00852 | -2.0696 | -2.0387 | -0.030895 | 0.00095449 | 98.06% | 0.17814 | 117.3709 | 1.4291 | 0.031208 | 0.96879
| 198 | 0.00567 | -2.2464 | -2.2554 | 0.0089439 | 7.9994e-05 | 98.7% | 0.12239 | 176.3668 | 1.4273 | 0.030991 | 0.96901
| 199 | 0.00411 | -2.3862 | -2.3765 | -0.009662 | 9.3353e-05 | 99.06% | 0.090296 | 243.309 | 1.4255 | 0.030781 | 0.96922
| 200 | 0.00324 | -2.4895 | -2.4834 | -0.006055 | 3.6663e-05 | 99.26% | 0.071889 | 308.642 | 1.4238 | 0.030578 | 0.96942
| 201 | 0.00264 | -2.5784 | -2.5761 | -0.0023239 | 5.4003e-06 | 99.39% | 0.058978 | 378.7879 | 1.4222 | 0.030382 | 0.96962
| 202 | 0.0023 | -2.6383 | -2.6545 | 0.016241 | 0.00026376 | 99.47% | 0.051582 | 434.7826 | 1.4206 | 0.030193 | 0.96981
| 203 | 0.00192 | -2.7167 | -2.7187 | 0.0020232 | 4.0934e-06 | 99.56% | 0.043247 | 520.8333 | 1.4191 | 0.03001 | 0.96999 | 0.029922
| 204 | 0.00168 | -2.7747 | -2.7687 | -0.0059912 | 3.5894e-05 | 99.61% | 0.037945 | 595.2381 | 1.4176 | 0.029833 | 0.97017
| 205 | 0.00156 | -2.8069 | -2.8044 | -0.0024299 | 5.9044e-06 | 99.64% | 0.035283 | 641.0256 | 1.4161 | 0.029662 | 0.97034
| 206 | 0.0015 | -2.8239 | -2.826 | 0.0020512 | 4.2073e-06 | 99.66% | 0.033949 | 666.6667 | 1.4147 | 0.029496 | 0.9705
| 207 | 0.0014 | -2.8539 | -2.8456 | -0.0082493 | 6.805e-05 | 99.68% | 0.031722 | 714.2857 | 1.4133 | 0.029335 | 0.97067
| 208 | 0.00134 | -2.8729 | -2.8638 | -0.0091231 | 8.323e-05 | 99.69% | 0.030383 | 746.2687 | 1.412 | 0.029179 | 0.97082 | 0.029103
| 209 | 0.00125 | -2.9031 | -2.8817 | -0.021378 | 0.00045702 | 99.71% | 0.028372 | 800 | 1.4107 | 0.029027 | 0.97097
| 210 | 0.00126 | -2.8996 | -2.8994 | -0.00018555 | 3.4431e-08 | 99.71% | 0.028596 | 793.6508 | 1.4095 | 0.02888 | 0.97112
| 211 | 0.00117 | -2.9318 | -2.917 | -0.014844 | 0.00022036 | 99.73% | 0.026581 | 854.7009 | 1.4082 | 0.028737 | 0.97126
| 212 | 0.00114 | -2.9431 | -2.9343 | -0.008804 | 7.751e-05 | 99.74% | 0.025908 | 877.193 | 1.4071 | 0.028598 | 0.9714
| 213 | 0.0011 | -2.9586 | -2.9514 | -0.0071973 | 5.1801e-05 | 99.75% | 0.02501 | 909.0909 | 1.4059 | 0.028462 | 0.97154 | 0.028396
| 214 | 0.00106 | -2.9747 | -2.9683 | -0.0063661 | 4.0528e-05 | 99.76% | 0.024112 | 943.3962 | 1.4048 | 0.028331 | 0.97167
| 215 | 0.00101 | -2.9957 | -2.985 | -0.010632 | 0.00011303 | 99.77% | 0.022988 | 990.099 | 1.4037 | 0.028202 | 0.9718
| 216 | 0.000988 | -3.0052 | -3.0016 | -0.0036746 | 1.3503e-05 | 99.77% | 0.022493 | 1012.1457 | 1.4026 | 0.028078 | 0.97192
| 217 | 0.000946 | -3.0241 | -3.0179 | -0.0062145 | 3.862e-05 | 99.78% | 0.021547 | 1057.0825 | 1.4015 | 0.027956 | 0.97204
| 218 | 0.000931 | -3.0311 | -3.034 | 0.0029761 | 8.8571e-06 | 99.79% | 0.021209 | 1074.1139 | 1.4005 | 0.027838 | 0.97216 | 0.02778
| 219 | 0.000918 | -3.0372 | -3.05 | 0.012809 | 0.00016407 | 99.79% | 0.020916 | 1089.3246 | 1.3995 | 0.027722 | 0.97228
| 220 | 0.000805 | -3.0942 | -3.0657 | -0.028488 | 0.00081157 | 99.81% | 0.018365 | 1242.236 | 1.3985 | 0.02761 | 0.97239
| 221 | 0.000786 | -3.1046 | -3.0813 | -0.0233 | 0.00054291 | 99.82% | 0.017936 | 1272.2646 | 1.3976 | 0.0275 | 0.9725
| 222 | 0.000737 | -3.1325 | -3.0967 | -0.035881 | 0.0012875 | 99.83% | 0.016827 | 1356.8521 | 1.3967 | 0.027393 | 0.97261
| 223 | 0.000714 | -3.1463 | -3.1118 | -0.034461 | 0.0011876 | 99.84% | 0.016306 | 1400.5602 | 1.3958 | 0.027288 | 0.97271 | 0.027237
| 224 | 0.0007 | -3.1549 | -3.1268 | -0.028056 | 0.00078713 | 99.84% | 0.015989 | 1428.5714 | 1.3949 | 0.027186 | 0.97281
| 225 | 0.00065 | -3.1871 | -3.1417 | -0.045416 | 0.0020626 | 99.85% | 0.014855 | 1538.4615 | 1.394 | 0.027086 | 0.97291
| 226 | 0.000669 | -3.1746 | -3.1563 | -0.018259 | 0.0003334 | 99.85% | 0.015286 | 1494.7683 | 1.3932 | 0.026989 | 0.97301
| 227 | 0.000643 | -3.1918 | -3.1708 | -0.021008 | 0.00044134 | 99.85% | 0.014697 | 1555.21 | 1.3923 | 0.026893 | 0.97311
| 228 | 0.000615 | -3.2111 | -3.1851 | -0.026054 | 0.00067881 | 99.86% | 0.014061 | 1626.0163 | 1.3915 | 0.0268 | 0.9732 | 0.026755
| 229 | 0.0006 | -3.2218 | -3.1992 | -0.022662 | 0.00051359 | 99.86% | 0.013721 | 1666.6667 | 1.3907 | 0.026709 | 0.97329
| 230 | 0.000595 | -3.2255 | -3.2131 | -0.012354 | 0.00015262 | 99.86% | 0.013607 | 1680.6723 | 1.3899 | 0.02662 | 0.97338
| 231 | 0.0005778 | -3.2382 | -3.2269 | -0.011322 | 0.00012819 | 99.87% | 0.013216 | 1730.7027 | 1.3892 | 0.026533 | 0.97347
| 232 | 0.000542 | -3.266 | -3.2405 | -0.025498 | 0.00065015 | 99.88% | 0.012402 | 1845.0185 | 1.3884 | 0.026448 | 0.97355
| 233 | 0.000534 | -3.2725 | -3.2539 | -0.018522 | 0.00034305 | 99.88% | 0.012221 | 1872.6592 | 1.3877 | 0.026364 | 0.97364 | 0.026323
| 234 | 0.000535 | -3.2716 | -3.2672 | -0.0044402 | 1.9715e-05 | 99.88% | 0.012243 | 1869.1589 | 1.387 | 0.026283 | 0.97372
| 235 | 0.000542 | -3.266 | -3.2803 | 0.01431 | 0.00020478 | 99.88% | 0.012402 | 1845.0185 | 1.3863 | 0.026203 | 0.9738
| 236 | 0.00051 | -3.2924 | -3.2933 | 0.00082348 | 6.7812e-07 | 99.88% | 0.011675 | 1960.7843 | 1.3856 | 0.026124 | 0.97388
| 237 | 0.000493 | -3.3072 | -3.306 | -0.0011179 | 1.2496e-06 | 99.89% | 0.011288 | 2028.3976 | 1.3849 | 0.026047 | 0.97395
| 238 | 0.000452 | -3.3449 | -3.3187 | -0.026203 | 0.00068661 | 99.9% | 0.010354 | 2212.3894 | 1.3842 | 0.025972 | 0.97403 | 0.025935
| 239 | 0.000471 | -3.327 | -3.3311 | 0.0041453 | 1.7184e-05 | 99.89% | 0.010787 | 2123.1423 | 1.3836 | 0.025899 | 0.9741
| 240 | 0.000483 | -3.3161 | -3.3434 | 0.027382 | 0.00074979 | 99.89% | 0.01106 | 2070.3934 | 1.383 | 0.025826 | 0.97417
| 241 | 0.000468 | -3.3298 | -3.3556 | 0.025838 | 0.00066762 | 99.89% | 0.010718 | 2136.7521 | 1.3823 | 0.025755 | 0.97424
| 242 | 0.000439 | -3.3575 | -3.3676 | 0.010062 | 0.00010125 | 99.9% | 0.010057 | 2277.9043 | 1.3817 | 0.025686 | 0.97431
| 243 | 0.000419 | -3.3778 | -3.3795 | 0.0016673 | 2.78e-06 | 99.9% | 0.0096014 | 2386.6348 | 1.3811 | 0.025618 | 0.97438 | 0.025584
| 244 | 0.000405 | -3.3925 | -3.3912 | -0.0013845 | 1.9169e-06 | 99.91% | 0.0092821 | 2469.1358 | 1.3805 | 0.025551 | 0.97445
| 245 | 0.000392 | -3.4067 | -3.4027 | -0.0039929 | 1.5943e-05 | 99.91% | 0.0089855 | 2551.0204 | 1.3799 | 0.025485 | 0.97451
| 246 | 0.000399 | -3.399 | -3.4141 | 0.01511 | 0.00022831 | 99.91% | 0.0091452 | 2506.2657 | 1.3794 | 0.025421 | 0.97458
| 247 | 0.000401 | -3.3969 | -3.4254 | 0.028554 | 0.00081534 | 99.91% | 0.0091909 | 2493.7656 | 1.3788 | 0.025358 | 0.97464
| 248 | 0.000392 | -3.4067 | -3.4365 | 0.029827 | 0.00088967 | 99.91% | 0.0089855 | 2551.0204 | 1.3783 | 0.025296 | 0.9747 | 0.025265
| 249 | 0.00039 | -3.4089 | -3.4475 | 0.038598 | 0.0014898 | 99.91% | 0.0089399 | 2564.1026 | 1.3777 | 0.025235 | 0.97477
| 250 | 0.000376 | -3.4248 | -3.4584 | 0.033575 | 0.0011273 | 99.91% | 0.0086203 | 2659.5745 | 1.3772 | 0.025175 | 0.97483
| 251 | 0.000336 | -3.4737 | -3.4691 | -0.004555 | 2.0748e-05 | 99.92% | 0.0077068 | 2976.1905 | 1.3767 | 0.025116 | 0.97488
| 252 | 0.000351 | -3.4547 | -3.4797 | 0.024997 | 0.00062484 | 99.92% | 0.0080495 | 2849.0028 | 1.3761 | 0.025059 | 0.97494
| 253 | 0.000359 | -3.4449 | -3.4901 | 0.045236 | 0.0020463 | 99.92% | 0.0082322 | 2785.5153 | 1.3756 | 0.025002 | 0.975 | 0.024982
| 253.7 | 0.00031814 | -3.4974 | -3.4974 | 0 | 1.9722e-31 | 99.93% | 0.0072987 | 3143.2527 | 1.3753 | 0.024963 | 0.97504
| 254 | 0.000348 | -3.4584 | -3.5005 | 0.042041 | 0.0017675 | 99.92% | 0.007981 | 2873.5632 | 1.3751 | 0.024946 | 0.97505
| 255 | 0.000326 | -3.4868 | -3.5107 | 0.023871 | 0.00056983 | 99.92% | 0.0074783 | 3067.4847 | 1.3746 | 0.024891 | 0.97511
| 256 | 0.000323 | -3.4908 | -3.5207 | 0.029921 | 0.00089524 | 99.93% | 0.0074098 | 3095.9752 | 1.3742 | 0.024838 | 0.97516
| 257 | 0.000323 | -3.4908 | -3.5307 | 0.03986 | 0.0015888 | 99.93% | 0.0074098 | 3095.9752 | 1.3737 | 0.024785 | 0.97522 | 0.024759
| 258 | 0.000306 | -3.5143 | -3.5405 | 0.026194 | 0.00068611 | 99.93% | 0.0070211 | 3267.9739 | 1.3732 | 0.024733 | 0.97527
| 259 | 0.000328 | -3.4841 | -3.5502 | 0.066039 | 0.0043612 | 99.92% | 0.007524 | 3048.7805 | 1.3728 | 0.024682 | 0.97532
| 260 | 0.000308 | -3.5114 | -3.5597 | 0.048289 | 0.0023318 | 99.93% | 0.0070669 | 3246.7532 | 1.3723 | 0.024631 | 0.97537
| 261 | 0.000331 | -3.4802 | -3.5692 | 0.089021 | 0.0079247 | 99.92% | 0.0075926 | 3021.148 | 1.3719 | 0.024582 | 0.97542
| 262 | 0.000279 | -3.5544 | -3.5785 | 0.024135 | 0.00058249 | 99.94% | 0.0064036 | 3584.2294 | 1.3714 | 0.024533 | 0.97547 | 0.024509
| 263 | 0.000285 | -3.5452 | -3.5878 | 0.042598 | 0.0018146 | 99.93% | 0.0065409 | 3508.7719 | 1.371 | 0.024485 | 0.97551
| 264 | 0.000273 | -3.5638 | -3.5969 | 0.033026 | 0.0010907 | 99.94% | 0.0062663 | 3663.0037 | 1.3706 | 0.024438 | 0.97556
| 265 | 0.000251 | -3.6003 | -3.6059 | 0.0055347 | 3.0633e-05 | 99.94% | 0.0057628 | 3984.0637 | 1.3702 | 0.024392 | 0.97561
| 266 | 0.000264 | -3.5784 | -3.6147 | 0.036353 | 0.0013216 | 99.94% | 0.0060604 | 3787.8788 | 1.3698 | 0.024346 | 0.97565
| 267 | 0.000236 | -3.6271 | -3.6235 | -0.0035583 | 1.2661e-05 | 99.95% | 0.0054194 | 4237.2881 | 1.3694 | 0.024301 | 0.9757 | 0.024279
| 268 | 0.000233 | -3.6326 | -3.6322 | -0.00044022 | 1.9379e-07 | 99.95% | 0.0053507 | 4291.8455 | 1.369 | 0.024257 | 0.97574
| 269 | 0.000239 | -3.6216 | -3.6408 | 0.019171 | 0.00036754 | 99.94% | 0.0054881 | 4184.1004 | 1.3686 | 0.024214 | 0.97579
| 270 | 0.000236 | -3.6271 | -3.6492 | 0.022153 | 0.00049073 | 99.95% | 0.0054194 | 4237.2881 | 1.3682 | 0.024171 | 0.97583
| 271 | 0.000234 | -3.6308 | -3.6576 | 0.026822 | 0.00071944 | 99.95% | 0.0053736 | 4273.5043 | 1.3678 | 0.024129 | 0.97587
| 272 | 0.000196 | -3.7077 | -3.6659 | -0.041871 | 0.0017532 | 99.95% | 0.0045029 | 5102.0408 | 1.3674 | 0.024087 | 0.97591 | 0.024067
| 273 | 0.000214 | -3.6696 | -3.674 | 0.0044564 | 1.986e-05 | 99.95% | 0.0049154 | 4672.8972 | 1.3671 | 0.024046 | 0.97595
| 274 | 0.000225 | -3.6478 | -3.6821 | 0.034299 | 0.0011764 | 99.95% | 0.0051674 | 4444.4444 | 1.3667 | 0.024006 | 0.97599
| 275 | 0.000216 | -3.6655 | -3.6901 | 0.02455 | 0.00060269 | 99.95% | 0.0049612 | 4629.6296 | 1.3663 | 0.023966 | 0.97603
| 276 | 0.000198 | -3.7033 | -3.698 | -0.0053514 | 2.8637e-05 | 99.95% | 0.0045487 | 5050.5051 | 1.366 | 0.023927 | 0.97607
| 277 | 0.000215 | -3.6676 | -3.7058 | 0.038219 | 0.0014607 | 99.95% | 0.0049383 | 4651.1628 | 1.3656 | 0.023889 | 0.97611 | 0.02387
| 278 | 0.000208 | -3.6819 | -3.7135 | 0.031552 | 0.00099554 | 99.95% | 0.0047779 | 4807.6923 | 1.3653 | 0.023851 | 0.97615
| 279 | 0.000209 | -3.6799 | -3.7211 | 0.041256 | 0.0017021 | 99.95% | 0.0048008 | 4784.689 | 1.3649 | 0.023813 | 0.97619
| 280 | 0.000222 | -3.6536 | -3.7286 | 0.074999 | 0.0056249 | 99.95% | 0.0050987 | 4504.5045 | 1.3646 | 0.023776 | 0.97622
| 281 | 0.000213 | -3.6716 | -3.7361 | 0.064479 | 0.0041575 | 99.95% | 0.0048925 | 4694.8357 | 1.3643 | 0.02374 | 0.97626
| 282 | 0.00018 | -3.7447 | -3.7435 | -0.0012572 | 1.5806e-06 | 99.96% | 0.0041361 | 5555.5556 | 1.364 | 0.023704 | 0.9763 | 0.023687
| 283 | 0.000177 | -3.752 | -3.7508 | -0.0012653 | 1.601e-06 | 99.96% | 0.0040673 | 5649.7175 | 1.3636 | 0.023669 | 0.97633
| 284 | 0.000185 | -3.7328 | -3.758 | 0.025146 | 0.00063233 | 99.96% | 0.0042507 | 5405.4054 | 1.3633 | 0.023634 | 0.97637
| 285 | 0.000178 | -3.7496 | -3.7651 | 0.015531 | 0.00024121 | 99.96% | 0.0040902 | 5617.9775 | 1.363 | 0.0236 | 0.9764
| 286 | 0.000175 | -3.757 | -3.7722 | 0.015211 | 0.00023137 | 99.96% | 0.0040214 | 5714.2857 | 1.3627 | 0.023566 | 0.97643
| 287 | 0.000164 | -3.7852 | -3.7792 | -0.0059943 | 3.5932e-05 | 99.96% | 0.0037691 | 6097.561 | 1.3624 | 0.023532 | 0.97647 | 0.023516
| 288 | 0.000168 | -3.7747 | -3.7861 | 0.011389 | 0.00012971 | 99.96% | 0.0038609 | 5952.381 | 1.3621 | 0.0235 | 0.9765
| 289 | 0.000159 | -3.7986 | -3.7929 | -0.0056749 | 3.2204e-05 | 99.96% | 0.0036544 | 6289.3082 | 1.3618 | 0.023467 | 0.97653
| 290 | 0.000163 | -3.7878 | -3.7997 | 0.011896 | 0.00014152 | 99.96% | 0.0037462 | 6134.9693 | 1.3615 | 0.023435 | 0.97657
| 291 | 0.00016 | -3.7959 | -3.8064 | 0.010543 | 0.00011117 | 99.96% | 0.0036774 | 6250 | 1.3612 | 0.023403 | 0.9766
| 292 | 0.000157 | -3.8041 | -3.8131 | 0.0089738 | 8.0529e-05 | 99.96% | 0.0036085 | 6369.4268 | 1.3609 | 0.023372 | 0.97663 | 0.023357
| 293 | 0.00016 | -3.7959 | -3.8197 | 0.023782 | 0.0005656 | 99.96% | 0.0036774 | 6250 | 1.3607 | 0.023341 | 0.97666
| 294 | 0.000151 | -3.821 | -3.8262 | 0.0051669 | 2.6697e-05 | 99.97% | 0.0034709 | 6622.5166 | 1.3604 | 0.023311 | 0.97669
| 295 | 0.000145 | -3.8386 | -3.8327 | -0.0059733 | 3.568e-05 | 99.97% | 0.0033332 | 6896.5517 | 1.3601 | 0.023281 | 0.97672
| 296 | 0.000144 | -3.8416 | -3.8391 | -0.0025673 | 6.5909e-06 | 99.97% | 0.0033102 | 6944.4444 | 1.3598 | 0.023251 | 0.97675
| 297 | 0.000123 | -3.9101 | -3.8454 | -0.064669 | 0.004182 | 99.97% | 0.0028282 | 8130.0813 | 1.3596 | 0.023222 | 0.97678 | 0.023208
| 298 | 0.000114 | -3.9431 | -3.8517 | -0.091366 | 0.0083478 | 99.97% | 0.0026215 | 8771.9298 | 1.3593 | 0.023193 | 0.97681
| 299 | 0.000124 | -3.9066 | -3.858 | -0.048599 | 0.0023618 | 99.97% | 0.0028511 | 8064.5161 | 1.359 | 0.023165 | 0.97684
| 300 | 0.000124 | -3.9066 | -3.8642 | -0.042398 | 0.0017976 | 99.97% | 0.0028511 | 8064.5161 | 1.3588 | 0.023137 | 0.97686
| 301 | 0.000146 | -3.8356 | -3.8703 | 0.034685 | 0.001203 | 99.97% | 0.0033561 | 6849.3151 | 1.3585 | 0.023109 | 0.97689
| 302 | 0.000108 | -3.9666 | -3.8764 | -0.090139 | 0.008125 | 99.98% | 0.0024837 | 9259.2593 | 1.3583 | 0.023082 | 0.97692 | 0.023068
| 303 | 0.000138 | -3.8601 | -3.8825 | 0.022377 | 0.00050074 | 99.97% | 0.0031725 | 7246.3768 | 1.358 | 0.023055 | 0.97695
| 304 | 0.000127 | -3.8962 | -3.8885 | -0.0076807 | 5.8992e-05 | 99.97% | 0.00292 | 7874.0157 | 1.3578 | 0.023028 | 0.97697
| 305 | 0.000112 | -3.9508 | -3.8945 | -0.05629 | 0.0031686 | 99.97% | 0.0025756 | 8928.5714 | 1.3576 | 0.023002 | 0.977
| 306 | 0.000122 | -3.9136 | -3.9004 | -0.013212 | 0.00017456 | 99.97% | 0.0028052 | 8196.7213 | 1.3573 | 0.022975 | 0.97702
| 307 | 0.000116 | -3.9355 | -3.9063 | -0.029215 | 0.00085353 | 99.97% | 0.0026674 | 8620.6897 | 1.3571 | 0.02295 | 0.97705 | 0.022937
| 308 | 0.000104 | -3.983 | -3.9122 | -0.070777 | 0.0050094 | 99.98% | 0.0023918 | 9615.3846 | 1.3568 | 0.022924 | 0.97708
| 309 | 0.000125 | -3.9031 | -3.918 | 0.014927 | 0.00022283 | 99.97% | 0.0028741 | 8000 | 1.3566 | 0.022899 | 0.9771
| 310 | 0.000112 | -3.9508 | -3.9238 | -0.026969 | 0.00072733 | 99.97% | 0.0025756 | 8928.5714 | 1.3564 | 0.022874 | 0.97713
| 311 | 0.00013 | -3.8861 | -3.9296 | 0.043521 | 0.0018941 | 99.97% | 0.0029889 | 7692.3077 | 1.3562 | 0.02285 | 0.97715
| 312 | 0.000119 | -3.9245 | -3.9353 | 0.01086 | 0.00011794 | 99.97% | 0.0027363 | 8403.3613 | 1.3559 | 0.022826 | 0.97717 | 0.022814
| 313 | 0.000113 | -3.9469 | -3.941 | -0.0059003 | 3.4813e-05 | 99.97% | 0.0025985 | 8849.5575 | 1.3557 | 0.022802 | 0.9772
| 314 | 0.000101 | -3.9957 | -3.9467 | -0.048975 | 0.0023985 | 99.98% | 0.0023229 | 9900.9901 | 1.3555 | 0.022778 | 0.97722
| 315 | 0.000105 | -3.979 | -3.9524 | -0.026637 | 0.00070955 | 99.98% | 0.0024148 | 9523.8095 | 1.3553 | 0.022755 | 0.97725
| 316 | 9.6e-05 | -4.0177 | -3.958 | -0.059729 | 0.0035676 | 99.98% | 0.002208 | 10416.6667 | 1.3551 | 0.022732 | 0.97727
| 317 | 0.000103 | -3.9872 | -3.9636 | -0.023547 | 0.00055447 | 99.98% | 0.0023689 | 9708.7379 | 1.3549 | 0.022709 | 0.97729 | 0.022698
| 318 | 9.3e-05 | -4.0315 | -3.9692 | -0.062304 | 0.0038818 | 99.98% | 0.0021391 | 10752.6882 | 1.3547 | 0.022686 | 0.97731
| 319 | 0.000108 | -3.9666 | -3.9748 | 0.008218 | 6.7535e-05 | 99.98% | 0.0024837 | 9259.2593 | 1.3545 | 0.022664 | 0.97734
| 320 | 0.0001 | -4 | -3.9804 | -0.01964 | 0.00038573 | 99.98% | 0.0022999 | 10000 | 1.3542 | 0.022642 | 0.97736
| 325 | 8.0562e-05 | -4.0939 | -4.0831 | -0.010803 | 0.00011671 | 99.98% | 0.0018533 | 12412.8576 | 1.3533 | 0.022535 | 0.97746
| 330 | 6.9487e-05 | -4.1581 | -4.1538 | -0.004259 | 1.8139e-05 | 99.98% | 0.0015987 | 14391.1568 | 1.3523 | 0.022435 | 0.97757
| 335 | 5.9933e-05 | -4.2223 | -4.2225 | 0.00011601 | 1.3458e-08 | 99.99% | 0.001379 | 16685.3992 | 1.3515 | 0.022339 | 0.97766
| 340 | 5.1681e-05 | -4.2867 | -4.2882 | 0.0015481 | 2.3965e-06 | 99.99% | 0.0011893 | 19349.4546 | 1.3506 | 0.022249 | 0.97775
| 345 | 4.5167e-05 | -4.3452 | -4.3504 | 0.0052475 | 2.7537e-05 | 99.99% | 0.0010395 | 22140.2413 | 1.3498 | 0.022163 | 0.97784
| 350 | 3.9521e-05 | -4.4032 | -4.4084 | 0.0052182 | 2.723e-05 | 99.99% | 0.00090959 | 25303.1329 | 1.3491 | 0.022082 | 0.97792
| 355 | 3.4744e-05 | -4.4591 | -4.4614 | 0.0022796 | 5.1964e-06 | 99.99% | 0.00079968 | 28782.3137 | 1.3484 | 0.022004 | 0.978
| 360 | 3.0835e-05 | -4.511 | -4.5088 | -0.0021885 | 4.7897e-06 | 99.99% | 0.00070975 | 32430.776 | 1.3477 | 0.021931 | 0.97807
| 365 | 2.8229e-05 | -4.5493 | -4.5498 | 0.00048141 | 2.3175e-07 | 99.99% | 0.00064979 | 35424.386 | 1.347 | 0.02186 | 0.97814
| 370 | 2.6058e-05 | -4.5841 | -4.5838 | -0.00031374 | 9.8432e-08 | 99.99% | 0.00059982 | 38376.4182 | 1.3464 | 0.021793 | 0.97821
| 375 | 2.4538e-05 | -4.6102 | -4.61 | -0.00019693 | 3.8781e-08 | 99.99% | 0.00056484 | 40753.7185 | 1.3458 | 0.021729 | 0.97827
| 380 | 2.3452e-05 | -4.6298 | -4.6277 | -0.0020787 | 4.3211e-06 | 99.99% | 0.00053985 | 42640.4647 | 1.3452 | 0.021668 | 0.97833
| 385 | 2.3018e-05 | -4.6379 | -4.6364 | -0.0015698 | 2.4644e-06 | 99.99% | 0.00052986 | 43445.0018 | 1.3447 | 0.021609 | 0.97839
| 390 | 2.3452e-05 | -4.6298 | -4.6352 | 0.0053294 | 2.8402e-05 | 99.99% | 0.00053985 | 42640.4647 | 1.3441 | 0.021553 | 0.97845
| 395 | 2.3669e-05 | -4.6258 | -4.6234 | -0.0024314 | 5.9118e-06 | 99.99% | 0.00054485 | 42249.2678 | 1.3436 | 0.021499 | 0.9785
| 400 | 2.5189e-05 | -4.5988 | -4.6004 | 0.0015923 | 2.5354e-06 | 99.99% | 0.00057983 | 39699.743 | 1.3432 | 0.021448 | 0.97855